## 표지
Park NX-Wafer
END USER: Samsung S4
SID No. : D160033-140725

## NX-Wafer
# | Module | Check Items | Min | Criteria | Max | Measurement | Unit | PASS/FAIL | Category | Trend | Remark
1 | Acoustic Enclosure | Check the specification of the nameplate (Refer to the drawing of layout and facility) |  | Yes |  | Yes | - | Pass | Configuration |  | 제원명판 제작 및 부착
2 | Acoustic Enclosure(Isolated AE) | AE type |  | Isolation |  | Isolation | - | Pass | Configuration |  | 분리형
2 | Acoustic Enclosure(Double Walled AE) | AE type |  | Double Walled |  |  | - |  | Configuration |  | 일체형
3 | AFM Controller | AFM Controller type |  | IND-M1 |  | IND-M1 | - | Pass | Configuration |  | Wafer/WLI : IND-M, IND-M1, Mask : IND-M1
4 | AFM Controller | Firmware version of Motor board |  | 0x6e31041e |  | 0x6e31041e | Version | Pass | Configuration
5 | AFM Controller | Firmware version of Base board |  | 0x6e11041e |  | 0x6e11041e | Version | Pass | Configuration
6 | AFM Controller | Line voltage selection |  | 230 |  | 230 | VAC | Pass | Configuration
7 | AFM Head | Check if the largest difference in PZT Alignment is less than 50um |  |  | 50 | 25 | um | Pass | Configuration |  | Head검사 성적서의 Z Scanner body & PZT Alignment 의 A/B, C/D, E/F, G/H의 차이가 50um이하인지 확인
8 | AFM Head | Check the warranty sticker |  | Yes |  | Yes | - | Pass | Configuration
9 | AFM Head | Z Detector offset |  |  | 200 | 19.75 | nm | Pass | Calibration
10 | AFM Head | AFM Head type |  | Auto standard |  | Auto standard | - | Pass | Configuration
11 | AFM Head | Probehand type |  | High Frequency |  | High Frequency | - | Pass | Configuration
12 | AFM Head | Beam intensity with AC160TS cantilever | 2.5 |  | 3.5 | 2.928 | V | Pass | Performance
13 | AFM Head | NCMDriveRangeLiquidMode |  | 1 |  | 1 | - | Pass | Configuration |  | Dsp/Info/NCMDriveRangeLiquidMode=1
14 | AFM Head | bDisableNCMDriveReduction |  | 1 |  | 1 | - | Pass | Configuration |  | Dsp/Info/bDisableNCMDriveReduction=1
15 | AFM Head | Check if the driving % changes proportionally as the NCM amplitude changes |  | Yes |  | Yes | - | Pass | Performance |  | Manual 참고, Amp. 1nm, 20nm, 100nm 일때 확인, 아래 값이 비례적으로 변하는지 확인한 후 Yes 진행
16 | AFM Head | NCM driving % at an amplitude of 1 nm |  | - |  | 0.06 | % | Pass | Performance |  | liquid mode 활성화 확인 필요 및 Amplitude 1nm로 일때의 Driving%, Capture 후 저장
17 | AFM Head | NCM driving % at an amplitude of 20 nm | 0.1 |  | 50 | 1.24 | % | Pass | Performance |  | liquid mode 활성화 확인 필요 및 Amplitude 20nm로 일때의 Driving%, Capture 후 저장
18 | AFM Head | NCM driving % at an amplitude of 100 nm |  | - |  | 6.18 | % | Pass | Performance |  | liquid mode 활성화 확인 필요 및 Amplitude 100nm로 일때의 Driving%, Capture 후 저장
19 | AFM Head | Specification of Step Height Sample for Z Linearity by Z scanner position test |  | - |  | 461.2 ± 2.8 | nm | Pass | Configuration |  | 아래 측정(Z Linearity by Z scanner position test) 에 사용하는 Sample(450nm) spec 기입 ex)450 ± 2.0nm
20 | AFM Head | Z Linearity by Z scanner position |  |  | 2 | 0.05 | % | Pass | Performance | Check item | Long head(2%) : Z scannner position -6,0,+6, Standard head(2%) : Z scannner position -3,0,+3
21 | AFM Head | Full range of Z Detector during Z Scan sweep | 10 |  | 18 | 12.72 | V | Pass | Performance | Check item | servo off
22 | AFM Head | Specification of Step Height Sample for Z Calibration | 458.4 | 461.2 | 464 | 461.2 | nm | Pass | Configuration |  | 아래 측정(Z Calibration) 에 사용하는 Sample(450nm) spec 값 min, max criteria에 기입
23 | AFM Head | Z Height calibration error |  |  | 0.5 | 0.073504 | % | Pass | Calibration | Check item
24 | AFM Head | - Z Height measurement value of the calibration sample |  | - |  | 461.539 | nm | Pass | Calibration
25 | AFM Head | Z Drive calibration error |  |  | 0.5 | 0.27342 | % | Pass | Calibration | Check item
26 | AFM Head | - Z Drive measurement value of the calibration sample |  | - |  | 462.461 | nm | Pass | Calibration
27 | AFM Head(NX-Wafer300mm) | Z Linearity by 4 different sample height | 0.98 |  | 1.02 | 0.99771 | - | Pass | Performance | Check item | 4 different Step Height Sample
27 | AFM Head(NX-Wafer300mm) | * 1800nm Step Height Sample (10 Repeats) | 1773 | 1783 | 1793 | 1782.584 | nm | Pass | Performance |  | 측정값 기입만 인쇄시 숨김처리, Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
28 | AFM Head(NX-Wafer300mm) | - 940nm Step Height Sample (10 Repeats) | 943 | 948.6 | 954.2 | 945.6755 | nm | Pass | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
29 | AFM Head(NX-Wafer300mm) | - 200~500nm Step Height Sample (10 Repeats) | 458.4 | 461.2 | 464 | 460.5411 | nm | Pass | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
30 | AFM Head(NX-Wafer300mm) | - 50~200nm Step Height Sample (10 Repeats) | 181.8 | 183.8 | 185.8 | 182.3668 | nm | Pass | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
31 | AFM Head(NX-Wafer300mm) | - Under 50nm Step Height Sample (10 Repeats) | 15.2 | 16.1 | 17 | 15.6328 | nm | Pass | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
31 | AFM Head(NX-Wafer200mm) | Z Linearity by 4 different sample height | 0.98 |  | 1.02 | #VALUE! | - | #VALUE! | Performance | Check item | 4 different Step Height Sample
31 | AFM Head(NX-Wafer200mm) | * 1800nm Step Height Sample (10 Repeats) |  |  |  |  | nm | - | Performance |  | 측정값 기입만 인쇄시 숨김처리, Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
31 | AFM Head(NX-Wafer200mm) | - 940nm Step Height Sample (10 Repeats) |  |  |  |  | nm |  | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
31 | AFM Head(NX-Wafer200mm) | - 200~500nm Step Height Sample (10 Repeats) |  |  |  |  | nm |  | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
31 | AFM Head(NX-Wafer200mm) | - 50~200nm Step Height Sample (10 Repeats) |  |  |  |  | nm |  | Performance |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입
31 | AFM Head(NX-Wafer200mm) | - Under 50nm Step Height Sample (10 Repeats) |  |  |  |  | nm |  | Performance
31 | AFM Head(Long) | Full range of Z Height during Z Scan sweep | 39 |  | 46 |  | um |  | Performance | Check item | servo off
31 | AFM Head(Long) | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 43 |  | 92 |  | - |  | Calibration | Check item
31 | AFM Head(Long) | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 24 |  | 28 |  | - |  | Calibration | Check item
31 | AFM Head(Long) | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) |  |  |  |  | - |  | Calibration | Check item
31 | AFM Head(Long) | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) |  |  |  |  | - |  | Calibration | Check item
31 | AFM Head(Long) | Z Height noise |  |  | 0.1 |  | nm |  | Performance | Check item
32 | AFM Head(Standard) | Full range of Z Height during Z Scan sweep | 15 |  |  | 17.8718 | um | Pass | Performance | Check item | servo off
33 | AFM Head(Standard) | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23 |  | 35 | 28.1004 | - | Pass | Calibration | Check item
34 | AFM Head(Standard) | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 9 |  | 14 | 11.6031 | - | Pass | Calibration | Check item
35 | AFM Head(Standard) | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) |  |  |  | 28.1227 | - | Pass | Calibration | Check item
36 | AFM Head(Standard) | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) |  |  |  | 11.671 | - | Pass | Calibration | Check item
37 | AFM Head(Standard) | Z Height noise | 0.02 |  | 0.05 | 0.03 | nm | Pass | Performance | Check item
38 | AFM Head(Auto Align) | A-B after Auto PSPD Alignment | -0.5 |  | 0.5 | 0.004 | V | Pass | Performance
39 | ATX | Check if the balls are attaching in each slot |  | Yes |  | Yes | - | Pass | Configuration
40 | ATX | Check if the magnet can move up and down using a software control |  | Yes |  | Yes | - | Pass | Configuration
41 | ATX | Check for CDA leakage |  | Yes |  | Yes | - | Pass | Configuration
42 | ATX | Check teaching position of all slots |  | Yes |  | Yes | - | Pass | Configuration
43 | ATX | Check Tip checker position |  | Yes |  | Yes | - | Pass | Configuration
44 | ATX | Check Probe hand checker |  | Yes |  | Yes | - | Pass | Configuration
45 | ATX(Auto Align Head) | A＋B difference during 3 times run test | -0.25 |  | 0.25 | 0.14 | V | Pass | Performance
46 | ATX(Auto Align Head) | A－B difference during 3 times run test | -1 |  | 1 | 0.97 | V | Pass | Performance
47 | Computer | Ethernet card type |  | Non-PoE |  | Non-PoE | - | Pass | Configuration
48 | Computer | Disable power saving mode in OS and BIOS |  | Yes |  | Yes | - | Pass | Configuration
49 | Computer | Enable RAID |  | Yes |  | Yes | - | Pass | Configuration
50 | Computer | RAID monitor software |  | Installed |  | Installed | - | Pass | Configuration
50 | Computer | Anti-virus software |  | Installed |  |  | - |  | Configuration
50 | Computer | Check for the latest update on anti-virus software |  | Yes |  |  | - |  | Configuration |  | anti virus software가 최신 업데이트 상태여야함
50 | Computer | Check for the number of viruses detected after the virus scan |  | 0 |  |  | - |  | Configuration |  | scan된 virus 개수 확인, virus가 0개여야 함
51 | Computer | Dongle for pattern recognition |  | Installed |  | Installed | - | Pass | Configuration
52 | Computer | Windows version |  | 24H2 |  | 24H2 | Version | Pass | Configuration |  | Window OS 선택하면 Version 알아서 바뀜, window10 -> 22h2, window11 -> 24h2
53 | Computer | Disable Windows update |  | Yes |  | Yes | - | Pass | Configuration
54 | Computer | Save computer configuration file |  | Yes |  | Yes | - | Pass | Configuration
55 | Computer | Check ethernet configuration for AFM Controller |  | Yes |  | Yes | - | Pass | Configuration |  | Address:192.168.0.1, Subnet mask: 255.255.255.0
56 | Computer | Check ethernet configuration for SlidingStage Controller |  | Yes |  | Yes | - | Pass | Configuration |  | Address:169.254.154.1, Subnet mask: 255.255.255.0
57 | Computer | Check ethernet configuration for on-axis camera |  | Yes |  | Yes | - | Pass | Configuration |  | Address:169.254.100.1, Subnet mask: 255.255.255.0
58 | Computer | OS(Operating System) type |  | Window11 64bit |  | Window11 64bit | - | Pass | Configuration
59 | Computer | OS / MS-office language |  | Korean |  | Korean | - | Pass | Configuration
60 | Computer | Check DotNET Framework Version | 4.6 |  |  | 4.8.1 | Version | Pass | Configuration |  | Version 4.6 이상인지 확인, ([PST] Environmental Check 프로그램 사용하여 확인)
61 | Computer | Check the installation status of the display refresh program |  | Yes |  | Yes | - | Pass | Configuration |  | Display refresh 프로그램 설치 확인
62 | Computer(HP) | Computer model |  | HP Z4 G5 Workstation |  | HP Z4 G5 Workstation | - | Pass | Configuration
62 | eAFM | Selectable amplifier gain |  | 106 to 109 |  |  | V/A |  | Configuration
62 | eAFM | Firmware version of FEM |  | 0x6E450421 |  |  | - |  | Configuration
62 | eAFM | Check LED lighting of FEM when switching modes (C-AFM, On / EFM, Off) |  | Yes |  |  | - |  | Configuration
62 | eAFM | Check add eAFM Mode to Head Mode |  | Yes |  |  | - |  | Configuration
62 | eAFM | Check the Bias Voltage (SampleChuck, EFM&C-AFM Tip Initialization sample at Sample Holder) |  | Yes |  |  | - |  | Configuration
62 | eAFM | C-AFM combo probehand for NX |  | Included |  |  | - |  | Configuration
62 | eAFM | EFM&C-AFM Tip Initialization sample (HOPG, Gold Tip 2ea) |  | Included |  |  | - |  | Configuration
62 | eAFM | AD-42-AS, 12M |  | Included |  |  | - |  | Configuration
62 | eAFM | CDT-FMR, 12M |  | Included |  |  | - |  | Configuration
62 | eAFM(Conductive AFM) | Noise level of current signal |  |  | 0.4 |  | pA |  | Performance | Check item
62 | eAFM(Conductive AFM) | Maximum value of current signal at 106 V/A (10V, Sample Bias) | 0.9 | 1 | 1.1 |  | μA |  | Performance | Check item
62 | eAFM(Conductive AFM) | Maximum value of current signal at 109 V/A (100mV, Sample Bias) | 9.9 | 10 | 10.1 |  | nA |  | Performance | Check item
62 | eAFM(Conductive AFM) | Resistance value of 10MΩ | 9 | 10 | 11 |  | MΩ |  | Performance | Check item
62 | eAFM(Conductive AFM) | Resistance between GND to Probehand |  |  | 0.5 |  | Ω |  | Performance | Check item
62 | eAFM(Conductive AFM) | Save Conductive AFM Image (Cobalt-Doped Iron Oxide Sample) |  | Yes |  |  | - |  | Configuration
62 | eAFM(EFM) | Check if EFM amplitude shows the magnitude of electrical signal |  | Yes |  |  | - |  | Configuration
62 | eAFM(EFM) | Check if the EFM phase shows 180° phase difference for the positive and negative bias |  | Yes |  |  | - |  | Configuration
62 | eAFM(EFM) | EFM Phase : For the positive bias | 89 | 90 | 91 |  | degree |  | Performance | Check item
62 | eAFM(EFM) | EFM Phase : For the negative bias | -91 | -90 | -89 |  | degree |  | Performance | Check item
62 | eAFM(EFM) | Lock-In configuration : Input channel |  | Vertical(A-B) |  |  | - |  | Configuration
62 | eAFM(EFM) | Lock-In configuration : Phase |  |  |  |  | degree | - | Performance | Check item
62 | eAFM(EFM) | Lock-In configuration : Time constant |  |  |  |  | msec | - | Performance | Check item
62 | eAFM(EFM) | Lock-In configuration : Sensitivity |  |  |  |  | V | - | Performance | Check item
62 | eAFM(EFM) | Lock-In configuration : Filter order |  | 2nd Order |  |  | - |  | Configuration
62 | eAFM(EFM) | Modulator Configuration : Output signal |  | Tip Bias |  |  | - |  | Configuration
62 | eAFM(EFM) | Modulator Configuration : Frequency |  |  |  |  | kHz | - | Performance | Check item
62 | eAFM(EFM) | Modulator Configuration : Drive |  |  |  |  | V | - | Performance | Check item
62 | eAFM(EFM) | Save EFM Image (EFM Stripe-patterned Surface Sample) |  | Yes |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Tip Initializer DB setting : TipInitialiizerSupport |  | 1 |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Tip Initializer stage position teach / XEA (EFM) |  | Teached |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Tip Initializer stage position teach / XEA (DC C-AFM) |  | Teached |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Tip Initializer stage position teach / XEA (AC C-AFM) |  | Teached |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Tip Initializer auto focus teach / XEA |  | Teached |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Check the Run Tip Initializer (EFM, Step1) / XEA |  | Yes |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Check the Run Tip Initializer (EFM, Step2) / XEA |  | Yes |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Check the Run Tip Initializer (DC, C-AFM) / XEA |  | Yes |  |  | - |  | Configuration
62 | eAFM(Tip Initializer) | Check the Run Tip Initializer (AC, C-AFM) / XEA |  | Yes |  |  | - |  | Configuration
63 | EFEM | Check if the SID on the EFEM QA report matches the SID on the equipment nameplate |  | Yes |  | Yes | - | Pass | Configuration |  | EFEM 업체 QA 리포트의 SID와 실제 장비의 명판 SID와 같은지 확인
64 | EFEM | Check if there is marking for robot initial position |  | Yes |  | Yes | - | Pass | Configuration
65 | EFEM | Check the safety interlock |  | Yes |  | Yes | - | Pass | Configuration
66 | EFEM | Check the load port indicator lamp |  | Yes |  | Yes | - | Pass | Configuration
67 | EFEM | Check clearance of robot and end effector against other module |  | Yes |  | Yes | - | Pass | Configuration |  | Loading 시 간섭여부 Check
68 | EFEM | Check the cassette ID reader |  | Yes |  | Yes | - | Pass | Configuration
69 | EFEM | Check the slot mapping |  | Yes |  | Yes | - | Pass | Configuration
70 | EFEM | Check load and unload from all slot |  | Yes |  | Yes | - | Pass | Configuration
71 | EFEM | Torque of fixing screws for End-Effector |  | 5 |  | 5 | kgf·cm | Pass | Configuration
72 | EFEM | Check the light curtain detection |  | Yes |  | Yes | - | Pass | Configuration
73 | EFEM | EFEM CDA pressure | 0.4 |  | 0.6 | 0.533 | Mpa | Pass | Configuration
74 | EFEM | EFEM vacuum pressure | -90 |  | -60 | -75 | kpa | Pass | Configuration
75 | EFEM | Indexer SW version |  |  |  | 5.2.4 Build54 | Version | Pass | Configuration
76 | EFEM | Check doubled wafer detection |  | Yes |  | Yes | - | Pass | Performance |  | Brooks robot인경우 안될수 있음
77 | EFEM | Check crossed wafer detection |  | Yes |  | Yes | - | Pass | Performance |  | Brooks robot인경우 안될수 있음
78 | EFEM | Wafer placement(Loading) error on the load port1 (X offset) |  |  | 100 | 18.45 | um | Pass | Performance | Check item | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고
79 | EFEM | Wafer placement(Loading) error on the load port1 (Y offset) |  |  | 100 | 31.32 | um | Pass | Performance | Check item | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고
80 | EFEM | Wafer placement(Loading) error on the load port2 (X offset) |  |  | 100 | 29.55 | um | Pass | Performance | Check item | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고, loadport2 없을시 삭제
81 | EFEM | Wafer placement(Loading) error on the load port2 (Y offset) |  |  | 100 | 43.52 | um | Pass | Performance | Check item | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고, loadport2 없을시 삭제
82 | EFEM | Cleanness of EFEM w/ airborne particle measurement |  | ISO1 |  | ISO1 | - | Pass | Performance
83 | EFEM | Check Aligner changes angle reference according to software configuration |  | Yes |  | Yes | - | Pass | Configuration |  | 0deg,90deg,180deg,270deg
84 | EFEM | Angle offset value of pre-aligner |  | - |  | 359.9 | degree | Pass | Performance |  | EFEM-Alinger 로딩 후 틀어진 각 조정하여 저장한 값 기입(XEA EFEM- Setup - Align Angle offset)
84 | EFEM(Bluecord aligner) | Centering accuracy of pre-aligner |  |  | 0.05 |  | degree |  | Performance
84 | EFEM(Logosol aligner) | Centering accuracy of pre-aligner |  |  | 0.04 |  | degree |  | Performance
85 | EFEM(Yaskawa aligner) | Centering accuracy of pre-aligner |  |  | 0.03 | 0.004 | degree | Pass | Performance
86 | EFEM | Air flow speed | 0.3 |  |  | 0.37 | m/s | Pass | Performance |  | 측정위치 : EFEM 로봇 origin position의 end effector 윗면으로 부터 50cm 위
87 | EFEM | Check enable pre-load |  | Yes |  | Yes | - | Pass | Configuration |  | 출하전, XEA에서 Test
88 | EFEM | Level to X axis on the load port 1 |  |  |  | 0.186 | degree | Pass | Configuration |  | 측정값기입
89 | EFEM | Level to Y axis on the load port 1 |  |  |  | -0.289 | degree | Pass | Configuration |  | 측정값기입
90 | EFEM | Level to X axis on the robot body | -0.1 |  | 0.1 | -0.01 | degree | Pass | Configuration |  | 검사성적서 값 기입
91 | EFEM | Level to Y axis on the robot body | -0.1 |  | 0.1 | 0 | degree | Pass | Configuration |  | 검사성적서 값 기입
92 | EFEM | Level to X axis on the robot upper arm | -0.1 |  | 0.1 | 0.04 | degree | Pass | Configuration |  | 검사성적서 값 기입
93 | EFEM | Level to Y axis on the robot upper arm | -0.1 |  | 0.1 | 0.08 | degree | Pass | Configuration |  | 검사성적서 값 기입
94 | EFEM | Level to X axis on the robot lower arm | -0.1 |  | 0.1 | 0.07 | degree | Pass | Configuration |  | 검사성적서 값 기입
95 | EFEM | Level to Y axis on the robot lower arm | -0.1 |  | 0.1 | 0.03 | degree | Pass | Configuration |  | 검사성적서 값 기입
96 | EFEM | Serial port number for the load port1 |  | 5 |  | 5 | - | Pass | Configuration
97 | EFEM | Measurement of surface resistance of the upper end-effector | 10 6 |  | 10 9 | 7.59 x 107 | Ω | Pass | Configuration |  | Pad type End-effector 만 PAD 표면저항 측정
98 | EFEM | Measurement of surface resistance of the lower end-effector | 10 6 |  | 10 9 | 3.51 x 108 | Ω | Pass | Configuration |  | Pad type End-effector 만 PAD 표면저항 측정
98 | EFEM(Single Load port) | Check the load button on the load port1 of the EFEM for any abnormalities |  | Yes |  |  | - |  | Configuration |  | EFEM Load port의 Load button 눌러서 indexer s/w에 event 발생하는 화면 캡쳐
98 | EFEM(Single Load port) | Check 250 cycle of wafer load and unload without error |  | Yes |  |  | - |  | Performance |  | Single LoadPort인 경우 250times
99 | EFEM(Dual Load port) | Check 500 cycle of wafer load and unload without error |  | Yes |  | Yes | - | Pass | Performance |  | YMTC 경우, LoadPort 별 500 times(기본은 LoadPort 별 250times)
100 | EFEM(Dual Load port) | Serial port number for the load port2 |  | 6 |  | 6 | - | Pass | Configuration
101 | EFEM(Dual Load port) | Level to X axis on the load port 2 |  |  |  | 0.135 | degree | Pass | Configuration |  | 측정값기입
102 | EFEM(Dual Load port) | Level to Y axis on the load port 2 |  |  |  | -0.298 | degree | Pass | Configuration |  | 측정값기입
103 | EFEM(Dual Load port) | Check the load button on the load port1 of the EFEM for any abnormalities |  | Yes |  | Yes | - | Pass | Configuration |  | EFEM Load port의 Load button 눌러서 indexer s/w에 event 발생하는 화면 캡쳐
104 | EFEM(Dual Load port) | Check the load button on the load port2 of the EFEM for any abnormalities |  | Yes |  | Yes | - | Pass | Configuration |  | EFEM Load port의 Load button 눌러서 indexer s/w에 event 발생하는 화면 캡쳐
104 | EFEM(FFU Louver type) | Air pressure @ 100% of FFU | 1 |  |  |  | Pa |  | Performance
105 | EFEM(FFU Open type) | Air pressure @ 60% of FFU | 1.25 |  |  | 1.72 | Pa | Pass | Performance
105 | EFEM(BCR) | Serial Port # |  | (선택) |  |  | - |  | Configuration |  | Dual loadport인 경우 2,4
105 | EFEM(BCR) | BCR(Bar Code Reader) teach |  | Installed |  |  | - |  | Configuration
105 | EFEM(OCR) | OCR(Sample ID Reader) Score | (선택) |  |  |  | - |  | Configuration |  | ioss인 경우 : 120 , cognex인 경우 : 198
106 | EFEM(RFID) | Serial Port # |  | 2&4 |  | 2&4 | - | Pass | Configuration |  | Dual loadport인 경우 2,4
107 | EFEM(RFID) | RFID type |  | LF |  | LF | - | Pass | Configuration
108 | EFEM(TDK loadport) | Install maintenance software for TDK loadport |  | Yes |  | Yes | - | Pass | Configuration |  | TDK loadport인 경우
109 | EFEM(Yaskawa) | Firmware backup for Yaskawa robot |  | Yes |  | Yes | - | Pass | Configuration |  | Yaskawa Robot인 경우
110 | EFEM(Yaskawa) | Firmware version of Yaskawa robot | 1.0.3 |  |  | 3.0.3 | Version | Pass | Configuration |  | Yaskawa Robot인 경우, 티칭팬던트에서 확인
111 | Facility Controller | Check facility power cable connection |  | Yes |  | Yes | - | Pass | Configuration
112 | Facility Controller | Check main Disconnect Switch On/Off function |  | Yes |  | Yes | - | Pass | Configuration
113 | Facility Controller | Check AE door sensor |  | Yes |  | Yes | - | Pass | Configuration
114 | Facility Controller | Check on and off of Ionizer |  | Yes |  | Yes | - | Pass | Configuration
115 | Facility Controller | Check on and off of Suction |  | Yes |  | Yes | - | Pass | Configuration
116 | Facility Controller | Check CCTV for EFEM and AFM |  | Yes |  | Yes | - | Pass | Configuration
117 | Facility Controller | Check vibration amplifier |  | Yes |  | Yes | - | Pass | Configuration
118 | Facility Controller | Check vibration sensors |  | Yes |  | Yes | - | Pass | Configuration
119 | Facility Controller | Check A-B interlock alarm |  | Yes |  | Yes | - | Pass | Configuration |  | Z stage 구동시, A-B값이 ±3V초과시 Alarm 발생확인
120 | Facility Controller | Check A+B Interlock alarm |  | Yes |  | Yes | - | Pass | Configuration |  | DB : ApproachSetPointRatio = 0.6 작을때 interlock이 걸리는지
121 | Facility Controller | Check sample sensing on the sample chuck by vacuum pressure |  | Yes |  | Yes | - | Pass | Configuration
122 | Facility Controller | Check A/E lamp on and off by door sensor |  | Yes |  | Yes | - | Pass | Configuration |  | Door sensor로 Lamp ON/OFF 확인
123 | Facility Controller | Check AVIS controller power on and off |  | Yes |  | Yes | - | Pass | Configuration
124 | Facility Controller | Check AVIS controller enable w/o red LED flashing (While XY stage stopped) |  | Yes |  | Yes | - | Pass | Configuration
125 | Facility Controller | Check signal tower for photo ionizer |  | Yes |  | Yes | - | Pass | Configuration
126 | Facility Controller | Check photo ionizer head toward to sample and AFM Head |  | Yes |  | Yes | - | Pass | Configuration
127 | Facility Controller | No power trip during power reset or usage |  | No |  | No | - | Pass | Configuration
128 | Facility Controller | EMO cover |  | NA |  | NA | - | Pass | Configuration |  | 중국향 장비는 Install, 아닌경우 NA
129 | Facility Controller | Signal Tower Color |  | RYG |  | RYG | - | Pass | Configuration |  | 사용색상 선택, 순서상관 없음
130 | Facility Controller | Ionizer Peak Power |  | 5 |  | 5 | KeV | Pass | Configuration |  | 사용 Power 선택
131 | Facility Controller | BaudRate |  | 9600 |  | 9600 | - | Pass | Configuration
132 | Facility Controller | Upper Temperature Limit |  | 80 |  | 80 | - | Pass | Configuration
133 | Facility Controller | Lower Temperature Limit |  | 0 |  | 0 | - | Pass | Configuration
134 | Facility Controller | PLC Version |  |  |  | PWC-SB 4Fa | Version | Pass | Configuration
135 | Facility Controller | Vacuum pressure on sample chuck | -80 |  | -60 | -75.1 | kPa | Pass | Configuration
136 | Facility Controller | FFU Check : XEA turn FFU ON / OFF function |  | Yes |  | Yes | - | Pass | Configuration
137 | Facility Controller | A/E FFU set on 100% and 20% then check RPM change |  | Yes |  | Yes | - | Pass | Configuration
138 | Facility Controller | EFEM FFU set on 100% and 20% then check RPM change |  | Yes |  | Yes | - | Pass | Configuration
139 | Facility Controller | AFM FFU speed |  | Normal=60%, Active=20% |  | Normal=60%, Active=20% | - | Pass | Configuration
140 | Facility Controller | Cabinet version |  |  |  | PWC-SB 4Fa | Version | Pass | Configuration
141 | Facility Controller | Check the pneumatic controller sensor settings (Nomal state : Green) |  | Yes |  | Yes | - | Pass | Configuration
142 | Facility Controller | Check the regulator vacuum pressure | -80 |  | -60 | -75.0 | kPa | Pass | Configuration |  | Cabinet에 있는 Regulator vacuum 압력 확인
143 | Facility Controller(PB) | Check the Enable Grip EMO switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
144 | Facility Controller(PB) | Check the Cabinet Rear EMO switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
145 | Facility Controller(PB) | Check the Cabinet Front Sub Box EMO switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
146 | Facility Controller(PB) | Check the Cabinet Plate EMO switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
147 | Facility Controller(PB) | Check the Cabinet Right EMO switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
148 | Facility Controller(PB) | Check the EFEM EMO-1 switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
149 | Facility Controller(PB) | Check the EFEM EMO-2 switch function test |  | Yes |  | Yes | - | Pass | Configuration |  | EMO Alarm 이미지 저장, PB인 경우 Warning창 + EMO 눌린 위치 표기 되야함
150 | Facility Controller(PB) | Check the cabinet front door alarm and buzzer |  | Yes |  | Yes | - | Pass | Configuration |  | [Cabinet door alarm DB setting] FailityController-Emergency-bUse : 1 FailityController-SignTower-bUseBuzzer : 1 위 DB 활성화 후 Cabinet door(Front/Rear/Side)를 열며, Alarm 확인(Signal tower Red/Buzzer on/XEA Alarm) 및 XEA 이미지 저장
151 | Facility Controller(PB) | Check the cabinet rear door alarm and buzzer |  | Yes |  | Yes | - | Pass | Configuration
152 | Facility Controller(PB) | Check the cabinet side door alarm and buzzer |  | Yes |  | Yes | - | Pass | Configuration
153 | Facility Controller(PB) | Swtich box power lamp color(Power Lamp Main, EFEM, AFM = Whilte) |  | White |  | White | - | Pass | Configuration
154 | Facility Controller(PB) | Photoionizer Color |  | White |  | White | - | Pass | Configuration
155 | Facility Controller(PB) | Swtich box lamp color (Power On = Green, Power Off = Black, Reset=Blue) |  | Green, Black, Blue |  | Green, Black, Blue | - | Pass | Configuration
156 | Facility Controller(PB) | Check if the power lamps(Main/EFEM/AFM) turn on when the main power is on |  | Yes |  | Yes | - | Pass | Configuration |  | power on 했을때 power lamp(main/efem/afm)가 점등 되는지 확인
156 | Facility Controller(T1) | Check the Cabinet EMO switches function test |  | Yes |  |  | - |  | Configuration |  | EMO Alarm 이미지 저장, Cabinet EMO Switch 위치 2군데 (Cabinet 오른쪽, Cabinet 안쪽)
156 | Facility Controller(T1) | Check the EFEM EMO switches function test |  | Yes |  |  | - |  | Configuration |  | EMO Alarm 이미지 저장, EFEM EMO Switch 위치 3군데 (Teaching pendant, EFEM Front, EFEM Monitor 쪽)
156 | Facility Controller(T1) | Serial Port # |  | 3 |  |  | - |  | Configuration
156 | Facility Controller(T1) | Power lamp color |  | White |  |  | - |  | Configuration
156 | Facility Controller(T1) | UPS Power lamp color |  | White |  |  | - |  | Configuration
156 | Facility Controller(T1) | Swtich box lamp color(Power On = Green, Power Off = White) |  | Green, White |  |  | - |  | Configuration
156 | Facility Controller(T1) | Photoionizer Tower Color |  | (선택) |  |  | - |  | Configuration |  | 사용색상 선택
156 | Facility Controller(T1) | Check if the power lamps(Power Cable/UPS Power) turn on when the main power is on |  | Yes |  |  | - |  | Configuration |  | power on 했을때 power lamp(Power Cable/UPS Power)가 점등 되는지 확인
156 | Facility Controller(UPS) | Check UPS capability for 5 minutes |  | Yes |  |  | - |  | Performance
157 | Focus Stage | Check the focus stage reset |  | Yes |  | Yes | - | Pass | Configuration |  | SW Limit 설정 후 Focus 맨위/맨아래 쪽에서 Reset 진행시 정상동작 여부 확인
158 | Focus Stage | Check the i-marking on the encoder bracket mounting bolts |  | Yes |  | Yes | - | Pass | Configuration
159 | Focus Stage | Check encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration
160 | Focus Stage | Encoder model |  | V2BCH |  | V2BCH | - | Pass | Configuration |  | XY Stage 상판/하판 사이의 encoder 육안으로 확인(품명적혀있음)
161 | Focus Stage | LowerLimit |  | H/W lower limit+500um |  | 1964.94 | um | Pass | Configuration |  | HW limit+500um
162 | Focus Stage | Check encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  | RGH24H : Good-Green, V2BCH : Good-Blue
163 | Focus Stage | UpperLimit |  | 9000 |  | 9000 | um | Pass | Configuration
164 | Focus Stage | AutoFocus Range (Tip Focus / Sample Focus) |  | 400 |  | 400 | um | Pass | Configuration
165 | Focus Stage | AutoFocus Speed (Tip Focus / Sample Focus) |  | 150 |  | 150 | um/s | Pass | Configuration
166 | Focus Stage | Fine Autofocus Range (Tip Focus / Sample Focus) |  | 50 |  | 50 | um | Pass | Configuration
167 | Focus Stage | Fine Autofocus Speed (Tip Focus / Sample Focus) |  | 30 |  | 30 | um/s | Pass | Configuration
168 | Focus Stage | UseEncoder |  | 1 |  | 1 | - | Pass | Configuration |  | FocusStage in DB
169 | Focus Stage | EncoderResolution |  | 0.05 |  | 0.05 | um | Pass | Configuration |  | Encoder 품명 육안 확인후 검색하여 Resolution 확인
170 | Focus Stage | Fine Autofocus Offset (Tip Focus / Sample Focus) |  | 0 |  | 0 | - | Pass | Configuration
171 | Focus Stage | OpticalAxisShiftPerFocusX | -0.01 |  | 0.01 | 0.0002 | - | Pass | Performance | Check item
172 | Focus Stage | OpticalAxisShiftPerFocusY | -0.01 |  | 0.01 | 0.0002 | - | Pass | Performance | Check item
173 | Focus Stage | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) |  |  | 0.0015 | 8e-05 | - | Pass | Performance | Check item
174 | Focus Stage | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) |  |  | 0.0015 | 5e-05 | - | Pass | Performance | Check item
174 | Laser Scattering | Laser intensity from laser source controller | 18 |  |  |  | mW |  | Performance
174 | Laser Scattering | Laser intensity from beam focusing module | 15 |  |  |  | mW |  | Performance
175 | Lift Pin | Lift pins synchronization offset check (X offset) |  |  | 50 | 30.17 | um | Pass | Performance | Check item | Lift pin 튜닝 후, up/down 5회 반복하여 offset 확인 guide 참고
176 | Lift Pin | Lift pins synchronization offset check (Y offset) |  |  | 50 | 10.38 | um | Pass | Performance | Check item | Lift pin 튜닝 후, up/down 5회 반복하여 offset 확인 guide 참고
177 | Main Body | Temperature sensor on Z stage |  | Installed |  | Installed | - | Pass | Configuration
178 | Main Body | Firmware version of Frame board |  | 0x6e25041e |  | 0x6e25041e | Version | Pass | Configuration
179 | Misc | Check if the mouse and keyboard are functioning properly after turning off the power and then turning it back on |  | Yes |  | Yes | - | Pass | Configuration |  | 전체 전원 Off 후 10분뒤 전원 On - 마우스,키보드 등 동작 이상없는지 확인
180 | Misc | Activate Thermal tune option in XEA |  | Activated |  | Activated | - | Pass | Configuration |  | Guide참고, xea.vmoption 활성화 및 XEA setup에서 use설정
181 | Misc | Resonance frequency of AC160 by thermal tune |  | - |  | 288.065 | kHz | Pass | Configuration |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용)
182 | Misc | Resonance frequency of AC160 by frequency sweep |  | - |  | 288.21 | kHz | Pass | Configuration |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용)
183 | Misc | Resonance frequency of AC240 by thermal tune |  | - |  | 88.068 | kHz | Pass | Configuration |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용)
184 | Misc | Resonance frequency of AC240 by frequency sweep |  | - |  | 88.1 | kHz | Pass | Configuration |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용)
185 | Misc | Delete unnecessary files and usage records from the computer according to internal guidelines |  | Yes |  | Yes | - | Pass | Configuration |  | PC 정리 가이드 참고
186 | Misc | Check SW & DB backup (Installed Software, XEA & XEService) |  | Yes |  | Yes | - | Pass | Configuration |  | CIM에 Software 및 DB Backup
187 | Misc | Skirt |  | Installed |  | Installed | - | Pass | Configuration
188 | Misc | Seismic bracket |  | Installed |  | Installed | - | Pass | Configuration
189 | Misc | Check if the bolts used for the AFM core module are assembled with standardized torque |  | Yes |  | Yes | - | Pass | Configuration |  | 옆에 표 참고
190 | Misc | Check if the cable holders are installed on both the AE and the Cabinet |  | Yes |  | Yes | - | Pass | Configuration
191 | Misc | Check the AE FFU cleaning |  | Yes |  | Yes | - | Pass | Configuration
192 | Misc | Check the jig bolts for the computer |  | Yes |  | Yes | - | Pass | Configuration
193 | Misc | Check the fixation of Cognex dongle cable on the rear of the PC |  | Yes |  | Yes | - | Pass | Configuration
193 | Misc(T1) | Check Facility Controller connect to A/E Cable holder |  |  | 0.5 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check Z Scanner body connect to A/E Cable holder |  |  | 1 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check XY Scanner connect to A/E Cable holder |  |  | 0.5 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check ATX Top plate connect to A/E Cable holder |  |  | 0.5 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check XY Stage connect to A/E Cable holder |  |  | 0.5 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check Focus Stage connect to A/E Cable holder |  |  | 1 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check Z Stage connect to A/E Cable holder |  |  | 0.5 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
193 | Misc(T1) | Check Frame Board connect to A/E Cable holder |  |  | 0.5 |  | Ω |  | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
194 | Misc(PB) | Check AFM Controller connect to Cabinet Earth Bar |  |  | 0.5 | 0.1 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
195 | Misc(PB) | Check Z Scanner body connect to A/E Earth Bar |  |  | 1 | 0.4 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
196 | Misc(PB) | Check XY Scanner connect to A/E Earth Bar |  |  | 0.5 | 0.3 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
197 | Misc(PB) | Check ATX Top plate connect to A/E Earth Bar |  |  | 0.5 | 0.3 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
198 | Misc(PB) | Check XY Stage connect to A/E Earth Bar |  |  | 0.5 | 0.3 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
199 | Misc(PB) | Check Focus Stage connect to A/E Earth Bar |  |  | 1 | 0.7 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
200 | Misc(PB) | Check Z Stage connect to A/E Earth Bar |  |  | 0.5 | 0.5 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
201 | Misc(PB) | Check Frame Board connect to A/E Earth Bar |  |  | 0.5 | 0.2 | Ω | Pass | Configuration |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인
202 | Misc(VLSI) | SHS12-1800 (Box) |  | - |  | 30907-008 | - | Pass | Configuration |  | VLSI샘플이 여러 개인 경우 항목 추가, 고객에게 나가는 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
203 | Misc(VLSI) | SHS12-1800 (Certification) |  | - |  | 30907-008 | - | Pass | Configuration |  | VLSI샘플이 여러 개인 경우 항목 추가, 고객에게 나가는 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
204 | Misc(VLSI) | SHS12-1800 (Sample) |  | - |  | 30907-008 | - | Pass | Configuration |  | VLSI샘플이 여러 개인 경우 항목 추가, 고객에게 나가는 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
205 | Misc(VLSI) | SHS12-4500 (Box) |  | - |  | 32039-004 | - | Pass
206 | Misc(VLSI) | SHS12-4500 (Certification) |  | - |  | 32039-004 | - | Pass
207 | Misc(VLSI) | SHS12-4500 (Sample) |  | - |  | 32039-004 | - | Pass
208 | Misc(VLSI) | SHS12-180 (Box) |  | - |  | 30658-023 | - | Pass
209 | Misc(VLSI) | SHS12-180 (Certification) |  | - |  | 30658-023 | - | Pass
210 | Misc(VLSI) | SHS12-180 (Sample) |  | - |  | 30658-023 | - | Pass
211 | Misc(VLSI) | STS3-12-180P (Box) |  | - |  | 13708-001-129 | - | Pass | Configuration |  | VLSI샘플이 여러 개인 경우 항목 추가, 고객에게 나가는 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
212 | Misc(VLSI) | STS3-12-180P (Certification) |  | - |  | 13708-001-129 | - | Pass | Configuration |  | VLSI샘플이 여러 개인 경우 항목 추가, 고객에게 나가는 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
213 | Misc(VLSI) | STS3-12-180P (Sample) |  | - |  | 13708-001-129 | - | Pass | Configuration |  | VLSI샘플이 여러 개인 경우 항목 추가, 고객에게 나가는 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
213 | Misc(Customer FAT) | Customer FAT |  | Yes |  |  | - |  | Performance |  | 고객 FAT 자료는 추가로 제작하여 제출
214 | Sample Chuck | Check chuck surface no damage |  | Yes |  | Yes | - | Pass | Configuration
215 | Sample Chuck(Sample holder) | Check sample holder |  | Yes |  | Yes | - | Pass | Configuration
216 | Sample Chuck | Wafer sample chuck type |  | Pin type |  | Pin type | - | Pass | Configuration
217 | Sample Chuck | Wafer sample chuck material |  | SiC |  | SiC | - | Pass | Configuration
218 | Sample Chuck | Tip position sample in XEA |  | Teached |  | Teached | - | Pass | Configuration
219 | Sample Chuck(w/ Sliding stage) | Check sample surface mapping |  |  | 30 | 23 | um | Pass | Configuration |  | Sim tape 작업 후 진행 (spec. <30um @300mm, <50um @200mm)
219 | Sample Chuck(C-AFM) | Check sample bias |  | Yes |  |  | - |  | Configuration
220 | Sample Chuck | Ground connection |  | Yes |  | Yes | - | Pass | Configuration |  | AL재질 및 eAFM Sample chuck 제외
220 | Sample Chuck(w/ AI) | Allowable max warpage, w/ Al sample chuck |  | N/A |  |  | - | - | Performance |  | 기준없음
220 | Sample Chuck(w/ SiC) | Allowable max warpage, w/ SiC sample chuck | 300 |  |  |  | um |  | Performance |  | Micron sample로 확인
221 | Sample Chuck(w/ SiC Anti-warpage) | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) |  |  | 800 | 755.93 | um | Pass | Performance |  | Micron sample로 확인
222 | Sample Chuck(w/ SiC Anti-warpage) | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Smile type) |  |  | -500 | -719.38 | um | Pass | Performance
223 | SlidingStage | Check sliding stage reset |  | Yes |  | Yes | - | Pass | Configuration
224 | SlidingStage | Maker(Motion Control Board) |  | Galil |  | Galil | - | Pass | Configuration
225 | SlidingStage | Galil Suite Version |  | 1.0.3.1088 |  | 1.0.3.1088 | Version | Pass | Configuration
226 | SlidingStage | FastScanPixelSize |  | 512 |  | 512 | - | Pass | Configuration |  | Info in Profiler DB
227 | SlidingStage | ProfileCountThreshold |  | 20 |  | 20 | - | Pass | Configuration |  | Info in Profiler DB
228 | SlidingStage | RecoveryCountThreshold |  | 20 |  | 20 | - | Pass | Configuration |  | Info in Profiler DB
229 | SlidingStage | PreventionFailureRate |  | 10 |  | 10 | - | Pass | Configuration |  | Info in Profiler DB
230 | SlidingStage | PreventionTorqueThreshold |  | 3 |  | 3 | - | Pass | Configuration |  | Info in Profiler DB
231 | SlidingStage | Position Based Profiler DB setting : ExtTriggerAUXSource |  | 4 |  | 4 | - | Pass | Configuration |  | Aux Cable Number : 4, 예외 YMTC(s/w version5.2.0B40) : 2
232 | SlidingStage | HomeMode |  | 4 |  | 4 | - | Pass | Configuration |  | Info in Profiler DB
233 | SlidingStage | Position Based Profiler DB setting : ExtTriggerThresholdVolt |  | 1.5 |  | 1.5 | - | Pass | Configuration |  | Default 1.5
234 | SlidingStage | PreventionFailAction |  | 1 |  | 1 | - | Pass | Configuration |  | Info in Profiler DB
235 | SlidingStage | UseRecovery |  | 1 |  | 1 | - | Pass | Configuration |  | Info in Profiler DB
236 | SlidingStage | Position Based Profiler DB setting : UsePositionSync |  | 1 |  | 1 | - | Pass | Configuration |  | Info in Profiler DB
237 | SlidingStage | Igain |  | 0.03 |  | 0.03 | - | Pass | Configuration |  | Info in Profiler DB, Tuning 방법 guide 참고
238 | SlidingStage | IGainWhileMoving |  | 0.03 |  | 0.03 | - | Pass | Configuration |  | Info in Profiler DB, Tuning 방법 guide 참고
239 | SlidingStage | OverProfileAfterMM |  | 0.2 |  | 0.2 | - | Pass | Configuration |  | Info in Profiler DB
240 | SlidingStage | OverProfileBeforeMM |  | 0.2 |  | 0.2 | - | Pass | Configuration |  | Info in Profiler DB
241 | SlidingStage | Mileage |  | 0 |  | 0 | - | Pass | Configuration |  | Reset 후 진행
242 | SlidingStage | Origin |  | 0 |  | 0 | - | Pass | Configuration |  | Info in Profiler DB
243 | SlidingStage | Position Based Profiler DB setting : ExtTriggerType |  | 0 |  | 0 | - | Pass | Configuration |  | 0 : Falling 1 : Raising 2 : Both
244 | SlidingStage | Dgain | 5 |  | 20 | 15 | - | Pass | Configuration |  | Info in Profiler DB, Tuning 방법 guide 참고
245 | SlidingStage | DGainWhileMoving | 5 |  | 20 | 15 | - | Pass | Configuration |  | Info in Profiler DB, Tuning 방법 guide 참고
246 | SlidingStage | Pgain | 5 |  | 20 | 15 | - | Pass | Configuration |  | Info in Profiler DB, Tuning 방법 guide 참고
247 | SlidingStage | PGainWhileMoving | 5 |  | 20 | 15 | - | Pass | Configuration |  | Info in Profiler DB, Tuning 방법 guide 참고
248 | SlidingStage | No Z scanner saturation during measurement |  | No |  | No | - | Pass | Performance |  | Horizontal line 측정 w/o fail
249 | SlidingStage | Profiler noise @200um |  |  | 0.25 | 0.118 | nm | Pass | Performance | Check item
250 | SlidingStage | Roll motion to X+ full scan |  |  | 100 | 9.9 | um | Pass | Performance | Check item
251 | SlidingStage | Position repeatability(10repeats, Position Based) |  |  | 1 | 0.52 | um | Pass | Performance | Check item | Position Base기준
252 | SlidingStage | OPM Repeatability @25mm |  |  | 12.9 | 4.037 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 Rep. 1σ 기입 (일체형은 Center값, 분리형은 Total RMS 값 기입)
253 | SlidingStage | OPM Repeatability @10mm |  |  | 5.6 | 1.739 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 Rep. 1σ 기입 (일체형은 Center값, 분리형은 Total RMS 값 기입)
254 | SlidingStage | OPM Repeatability @5mm |  |  | 3.3 | 1.322 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 Rep. 1σ 기입 (일체형은 Center값, 분리형은 Total RMS 값 기입)
255 | SlidingStage | OPM Repeatability @1mm |  |  | 1.6 | 0.657 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 Rep. 1σ 기입 (일체형은 Center값, 분리형은 Total RMS 값 기입)
255 | SlidingStage(Double Walled AE) | Max OPM @25mm |  |  | 250 |  | nm |  | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (일체형은 Center값 기입)
255 | SlidingStage(Double Walled AE) | Max OPM @10mm |  |  | 100 |  | nm |  | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (일체형은 Center값 기입)
255 | SlidingStage(Double Walled AE) | Max OPM @5mm |  |  | 50 |  | nm |  | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (일체형은 Center값 기입)
255 | SlidingStage(Double Walled AE) | Max OPM @1mm |  |  | 18 |  | nm |  | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (일체형은 Center값 기입)
256 | SlidingStage(Isolated AE) | Max OPM @25mm |  |  | 200 | 133.469 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (분리형은 Total Max 값 기입)
257 | SlidingStage(Isolated AE) | Max OPM @10mm |  |  | 80 | 38.704 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (분리형은 Total Max 값 기입)
258 | SlidingStage(Isolated AE) | Max OPM @5mm |  |  | 40 | 24.385 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (분리형은 Total Max 값 기입)
259 | SlidingStage(Isolated AE) | Max OPM @1mm |  |  | 13 | 10.175 | nm | Pass | Performance | Check item | AFP Repeatability 프로그램을 이용해 분석하여 OPM Max 기입 (분리형은 Total Max 값 기입)
259 | SlidingStage(10mm) | Torque Profile (10mm, 200um/s) |  |  | 2 |  | - |  | Performance | Check item | 평균값 기입, Sliding Stage Recipe 사용, Sample : Bare wafer, Pixel : 8,192pxl
259 | SlidingStage(10mm) | Torque Profile (10mm, 20um/s) |  |  | 2 |  | - |  | Performance | Check item | 평균값 기입, Sliding Stage Recipe 사용, Sample : Bare wafer, Pixel : 8,192pxl
260 | SlidingStage(25mm,50mm) | Torque Profile (50mm, 200um/s) |  |  | 2 | 0.365 | - | Pass | Performance | Check item | 평균값 기입, Sliding Stage Recipe 사용, Sample : Bare wafer, Pixel : 8,192pxl
261 | SlidingStage(25mm,50mm) | Torque Profile (50mm, 20um/s) |  |  | 2 | 0.512 | - | Pass | Performance | Check item | 평균값 기입, Sliding Stage Recipe 사용, Sample : Bare wafer, Pixel : 8,192pxl
261 | SlidingStage(10mm) | FullRecoveryProfileLength |  | 10 |  |  | - |  | Configuration |  | Info in Profiler DB
261 | SlidingStage(10mm) | PosMax | 5 |  |  |  | - |  | Configuration |  | Info in Profiler DB
261 | SlidingStage(10mm) | PosMin |  |  | -5 |  | - |  | Configuration |  | Info in Profiler DB
261 | SlidingStage(10mm) | MaxLength | 10 |  |  |  | mm |  | Configuration |  | Info in Profiler DB
262 | SlidingStage(25mm,50mm) | FullRecoveryProfileLength |  | 50 |  | 50 | - | Pass | Configuration |  | Info in Profiler DB
263 | SlidingStage(25mm,50mm) | PosMax | 25 |  |  | 25.5 | - | Pass | Configuration |  | Info in Profiler DB
264 | SlidingStage(25mm,50mm) | PosMin |  |  | -25 | -25.5 | - | Pass | Configuration |  | Info in Profiler DB
265 | SlidingStage(25mm,50mm) | MaxLength | 50 |  |  | 51 | mm | Pass | Configuration |  | Info in Profiler DB
266 | Smoke Detector | Check the smoke detector setting |  | Yes |  | Yes | - | Pass | Configuration
267 | SW | 'Xms4096m' in xea.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  | Park systems - XEA - Bin - xea.vmoptions
268 | SW | 'Xmx16384m' in xea.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  | Park systems - XEA - Bin - xea.vmoptions
269 | SW | 'Xms128m' in xei.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  | Park systems - XEI - Bin - xei.vmoptions
270 | SW | 'Xmx2048m' in xei.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  | Park systems - XEI - Bin - xei.vmoptions
271 | SW | 'Dcom.psia.xei.profileSupported=true' in xei.vmoptions |  | Yes |  | Yes | - | Pass | Configuration
272 | SW | 'Dcom.psia.xea.useInDieAlign=true' in xea.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  | Park systems - XEA - Bin - xea.vmoptions
273 | SW | JAVA JRE version |  | jre-8u191 |  | jre-8u191 | Version | Pass | Configuration
274 | SW | Head Mode in XEA |  | C-AFM, NC-AFM |  | C-AFM, NC-AFM | - | Pass | Configuration
275 | SW | SmartAnalysis version |  | 1.4.3.1480 |  | 1.4.3.1480 | Version | Pass | Configuration |  | 최신 RTM 버전 사용
276 | SW | Z scale in XEA |  | 0.7 |  | 0.7 | - | Pass | Configuration
277 | SW | XEA version |  | 5.2.6 RTMx |  | 5.2.6 RTM5 | Version | Pass | Configuration |  | 5.2.4 빌드버전, YMTC 제외
278 | SW | XEService version |  | 5.2.6 RTMx |  | 5.2.6 RTM5 | Version | Pass | Configuration |  | 5.2.4 빌드버전, YMTC 제외
279 | SW | XEI version |  | 5.2.4 B105 |  | 5.2.4 B105 | Version | Pass | Configuration
280 | SW | XEPi version |  | 5.2.0 B28 |  | 5.2.0 B28 | Version | Pass | Configuration
281 | SW | Cantilever DB version |  | 1.3 |  | 1.3 | Version | Pass | Configuration
282 | SW | Simulated Automation Test |  | Yes |  | Yes | - | Pass | Performance
283 | SW | Head Crash DB setting: A+B Interlock |  | 0.5 |  | 0.5 | V | Pass | Configuration |  | DB: DSP>Cantilever>MinimalIntensity (Cantiliver: AC160TS)
284 | SW | Head Crash DB setting: A-B Interlock |  | 6 |  | 6 | V | Pass | Configuration |  | DB: DSP>Cantilever>SafeErrorSignal (Cantiliver: AC160TS)
285 | SW | Head Crash DB setting: Z Detector Interlock |  | 0.15 |  | 0.15 | - | Pass | Configuration |  | DB: DSP/ZStage/HeadCrashThreshold
286 | SW | Head Crash DB setting: XY Stage Interlock |  | 1 |  | 1 | - | Pass | Configuration |  | DB: XYStage/Info/bUseInterlock
287 | System Noise | # of AVIS Element |  | 8 |  | 8 | ea | Pass | Configuration
288 | System Noise | Lateral(C-D) sensitivity |  |  |  | 15 | mV/nm | Pass | Configuration | Check item | B100-NCH A-B Calibration(F/D Curve)
289 | System Noise | Lateral(C-D) noise |  |  | 0.25 | 0.174 | nm | Pass | Performance | Check item | B100-NCH Calibration 후 측정
290 | System Noise | Min. Vibration Transmissbility (~100Hz) |  |  | 0.1 | 0.026 | - | Pass | Performance | Check item | Floor vs AFM Base
291 | System Noise(NX-Wafer 300mm, W/ Sliding Stage) | RMS of ZeroScan |  |  | 0.1 | 0.055 | nm | Pass | Performance | Check item | Z Drive(Long head), Z Height(standard head)
291 | System Noise(NX-Wafer 200mm, W/ Sliding Stage) | RMS of ZeroScan |  |  | 0.1 |  | nm |  | Performance | Check item | Z Drive(Long head), Z Height(standard head)
291 | System Noise(NX-Wafer 300mm) | RMS of ZeroScan |  |  | 0.07 |  | nm |  | Performance | Check item | Z Drive(Long head), Z Height(standard head)
291 | System Noise(NX-Wafer 200mm) | RMS of ZeroScan |  |  | 0.05 |  | nm |  | Performance | Check item | Z Drive(Long head), Z Height(standard head)
292 | Tip Characterizer Sample | Tip Characterizer stage position teach / XEA |  | Teached |  | Teached | - | Pass | Configuration |  | Guide 참고
293 | Tip Characterizer Sample | Tip Characterizer DB setting : TipCharacterizerSupport |  | 1 |  | 1 | - | Pass | Configuration |  | Guide 참고
294 | Tip Characterizer Sample | Tip Characterizer auto focus teach / XEA |  | Teached |  | Teached | - | Pass | Configuration |  | Guide 참고
295 | Tip Characterizer Sample | Tip Characterizer sample width |  | - |  | 193.1 ± 1.9 | nm | Pass | Configuration |  | Guide 참고
296 | Vision | Check the Vision brightness calibration |  | Yes |  | Yes | - | Pass | Performance
297 | Vision | LightStrengthGain |  |  |  | 0.53 | - | Pass | Performance | Check item | Vision brightness cal 후 DB 값 기입 Vision-Camera-LightStrengthGain, Cal 값 기입 후 출고시 1로 설정해서 출하
298 | Vision | Check the white balance |  | Yes |  | Yes | - | Pass | Configuration |  | Tip-Sample Distance가 어느정도 떨어져 있는 상태에서 White Paper에 Focusing -Brightness 80/Contrast 0/Light strength 50 --> White Balance를 여러 번 클릭해서 WB 맞추기(Vision 밝기가 변화 없을 때까지 -> Gray Color)
299 | Vision | Field Of View in large view | 800 |  |  | 839.16 | - | Pass | Calibration |  | FieldOfView in Vision Carmera DB
300 | Vision | Check Camera illumination function |  | Yes |  | Yes | - | Pass | Configuration |  | Tip이나 Sample focus 후 1. Light Strength100%(Capture) 2. Light strength 0%(Capture) (Contrast 50%, Brightness는 50%)
301 | Vision | Check sample autofocus function |  | Yes |  | Yes | - | Pass | Configuration
302 | Vision | Check tip auto focus function |  | Yes |  | Yes | - | Pass | Configuration
303 | Vision | Check tip pattern recognition |  | Yes |  | Yes | - | Pass | Configuration
304 | Vision | Check LightSource ON |  | Yes |  | Yes | - | Pass | Configuration
305 | Vision | LightSource(LED) Color |  | White |  | White | - | Pass | Configuration
306 | Vision | Supply power of the on-axis camera through PoE |  | No |  | No | - | Pass | Configuration
307 | Vision | Supply power of the on-axis camera through frame board |  | Yes |  | Yes | - | Pass | Configuration
308 | Vision | Camera model |  | Mako G-507C |  | Mako G-507C | - | Pass | Configuration
309 | Vision | Camera maker |  | AVT_GIGECAM |  | AVT_GIGECAM | - | Pass | Configuration |  | Carmera in Vision DB
310 | Vision | Camera alignment |  |  | 5 | 4.5 | um | Pass | Configuration |  | Vision상에 보여지는 Chip carrier의 상단 축과(비전 상단끝에 걸치도록 조정) 비전 상단 끝부분 영역이 Rotate 없이 수평으로 맞추어져 있는지 Check, 필요시 Camera Angle을 조정
311 | Vision | Vision software version |  | 7.0.2 |  | 7.0.2 | Version | Pass | Configuration
312 | Vision | FullPixel |  | 2440 |  | 2440 | - | Pass | Configuration |  | Carmera in Vision DB
313 | Vision | FullPixelHeight |  | 1828 |  | 1828 | - | Pass | Configuration |  | GT2450C : 1830, Mako G-507C : 1828
314 | Vision | FPS(FramePerSecond) of large view | 15 |  |  | 25 | FPS | Pass | Configuration |  | Large view 일때의 FPS 기입
315 | XY Scanner | Check the warranty sticker |  | Yes |  | Yes | - | Pass | Configuration
316 | XY Scanner | Full range of X Detector during X Scan sweep | 150 |  | 240 | 220.8067 | um | Pass | Performance | Check item | servo off
317 | XY Scanner | Full range of X1 Detector during X Scan sweep | 150 |  | 240 | 218.7191 | um | Pass | Performance | Check item | servo off
318 | XY Scanner | Full range of X2 Detector during X Scan sweep | 150 |  | 240 | 222.8943 | um | Pass | Performance | Check item | servo off
319 | XY Scanner | Full range of Y Detector during Y Scan sweep | 150 |  | 240 | 218.8544 | um | Pass | Performance | Check item | servo off
320 | XY Scanner | Full range of Y1 Detector during Y Scan sweep | 150 |  | 240 | 220.8703 | um | Pass | Performance | Check item | servo off
321 | XY Scanner | Full range of Y2 Detector during Y Scan sweep | 150 |  | 240 | 216.8386 | um | Pass | Performance | Check item | servo off
322 | XY Scanner | Full range of X1 Detector Raw during X Scan sweep | 11 |  | 17 | 14.26 | V | Pass | Performance | Check item | servo off
323 | XY Scanner | Full range of X2 Detector Raw during X Scan sweep | 11 |  | 17 | 14.32 | V | Pass | Performance | Check item | servo off
324 | XY Scanner | Full range of Y1 Detector Raw during Y Scan sweep | 11 |  | 17 | 14.51 | V | Pass | Performance | Check item | servo off
325 | XY Scanner | Full range of Y2 Detector Raw during Y Scan sweep | 11 |  | 17 | 14.41 | V | Pass | Performance | Check item | servo off
326 | XY Scanner | X Detector stroke calibration error (nominal scale error X) |  |  | 0.5 | 0.068 | % | Pass | Calibration | Check item | 신규 분석 프로그램인 XY Scanner Calibration Sample Analyzer 사용
327 | XY Scanner | Y Detector stroke calibration error (nominal scale error Y) |  |  | 0.5 | 0.215 | % | Pass | Calibration | Check item | 신규 분석 프로그램인 XY Scanner Calibration Sample Analyzer 사용
328 | XY Scanner | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 200 |  | 300 | 273.0679 | - | Pass | Calibration | Check item
329 | XY Scanner | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 200 |  | 300 | 287.0047 | - | Pass | Calibration | Check item
330 | XY Scanner | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 200 |  | 300 | 279.0944 | - | Pass | Calibration | Check item
331 | XY Scanner | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 200 |  | 300 | 279.4835 | - | Pass | Calibration | Check item
332 | XY Scanner | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 150 |  | 220 | 183.612 | um | Pass | Calibration | Check item
333 | XY Scanner | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 150 |  | 220 | 187.701 | um | Pass | Calibration | Check item
334 | XY Scanner | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) |  | 101.01 |  | 101.01 | um | Pass | Calibration
335 | XY Scanner | XY Orthogonality | 89.7 |  | 90.3 | 89.962 | degree | Pass | Performance
336 | XY Scanner | X Non-linearity error | 0 |  | 4 | 3.48 | % | Pass | Performance | Check item
337 | XY Scanner | Y Non-linearity error | 0 |  | 4 | 2.89 | % | Pass | Performance | Check item
338 | XY Scanner | X axis OPM |  |  | 4 | 0.72 | nm | Pass | Performance | Check item | @80x80um
339 | XY Scanner | Y axis OPM |  |  | 4 | 0.618 | nm | Pass | Performance | Check item | @80x80um
340 | XY Scanner | Check the full stroke image(servo on) |  | Yes |  | Yes | - | Pass | Performance |  | scan size : 100x100um, overscan : 0%, 이미지 측정하여 왜곡 확인 (orthogonality 측정시랑 같은 조건으로 scan size 및 overscan 만 변경)
341 | XY Scanner | X Ringing(Rpv) |  |  | 1 | 0.099 | nm | Pass | Performance | Check item | Recipe사용
342 | XY Scanner | Y Ringing(Rpv) |  |  | 1 | 0.151 | nm | Pass | Performance | Check item | Recipe사용
343 | XY Scanner(w/ SiC Anti-warpage chuck) | X Detector noise (XY Servo on) |  |  | 0.5 | 0.345 | nm | Pass | Performance | Check item | Anti-warpage 샘플척만
344 | XY Scanner(w/ SiC Anti-warpage chuck) | Y Detector noise (XY Servo on) |  |  | 0.6 | 0.351 | nm | Pass | Performance | Check item | Anti-warpage 샘플척만
345 | XY Scanner(w/ SiC Anti-warpage chuck) | X1 Detector noise (XY Servo on) |  |  | 0.6 | 0.47 | nm | Pass | Performance | Check item | Anti-warpage 샘플척만
346 | XY Scanner(w/ SiC Anti-warpage chuck) | Y1 Detector noise (XY Servo on) |  |  | 0.7 | 0.476 | nm | Pass | Performance | Check item | Anti-warpage 샘플척만
347 | XY Scanner(w/ SiC Anti-warpage chuck) | X2 Detector noise (XY Servo on) |  |  | 0.6 | 0.464 | nm | Pass | Performance | Check item | Anti-warpage 샘플척만
348 | XY Scanner(w/ SiC Anti-warpage chuck) | Y2 Detector noise (XY Servo on) |  |  | 0.7 | 0.474 | nm | Pass | Performance | Check item | Anti-warpage 샘플척만
348 | XY Scanner(w/ SiC Fork type chuck) | X Detector noise (XY Servo on) |  |  | 0.4 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Fork type chuck) | Y Detector noise (XY Servo on) |  |  | 0.4 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Fork type chuck) | X1 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Fork type chuck) | Y1 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Fork type chuck) | X2 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Fork type chuck) | Y2 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Bar type chuck) | X Detector noise (XY Servo on) |  |  | 0.4 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Bar type chuck) | Y Detector noise (XY Servo on) |  |  | 0.4 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Bar type chuck) | X1 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Bar type chuck) | Y1 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Bar type chuck) | X2 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ SiC Bar type chuck) | Y2 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ AL Bar type chuck) | X Detector noise (XY Servo on) |  |  | 0.4 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ AL Bar type chuck) | Y Detector noise (XY Servo on) |  |  | 0.4 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ AL Bar type chuck) | X1 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ AL Bar type chuck) | Y1 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ AL Bar type chuck) | X2 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(w/ AL Bar type chuck) | Y2 Detector noise (XY Servo on) |  |  | 0.5 |  | nm |  | Performance | Check item
348 | XY Scanner(NX-Wafer 200mm) | ResonanceFrequencyHz |  | 200 |  |  | Hz |  | Calibration |  | X/YScanner in DB
348 | XY Scanner(NX-Wafer 200mm) | ServoCutoffFrequencyHz |  | 80 |  |  | Hz |  | Calibration |  | X/YScanner in DB
349 | XY Scanner(NX-Wafer 300mm) | ResonanceFrequencyHz |  | 160 |  | 160 | Hz | Pass | Calibration |  | X/YScanner in DB
350 | XY Scanner(NX-Wafer 300mm) | ServoCutoffFrequencyHz |  | 50 |  | 50 | Hz | Pass | Calibration |  | X/YScanner in DB
351 | XY Stage | Check the i-marking on the XY stage mounting bolts |  | Yes |  | Yes | - | Pass | Configuration
352 | XY Stage | Align the XY stage origin and sample origin |  | Yes |  | Yes | - | Pass | Calibration
353 | XY Stage | RMS error after Stage Calibration |  |  | 5 | 2.87 | um | Pass | Calibration | Check item | 2번 진행, 2번째 값 기입
354 | XY Stage | Check X axis (+) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration
355 | XY Stage | Check X axis (-) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration
356 | XY Stage | Check Y axis (+) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration
357 | XY Stage | Check Y axis (-) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration
358 | XY Stage | Check XY stage reset |  | Yes |  | Yes | - | Pass | Configuration |  | SW Limit 설정 후 XY Stage 상/하/좌/우 끝쪽에서 Reset 진행시 정상동작 여부 확인
359 | XY Stage | Check X axis encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration
360 | XY Stage | Check Y axis encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration
361 | XY Stage | Save XY stage setup file |  | Yes |  | Yes | - | Pass | Configuration
362 | XY Stage | Suction |  | Enabled |  | Enabled | - | Pass | Configuration
363 | XY Stage | Check the circle align |  | Yes |  | Yes | - | Pass | Configuration
364 | XY Stage | Encoder model |  | V2BCH |  | V2BCH | - | Pass | Configuration |  | XY Stage 상판/하판 사이의 encoder 육안으로 확인(품명적혀있음)
365 | XY Stage | Maker of the motion control board |  | HCMNEW |  | HCMNEW | - | Pass | Configuration |  | Info in XY Stage DB
366 | XY Stage | Check X axis encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  | RGH24H : Good-Green, V2BCH : Good-Blue
367 | XY Stage | Check Y axis encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  | RGH24H : Good-Green, V2BCH : Good-Blue
368 | XY Stage | Encoder Resolution |  | 0.05 |  | 0.05 | um | Pass | Configuration |  | Encoder 품명 육안 확인후 검색하여 Resolution 확인
369 | XY Stage | BaudRate |  | 38400 |  | 38400 | - | Pass | Configuration
370 | XY Stage | Serial Port # |  | 1 |  | 1 | - | Pass | Configuration
371 | XY Stage | Image offset after moving from -X to +X (X scan) |  |  | 200 | 11.7188 | nm | Pass | Performance | Check item | Plus direction move Drift, X move X Scan(XX) 700sec (-X to +X move, image offset
372 | XY Stage | Image offset after moving from -Y to +Y (Y scan) |  |  | 200 | 42.4805 | nm | Pass | Performance | Check item | Plus direction move Drift, Y move Y Scan(YY) 700sec
373 | XY Stage | Image offset after moving from -X to +X (Y scan) |  |  | 200 | 34.4238 | nm | Pass | Performance | Check item | Plus direction move Drift, X move Y Scan(XY) 700sec
374 | XY Stage | Image offset after moving from -Y to +Y (X scan) |  |  | 200 | 12.4512 | nm | Pass | Performance | Check item | Plus direction move Drift, Y move X Scan(YX) 700sec
375 | XY Stage | Image offset after moving from +X to -X (X scan) |  |  | 200 | 18.3105 | nm | Pass | Performance | Check item | Minus direction move Drift, X move X Scan(XX) 700sec
376 | XY Stage | Image offset after moving from +Y to -Y (Y scan) |  |  | 200 | 24.9023 | nm | Pass | Performance | Check item | Minus direction move Drift, Y move Y Scan(YY) 700sec
377 | XY Stage | Image offset after moving from +X to -X (Y scan) |  |  | 200 | 81.2988 | nm | Pass | Performance | Check item | Minus direction move Drift, X move Y Scan(XY) 700sec
378 | XY Stage | Image offset after moving from +Y to -Y (X scan) |  |  | 200 | 43.2129 | nm | Pass | Performance | Check item | Minus direction move Drift, Y move X Scan(YX) 700sec
379 | XY Stage | XY Stage positioning X offset (Vision Pattern Matching) |  |  | 4 | 0.6299 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
380 | XY Stage | XY Stage positioning Y offset (Vision Pattern Matching) |  |  | 4 | 1.792 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
381 | XY Stage | XY Stage positioning X 1σ (Vision Pattern Matching) |  |  | 0.8 | 0.1246 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
382 | XY Stage | XY Stage positioning Y 1σ (Vision Pattern Matching) |  |  | 0.8 | 0.342 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
383 | XY Stage | XY Stage positioning X offset (In-Die Align) |  |  | 4 | 1.6943 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
384 | XY Stage | XY Stage positioning Y offset (In-Die Align) |  |  | 4 | 2.4268 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
385 | XY Stage | XY Stage positioning X 1σ (In-Die Align) |  |  | 0.8 | 0.27 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
386 | XY Stage | XY Stage positioning Y 1σ (In-Die Align) |  |  | 0.8 | 0.4755 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
387 | XY Stage | XY Stage positioning X offset (LLC Translation) |  |  | 4 | 2.0215 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
388 | XY Stage | XY Stage positioning Y offset (LLC Translation) |  |  | 4 | 2.2217 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
389 | XY Stage | XY Stage positioning X 1σ (LLC Translation) |  |  | 0.8 | 0.3857 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
390 | XY Stage | XY Stage positioning Y 1σ (LLC Translation) |  |  | 0.8 | 0.4503 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
391 | XY Stage | XY Stage positioning X offset (Global Align) |  |  | 7.5 | 0.6885 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
392 | XY Stage | XY Stage positioning Y offset (Global Align) |  |  | 7.5 | 1.8604 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
393 | XY Stage | XY Stage positioning X 1σ (Global Align) |  |  | 2.2 | 0.1465 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
394 | XY Stage | XY Stage positioning Y 1σ (Global Align) |  |  | 2.2 | 0.4678 | um | Pass | Performance | Check item | Recipe/분석프로그램 사용, Metro cal 사용
394 | XY Stage(275mmx200mm) | Sample Loading Position to X |  | - |  |  | - | - | Configuration |  | DB값 기입, XY Stage 체결위치 check
394 | XY Stage(275mmx200mm) | Sample Loading Position to Y |  | - |  |  | - | - | Configuration |  | DB값 기입, XY Stage 체결위치 check
394 | XY Stage(275mmx200mm) | StrokeX | 275000 | ¦StrokeXMax¦+¦StrokeXMin¦≥275000um |  | 0 | um | Fail | Configuration |  | StrokeXMax와 StrokeYMin 기입하면 자동
394 | XY Stage(275mmx200mm) | StrokeXMax | 137500 |  |  |  | um |  | Configuration
394 | XY Stage(275mmx200mm) | StrokeXMin |  |  | -137500 |  | um |  | Configuration
394 | XY Stage(275mmx200mm) | StrokeY | 200000 | ¦StrokeYMax¦+¦StrokeYMin¦≥200000um |  | 0 | um | Fail | Configuration |  | StrokeXMax와 StrokeYMin 기입하면 자동
394 | XY Stage(275mmx200mm) | StrokeYMax | 100000 |  |  |  | um |  | Configuration
394 | XY Stage(275mmx200mm) | StrokeYMin |  |  | -100000 |  | um |  | Configuration
394 | XY Stage(300mmx200mm) | Sample Loading Position to X |  | - |  |  | - | - | Configuration |  | DB값 기입, XY Stage 체결위치 check
394 | XY Stage(300mmx200mm) | Sample Loading Position to Y |  | - |  |  | - | - | Configuration |  | DB값 기입, XY Stage 체결위치 check
394 | XY Stage(300mmx200mm) | StrokeX | 300000 | ¦StrokeYMax¦+¦StrokeYMin¦≥300000um |  | 0 | um | Fail | Configuration |  | StrokeXMax와 StrokeYMin 기입하면 자동
394 | XY Stage(300mmx200mm) | StrokeXMax | 150000 |  |  |  | um |  | Configuration
394 | XY Stage(300mmx200mm) | StrokeXMin |  |  | -150000 |  | um |  | Configuration
394 | XY Stage(300mmx200mm) | StrokeY | 200000 | ¦StrokeYMax¦+¦StrokeYMin¦≥200000um |  | 0 | um | Fail | Configuration |  | StrokeXMax와 StrokeYMin 기입하면 자동
394 | XY Stage(300mmx200mm) | StrokeYMax | 100000 |  |  |  | um |  | Configuration
394 | XY Stage(300mmx200mm) | StrokeYMin |  |  | -100000 |  | um |  | Configuration
395 | XY Stage(400mmx300mm) | Sample Loading Position to X |  | - |  | -200078 | - | Pass | Configuration |  | DB값 기입, XY Stage 체결위치 check
396 | XY Stage(400mmx300mm) | Sample Loading Position to Y |  |  | -20000 | -39412.2 | - | Pass | Configuration |  | DB값 기입, XY Stage 체결위치 check, 분리형 AE인 경우만 Spec있음 일체형 AE는 Spec 없음
397 | XY Stage(400mmx300mm) | StrokeX | 400000 | ¦StrokeXMax¦+¦StrokeXMin¦≥400000um |  | 406838.4 | um | Pass | Configuration |  | StrokeXMax와 StrokeYMin 기입하면 자동
398 | XY Stage(400mmx300mm) | StrokeXMax | 200000 |  |  | 203220 | um | Pass | Configuration
399 | XY Stage(400mmx300mm) | StrokeXMin |  |  | -200000 | -203618.4 | um | Pass | Configuration
400 | XY Stage(400mmx300mm) | StrokeY | 300000 | ¦StrokeYMax¦+¦StrokeYMin¦≥300000um |  | 302279.7 | um | Pass | Configuration |  | StrokeXMax와 StrokeYMin 기입하면 자동
401 | XY Stage(400mmx300mm) | StrokeYMax | 150000 |  |  | 151767.2 | um | Pass | Configuration
402 | XY Stage(400mmx300mm) | StrokeYMin |  |  | -150000 | -150512.5 | um | Pass | Configuration
402 | XY Stage(NX-Wafer 200mm) | StageSafetyX | 200000 | ¦StageSafetyXMax¦+¦StageSafetyXMin¦≥200000um |  | 0 | um | Fail | Configuration
402 | XY Stage(NX-Wafer 200mm) | StageSafetyXMax | 100000 |  | 105000 |  | um |  | Configuration
402 | XY Stage(NX-Wafer 200mm) | StageSafetyXMin | -105000 |  | -100000 |  | um |  | Configuration
402 | XY Stage(NX-Wafer 200mm) | StageSafetyY | 200000 | ¦StageSafetyXMax¦+¦StageSafetyXMin¦≥200000um |  | 0 | um | Fail | Configuration
402 | XY Stage(NX-Wafer 200mm) | StageSafetyYMax | 100000 |  | 105000 |  | um |  | Configuration
402 | XY Stage(NX-Wafer 200mm) | StageSafetyYMin | -105000 |  | -100000 |  | um |  | Configuration
403 | XY Stage(NX-Wafer 300mm) | StageSafetyX | 300000 | ¦StageSafetyXMax¦+¦StageSafetyXMin¦≥300000um |  | 300274.64 | um | Pass | Configuration
404 | XY Stage(NX-Wafer 300mm) | StageSafetyXMax | 150000 |  | 155000 | 150204.79 | um | Pass | Configuration
405 | XY Stage(NX-Wafer 300mm) | StageSafetyXMin | -155000 |  | -150000 | -150069.85 | um | Pass | Configuration
406 | XY Stage(NX-Wafer 300mm) | StageSafetyY | 300000 | ¦StageSafetyXMax¦+¦StageSafetyXMin¦≥300000um |  | 300504.41 | um | Pass | Configuration
407 | XY Stage(NX-Wafer 300mm) | StageSafetyYMax | 150000 |  | 155000 | 150129.72 | um | Pass | Configuration
408 | XY Stage(NX-Wafer 300mm) | StageSafetyYMin | -155000 |  | -150000 | -150374.69 | um | Pass | Configuration
409 | Z Stage | Check upper limit sensor function |  | Yes |  | Yes | - | Pass | Configuration
410 | Z Stage | Check lower limit sensor function |  | Yes |  | Yes | - | Pass | Configuration
411 | Z Stage | Check the Z stage reset |  | Yes |  | Yes | - | Pass | Configuration |  | SW Limit 설정 후 Z Stage 맨위/맨아래 쪽에서 Reset 진행시 정상동작 여부 확인
412 | Z Stage | Check encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration
413 | Z Stage | Encoder model |  | V2BCH |  | V2BCH | - | Pass | Configuration |  | XY Stage 상판/하판 사이의 encoder 육안으로 확인(품명적혀있음)
414 | Z Stage | Check encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  | RGH24H : Good-Green, V2BCH : Good-Blue
415 | Z Stage | Check the i-marking on the encoder bracket mounting bolts |  | Yes |  | Yes | - | Pass | Configuration
416 | Z Stage | LowerLimit |  | 1Port, 1Slot Z Stage Position -300um |  | 4648 | um | Pass | Configuration |  | 1Port 1Slot -300um, ATX 없는 경우 Samplechuck approach position -500um
417 | Z Stage | EncoderResolution |  | 0.05 |  | 0.05 | um | Pass | Configuration |  | Encoder 품명 육안 확인후 검색하여 Resolution 확인
418 | Z Stage | UpperLimit |  | 29000 |  | 29000 | um | Pass | Configuration
419 | Z Stage | UseEncoder |  | 1 |  | 1 | - | Pass | Configuration |  | ZStage in DB
420 | Z Stage | HeadCrashThreshold |  | 0.15 |  | 0.15 | - | Pass | Configuration |  | ZStage in DB
421 | Z Stage | Initial Z | 12000 |  | 15000 | 15000 | um | Pass | Configuration
422 | Z Stage | Z Stage material |  | S45C |  | S45C | - | Pass | Configuration |  | 옆 사진 참고하여 해당 두가지 파트에 자석이 각각 붙는지 확인한다. (붙으면 s45c, 아니면 Fail)
423 | Z Stage | Z Drift |  |  | 20 | 14.379 | nm | Pass | Performance | Check item | Zeroscan recipe와 함께 측정되게 변경(v3.13.0), ZeroScan_Zdrift Recipe 사용
424 | Performance | Specification of Step Height Sample for repeatability test |  | - |  | 183.8 ± 2.0 | nm | Pass | Performance |  | 아래 측정(AFM,AFP Repeatability) 에 사용하는 Sample(180nm) spec 기입 ex)183.8 ± 2.0nm
425 | Performance | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) |  |  | 1 | 0.33 | nm | Pass | Performance | Check item | 측정시 Sample EFEM으로 로딩하는 Job으로 진행 필요 3.13.0 부터 해당레시피 변경됨
426 | Performance | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) |  |  | 0.5 | 0.06 | % | Pass | Performance | Check item | 측정시 Sample EFEM으로 로딩하는 Job으로 진행 필요 3.13.0 부터 해당레시피 변경됨
427 | Performance | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) |  |  | 1 | 0.42 | nm | Pass | Performance | Check item | 1day로 변경
428 | Performance | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) |  |  | 0.5 | 0.08 | % | Pass | Performance | Check item | 1day로 변경
429 | Performance | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) |  |  | 1 | 0.63 | nm | Pass | Performance | Check item | 1day로 변경
430 | Performance | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) |  |  | 0.5 | 0.12 | % | Pass | Performance | Check item | 1day로 변경
431 | Performance | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) |  |  | 1 | 0.28 | nm | Pass | Performance | Check item | 1day로 변경
432 | Performance | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) |  |  | 0.5 | 0.05 | % | Pass | Performance | Check item | 1day로 변경
433 | Performance | Cycle time for 15-point measurement |  | - |  | 1723.196 | sec | Pass | Performance | Check item | Static Repeatability w/ 180nm Step Height Sample 측정 Data log로 확인, 가이드 참고하여 분석시트 사용
434 | Performance | Specification of Step Height Sample for short-term stability test |  | - |  | 462.3 ± 2.7 | nm | Pass | Performance |  | 아래 측정(Short-term stability) 에 사용하는 Sample spec 기입 ex)450 ± 2.0nm, wafer sample 사용(180nm or 450nm)
435 | Performance | Short-term stability using Step Height Sample (3σ) |  |  | 1 | 0.44724 | nm | Pass | Performance | Check item | Cal 450nm, 450nm or 180nm의 Sample을 활용하여 72시간(3days) 측정 진행 1recipe(100 repeats)를 21번 반복 측정 - 총 2100 point의 3시그마 사용, Guide 및 분석시트 사용
436 | Performance | Roughness Repeatability Test 1x1um(10Repeats) |  | - |  | - | - | Pass
437 | Performance | Roughness Repeatability 1x1um(RSD%) |  |  | 5 | 0.798 | % | Pass | Performance | Check item | 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
438 | Performance | Roughness Repeatability 1x1um(1σ) |  |  | 0.01 | 0.001 | nm | Pass | Performance | Check item | 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
439 | Performance | Roughness Repeatability Test 80x80um(10Repeats) |  | - |  | - | - | Pass
440 | Performance | Roughness Repeatability 80x80um(RSD%) |  |  | 5 | 1.254 | % | Pass | Performance | Check item | 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
441 | Performance | Roughness Repeatability 80x80um(1σ) |  |  | 0.01 | 0.002 | nm | Pass | Performance | Check item | 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
441 | Customer | Check the Z Scanner Range | 15 |  |  |  | um |  | Performance |  | On-site Sample Test
441 | Customer | Check the Resolution of 1line (Max Data Length in Profile mode) | 25000 |  |  |  | pts/line |  | Performance |  | On-site Sample Test
441 | Customer | Check the Profiling Speed (Max Speed in Profile mode) | 0.5 |  |  |  | mm/sec |  | Performance |  | On-site Sample Test
441 | Customer | Check the Profiling Length (Max Profiling Length in Profile mode) | 6 |  |  |  | mm |  | Performance |  | On-site Sample Test
441 | Customer | Check the Max Scan Speed (Max Scan speed in AFM mode) | 3 |  |  |  | Hz |  | Performance |  | On-site Sample Test
441 | Customer | Vertical Resolution Check |  |  | 0.1 |  | nm |  | Performance |  | On-site Sample Test
441 | Customer | AFM Static Repeatability Test (3σ, 30 Repeats) |  |  | 0.6 |  | % |  | Performance |  | Bidding Spec Test, 30 Repeats에 대한 3sigma 값 입력
441 | Customer | AFM Dynamic Repeatability Test (3σ, 30 Repeats) |  |  | 1 |  | % |  | Performance |  | Bidding Spec Test, 30 Repeats에 대한 3sigma 값 입력
441 | Customer | AFP Static Repeatability Test (3σ, 30 Repeats) |  |  | 0.6 |  | % |  | Performance |  | Bidding Spec Test, 30 Repeats에 대한 3sigma 값 입력
441 | Customer | AFP Dynamic Repeatability Test (3σ, 30 Repeats) |  |  | 1 |  | % |  | Performance |  | Bidding Spec Test, 30 Repeats에 대한 3sigma 값 입력
441 | Customer | Variation range of start-point location (30 Repeats) |  |  | 8 |  | um |  | Performance |  | Measurement Location Test : By dynamic repeat 30 times using VLSI sample of which height is 180nm, and its variation range of start-point location should be less than 8um | Recipe 필요

## Last
Product Quality Conformity Certificate
This is to certify the product listed below and associated components have been manufactured, assembled, and tested by Park Systems, in accordance with established manufacturing procedures and ISO9001 and ISO14001 guidelines.
|  | Product Model | [5 empty] | NX-Wafer
|  | SID(System ID) Number | [5 empty] | D160033-140725
|  | Reference Document | [5 empty] | Industrial Check List v3.21.1
|  | Date of Final Test | [5 empty] | 2025-11-24
|  | End User | [5 empty] | Samsung S4
|  | Manufacturing Engineer | [5 empty] | Myeongjin Baek
|  | Production QC Engineer | [5 empty] | Levi Baek
|  | Team Manager of Production QC | [5 empty] | Jayden Lee
-
|  | Teached |  | Teached |  | Pass
5.2.4 빌드버전, YMTC 제외
5.2.4 빌드버전, YMTC 제외
|  |  |  | 1 |  | Fail
um
Resonance Frequency in DB
Check XY stage reset
Save XY Stage setup file
Suction |  | Enabled |  | Enabled |  | Pass
Check the circle align
Maker of the motion control board(Maker in XY Stage DB)
Check X axis encoder reader LED(Green: Good, Red: No Good)
Check Y axis encoder reader LED(Green: Good, Red: No Good)
Encoder Resolution
Sample Loading Position to X
Sample Loading Position to Y
Image offset after moving from -X to +X | [6 empty] | Plus direction move Drift, X move X Scan(XX) 700sec (-X to +X move, image offset
Image offset after moving from -Y to +Y
Image offset after moving from -X to +Y
Image offset after moving from -Y to +X
Image offset after moving from +X to -X
Image offset after moving from +Y to -Y
Image offset after moving from +X to -Y
Image offset after moving from +Y to -X
X axis, Repeatability @1st position
X axis, Repeatability @ 2nd position
X axis, Repeatability @ 3rd position
X axis, Repeatability @ 4th position
X axis, Accuracy @1st position | [6 empty] | Y axis, Accuracy @1st position
X axis, Accuracy @ 2nd position | [6 empty] | Y axis, Accuracy @ 2nd position
X axis, Accuracy @ 3rd position | [6 empty] | Y axis, Accuracy @ 3rd position
X axis, Accuracy @ 4th position | [6 empty] | Y axis, Accuracy @ 4th position
Y axis, Repeatability @1st position
Y axis, Repeatability @ 2nd position
Y axis, Repeatability @ 3rd position
Y axis, Repeatability @ 4th position
Y axis, Accuracy @1st position
Y axis, Accuracy @ 2nd position
Y axis, Accuracy @ 3rd position
Y axis, Accuracy @ 4th position
Check the Z stage reset
Z Drift (Short Traveling)
Check the Z Scanner Range
Check the Resolution of 1line (Max Data Length in Profile mode) |  |  |  | 524288
Check the Profiling Speed (Max Speed in Profile mode)
Check the Profiling Length (Max Profiling Length in Profile mode)
Check the Max Scan Speed (Max Scan speed in AFM mode)

## v3.21.1
Version | 작성자 | 적용시기 | Check 항목 | 상세내용 | 비고
v3.00 | 장순영 | 2023-03-14 | new 양식 생성 | v3.00 통합 양식으로 새롭게 제작 함.
v3.01 | 장순영 | 2023-03-16 | EFEM (Wafer, WLI, Mask)
|  |  | Air pressure test 항목 세분화 | 1. Air pressure @ 60% of FFU 변경
⑴ Module name 변경 : EFEM ⇒ EFEM(FFU Open type)
⑵ Spec 변경 1.2Pa ⇒ 1Pa
2. Air pressure @ 100% of FFU 추가
XY Scanner (Wafer, WLI, Mask)
|  |  | Rining test | Spec 기입 : 1nm
Sliding Stage (Wafer, WLI)
|  |  | Profile noise | Spec 기입 : 0.2nm
Rq값 기입으로 변경
v3.1.1 | 장순영 | 2023-03-20 | 양식 변경 | 버전 관리 변경 v3.xx ⇒ v3.x.x (두번째 자리 변경은 spec 및 내용 변경관련 주요한 사항 수정, 세번째 자리는 오타 수정 등 간단한 사항 수정)
"Serial Number" 탭 삭제
"Last" 탭에 Product Quality Conformity Certificate 페이지 변경
필터링 후 순번 오류 수정(함수변경)
AFM Head(Standard/Long) (Wafer, WLI)
|  |  | Module 명 수정 | Auto PSPD Alignment (A-B)의 module 명 수정 : AFM Head ⇒ AFM Head(Auto Align)
|  |  | Check item 명 수정 | MicrometerPerAdcHigh of Z Height in DB ⇒ Z Height Calibration(Large Mode) : ZDetector MicrometerPerAdcHigh
MicrometerPerDacHigh of Z Height in DB ⇒ Z Drive Calibration(Large Mode) : ZSanner MicrometerPerDacHigh
|  |  | Spec 수정 | Z Height Calibration(Large Mode) : ZDetector MicrometerPerAdcHigh : min 25 ⇒ min 23
Z Drive Calibration(Large Mode) : ZSanner MicrometerPerDacHigh : max 12 ⇒ max 14
ATX(Wafer, WLI, Mask, Spare Part)
|  |  | Module 명 수정 | A ± B difference during 3times(FT2) run test module 명 수정 : ATX ⇒ ATX(Auto Align Head)
System Noise(Long/Standard)
|  |  | Check item 명 오타 수정 | sliding stage 오타 수정
|  |  | Module 명 수정 | System Noise(Long-S/S), System Noise(Standard-S/S) 로 변경
Sliding Stage (Wafer, WLI)
|  |  | Spec 오류 수정 | HomeMode Spec 수정 : 3 ⇒ 4
OverProfileAfterMM Spec 수정 : 0.1 ⇒ 0.2
OverProfileBeforeMM Spec 수정 : 0.1 ⇒ 0.2
TSH(Mask)
|  |  | Check item 명 수정 | Z Height MicrometerPerAdcHigh ⇒ Z Height Calibration : ZDetector MicrometerPerAdcHigh
Z Drive MicrometerPerDacHigh ⇒ Z Drive Calibration : ZSanner MicrometerPerDacHigh
|  |  | Spec 수정 | Z Height Calibration : ZDetector MicrometerPerAdcHigh : min 25 ⇒ min 23
XY Scanner(Wafer, WLI)
|  |  | Check item 명 수정 | X1Scanner Calibration(Closed loop) : XDetector MicrometerPerAdc (X1) 로 변경
X2Scanner Calibration(Closed loop) : XDetector AuxMicrometerPerAdc (X2) 로 변경
Y1Scanner Calibration(Closed loop) : YDetector MicrometerPerAdc (Y1) 로 변경
Y2Scanner Calibration(Closed loop) : YDetector AuxMicrometerPerAdc (Y2) 로 변경
Xscanner Calibration(Open loop) : XScanner MicrometerPerDacHigh 로 변경
Yscanner Calibration(Open loop) : YScanner MicrometerPerDacHigh 로 변경
Closed loop scanning range : X/YScanner ServoScanRangeMicrometer 로 변경
EFEM(RFID) (Wafer, WLI, Mask)
|  |  | Spec 수정 | LW ⇒ LF
v3.2.1 | 장순영 | 2023-03-28 | 양식 변경 | 글꼴 통일
바닥글 오류 수정
AFM Controller (WLI)
|  |  | Spec 수정 | 기준 추가 : IND-M1
Facility Controller (Wafer, WLI, Mask)
|  |  | Check item 명 오타 수정 | Up and Down of ATX
Sample Chuck (Wafer, WLI)
|  |  | Check item 명 수정 | Check Lift Pin CDA Leak 으로 변경
Sliding Stage (Wafer, WLI)
|  |  | Spec 오류 수정 | FullRecoveryProfileLength Spec 수정 : 20 ⇒ 25
|  |  | 항목 세분화 | Full Stroke 10mm, 25mm/50mm 로 항목 세분화
|  |  | 항목 세분화 | POSMAX, POSMIN (10mm) 체크항목 생성
Vision (Wafer, WLI, Mask)
|  |  | Spec 수정 | LightSorce(LED) Color : 선택 가능하게끔 변경 (White, Yellow)
XY Scanner(Wafer, WLI)
|  |  | 항목 세분화 | XY Linearity : X, Y 분리
AFM Head(Auto Align)
|  |  | Spec 오류 수정 | Auto PSPD Alignment (A-B) : ±0.5V
System Noise(Long/Standard)
|  |  | Module 명 수정 | System Noise(Long head W/Sliding Stage), System Nois(Standard head W/Sliding Stage) 로 변경
Facility Controller(Wafer, WLI, Mask)
|  |  | Module 명 수정 및 삭제 | Facility Cotroller Serial port# : T1에만 적용, Mask는 PB이기에 삭제
Vision (Wafer, WLI, Mask)
|  |  | Check item 추가 | White Balance 설정 추가
v3.3.1 | 장순영 | 2023-04-17 | EFEM
|  |  | (Wafer, WLI, Mask)항목 세분화 | Check the robot leveling (Robot body), Check the robot leveling (Upper arm), Check the robot leveling (Lower arm)
|  |  | (Wafer, WLI)Check item 추가 | Check Aligner changes angle reference according to software configuration
System Noise
|  |  | (Wafer, WLI,Mask)Criteria 수정 | # of AVIS Element : 6ea 추가
[YMTC] FAT
|  |  | 항목추가 | 1. Check the Z Scanner Range
2. Check the Resolution of 1line
3. Check the Profiling Speed
4. Check the Profiling Length
5. Check the Max Scan Speed
6. Vertical Resolution Check
7. AFP Static Repeatability Test (3σ, 30 Repeats)
8. AFM Dynamic Repeatability Test (3σ, 30 Repeats)
9. AFP Static Repeatability Test (3σ, 30 Repeats)
10. AFP Dynamic Repeatability Test (3σ, 30 Repeats)
11. Variation range of start-point location (30 Repeats)
--총 11가지 추가--
v3.4.1 | 장순영 | 2023-05-08 | 양식 변경 | 필터링 후 순번 오류 수정(함수변경)
오타 수정
Check item 명 수정 : Check~시작되는 test에 대해서 기준 Yes로 변경 및 전체적인 Check item명 수정
|  |  | AE(Double Walled AE) | 일체형 AE 명칭 수정
Sliding Stage (Wafer, WLI)
|  |  | Spec 수정 | Profile noise @200um spec 수정 : 0.2nm ⇒ 0.25nm
|  |  | 측정값 변경 | OPM 측정값 변경 : Center 측정값으로 기입
|  |  | Test 변경 | Toque profile test 간소화(50mm의 고속/저속만 측정)
|  |  | Repeatability(Wafer, WLI) -> Performance(Wafer, WLI) | Performance로 변경
|  |  | 측정 조건 변경 | 1. AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
2. AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
3. AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
4. AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
5. AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
6. AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
7. AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
8. AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
-- 위 8가지 항목에 대해 측정시간 변경 : 3days ⇒ 1day --
|  |  | 측정 조건 변경 | Short-term stability w/ 450nm Step Height Sample (3σ)
-- 위 항목에 대해 측정조건 변경 : 하루에 아침5번, 점심5번, 저녁5번씩 총 3일 측정(총 데이터 45개)
AFM Head(Standard/Long) (Wafer, WLI, Mask)
|  |  | Test 추가 | Z Linearity by Z scanner position 추가
Performance
|  |  | Spec 수정 | Short-term stability w/ 450nm Step Height Sample (3σ) <1nm
|  |  | NX-Mask | 전반적인 내용 수정
Mask FAT 내용추가
v3.5.1 | 장순영 | 2023-05-15 | 수식오류 변경 | Pass/Fail 수식오류 변경
|  |  | EFEM | Check item명 및 Spec 수정
|  |  | Check item 명 수정 | Minimun adjustable angle of pre-aligner -> Angle offset value of pre-aligner
|  |  | Spec 수정 | 값 기입으로 Spec 변경
|  |  | Sliding Stage (Wafer, WLI) | AE Type에 따른 Max OPM spec 세분화
|  |  | Spec 수정 | 일체형 : 250,100,50,18 / 분리형 : 200,80,40,13
|  |  | XY Scanner(w/ SiC Anti-warpage) | X,Y,X1,X2,Y1,Y2 Detector noise spec. 세분화 (anti warp/그 외)
Anti warpage(X:0.5, Y:0.8, X1/X2:0.6, Y1/Y2:0.9)
EFEM(Wafer, WLI, Mask)
|  |  | Spec 수정 | EFEM CDA pressure : 0.4~0.6 으로 spec 변경
EFEM vacuum pressure : -90~-60 으로 spec 변경
AFM Head(Standard/Long) / TSH
|  |  | Spec 수정 | Z Linearity by Z scanner position : 2% ->1.5% spec 변경
v3.6.1 | 장순영 | 2023-05-31 | 양식 변경
|  |  | System 추가 | NX-HDM EFEM 시스템 추가
Facility controller
|  |  | Check item 명 수정 | Alarm tower color -> Photoionizer Tower Color로 변경
EFEM
|  |  | 단위 변경 | Air flow 단위 변경 mm/s -> m/s
|  |  | 항목 세분화 | Pre-aligner 종류에 따른 spec 변경 : Centering accuracy of pre-aligner
AFM Head(long)
|  |  | Spec 수정 | Full range in um during Z Scan sweep Spec 수정 : min 34 max 42um
|  |  | Check item 삭제 | (long head) Z Height noise 삭제
Vision
|  |  | Spec 오류 수정 | Field of view Spec 수정 : min 800
Facility Controller
|  |  | Spec 수정 | Signal Tower Color 추가 : RYGW
Performance(mask)
|  |  | Check item 추가 | Mask Roughness Repeatability 추가
v3.7.1 | 장순영 | 2023-06-09 | Performance(mask)
|  |  | 항목 세분화 | Mask Roughness Repeatability (RSD%, 1σ) 로 세분화
Sliding Stage (Wafer, WLI)
|  |  | Spec 수정 | Position Based Profiler DB setting : ExtTriggerAUXSource : 2-->4 수정
SW
|  |  | Spec 수정 | SmartAnalysis version : Basic -> 1.2.1.0
System Noise
|  |  | Spec 수정 | Check item 세분화 변경, NX-Wafer300mm (Sliding stage X) Zeroscan 0.07nm로 변경
System Noise(Mask)
|  |  | Spec 수정 | Zeroscan spec변경 0.12 -->0.1nm (Rq)
EFEM
|  |  | Check item 명 수정 | Level to X axis on the robot body
Level to Y axis on the robot body
Level to X axis on the robot upper arm
Level to Y axis on the robot upper arm
Level to X axis on the robot lower arm
Level to Y axis on the robot lower arm
SlidingStage
|  |  | 항목 세분화 | FullRecoveryProfileLength 세분화
|  |  | Spec 수정 | 10mm경우 spec 10, 25mm/50mm인 경우 spec 50
WLI Z scanner
|  |  | Spec 수정 | Z Detector max/min spec 수정 (0.3 -> 0.2)
WLI Turret Stage
|  |  | Spec 수정 | Limit range spec 수정(41000->-41000)
WLI SW
|  |  | Spec 표기 수정 | XIMEA API Software install PKG spec표기 수정 (4.23.03)
WLI Misc
|  |  | Spec 수정 | TestBench version spec 수정(4.1.0 -> 4.0.28)
v3.7.2 | 장순영 | 2023-06-22 | Vision
|  |  | Spec 표기 수정 | Camera alignment spec표기 수정 : Degree -> um
Sliding Stage (Wafer, WLI)
|  |  | 오타수정 | Position Based Profiler DB setting : ExtTriggerAUXSource 6으로 오류 기입 -> 4로 수정
|  |  | 항목 세분화 | OPM Repeatability 기입값 분리형, 일체형 세분화
Sample Chuck (Mask)
|  |  | Check item 명 수정 | Check if the lock stage holds the mask
Sample Chuck (Mask)
|  |  | Check item 삭제 | Ground connection
Facility Controller
|  |  | Check item 추가 | Check the all EMO switches function test
XY Stage(Mask)
|  |  | 항목 세분화 | Scan offset accuracy (Precision) : SPEC max precision 5um
Aligned stage accuracy : X offset, Y offset max 1um
Computer
|  |  | Check item 추가 | MS-office product key
v3.8.2 | 장순영 | 2023-07-26 | TSH(50um, 100um)
|  |  | Check item 추가 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
Y Detector calibration value (MicrometerPerAdc in YDetector DB)
TSH
|  |  | Check item 명 수정 | Full range of X detector raw 로 변경
Full range of Y detector raw 로 변경
Vision
|  |  | Check item 명 수정 | FPS(FramePerSecond) -> FPS(FramePerSecond) of large view
Performance
|  |  | Check item 명 수정 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
%GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
AFM Head
|  |  | Spec 수정 | Full range in voltage during Z Scan sweep : 11V~15V -> 10V~18V 로 변경
Facility Controller
|  |  | Check item 추가 | Cabinet version 기입 추가
양식 변경
|  |  | System 추가 | NX-HDM 시스템 추가
v3.9.2 | 장순영 | 2023-08-10 | AFM Head(Long)
|  |  | Spec 수정 | Full range in um during Z Scan sweep Spec 수정 : min 39 max 46um
AFM Controller
|  |  | Spec 수정 | Firmware verision 추가
NX-HDM
|  |  | Check item 수정 | XY Stage : Serial Port #, BaudRate 삭제
Facility Controller : Serial Port # 1로 변경
Facility Controller : Ionizer 세분화
Sample chuck : Sample holder 세분화
EFEM(NX-Wafer)
|  |  | 오타수정 | Logosol 수정
Sample Chuck(w/ SiC Anti-warpage)
|  |  | 항목 세분화 | Allowable max warpage, w/ SiC Anti-warpage sample chuck Smile type/Crying type 세분화 스펙 지정
AFM Head/TSH
|  |  | Check item 추가 | Z Linearity by 5 different sample height : 300mm 인경우
Z Linearity by 4 different sample height : 200mm/150mm 인경우
v3.10.2 | 장순영 | 2023-09-01 | XY Scanner
|  |  | Check item 추가 | Check the full stroke image(servo on) 추가
|  |  | Check item 추가 | Full range in um during X1/X2/Y1/Y2 Scan sweep
양식 변경
|  |  | System 추가 | NX-PSS 시스템 추가(AFM1+EFEM, AFM2)
SW
|  |  | Spec 수정 | SmartAnalysis verision 수정 : 1.2.1.0 -> 1.4.3.1480 (최신 RTM)
Misc(VLSI)
|  |  | Check item 추가 | Check the VLSI sample model number (Box, Certification, Sample)
Check the VLSI sample serial number (Box, Certification, Sample)
v3.11.2 | 장순영 | 2023-10-16 | WLI Z scanner
|  |  | Check item 명 수정 | WLI_Scale -> WLI_Scale(Z Calibration)
Vision
|  |  | Check item 명 수정 | Check LightSource ON/LightSource(LED) Color 오타 수정
XY Scanner,TSH 50/100um
|  |  | Check item 명 수정 | 아래 8가지에 대해 XY Servo off 명칭 추가
Full range in um during X1 Scan sweep (XY Servo off)
Full range in um during X2 Scan sweep (XY Servo off)
Full range in um during Y1 Scan sweep (XY Servo off)
Full range in um during Y2 Scan sweep (XY Servo off)
Full range in voltage during X1 Scan sweep (XY Servo off)
Full range in voltage during X2 Scan sweep (XY Servo off)
Full range in voltage during Y1 Scan sweep (XY Servo off)
Full range in voltage during Y2 Scan sweep (XY Servo off)
|  |  | Lift Pin | : check module 추가
|  |  | Check item spec 변경 | Lift pins synchronization offset check : spec. ≤ 50um
|  |  | Check item 명 수정 | Check lift pins synchronization -> Lift pins synchronization offset check (spec. ~50um) : X/Y offset
EFEM
|  |  | Check item spec 변경 | OCR score : Ioss 120, Cognex 198
|  |  | Check item 추가 | Wafer placement(Loading) error on the load port1 (X offset)
Wafer placement(Loading) error on the load port1 (Y offset)
Wafer placement(Loading) error on the load port2 (X offset)
Wafer placement(Loading) error on the load port2 (Y offset)
XY Scanner
|  |  | Check item 단위 수정 | ResonanceFrequencyHz/ServoCutoffFrequencyHz : 단위 Hz 추가
System Noise
|  |  | Check item 단위 수정 | 단위 제거
|  |  | XY Stage | *XY Stage controller 없는 시스템
|  |  | Check item spec 변경 | Maker of the motion control board : HCMNEW -> PSC
SW
|  |  | Check item spec 변경 | Cantilever DB Version : 1.0 -> 1.1
AFM Head(NX-Wafer300/200mm)
|  |  | Check item 변경 | Z Linearity by 5 different sample height / *1800nm Step Height Sample (10 Repeats) test는 측정값만 기입, linearity값에 포함X
XY Stage
|  |  | Check item spec 변경 | Encoder model : RGH24H30D30A -> RGH24H (Cable 길이 정보 삭제)
|  |  | Check item 추가 | Align the XY stage origin and sample origin
|  |  | Tip Characterizer Sample | : check module 추가
|  |  | Check item 추가 | Tip Characterizer stage position teach / XEA
Tip Characterizer DB setting : TipCharacterizerSupport
Tip Characterizer auto focus teach / XEA
Tip Characterizer sample width
Misc
|  |  | Check item 추가 | Activate Thermal tune option in XEA
Resonance frequency of AC160 by thermal tune
Resonance frequency of AC160 by frequency sweep
Resonance frequency of AC240 by thermal tune
Resonance frequency of AC240 by frequency sweep
AFM Head
|  |  | Check item 추가 | Check if the largest difference in PZT Alignment is less than 50um
Computer
|  |  | Check item 추가 | Check for the latest update on anti-virus software
Check for the number of viruses detected after the virus scan.
Performance
|  |  | Check item 추가 | Test 샘플 정보 기입란 추가 : Specification of Step height sample for Repeatability test (180nm)
Test 샘플 정보 기입란 추가 :Specification of Step Height Sample for short-term stability test (450nm)
AFM Head
|  |  | Check item 추가 | Test 샘플 정보 기입란 추가 : Specification of Step Height Sample for Z Linearity by Z scanner position test (180nm)
|  |  | Customer | : check module 추가
[YMTC]FAT -> Customer로 변경
v3.12.2 | 장순영 | 2023-12-12 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
EFEM(FFU Open type)
|  |  | Check item spec 변경 | Air pressure @ 60% of FFU : Spec 1.0 -> 1.25
AFM Controller
|  |  | Check item spec 추가 | Firmware version of Motor board : 0x6E31041C 추가
Firmware version of Base board : 0x6E11041C 추가
[Spare Part] AFM Head(Long)
|  |  | Check item 추가 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
Z Height noise : check list 인쇄시 숨김처리
[Spare Part] AFM Head(Standard)
|  |  | Check item 추가 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
Computer
|  |  | Check item 추가 | Check DotNET Framework Version : Spec 4.6이상
TSH
|  |  | Check item spec 변경 | Full range in voltage during Z Scan sweep : 15V이상 -> 11.8V~15.4V(이전 Spec오류)
AFM Head(Wafer,HDM,PSS)
|  |  | Check item 추가 | AFM Head(NX-Wafer200mm) : 4) Under 50nm Step Height Sample (10 Repeats) 추가 |  | 샘플 구입 완료되어 추가됨
XY Stage(Mask)
|  |  | Check item 추가 | Aligned stage accuracy (X/Y offset-Point Align1/2/3) : Aligned stage accuracy 값에서 추출하는 테스트
AFM Head(Wafer,HDM,PSS,WLI)
|  |  | Check item spec 추가 | AFM Head type 4가지로 관리 : Standard/Long -> Auto standard/Auto long/Standard/Long
EFEM(Wafer, WLI, Mask)
|  |  | Check item 명 수정 | RFID Type 오타 수정 : LW -> LF
EFEM
|  |  | 항목 세분화 | Check 250 cycle of wafer load and unload without error (Single loadport)
Check 500 cycle of wafer load and unload without error (Dual loadport)
XY Stage(Wafer/WLI)
|  |  | Check item 추가 | XY Stage positioning X/Y offset (Vision Pattern Matching/In-Die Align/LLC Translation/Global Align)
EFEM(Wafer, WLI, Mask)
|  |  | Spec수정 | Cleanness of EFEM w/ airborne particle measurement SPEC : ISO3 -> ISO1 |  | Price list와 매칭
Laser Scattering(Wafer/HDM)
|  |  | Check item 추가 | Laser intensity from laser source controller
Laser intensity from beam focusing module
v3.13.0 | 장순영 | 2024-01-30 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
Performance(Wafer,HDM,WLI)
|  |  | Check item 추가 | Cycle time for 15-point measurement 추가 |  | AFM, Static Repeatability w/ 180nm Step Height Sample측정시 log 확인 필요, 해당 Recipe 변경됨(Recipe v3.13.0)
XY Stage(Wafer,WLI,Mask)
|  |  | Check item spec 변경 | Sample Loading Position to Y : spec max -20000 추가 |  | XY Stage 체결 문제로 인한 스펙 추가
Z Stage
|  |  | Check item 방법 변경 | Z Drift Recipe 변경, zeroscan recipe와 함께 측정되게 변경
|  |  | Check item spec 단위 변경 | Z Drift spec 단위 변경, 0.3nm/s -> 20nm
AFM Head(Standard)/TSH
|  |  | Check item spec 변경 | Full range in um during Z Scan sweep : Max spec 삭제
AFM Head(Long)
|  |  | Check item spec 변경 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) : 60~70 ->43~92
|  |  | Check item 추가 | Z Height noise spec 확정으로 테스트 추가됨 : spec < 0.1nm
양식 변경
|  |  | 서명란 변경 | 조직개편으로 인한 서명란 수정
|  |  | Check item 문구 변경 | 연구용 Checklist와 명칭 통일(NX-Wafer Sheet에 파란색 음영으로 표시 되어있음)
XY Scanner
|  |  | Check item spec 변경 | X/Y Non-linearity error : ±4%, 0~4%로 수정 |  | 분석시트상 절대값으로 되어있어 수정
TSH
|  |  | Check item spec 변경 | Z Linearity by Z scanner position SPEC 변경 : 1.5% -> 2%
v3.14.0 | 장순영 | 2024-02-26 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
|  |  | 양식 변경 | Category 분류 위치 변경 및 Trend 분류 추가
Core Module의 Module No. 및 Revision 번호 기록 추가
AFM Head(Wafer,HDM,PSS,WLI,Spare)
|  |  | Check item spec 변경 | Z Linearity by Z scanner position SPEC 변경 : 1.5% -> 2% |  | 통일화 진행 (Long/Standard)
|  |  | Check item test 방법 변경 | Long head(2%) : Z scannner position -6,0,+6, Standard head(2%) : Z scannner position -3,0,+3
v3.15.0 | 장순영 | 2024-03-29 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
TSH
|  |  | Check item spec 변경 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) : MAX 12 -> X (Max Spec. 제거)
|  |  |  | TSH Head를 사용하는 장비에서 Z Drift Spec 변경 20 ->45 | 품질개선 24-036
Facility Controller(PB)
|  |  | Check item 추가 | Check the cabinet front/rear/side door alarm 3가지 추가 | 품질개선 24-037 | OBQ 발생
Misc
|  |  | Check item 추가 | Check if the bolts used for the AFM core module are assembled with standardized torque
EFEM(Mask)
|  |  | Check item 추가 | UsebreakInPrepare : 1 | 품질개선 24-054 | Hoya NX-Mask에서 Mask 샘플 손상발견하여, UsebreakInprepare 활성화 필요
SW
|  |  | Check item 추가 | Xms4096m in xea.vmoptions | 품질개선 24-055 | 산업용 장비 PC에서 고용량 Recipe 사용 시 SW Freezing 되는 현상 발생 (Hoya Mask)
Xmx16384m in xea.vmoptions
WLI Repeatability(WLI)
|  |  | Check item test 방법 변경 | 아래 두 테스트 모두 출고되는 가장 큰 렌즈 사용하여 측정 진행
Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
WLI Vision(WLI)
|  |  | Check item 추가 | Port1 Lens magnification
Port2 Lens magnification
|  |  | Check item 명 및 Check item test 방법 변경 | 아래 테스트 모두 출고되는 가장 큰 렌즈 사용하여 측정 진행
Angle Align Between Port1 Lens and Sample Chuck: X
Angle Align Between Port1 Lens and Sample Chuck: Y
Angle Align Between Port2 Lens and Sample Chuck: X
Angle Align Between Port2 Lens and Sample Chuck: Y
Angle Align Between Port2 Lens and VLSI SHS: X
Angle Align Between Port2 Lens and VLSI SHS: Y
Lens Port1 (to OnAxis) Offset
Lens Port2 (to OnAxis) Offset
Port1 Lens Background Effect Remove Setup
Port2 Lens Background Effect Remove Setup
Port1 Lens Resolution
Port2 Lens Resolution
|  |  | AFM Head, TSH | 아래 항목들 앞에 순번 제거 |  | 관리도 항목으로 통일화 목적
|  |  | Check item 명 변경 | - 940nm Step Height Sample (10 Repeats)
- 200~500nm Step Height Sample (10 Repeats)
- 50~200nm Step Height Sample (10 Repeats)
- Under 50nm Step Height Sample (10 Repeats)
|  |  | Elec-regulator(Spare Part) | 아래 DB 설정 항목들 추가 및 기능 확인 내용 추가
XEA version
XEService version
Check the serial communication
Detector - bUse
Detector - bUseRegulator
Regulator - Info - bSimulation
Regulator - Info - bUse
Regulator - Info - bUseLog
Regulator - Info - Delay
Regulator - Info - DetectThreshold
Regulator - Info - FS
Regulator - Info - HighVacuum
Regulator - Info - LowVacuum
Regulator - Info - Maker
Regulator - Info - UpdataDelayMillisec
Regulator - Info - ValidThreshold
Regulator - SerialPort - PortNo
Regulator - SerialPort - BaudRate
Check if the vacuum changes from -5kPa to -50kPa
Check if the vacuum changes from -50kPa to -5kPa
Check if the system transitions to a high vacuum state after the job is completed
Misc
|  |  | Check item 추가 | Delete unnecessary files and usage records from the computer according to internal guidelines | 품질개선 24-059 | OBQ 발생
|  | 2024-04-05 | System noise | 아래 두 테스트에 사용하는 Tip 변경 (HSC60 -> B100-NCH)
|  |  | Remark 변경 | Lateral(C-D) sensitivity
Lateral(C-D) noise
v3.16.0 | 장순영 | 2024-05-31 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
WLI Z scanner(WLI)
|  |  | Check item spec 변경 | WLI_Scale(Z Calibration) : min30 -> min25~max39 (출고장비 데이터의 3sigma 값으로 선정) | 품질개선 24-079 | Spec out 발생
Elec-regulator(WLI)
|  |  | Check item 추가 | Check the serial communication | WLI System에 Elec-regulator default 적용/Test 가이드 참고
Check XEA DB configuration for elec-regulator
Check if the vacuum changes from -5kPa to -50kPa
Check if the vacuum changes from -50kPa to -5kPa
Check if the system transitions to a high vacuum state after the job is completed
Conductive AFM(Wafer)
|  |  | Check item 추가 | Selectable amplifier gain | C-AFM 산업용적용에 의한 Checklist 내용추가
Noise level of current signal
1) Maximum value of current signal at 106 V/A
2) Maximum value of current signal at 109 V/A
Resistance value of 1kΩ
Resistance value of 1MΩ
Firmware version of FEM
Resistance between GND to Probehand
C-AFM combo probehand for NX
Conductive AFM test sample (Cobalt-Doped Iron Oxide)
CONTSCPt, 3M
25Pt300B, 3M
CDT-FMR, 3M
Computer
|  |  | Check item 추가 | Ethernet card type 추가
|  |  | Vision |  | Vision power cable 적용에 따른 내용 변경
|  |  | Check item 삭제 | Ethernet card type for on-axis camera 항목 삭제
|  |  | Check item 추가 | Supply power of the on-axis camera through PoE
Supply power of the on-axis camera through frame board
AFM Controller
|  |  | Check item spec 변경 | Firmware version of Motor board (0x6E300418, 0x6E310418, 0x6E31041C -> 6e30041e, 6e31041e) | firmware 버전(4.28 -> 4.30)
Firmware version of Base board (0x6E110418, 0x6E100418, 0x6E11041C -> 6e10041e, 6e11041e)
Main body
|  |  | Check item spec 변경 | Firmware version of Frame board (6E250418 -> 6e24041e,6e25041e) | firmware 버전(4.28 -> 4.30)
Facility Controller(mask)
|  |  | Check item 추가 | Check the communication | 품질개선 24-042 | NX-Mask 온습도 센서 Test 추가건/Test 가이드 참고
Check XEA DB configuration for temperature and humidity sensor
v3.17.0 | 장순영 | 2024-07-05 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
Sliding Stage(Wafer,WLI,Spare Part)
|  |  | Check item spec 변경 | Torque Profile (10mm, 200um/s) spec. 6-->2로변경 | Sliding stage Motor current 증가 제품 평가 결과 근거
|  |  |  | Torque Profile (10mm, 20um/s) spec. 6-->2로변경 | NX-STC-SS 1C 버전 사용 시 적용 필수
Torque Profile (50mm, 200um/s) spec. 6-->2로변경
Torque Profile (50mm, 20um/s) spec. 6-->2로변경
PreventionTorqueThreshold spec. 6-->3로변경
Igain 0.1 -> 0.03 로 변경
IGainWhileMoving 0.1 -> 0.03 로 변경
Dgain 40~100 -> 5~20 로 변경
DGainWhileMoving 30~110 -> 5~20 로 변경
Pgain 40~100 -> 5~20 로 변경
PGainWhileMoving 30~110 -> 5~20 로 변경
|  |  | Misc(Wafer,WLI) | Check if the bolts used for the AFM core module are assembled with standardized torque의 Guide 표 변경 | 기술미팅 240327, Sliding stage OPM 측정의 최적 토크 적용 및 ATX Bolt 풀림 방지로 인해 spring washer 추가
XY Stage(Wafer, WLI)
|  |  | Check item spec 변경 | Encoder model 추가 (RGH24H, V2BCH) | Encoder 단종으로 인한 변경
Check encoder reader LED (RGH24H : Good-Green, V2BCH : Good-Blue)
Focus Stage(Wafer, HDM, PSS, WLI)
|  |  | Check item spec 변경 | Encoder model 추가 (RGH24H, V2BCH)
Check encoder reader LED (RGH24H : Good-Green, V2BCH : Good-Blue)
Z Stage(Wafer, HDM, PSS, WLI)
|  |  | Check item spec 변경 | Encoder model 추가 (RGH24H, V2BCH)
Check encoder reader LED (RGH24H : Good-Green, V2BCH : Good-Blue)
양식 변경
|  |  |  | Core Module의 Module No. 및 Revision 번호 기록 삭제 | Packinglist에 추가됨
v3.18.0 | 장순영 | 2024-08-13 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
XY Stage(Wafer, WLI)
|  |  | Check item spec 추가 | 아래 두 항목에 대한 spec 추가 | 품질개선 24-144 | XY stage positioning test 中 LLC Translation test spec 관리
XY Stage positioning X offset (LLC Translation)
XY Stage positioning Y offset (LLC Translation)
|  |  | Check item 추가 | 아래 두 항목에 대한 Check item 추가 | 품질개선 24-144 | XY stage positioning test 中 LLC Translation test spec 관리
XY Stage positioning X 1σ (LLC Translation)
XY Stage positioning Y 1σ (LLC Translation)
Performance(Wafer, HDM, PSS, WLI)
|  |  | Check item 삭제 | 아래 %GRR 3가지 항목 삭제
%GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
%GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
%GRR w/ roughness of non-patterened wafer
Performance(Mask)
|  |  | Check item 추가 | Z-XY Orthogonality (X direction-Left) | Recipe(11.Mask - 01.TSH - TSH_Z-XY Orthogonality) 의 TGF11 Pameter 사용하여 10번 Repeat후 해당 Method로 Analysis 진행한 값 분석시트에 기입
Z-XY Orthogonality (X direction-Right)
Z-XY Orthogonality (Y direction-Left)
Z-XY Orthogonality (Y direction-Right)
Misc(Wafer,WLI,Mask)
|  |  | Check item 추가 | NX-Wafer Ground map 추가 | [MC] NX-Wafer Ground Map Guide 참고
|  |  | T1 경우 | 아래 주요 모듈에 대한 Ground 확인 (spec < 0.5Ω)
Check Facility Controller connect to A/E Cable holder
Check Z Scanner body connect to A/E Cable holder
Check XY Scanner connect to A/E Cable holder
Check ATX Top plate connect to A/E Cable holder
Check XY Stage connect to A/E Cable holder
Check Focus Stage connect to A/E Cable holder
Check Z Stage connect to A/E Cable holder
Check Frame Board connect to A/E Cable holder
|  |  | PB 경우 | 아래 주요 모듈에 대한 Ground 확인 (spec < 0.5Ω)
Check Cabinet Earth Bar connect to A/E Earth Bar
Check AFM Controller connect to Cabinet Earth Bar
Check Z Scanner body connect to A/E Earth Bar
Check XY Scanner connect to A/E Earth Bar
Check ATX Top plate connect to A/E Earth Bar
Check XY Stage connect to A/E Earth Bar
Check Focus Stage connect to A/E Earth Bar
Check Z Stage connect to A/E Earth Bar
Check Frame Board connect to A/E Earth Bar
v.3.18.1 | 장순영 | 2024-08-14 | Computer(HP)
|  |  | Check item spec 변경 | Computer model 명 추가 (Z4G5)
v3.19.0 | 장순영 | 2024-09-23 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
|  |  | Facility Controller(PB) | EMO 위치별 동작 여부 확인 세분화 (기존에 All EMO Switches 확인으로 되어있던 항목을 세부적으로 나눔) | 품질개선 24-159 | Enable Grip EMO Switch 문제 개선 건
|  |  | Check item 추가 | Check the Enable Grip EMO switch function test
Check the Cabinet Rear EMO switch function test
Check the Cabinet Front Sub Box EMO switch function test
Check the Cabinet Plate EMO switch function test
Check the Cabinet Right EMO switch function test
Check the EFEM EMO-1 switch function test
Check the EFEM EMO-2 switch function test
Check the EFEM EMO-3 switch function test
|  |  | Facility Controller(T1/T2) | EMO 위치별 동작 여부 확인 세분화 (기존에 All EMO Switches 확인으로 되어있던 항목을 세부적으로 나눔)
|  |  | Check item 추가 | Check the Cabinet EMO switches function test
Check the EFEM EMO switches function test
Facility Controller(T3)
|  |  | Check item 추가 | Check the Cabinet EMO switch function test
|  |  | Performance(Wafer, HDM, PSS, WLI) |  | 품질개선 24-160 | 산업, 연구장비 Roughness Test Checklist 추가 요청 건
|  |  | Check item 추가 | Roughness Repeatability 1x1um(10Repeats, RSD%)
Roughness Repeatability 80x80um(10Repeats, RSD%)
Vision
|  |  | Check item spec 변경 | Camera model 명 추가(Mako G-507C)
FullPixelHeight 값 추가(1828) : Mako G-507C 추가로 인한 Pixel값 추가
Computer
|  |  | Check item spec 변경 | MS-office product key (MS Office 365 포함 중단으로 인해 포함 유/무 선택할 수 있도록 변경)
|  |  | AFM Head/TSH |  | 품질개선 24-165 | 산업용 Checklist에 QC Checklist 항목 반영 (총20가지)
|  |  | Check item 추가 | Check the warranty sticker
EFEM
|  |  | Check item 추가 | Check if the SID on the EFEM QA report matches the SID on the equipment nameplate
Facility Controller(Wafer,HDM,PSS, WLI, Mask)
|  |  | Check item 추가 | Check the regulator vacuum pressure
Check the pneumatic controller sensor settings (Nomal state : Green)
Focus Stage
|  |  | Check item 추가 | Check the i-marking on the encoder bracket mounting bolts
Misc
|  |  | Check item 추가 | Check if the mouse and keyboard are functioning properly after turning off the power and then turning it back on
Check if the cable holders are installed on both the AE and the Cabinet
Check the AE FFU cleaning
Check the jig bolts for the computer
Check the fixation of Cognex dongle cable on the rear of the PC
XY Scanner
|  |  | Check item 추가 | Check the warranty sticker
Z Stage
|  |  | Check item 추가 | Check the i-marking on the encoder bracket mounting bolts
Z Stage(WLI)
|  |  | Check item 추가 | Check for interference between the Z and F encoder brackets at the Z stage hardware limit position.
Focus Stage(Mask/TSH)
|  |  | Check item 추가 | Check for interface between the focus stage and head cover when lowering the focus stage
SW
|  |  | Check item 추가 | Head Crash DB setting: A+B Interlock
Head Crash DB setting: A-B Interlock
Head Crash DB setting: Z Detector Interlock
Head Crash DB setting: XY Stage Interlock
XY Stage(Wafer,HDM,PSS, WLI, Mask)
|  |  | Check item 추가 | Check the i-marking on the XY stage mounting bolts
Smoke Detector
|  |  | Check item 추가 | Check the smoke detector setting
양식 변경
|  |  | System 추가 | NX-TSH600, NX-TSH1518, NX-TSH2326 추가
v3.19.0 | 장순영 | 2024-09-25 | Misc(PB)
|  |  | Check item 삭제 | Check Cabinet Earth Bar connect to A/E Earth Bar
|  | 2025-01-15 | Computer
|  |  | Check item spec 변경 | OS(Operating System) type(Window11 추가)
Windows version(24H2 추가)
|  | 2025-01-16 | Misc(Wafer)
|  |  | Check item spec 변경 | Check Z Scanner body connect to A/E Earth Bar (0.5Ω ->1Ω)
Check Focus Stage connect to A/E Earth Bar(0.5Ω ->1Ω)
v3.20.0 | 장순영 | 2025-02-21 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
Z Stage
|  |  | Check item 추가 | Z Stage material : S45C | Z Stage 재질변경
AFM Head/TSH
|  |  | Check item spec 변경 | NCM driving% with AC160TS cantilever spec. 변경 1~50% -> 0.1~50%, Amplitude 20nm | 품질개선25-021, 24-071(TSB)
|  |  | Check item 추가 | NCMDriveRangeLiquidMode =1 | 품질개선25-021, 24-071(TSB)
|  |  |  | bDisableNCMDriveReduction =1 | 품질개선25-021, 24-071(TSB)
Computer
|  |  | Check item 삭제 | MS-office product key 삭제 | 24-044(TSB) : MS Office ->Libre Office로 변경
Facility Controller
|  |  | Check item spec 변경 | EMO cover : Cover -> Installed
Conductive AFM
|  |  | Check item 삭제 | 아래 13건 Check Item 삭제
Selectable amplifier gain
Noise level of current signal
1) Maximum value of current signal at 106 V/A
2) Maximum value of current signal at 109 V/A
Resistance value of 1kΩ
Resistance value of 1MΩ
Firmware version of FEM
Resistance between GND to Probehand
C-AFM combo probehand for NX
Conductive AFM test sample (Cobalt-Doped Iron Oxide)
CONTSCPt, 3M
25Pt300B, 3M
CDT-FMR, 3M
|  |  | eAFM(Wafer) | 아래 9건 Check item 추가 | eAFM 초도 완료건
|  |  | Check item 추가 | Selectable amplifier gain
Firmware version of FEM
Check LED lighting of FEM when switching modes (C-AFM, On / EFM, Off)
Check add eAFM Mode to Head Mode
Check the Bias Voltage (SampleChuck, EFM&C-AFM Tip Initialization sample at Sample Holder)
C-AFM combo probehand for NX
EFM&C-AFM Tip Initialization sample (HOPG, Gold Tip 2ea)
AD-42-AS, 12M
CDT-FMR, 12M
|  |  | eAFM(Conductive AFM) (Wafer) | 아래 6건 Check Item 추가
|  |  | Check item 추가 | Noise level of current signal
Maximum value of current signal at 106 V/A (10V, Sample Bias)
Maximum value of current signal at 109 V/A (100mV, Sample Bias)
Resistance value of 10MΩ
Resistance between GND to Probehand
Save Conductive AFM Image (Cobalt-Doped Iron Oxide Sample)
|  |  | eAFM(EFM) (Wafer) | 아래 13건 Check Item 추가
|  |  | Check item 추가 | Check if EFM amplitude shows the magnitude of electrical signal
Check if the EFM phase shows 180° phase difference for the positive and negative bias
EFM Phase : For the positive bias
EFM Phase : For the negative bias
Lock-In configuration : Input channel
Lock-In configuration : Phase
Lock-In configuration : Time constant
Lock-In configuration : Sensitivity
Lock-In configuration : Filter order
Modulator Configuration : Output signal
Modulator Configuration : Frequency
Modulator Configuration : Drive
Save EFM Image (EFM Stripe-patterned Surface Sample)
|  |  | eAFM(Tip Initializer) | 아래 9건 Check Item 추가
|  |  | Check item 추가 | Tip Initializer DB setting : TipInitialiizerSupport
Tip Initializer stage position teach / XEA (EFM)
Tip Initializer stage position teach / XEA (DC C-AFM)
Tip Initializer stage position teach / XEA (AC C-AFM)
Tip Initializer auto focus teach / XEA
Check the Run Tip Initializer (EFM, Step1) / XEA
Check the Run Tip Initializer (EFM, Step2) / XEA
Check the Run Tip Initializer (DC, C-AFM) / XEA
Check the Run Tip Initializer (AC, C-AFM) / XEA
양식 변경
|  |  | 인쇄 양식 변경 | 글자 크기변경(10->15) | 품질개선25-014, 인쇄시 시인성 향상
Performance
|  |  | Check item spec 추가 | Roughness Repeatability spec 수립, RSD < 5% or 1σ <0.01nm (둘 중 하나 만족하면 Pass) | 품질개선25-027
Facility Controller(PB)
|  |  | Check item 문구 수정 | Check the cabinet front door alarm -> Check the cabinet front door alarm and buzzer | 품질개선25-026
Sample Chuck(w/ SiC Anti-warpage)
|  |  | Check item spec 변경 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) : min 900 -> max 800 | 품질개선25-036 | Warpage sample 분실로 인해 spec 변경
Sample Chuck(w/ Sliding stage)
|  |  | Check item spec 변경 | Check sample surface mapping : 50um -> 30um@300mm, 50um@200mm | 품질개선25-034 | Sim tape 작업으로 인한 평탄도 개선, 300mm system
XY Scanner(w/ SiC Anti-warpage chuck)
|  |  | Check item spec 변경 | Y Detector noise (XY Servo on) : 0.8nm -> 0.6nm | 품질개선25-058 | 경량화 Anti-warpage sample chuck 투입으로 인한 XY Detector noise spec 변경 건
Y1 Detector noise (XY Servo on) : 0.9nm -> 0.7nm
Y2 Detector noise (XY Servo on) : 0.9nm -> 0.7nm
양식 변경
|  |  | System 추가 | NX-TSH300 Overlay | 초도평가 완료 후 추가
v3.20.1 | 장순영 | 2025-03-06 | *즉시 적용건
XY Stage(Wafer, HDM, PSS, WLI, Mask)
|  |  | Check item 삭제 | 아래 16가지 항목 삭제 | 품질개선 25-072 | XY Stage Repeat Test Repeat 횟수 조정 件. (삭제로 변경)
X axis, Repeatability @1st position
X axis, Repeatability @ 2nd position
X axis, Repeatability @ 3rd position
X axis, Repeatability @ 4th position
X axis, Accuracy @1st position
X axis, Accuracy @ 2nd position
X axis, Accuracy @ 3rd position
X axis, Accuracy @ 4th position
Y axis, Repeatability @1st position
Y axis, Repeatability @ 2nd position
Y axis, Repeatability @ 3rd position
Y axis, Repeatability @ 4th position
Y axis, Accuracy @1st position
Y axis, Accuracy @ 2nd position
Y axis, Accuracy @ 3rd position
Y axis, Accuracy @ 4th position
XY Scanner
|  |  | Check item 추가 | Calibration 오차 기준 설정에 따른 Check item 추가 | 품질개선 25-075 | X,Y,Z Calibration QC Spec Checklist 반영 요청 건
X Detector stroke calibration error (nominal scale error X) : <spec.0.5%
Y Detector stroke calibration error (nominal scale error Y) : <spec.0.5%
TSH
|  |  | Check item 추가 | Calibration 오차 기준 설정에 따른 Check item 추가
X Detector stroke calibration error (nominal scale error X) : <spec.0.5%
Y Detector stroke calibration error (nominal scale error Y) : <spec.0.5%
AFM Head/TSH
|  |  | Check item 추가 | Specification of Step Height Sample for Z Calibration
Z Height calibration error : <spec.0.5%
- Z Height measurement value of the calibration sample
Z Drive calibration error : <spec.0.5%
- Z Drive measurement value of the calibration sample
Performance(TSH600)
|  |  | Check item 추가 | Z-XY Orthogonality (X direction-Left) | TSH600만 누락된 항목으로 추가
Z-XY Orthogonality (X direction-Right)
Z-XY Orthogonality (Y direction-Left)
Z-XY Orthogonality (Y direction-Right)
v3.20.2 | 장순영 | 2025-03-14 | Computer
|  |  | Check item 추가 | Check the installation status of the display refresh program | 품질개선 25-087 | 산업용 장비 Display Refresh Setting 관리 방법 논의 요청 건
EFEM(Wafer, WLI)
|  |  | Check item 추가 | Measurement of surface resistance of the upper end-effector : SPEC. 10^6~10^9 | 품질개선 25-084 | End Effector 개선 PAD 확산 적용 검토 요청의 건
Measurement of surface resistance of the lower end-effector : SPEC. 10^6~10^9
|  | 2025-03-24 | Misc
|  |  | Check item 기준 변경 | Check if the bolts used for the AFM core module are assembled with standardized torque | 품질개선 25-096 | 산업용 System XY Scanner 체결 Bolt 변경 요청 건
Bolt Size 변경(XY Scanner 체결 볼트)
|  | 2025-03-26 | XY Stage | 아래 항목들 중 1σ 항목 추가 및 XY Stage positioning Spec 추가
|  |  | Check item 추가 및 Spec 추가 | XY Stage positioning X offset (Vision Pattern Matching) : spec. <4um | 품질개선 25-103 | 산업용 XY Stage positioning test spec 수립
XY Stage positioning Y offset (Vision Pattern Matching) : spec. <4um
XY Stage positioning X 1σ (Vision Pattern Matching) : spec. <0.8um
XY Stage positioning Y 1σ (Vision Pattern Matching) : spec. <0.8um
XY Stage positioning X offset (In-Die Align) : spec. <4um
XY Stage positioning Y offset (In-Die Align) : spec. <4um
XY Stage positioning X 1σ (In-Die Align) : spec. <0.8um
XY Stage positioning Y 1σ (In-Die Align) : spec. <0.8um
XY Stage positioning X offset (Global Align) : spec. <7.5um
XY Stage positioning Y offset (Global Align) : spec. <7.5um
XY Stage positioning X 1σ (Global Align) : spec. <2.2um
XY Stage positioning Y 1σ (Global Align) : spec. <2.2um
|  | 2025-05-30 | SW
|  |  | Check item spec 변경 | Cantilever DB version : 1.1 ->1.3 | eAFM Default cantilever DB 추가
v3.21.0 | 장순영 | 2025-06-17 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
SW(Wafer, WLI, Overlay)
|  |  | Check item 추가 | 'Dcom.psia.xea.useInDieAlign=true' in xea.vmoptions | 품질개선 25-134 | XY stage positioning test를 위한 In-Die align VMOPTION 관련 Checklist 항목 추가 요청의 件
양식 변경
|  |  | System 추가 | NX-Hybrid WLI(Interferom) | 초도평가 완료 후 추가
Misc(Mask, Overlay)
|  |  | Check item 삭제 | Check XY Scanner connect to A/E Earth Bar | TSH만
Misc(Mask, Overlay)
|  |  | Check item spec 변경 | Check Z Scanner body connect to A/E Earth Bar : Spec.<1.5Ω | TSH만
Facility(PB, Overlay, TSH600)
|  |  | Check item 추가 | Check if the power lamps(Main/EFEM/AFM) turn on when the main power is on | 품질개선 25-133
Facility(T1)
|  |  | Check item 추가 | Check if the power lamps(Power Cable/UPS Power) turn on when the main power is on | 품질개선 25-133
AFM Head, TSH
|  |  | Check item 추가 | Check if the driving % changes proportionally during NCM modulation | 품질개선 25-160
XY Scanner
|  |  | Check item spec 변경 | 아래 다섯개의 항목 spec. 110≤ -> 150~240 으로 변경 | 품질개선 25-178
Full range of X Detector during X Scan sweep
Full range of X1 Detector during X Scan sweep
Full range of X2 Detector during X Scan sweep
Full range of Y Detector during Y Scan sweep
Full range of Y1 Detector during Y Scan sweep
Full range of Y2 Detector during Y Scan sweep
|  |  | Sample chuck(mask) |  | 품질개선 25-164
|  |  | Check item 추가 | Check if the lock stage has been taught
Check if any alarm occurs during the lock stage repeat
|  |  | EFEM(single, dual) |  | 품질개선 24-216
|  |  | Check item 추가 | Check the load button on the load port1 of the EFEM for any abnormalities
Check the load button on the load port2 of the EFEM for any abnormalities
|  |  | Vision |  | 품질개선 25-180
|  |  | Check item 추가 | Check the Vision brightness calibration
LightStrengthGain
SW
|  |  | Check item spec 변경 | Cantilever DB version : 1.1 ->1.3
v3.21.1 | 장순영 | 2025-06-24 | *Checklist에 음영 표시건이 개정내용
*즉시 적용건은 노란색 음영으로 표기
AFM Head, TSH
|  |  | Check item 삭제 | NCM driving% with AC160TS cantilever
|  |  | Check item 변경 | Check if the driving % changes proportionally during NCM modulation을 아래 문장으로 변경 | 품질개선 25-160
Check if the driving % changes proportionally as the NCM amplitude changes
|  |  | Check item 추가 | 위 테스트를 세분화하여 수정, 아래 체크 item 3가지 추가 | 품질개선 25-160
NCM driving % at an amplitude of 1 nm
NCM driving % at an amplitude of 20 nm
NCM driving % at an amplitude of 100 nm
Performance(Mask, Overlay)
|  |  | Check item 추가 | Mask Roughness Repeatability Test 1x1um(10Repeats)
Mask Roughness Repeatability 1x1um(RSD%)
Mask Roughness Repeatability 1x1um(1σ)
Mask Roughness Repeatability Test 40x40um(10Repeats)
Mask Roughness Repeatability 40x40um(RSD%)
Mask Roughness Repeatability 40x40um(1σ)
|  |  | Check item spec 추가 | Roughness Repeatability spec 수립, RSD < 5% or 1σ <0.01nm (둘 중 하나 만족하면 Pass)
|  | 2025-07-01 | Performance(Wafer, HDM, WLI)
|  |  | Check item 문구 수정 | Short-term stability using Step Height Sample (3σ) | 450nm 샘플과180nm 샘플 둘 중 하나 사용으로 변경됨에 따라 문구 변경
|  | 2025-08-25 | Z Stage(HDM, HDM EFEM, PSS)
|  |  | Check item spec 변경 | Z Stage material S45C -> AL6061 | 해당 장비들의 z stage의 재질변경 제외 건들
|  | 2025-08-29 | EFEM(HDM EFEM)
|  |  | Check item 추가 | Aligner maker model에 추가 (기존에는 bluecord만 있었음 logosol 추가함) : logosol spec. 0.04