## Equipments
No | SID | 장비명 | 종료일 | R/I | Model | XY Scanner | Head Type | MOD/VIT | Sliding Stage | Sample Chuck | AE
1 | C160107-030225 | BYD Semiconductor(Jinan) | 2025-11-30 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units | 10mm | SiC Anti-warpage chuck | Double Walled
2 | J150021-250625 | LAPIS Semiconductor | 2025-10-29 | Research | NX20 300mm | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
3 | C180006-180625 | ZJEAGLES | 2025-10-28 | Industrial | NX-HDM | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck
4 | BR10007-190825 | Brazilian Center for Research in Physics | 2025-10-27 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
5 | C160110-150525 | OEINO | 2025-10-27 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
6 | D280037-040825 | 램테크놀러지㈜ | 2025-10-21 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
7 | TH21005-261124 | WD Storage Technologies Ltd #3 | 2025-10-21 | Industrial | NX-PTR | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  |  | Double Walled
8 |  | LG Display #3 | 2025-10-21 | Industrial | NX-TSH1518 | Single 100µm | TSH 100µm |  |  |  | Isolated
9 | TH21006-261124 | WD Storage Technologies Ltd #4 | 2025-10-21 | Industrial | NX-PTR | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  |  | Double Walled
10 | TH21007-261124 | WD Storage Technologies Ltd #5 | 2025-10-21 | Industrial | NX-PTR | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  |  | Double Walled
11 |  | PSI EAST OFFICE | 2025-10-16 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
12 |  | Kitagawa Gresstech Co., Ltd | 2025-10-16 | Research | NX20 300mm | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
13 |  | Zhejiang SDS | 2025-10-14 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
14 | TH21003-261124 | WD Storage Technologies Ltd #1 | 2025-10-14 | Industrial | NX-PTR | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  |  | Double Walled
15 | TH21004-261124 | WD Storage Technologies Ltd #2 | 2025-10-14 | Industrial | NX-PTR | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  |  | Double Walled
16 |  | Fujian Jingan Optoelectronics | 2025-10-13 | Industrial | NX-PSS | Dual 100µm | Standard | Single MOD 2 units |  | AL Bar type chuck | Double Walled
17 |  | Micron F6 | 2025-10-02 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
18 |  | 포항공과대학교_화학공학과 | 2025-10-01 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
19 |  | Dalian Institute of Chemical Physics | 2025-10-01 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
20 |  | CNR-IFN | 2025-09-29 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
21 |  | Intel Hillsboro #5 | 2025-09-26 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
22 |  | LMU Munich Nanophotonics Group | 2025-09-24 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
23 |  | University of Electronic Science and Technology of China | 2025-09-23 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE203
24 |  | Analog Devices GEN, Trias INC. | 2025-09-22 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
25 |  | 한국에너지기술연구원 | 2025-09-19 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | Glove Box
26 |  | UNIST_신소재공학과 | 2025-09-18 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
27 |  | Intel Hillsboro#6 | 2025-09-17 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
28 |  | PSS Demo | 2025-09-12 | Research | FX40 | Single 100µm | FX Standard
29 |  | Clemson University | 2025-09-12 | Research | NX-Hivac | Single 100µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
30 |  | Lumentum | 2025-09-12 | Industrial | NX-HDM | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck | Double Walled
31 |  | Uzkimyosanoat Joint Stock Company | 2025-09-10 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
32 |  | Momentive Performance Materials (India) | 2025-09-08 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
33 |  | 서울대학교_에너지자원공학과 | 2025-09-03 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | Glove Box
34 |  | 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
35 |  | 한국나노마이스터고등학고 | 2025-08-29 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
36 |  | IOE CAS (Chengdu) | 2025-08-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Bar type chuck | Double Walled
37 |  | Samsung(SEC) P2D | 2025-08-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
38 |  | Sotacarbo Spa | 2025-08-20 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
39 | D10200-200325 | 고려대학교_융합대학원 | 2025-08-19 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
40 |  | 공주대학교 공동실험실습관 | 2025-08-11 | Research | FX40 | Single 100µm | Auto Align Standard | Mini450F |  |  | FX40 AE
41 |  | 한국생산기술연구소 | 2025-08-11 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
42 |  | Galaxycore(2) | 2025-08-11 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
43 |  | GRANDIT#2 | 2025-08-08 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
44 |  | 강원대학교 전기전자공학과 | 2025-08-06 | Research | NX7 | Single 50µm | Standard | Accurion i4 | 선택하세요 | 선택하세요 | AE101
45 | T150014-050325 | Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE401
46 |  | National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  |  | AE202
47 |  | RATC Demo | 2025-08-05 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  | Vacuum Sample Chuck | Chamber
48 |  | Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
49 |  | University of Padova | 2025-08-01 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
50 |  | 강원대학교_공동실험실습관 | 2025-07-30 | Research | FX40 | Single 100µm | Standard | Mini450F |  |  | FX40 A/E
51 |  | Intel Hillsboro#4 | 2025-07-29 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
52 |  | University of North Carolina | 2025-07-25 | Research | NX10 | Single 50µm | SICM Long Head | Accurion i4 |  |  | AE101
53 | U10094-020625 | University of California at Berkeley | 2025-07-25 | Research | NX10 | Single 50µm | Standard | Accurion i4
54 | U28013-270525 | Analog Devices Inc | 2025-07-24 | Research | FX40 | Single 100µm | Standard | Mini450F |  |  | FX40 A/E
55 | D10204-150525 | HAESUNG DS | 2025-07-23 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
56 |  | Samsung NX-TSH 400TF | 2025-07-23 | Industrial | NX-TSH400TF | 선택하세요 | TSH 100µm | Dual MOD 8 units | 50mm | Porous Sample Chuck | Isolated
57 | D30026-300525 | 아주대학교_기계공학과 | 2025-07-17 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
58 | IN10024-090525 | Vellore Institute of Technology | 2025-07-10 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
59 | U33001-120324 | Micron | 2025-07-10 | Industrial | NX-Wafer Plus | Dual 100µm(300mm) | Auto Align Standard |  | 50mm | SiC Anti-warpage chuck | Isolated
60 | F10009-220525 | Neel Institute | 2025-07-08 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
61 | J10042-080425 | Tomoegawa Corporation | 2025-07-07 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
62 | U10092-310325 | University of Utah | 2025-07-07 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
63 | D30025-070525 | 한양대학교 융합전자공학과 | 2025-07-04 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
64 | S22001-180325 | Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Research | NX-Hivac | Single 100µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
65 | C28008-070425 | University of Chinese Academy of Sciences | 2025-06-30 | Research | FX40 | Single 100µm | Auto Align Standard | Mini450F |  |  | FX40 AE
66 | T30002-230425 | Chung Yuan Christian University | 2025-06-27 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
67 | C26007-050824 | Wuxi DIS Microelectronics | 2025-06-25 | Industrial | NX-Mask | 선택하세요 | TSH 50µm | Dual MOD 8 units | 선택하세요 | Mask | Isolated
68 | IC10001-050525 | University of Iceland | 2025-06-24 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
69 | C160101-271124 | United Nova Technology | 2025-06-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units |  | AL Bar type chuck | Double Walled
70 | U28012-090425 | Silicon Labs | 2025-06-23 | Research | FX40 | Single 100um | Auto Align Standard | Mini450F |  |  | FX40 AE
71 | C10103-070325 | Henkel Corporation, Ltd. | 2025-06-17 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
72 | D150027-130325 | 에이티아이 | 2025-06-17 | Research | NX20 300mm | Dual 100µm | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
73 | C30012-280425 | Xidian University Guangzhou Institute of Technology | 2025-06-16 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
74 | D10202-300425 | 도레이첨담소재 | 2025-06-16 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
75 | T25011-210125 | TSMC AP3(400TF) | 2025-06-15 | Industrial | NX-TSH400TF |  | TSH 100µm | Dual MOD 8 units +Passive MOD 2 units | 50mm | Porous Sample Chuck | Isolated
76 | U27005-100225 | Intel Hillsboro#3 | 2025-06-13 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
77 | G27001-221024 | PSE demotool | 2025-06-13 | Industrial | NX-Hybrid WLI(Interferom Ver.) | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
78 | F28007-240325 | PSF Demo Lab | 2025-06-12 | Research | FX40 | Single 100um | Auto Align Standard | Mini450F |  |  | FX40 AE
79 |  | Dupont Taiwan | 2025-06-11 | Research | NX20 | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
80 |  | 울산과학기술원 물리학과 | 2025-06-10 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
81 | J28003-170225 | PSJ DEMO | 2025-06-04 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
82 | N34001-171224 | National Metrology Institute | 2025-06-04 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
83 | C160103-041224 | IME CAS(Beijing) | 2025-06-04 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Bar type chuck | Double Walled
84 | D10201-310325 | 단국대학교_융합반도체공학과 | 2025-06-02 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
85 | E34001-100325 | PS UK Demo Lab | 2025-06-02 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
86 | C26008-300924 | QMOT(Quanyi)(3) | 2025-06-02 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
87 | C160099-250924 | Dongguan Guangmao Technology(3) | 2025-06-02 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
88 | D10200-200325 | 고려대학교 융합대학원 | 2025-05-29 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
89 | C10102-060325 | The Second Research Institute of CAAC | 2025-05-29 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
90 | G160013-060125 | Infineon Dresden(4) | 2025-05-29 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
91 | C30011-200325 | Shanxi Institute of Technology | 2025-05-28 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
92 | D10199-240325 | 광주과학기술원_반도체공학과 | 2025-05-20 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
93 | I30001-250325 | University of Siena | 2025-05-19 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
94 | SW30002-010425 | University of Zurich | 2025-05-19 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE203
95 | T150014-050325 | Hon Hai Technology Group (FOXCONN) | 2025-05-16 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
96 | C150145-200225 | Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
97 | G34003-171224 | Walther-Meissner-Institute | 2025-05-15 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
98 | 043-0000-005 | ATC Demo | 2025-05-15 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
99 | PL30002-240325 | AGH University | 2025-05-12 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
100 | G30004-120325 | University of Siegen, Germany | 2025-05-09 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE203
101 | C31008-130225 | Wuhan Cyanopto Technology | 2025-05-07 | Research | NX15 | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
102 | D150026-020125 | 아이큐랩 | 2025-05-07 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
103 | U10091-240225 | Morgan State University | 2025-05-02 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
104 | D30023-080225 | 울산과학기술 신소재공학과(UNIST) | 2025-04-29 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
105 | T10015-050325 | Taiwan Instrument Research Institute(TIRI) | 2025-04-29 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
106 | I10020-100125 | CNR-Nanotec(Lecce) | 2025-04-28 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
107 | C150143-200125 | BestBay Co.,Limited | 2025-04-28 | Research | NX20 300mm | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
108 | D280032-080125 | 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
109 | T34001-050325 | Taiwan Instrument Research Institute | 2025-04-22 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
110 | U30009-200125 | Purdue University | 2025-04-21 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
111 | G28005-310125 | PSE demo lab (1/2) | 2025-04-21 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
112 | RO10002-200225 | Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
113 |  | Samsung (SEC) SRD | 2025-04-16 | Industrial | NX-Wafer Plus | Dual 100µm(300mm) | Auto Align Standard |  | 50mm | SiC Anti-warpage chuck | Isolated
114 |  | UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
115 | C32018-301224 | Guangdong Huicheng Vacuum Technology | 2025-04-15 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
116 | C10099-070125 | North China University of Technology | 2025-04-14 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
117 |  | University of California, Davis | 2025-04-14 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
118 | J150018-301224 | Kyocera Corporation | 2025-04-14 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
119 | C160100-291024 | JCET Microelectronics | 2025-04-14 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
120 | C150144-220125 | Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
121 | C10098-030125 | East China Normal University | 2025-04-10 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
122 |  | Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
123 | U150051-301224 | Accretech America Inc. | 2025-04-08 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
124 | C160106-220125 | Hubei Dinghui | 2025-04-07 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
125 | U22008-200125 | PSI Demo | 2025-04-04 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
126 | C160105-060125 | Zhejiang ICTruth | 2025-04-04 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
127 |  | Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
128 | U23013-291124 | University of Texas at Austin | 2025-04-01 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE203
129 | C10097-020125 | Zhejiang Sci-tech University | 2025-04-01 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
130 | 012-0020-002 | App Team Demo | 2025-04-01 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
131 |  | SZMTC#4 | 2025-04-01 | Industrial | NX-PSS | Dual 100µm | Standard | Single MOD 2 units |  | AL Bar type chuck | 선택하세요
132 |  | SZMTC#5 | 2025-04-01 | Industrial | NX-PSS | Dual 100µm | Standard | Single MOD 2 units |  | AL Bar type chuck | 선택하세요
133 | D10198-231224 | 가천대학교 전자공학부 | 2025-03-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
134 |  | SZMTC | 2025-03-28 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
135 |  | 가천대학교_전자공학부 | 2025-03-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
136 | C30010-310125 | Northwest University | 2025-03-27 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
137 | C32017-251124 | Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
138 | C33001-050824 | SwaySure Technology | 2025-03-26 | Industrial | NX-Wafer Plus | Dual 100µm(300mm) | Auto Align Standard |  | 50mm | SiC Anti-warpage chuck | Isolated
139 |  | Samsung NRD-K eAFM | 2025-03-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Fork type chuck | Isolated
140 | U30008-141124 | University of California-San Diego | 2025-03-21 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
141 | C150142-101224 | Xinping Semiconductor (Shenzhen) | 2025-03-20 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
142 | C32016-211024 | Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
143 |  | University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
144 | D10197-170125 | 서울대학교_농림생물자원학부 | 2025-03-13 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
145 | TH30001-241224 | Thammasat University | 2025-03-13 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
146 |  | Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
147 |  | TSMC F22(1) | 2025-03-09 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Fork type chuck | Isolated
148 | C160098-250924 | Dongguan Guangmao Technology(2) | 2025-03-07 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
149 |  | TSMC AP2 Mini Line (4) | 2025-02-28 | Industrial | NX-TSH |  | TSH 100µm | Dual MOD 6 units |  | Coreflow customized | Isolated
150 |  | TSMC F22(2) | 2025-02-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | Fork type | Isolated
151 |  | Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
152 | D30022-031224 | 충북대학교 신소재공학과 | 2025-02-25 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
153 |  | Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
154 | C28007-181124 | The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40 AE
155 |  | BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
156 | C10095-111024 | Beijing Institute of Technology | 2025-02-14 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
157 |  | 한양대학교_물리학과 | 2025-02-13 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
158 | G34002-030924 | Bosch Reutilngen | 2025-02-13 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
159 | I10019-181224 | INRIM Torino | 2025-02-13 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
160 | C150137-021024 | Huntersun Huzhou | 2025-02-12 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
161 | BR30001-141124 | Federal University of Minas Gerais | 2025-02-11 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
162 | D160031-300824 | Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Industrial | NX-Wafer | Dual 100µm | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
163 | G150036-121124 | Innolume GmbH | 2025-02-07 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
164 | U34003-130924 | TIE Cerium Laboratories | 2025-02-05 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
165 |  | PSE demo lab | 2025-02-03 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
166 | D160028-160724 | SK하이닉스 M11 (Flash) | 2025-02-03 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
167 |  | TSMC AP2 Mini Line(3) | 2025-01-31 | Industrial | NX-TSH |  | TSH 100µm | Dual MOD 6 units |  | Coreflow customized | Isolated
168 |  | TSMC AP2 Mini Line(3) | 2025-01-31 | Industrial | NX-TSH |  | TSH 100µm | Dual MOD 6 units |  | Coreflow customized | Isolated
169 | C23006-211024 | Tianjin Medical University | 2025-01-21 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
170 | C10096-061124 | China Aluminum Science and Technology Research Institute | 2025-01-13 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
171 | MY150002-091024 | NILT Malaysia Sdn. Bhd. | 2025-01-13 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
172 | G10036-101224 | Free University of Berlin | 2025-01-13 | Research | NX10 | Single 10µm | Standard | Accurion i4 |  |  | AE203
173 |  | Xuzhou Meixing OE Technology | 2025-01-10 | Industrial | NX-PSS | Dual 100µm | Standard | Single MOD 2 units |  | AL Bar type chuck | Double Walled
174 | C24005-190924 | Xuzhou Meixing OE Technology(NX-PSS) | 2025-01-10 | Industrial | NX-PSS | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | AL Bar type chuck
175 |  | HEPIA | 2025-01-09 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
176 | S10012-060623 | UMC Singapore | 2025-01-03 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
177 | C30007-290824 | Shanghai Polytechnic University | 2024-12-31 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
178 |  | Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
179 |  | Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
180 | C160091-290424 | Huahong Wuxi #6 | 2024-12-27 | Industrial | NX-Wafer Basic | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
181 |  | Shenzhen JWKW(2) | 2024-12-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
182 | C160095-300724 | Shenzhen JWICW #1 | 2024-12-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units | 10mm | SiC Anti-warpage chuck | Double Walled
183 |  | Shenzhen JWKW (NX-HDM) | 2024-12-24 | Industrial | NX-HDM | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | AL Bar type chuck
184 | C31007-300824 | Fudan University | 2024-12-23 | Research | NX15 | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
185 |  | NASA Kennedy Space Center KSC | 2024-12-23 | Research | NX-Hivac | Single 100µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
186 |  | 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
187 |  | Micron F16(3) (Taiwan) | 2024-12-23 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
188 | T25004-160624 | TSMC AP2 Mini Line #2 | 2024-12-20 | Industrial | NX-TSH |  | TSH 100µm | Dual MOD 6 units |  | Coreflow customized | Isolated
189 |  | 서강대학교_화공생명공학과 | 2024-12-19 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
190 |  | 인하대학교_물리학과 | 2024-12-17 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
191 |  | University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
192 | G160011-260923 | Fraunhofer IPMS Dresden | 2024-12-13 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Single MOD 6 units | 50mm | SiC Flat type chuck | Double Walled
193 |  | EVGroup | 2024-12-13 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Bar type chuck | Isolated
194 | J150017-260924 | Omron | 2024-12-12 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
195 |  | CXT #9 | 2024-12-10 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Double Walled
196 |  | CXT #8 | 2024-12-10 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Double Walled
197 |  | University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
198 | J160026-120624 | JSR Corp. | 2024-12-06 | Industrial | NX-Wafer Basic | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 10mm | SiC Anti-warpage chuck | Double Walled
199 |  | NVIDIA (Chuan Zhang) | 2024-12-02 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
200 |  | Crystaland Co.,Ltd. | 2024-11-25 | Research | NX15 | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
201 |  | Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
202 |  | 전남대학교_고분자융합소재공학부 | 2024-11-20 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
203 |  | HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
204 | C160092-290424 | BYD Chengdu | 2024-11-18 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
205 |  | Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
206 |  | 충남대학교_기계공학부 | 2024-11-13 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
207 |  | Indian Institute of Technology-HYD | 2024-11-08 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
208 |  | Istanbul Technology University (ITU) | 2024-11-08 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
209 | C10093-160724 | Zhejiang Sci-Tech University | 2024-11-04 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
210 |  | Hauhong Wuxi #5 | 2024-11-01 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
211 |  | Dongguan Guangmao Technology | 2024-11-01 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
212 |  | 강원대학교_물리학과 | 2024-10-30 | Research | FX40 | Single 100µm | Auto Align Standard | Mini450F |  |  | FX40AE
213 |  | 한양대학교_화학과 | 2024-10-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
214 |  | Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
215 |  | Nanyang Technological University_CEE | 2024-10-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
216 |  | Huawei Technologies | 2024-10-26 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Pin Bar type chuck | Double Walled
217 |  | Hsingchen Technology | 2024-10-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
218 |  | NUS CA2DM | 2024-10-24 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
219 |  | Arizona State University (Umberto Celano) | 2024-10-23 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
220 |  | Black Semiconductor GmbH | 2024-10-23 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
221 | C150135-290424 | Shenzhen Technology University | 2024-10-23 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
222 |  | 금오공과대학교_공동실험실습관 | 2024-10-23 | Research | FX40 | Single 100µm | Auto Align Standard | Mini450F |  |  | FX40AE
223 |  | 고려기연 | 2024-10-22 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
224 | T150013-310724 | Kanto Chemical | 2024-10-18 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
225 |  | 안동대학교_공동실험실습관 | 2024-10-16 | Research | FX40 | Single 100µm | Auto Align Standard | Mini450F |  |  | FX40AE
226 |  | SJ Semi Wuxi | 2024-10-15 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
227 | U10088-080724 | Genomics | 2024-10-10 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
228 |  | Quanzhou San'an Semiconductor(2) | 2024-10-10 | Industrial | NX-HDM | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | AL Bar type chuck
229 | C26006-101123 | heguang photomask | 2024-10-07 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
230 | U10089-160724 | University of the District of Columbia | 2024-10-04 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
231 | G30003-080824 | Technical University of Dresden (Niese) | 2024-10-02 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
232 | U28011-020924 | Stanford University | 2024-10-02 | Research | FX40 | Single 100µm | Auto Align Standard | Mini450F |  |  | FX40AE
233 | J30005-240624 | Nagoya University(Prof.Taniyama) | 2024-09-30 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE203
234 |  | ITB Indonesia | 2024-09-27 | Research | FX40 | Single 100µm | Auto Align Standard | Accurion i4 |  |  | FX40AE
235 | D10183-260624 | DGIST | 2024-09-26 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
236 | J160025-020424 | TSMC F23(Kumamoto) | 2024-09-26 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 7 units | 50mm | SiC Fork type chuck | Double Walled
237 |  | Qorvo #3 | 2024-09-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 6 units | 50mm | SiC Bar type chuck | Double Walled
238 |  | Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
239 | C31004-290324 | Wuhan Epiflag | 2024-09-23 | Research | NX15 | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
240 | C32014-120624 | Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 medium |  | Vacuum Sample Chuck | AE201
241 | T160041-190424 | Winbond | 2024-09-20 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
242 | S10013-210624 | National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
243 | U10087-260624 | Harvard University (Nocera) | 2024-09-19 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
244 |  | Rolex (Dudon) | 2024-09-19 | Research | FX40 | Single 100µm | Auto Align Standard | Accurion i4 |  |  | FX40AE
245 | GR22001-200524 | INN Demokritos, (El Sachat) | 2024-09-13 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
246 | E150008-190624 | University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
247 |  | Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
248 |  | Zhejiang University_Physics | 2024-09-10 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
249 |  | Sant'Anna Pisa (Poliziani) | 2024-09-09 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
250 |  | Alba Lulia University (SchaeferSE) | 2024-09-09 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
251 |  | San'an-ST Microelectronics | 2024-09-06 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
252 | J30007-280624 | Hirosaki University | 2024-09-03 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE203
253 |  | TSMC AP3 RD | 2024-09-03 | Industrial | NX-TSH |  | TSH 100µm | Dual MOD 6 units |  | Coreflow customized | Isolated
254 | I10015-070624 | Polimi, Italy | 2024-09-02 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
255 |  | 고려대학교 융합에너지공학과 | 2024-09-02 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
256 |  | Hunan San'an Semiconductor | 2024-08-30 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
257 | D160027-060324 | Samsung LED | 2024-08-30 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
258 | I23001-280923 | University of Bologna | 2024-08-29 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
259 | D280017-290724 | 주성엔지니어링 | 2024-08-28 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
260 | U160027-260224 | Dupont US | 2024-08-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Bar type chuck | Isolated
261 |  | TSMC F20 EBO | 2024-08-28 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
262 |  | Hauhong Wuxi #4 | 2024-08-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
263 |  | University of Arkansas (Hugh Churchill) | 2024-08-23 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
264 | D30019-210524 | 서강대학교 물리학과 | 2024-08-23 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
265 |  | 서강대학교_물리학과 | 2024-08-23 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
266 |  | University of Nevada(Ben King) | 2024-08-21 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
267 | C10090-010424 | Shandong University | 2024-08-20 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
268 | D10181-310524 | 한양대학교 화학과 | 2024-08-20 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
269 | RO150001-120624 | Alba Iulia University, Romania | 2024-08-20 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
270 | D10175-110324 | 서울대학교 화학생물공학부 | 2024-08-16 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | Glove Box
271 | J30006-270624 | Tokyo Metropolitan University | 2024-08-14 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
272 | U160029-080324 | Global Wafers America | 2024-08-12 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Bar type chuck | Double Walled
273 |  | Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
274 |  | University of Toronto(De La Barrera) | 2024-08-08 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
275 |  | Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Research | NX7 | Single 50µm | Standard | Accurion Nano30 |  |  | AE101
276 |  | University of Limerick (PSE) | 2024-08-08 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
277 |  | 목포해양대학교 | 2024-08-08 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
278 |  | Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
279 |  | University of Limerick | 2024-08-05 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
280 |  | Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
281 | J23001-230524 | Osaka University | 2024-08-02 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
282 |  | Sincere International Trade Co., Ltd | 2024-07-31 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
283 |  | ATC(Demoroom) | 2024-07-30 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
284 |  | Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
285 | F150008-300623 | Park Systems Europe GmbH (PSE) | 2024-07-26 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
286 |  | Henan Academy of Sciences | 2024-07-25 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
287 |  | Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
288 | D10178-080524 | 목포해양대학교 기관시스템공학부 | 2024-07-19 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
289 |  | Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
290 | G160012-271223 | Bosch | 2024-07-12 | Industrial | NX-Wafer | Dual 100µm | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
291 | C160086-200224 | Huahong Chengdu | 2024-07-12 | Industrial | NX-Wafer | Dual 100µm(300mm) | Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
292 | U27003-260324 | Intel Hillsboro#2 | 2024-07-12 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
293 |  | 엘티씨에이엠㈜ | 2024-07-11 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
294 |  | Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
295 | D23006-290124 | 충남대학교 분석과학기술대학원 | 2024-07-09 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
296 |  | Kyocera SLD Laser, Inc | 2024-06-28 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
297 | D150024-080524 | 에이앤아이 | 2024-06-28 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
298 |  | University of Bath (Prof. Kamal) | 2024-06-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
299 | T160040-080124 | TSMC F20 #3 (Hsinchu) | 2024-06-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Fork type chuck | Double Walled
300 | D10177-080524 | 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
301 | T160039-291223 | TSMC F20 #2 (Hsinchu) | 2024-06-27 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Fork type chuck | Double Walled
302 | C160070-170223 | Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Bar type chuck | Double Walled
303 |  | Qorvo #2(Thomas Russel) | 2024-06-26 | Industrial | NX-Wafer | Dual 100µm | Auto Align Long | Dual MOD 6 units | 50mm | SiC Bar type chuck | Double Walled
304 |  | National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
305 | C150129-200224 | Konfoong Materials International Co., Ltd. | 2024-06-25 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
306 | C160081-171123 | Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Bar type chuck | Double Walled
307 |  | Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
308 |  | Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Research | FX200 | Dual 100µm | FX Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | FX200 AE
309 | C10088-040124 | Tsinghua University | 2024-06-20 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
310 | C150131-110324 | Sanya Research Institute of Hainan University | 2024-06-18 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
311 |  | TSMC F20 #1 (Hsinchu) | 2024-06-17 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Fork type chuck | Double Walled
312 |  | PSCN Guangzhou Demo | 2024-06-14 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
313 |  | 서울과학기술대학교 | 2024-06-14 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
314 |  | Chongqing San'an Semiconductor | 2024-06-11 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
315 |  | University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Research | NX7 | Single 50µm | Standard | Minus-K |  |  | AE101
316 |  | GRG Test Group | 2024-06-05 | Research | NX10 | Single 100µm | Long | Accurion i4 |  |  | AE203
317 |  | University of Texas (Austin) | 2024-06-05 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
318 | S160007-071223 | Siltronic Singapore | 2024-05-31 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units |  | AL Bar type chuck | Double Walled
319 | D10176-250324 | 한국표준과학연구원 | 2024-05-30 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
320 |  | National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
321 |  | CXT #5 | 2024-05-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Bar type chuck | Double Walled
322 | C150124-081123 | Linton Kayex Technology | 2024-05-22 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
323 |  | Jiangsu Leadmicro | 2024-05-22 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
324 |  | NAURA, Beijing | 2024-05-20 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Bar type chuck | Isolated
325 |  | Getinge Atrium Medical(Dao) | 2024-05-17 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
326 | D280013-271123 | 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
327 |  | Park Systems Inc. (PSI) | 2024-05-09 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
328 |  | Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
329 | U150048-130224 | University of Texas, Arlington | 2024-05-08 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
330 |  | 충남테크노파크-디스플레이센터 | 2024-05-08 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
331 |  | Cenimat (Tomas) | 2024-05-08 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
332 |  | Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Research | NX7 | Single 100µm | Standard | Accurion Nano30 |  |  | AE101
333 |  | Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
334 | PL30001-290224 | Technolutions | 2024-05-07 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
335 |  | SMSC TD | 2024-04-30 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Pin Bar type chuck | Isolated
336 |  | International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
337 |  | Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | AL Bar type chuck | Double Walled
338 |  | Aixtron SE | 2024-04-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units |  | AL Bar type chuck | Double Walled
339 | C150123-011123 | Shenzhen Pengjin High-Tech | 2024-04-25 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
340 |  | University of Illinois at Urbana Champaign | 2024-04-24 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
341 |  | IQS Nanoptiqs | 2024-04-24 | Research | NX20 | Dual 100µm | Long | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
342 | P150002-260523 | HK SAE co.,ltd | 2024-04-19 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
343 |  | Jiangsu Jixin Advance Materials | 2024-04-11 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
344 |  | SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
345 | U30003-240124 | Southern Illinois University(Shamsi) | 2024-04-08 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
346 | U23012-091023 | University of Delaware(Rachel) | 2024-04-05 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
347 |  | International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
348 | D160025-230523 | Samsung Display A1 | 2024-03-29 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units |  | SiC Anti-warpage chuck | Isolated
349 | U10084-050124 | University of Washington | 2024-03-29 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
350 | C26005-040823 | SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
351 | U160025-041023 | Qorvo (Thomas Russel) | 2024-03-25 | Industrial | NX-Wafer | Dual 100µm | Auto Align Long | Dual MOD 6 units | 50mm | SiC Pin Bar type chuck | Double Walled
352 | D10173-120124 | 성균관대학교(기계공학부) | 2024-03-25 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
353 |  | Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
354 | ES28004-241123 | Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
355 |  | Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
356 | D30017-120124 | 한국나노마이스터고등학교 | 2024-03-12 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
357 |  | 단국대학교(전기전자공학부) | 2024-03-08 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
358 |  | NIPER Raebareli | 2024-03-08 | Research | NX12 | Single 100µm | Standard | Accurion i4 medium |  |  | AE204
359 | C30003-211223 | China Academy of Engineering Physics | 2024-03-07 | Research | NX7 | Single 50µm | Standard | Minus-K |  |  | AE101
360 | C30004-010124 | State Grid Tianjin | 2024-03-04 | Research | NX7 | Single 10µm | Standard | Minus-K |  |  | AE101
361 | C32013-281123 | Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
362 | D10170-071223 | 한국전자기술연구원 | 2024-02-29 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
363 | C26004-220623 | QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
364 |  | Nagoya Univ(Prof. Kato) | 2024-02-16 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
365 |  | Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
366 | D280012-251023 | 김해의생명산업진흥원 | 2024-02-16 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
367 | I10012-081123 | National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
368 | LU22001-040823 | University of Luxembourg (ST Instruments) | 2024-02-02 | Research | NX-Hivac | Single 50µm | NX-Hivac Auto Align | Single MOD 2 units |  |  | Chamber
369 |  | 경남테크노파크 | 2024-01-31 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
370 |  | 베스트그래핀 | 2024-01-31 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
371 |  | 한국해양대학교 나노반도체공학과 | 2024-01-31 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
372 |  | Tianjin De Quan Xing Ye technology | 2024-01-29 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
373 | C160071-170223 | Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
374 | C160072-170223 | Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
375 |  | Xidian University | 2024-01-24 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
376 | D30016-191023 | 중앙대학교 전기전자공학부 | 2024-01-18 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
377 |  | PSI(Mexico) Demo | 2024-01-17 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
378 |  | Intel Hillsboro #1 | 2024-01-15 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
379 |  | 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
380 |  | 전북대학교 물리학과 | 2023-12-28 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | FX40AE
381 |  | QST | 2023-12-28 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | FX40AE
382 |  | 한국폴리텍대학교_인천 | 2023-12-27 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
383 |  | Missouri State University(Rishi Patel) | 2023-12-27 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
384 |  | Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
385 |  | IIPE Visakhapatnam | 2023-12-22 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
386 |  | PowerEpi Semiconductor | 2023-12-22 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
387 |  | Hon Young Semiconductor | 2023-12-22 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
388 |  | Google(Melissa Archer) | 2023-12-20 | Industrial | NX-Wafer | Dual 100µm | Auto Align Standard | Dual MOD 6 units | 10mm | AL Bar type chuck | Double Walled
389 |  | National Taiwan Normal University | 2023-12-20 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
390 |  | China University of Petroleum | 2023-12-20 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | Glove Box
391 |  | Chulalongkorn University | 2023-12-19 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
392 |  | XFAB France #2 | 2023-12-15 | Industrial | NX-Wafer | Dual 100µm | Auto Align Standard | Dual MOD 6 units | 50mm | SiC Anti-warpage chuck | Double Walled
393 |  | 한양대학교 나노광전자학과 | 2023-12-15 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
394 |  | Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
395 |  | Tokyo Electron Kyushu | 2023-12-14 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units |  | SiC Anti-warpage chuck | Isolated
396 |  | 서울대학교 전기정보공학과 | 2023-12-14 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
397 |  | Harland Medical Systems | 2023-12-12 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
398 |  | Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
399 |  | Brock University(Kaur) | 2023-12-11 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
400 |  | EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
401 |  | EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
402 |  | EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
403 |  | Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
404 |  | Zhongke Dexin Sensing Technology | 2023-12-08 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
405 |  | 충북대학교 물리학과 | 2023-12-05 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
406 |  | Sabaragamuwa University | 2023-12-05 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
407 |  | Jinan Guangwei Semiconductor Technology | 2023-12-04 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
408 |  | Samsung P3 Foundry S5 | 2023-12-01 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
409 |  | Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units |  | AL Bar type chuck | Double Walled
410 |  | 부산대학교 물리학과 | 2023-11-30 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
411 |  | Sincere Internation | 2023-11-30 | Research | NX15 | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
412 |  | University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | FX40AE
413 |  | Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units |  | SiC Bar type chuck | Double Walled
414 |  | Khalifa University | 2023-11-24 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
415 |  | PrimeNano(Eduard) | 2023-11-22 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
416 |  | 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
417 |  | Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Research | NX15 | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
418 |  | Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Research | NX15 | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
419 |  | Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
420 |  | Anhui University | 2023-11-15 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
421 |  | University of Wuerzburg | 2023-11-13 | Research | NX7 | Single 100µm | Standard | Accurion i4 |  |  | AE101
422 |  | DESY (Dr.Chumakov) | 2023-11-13 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
423 |  | Fukoku | 2023-11-10 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE203
424 |  | Exquisite-Automotive Systems | 2023-11-10 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
425 |  | 인하대학교 물리학과 | 2023-11-10 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | Glove Box
426 |  | Shanghai Chisin Instruments (Demo) | 2023-11-08 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
427 |  | 전북대학교 공동실험실습관 | 2023-11-08 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
428 |  | Trumpf Photonics (PSE) | 2023-11-07 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
429 |  | Universidad de Cuenca | 2023-11-06 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
430 |  | Longzhong Laboratory | 2023-11-03 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
431 |  | CXT #4 | 2023-11-01 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
432 |  | AI IMAGE TECHNOLOGY | 2023-10-31 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
433 |  | Qingfang (Qingdao) Technology | 2023-10-31 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
434 |  | 한국교통대학교 화학산업연구소 | 2023-10-30 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
435 |  | University of Aveiro (PSE) | 2023-10-30 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
436 |  | 경상국립대학교 화학과 | 2023-10-27 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
437 |  | 부경대학교 고분자공학과 | 2023-10-27 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
438 |  | Samsung P3D | 2023-10-26 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
439 |  | Guangzhou University | 2023-10-20 | Research | NX7 | Single 100µm | Standard | Accurion i4 |  |  | AE101
440 |  | Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE203
441 |  | TYSiC#8 | 2023-10-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
442 |  | TYSiC#9 | 2023-10-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
443 |  | TYSiC#10 | 2023-10-19 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
444 |  | TYSiC#11 | 2023-10-19 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
445 |  | TYSiC#12 | 2023-10-19 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
446 |  | RMUTR University | 2023-10-17 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
447 |  | University of Mississippi(Vig Sundaresan) | 2023-10-16 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
448 |  | IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
449 |  | Cornell University(Liu) | 2023-10-13 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE101
450 |  | NovaSic (PSE) | 2023-10-13 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
451 |  | IDEADED (PSE) | 2023-10-13 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
452 |  | 삼성전자 (PE) | 2023-10-13 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
453 |  | TDK | 2023-10-12 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
454 |  | National Chung Hsing University | 2023-10-12 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
455 |  | Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Industrial | NX-HDM | Dual 100µm | Standard | Single MOD 2 units |  | AL Bar type chuck | AE401
456 |  | Shenzhen JWKW | 2023-10-10 | Industrial | NX-HDM | Dual 100µm | Standard | Dual MOD 4 units |  | AL Bar type chuck | AE401
457 |  | Brookhaven National Lab(BNL) | 2023-10-06 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
458 |  | SUMCO | 2023-10-05 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units |  | SiC Fork type chuck | Isolated
459 |  | SUMCO #2 | 2023-10-04 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units |  | SiC Fork type chuck | Isolated
460 |  | Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
461 |  | Hebei Synlight Semiconductor #7 | 2023-10-04 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
462 |  | 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
463 |  | Zhejiang University_Quzhou | 2023-09-26 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
464 |  | East China University of Science and Technology | 2023-09-25 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
465 |  | TYSiC#3 | 2023-09-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
466 |  | TYSiC#4 | 2023-09-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
467 |  | TYSiC#5 | 2023-09-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
468 |  | TYSiC#6 | 2023-09-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
469 |  | TYSiC#7 | 2023-09-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
470 |  | 경상국립대학교 반도체공학과 | 2023-09-20 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
471 |  | TSMC F23 (Kumamoto) | 2023-09-19 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Fork type chuck | Double Walled
472 |  | Swegan AB (PSE) | 2023-09-19 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
473 |  | Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
474 |  | 부산대학교 재료공학부 (세원) | 2023-09-18 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
475 |  | Wuhan Aroptics-tech | 2023-09-18 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
476 |  | Infineon Villach (PSE) #7 | 2023-09-14 | Industrial | NX-Wafer | Dual 100µm | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
477 |  | Infineon Villach (PSE) #8 | 2023-09-14 | Industrial | NX-Wafer | Dual 100µm | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
478 |  | Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Research | NX7 | Single 10µm | Standard | Mini450F |  |  | AE101
479 |  | Infineon Kulim #2 | 2023-09-08 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
480 |  | Infineon Kulim #2_1 | 2023-09-08 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
481 |  | Xavier University | 2023-09-08 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE101
482 |  | Rebeca Ribeiro | 2023-09-08 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
483 |  | SweGan, Sweden | 2023-09-07 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
484 |  | 부산대학교_공과대학 | 2023-08-31 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
485 |  | Sincere International Trading Company Limited #2 | 2023-08-31 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
486 |  | University of Texas at Arlington | 2023-08-31 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
487 |  | Westlake University_Jia Xing Huang Lab | 2023-08-31 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
488 |  | SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Industrial | NX-HDM | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck | AE401
489 |  | SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Industrial | NX-HDM | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck | AE401
490 |  | Nexchip #3 | 2023-08-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 10mm | SiC Anti-warpage chuck | Double Walled
491 |  | Hisense Broadband,Inc. | 2023-08-28 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
492 |  | Park Systems Singapore(demo 장비) | 2023-08-28 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
493 |  | Columbia Univ | 2023-08-25 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
494 |  | 주식회사 에이치앤에스 | 2023-08-22 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
495 |  | SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Industrial | NX-HDM | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck | AE401
496 |  | SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Industrial | NX-HDM | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck | AE401
497 |  | Dyson UK | 2023-08-21 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
498 |  | GRANDIT | 2023-08-18 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
499 |  | Infineon Villach #7 | 2023-08-08 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
500 |  | Infineon Villach #8 | 2023-08-08 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | SiC Bar type chuck | Double Walled
501 |  | George Mason University | 2023-08-07 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
502 |  | ICMAB-CSIC | 2023-08-07 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE101
503 |  | SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Industrial | NX-HDM | Dual 100µm | Auto Align Standard | Dual MOD 4 units |  | AL Bar type chuck | AE401
504 |  | 성균관대학교 신소재공학부 | 2023-08-04 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
505 |  | ICFO | 2023-08-01 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
506 |  | School of Physical and Mathematical Sciences | 2023-07-31 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
507 |  | Shanghai GTX Semiconductor | 2023-07-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
508 |  | Shanghai GTX Semiconductor_1 | 2023-07-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
509 |  | Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
510 |  | Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
511 |  | Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units |  | AL Bar type chuck | Double Walled
512 |  | 부경대학교 화학과 | 2023-07-26 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
513 |  | CXMT Hefei #7 | 2023-07-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
514 |  | Tokuyama Corporation | 2023-07-21 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
515 |  | 제주대학교 화학코스메틱스학과 | 2023-07-21 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE203
516 |  | TYSiC #2 | 2023-07-20 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
517 |  | Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
518 |  | OLED AOMORI | 2023-07-17 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
519 |  | Nanjing University_Electronic Science & Engineering | 2023-07-14 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
520 |  | TYSiC #1 | 2023-07-14 | Research | NX20 Lite | Single 100µm | Standard | Accurion i4 |  | Vacuum Sample Chuck | AE201
521 |  | PS UK Demo | 2023-07-13 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
522 |  | SMOC | 2023-07-10 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
523 |  | Samsung P3 | 2023-06-30 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
524 |  | Sicentury Semiconductor Technology | 2023-06-30 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
525 |  | Zhejiang Jinray Electronic Technology | 2023-06-30 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
526 |  | 전남대학교_석유화학신소재공학과 | 2023-06-30 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
527 |  | Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
528 |  | Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
529 |  | OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units |  | SiC Fork type chuck | Double Walled
530 |  | CXMT Beijing #3 | 2023-06-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
531 |  | OmniVision Semiconductor Shanghai #2_1 | 2023-06-28 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units |  | SiC Fork type chuck | Double Walled
532 |  | Polytechnic University of Milan #2 | 2023-06-28 | Research | NX10 | Single 50µm | Standard | Accurion i4 |  |  | AE101
533 |  | 목포대학교_첨단재료공학과 | 2023-06-28 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
534 |  | 삼성전자 | 2023-06-27 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
535 |  | 안동대학교_첨단재료공학과 | 2023-06-26 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
536 |  | Eindhoven University of Technology | 2023-06-22 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
537 |  | 에이프로세미콘 | 2023-06-20 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
538 |  | University of Bergamo | 2023-06-15 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE101
539 |  | AIST | 2023-06-15 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE203
540 |  | Ruhr-University of Bochum | 2023-06-14 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
541 |  | PSI DEMO Replacement | 2023-06-14 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
542 |  | Teknotip Cankiri University | 2023-06-13 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | Glove Box
543 |  | Universitas Andalas | 2023-06-13 | Research | FX40 | Single 100µm | FX Standard | Accurion i4 |  |  | FX40AE
544 |  | SPTS | 2023-06-09 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
545 |  | Westlake University_School of Science | 2023-06-01 | Research | NX10 | Single 100µm | Standard | Accurion i4 |  |  | AE203
546 |  | Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
547 |  | HC SemiTek | 2023-05-30 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
548 |  | PSE Demo | 2023-05-26 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
549 |  | PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
550 |  | PXW SEMICONDUCTOR MANUFACTURING_1 | 2023-05-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
551 |  | Mahidol University | 2023-05-24 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
552 |  | 성균관대학교 정보통신대학 | 2023-05-24 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
553 |  | KAIST(물리학과) | 2023-05-24 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
554 |  | Hangzhou HFC Semiconductor | 2023-05-23 | Research | NX20 | Dual 100µm(300mm) | Standard | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402
555 |  | UNC | 2023-05-23 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
556 |  | Swansea University | 2023-05-19 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE101
557 |  | Quantlase Lab | 2023-05-18 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
558 |  | A Inc._Aurora | 2023-05-16 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 6 units | 10mm | SiC Bar type chuck | Double Walled
559 |  | Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
560 |  | CXT #2 | 2023-05-11 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
561 |  | CXT #3 | 2023-05-11 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
562 |  | Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
563 |  | 동아대학교_신소재공학과 | 2023-05-10 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
564 |  | Intel Chandler | 2023-05-08 | Industrial | NX-TSH |  | TSH 100µm
565 |  | Lanzhou University | 2023-05-03 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
566 |  | Tianjin Normal University | 2023-05-03 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE101
567 |  | Jincheng industrial research institute | 2023-04-28 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
568 |  | 한국화학연구원 | 2023-04-28 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | Glove Box
569 |  | Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
570 |  | Tel Aviv University | 2023-04-26 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
571 |  | Tel Aviv University_NX12 | 2023-04-26 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
572 |  | Hangzhou Zeda Semiconductor | 2023-04-25 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
573 |  | TSMC AP6 | 2023-04-24 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
574 |  | Hacettepe University | 2023-04-24 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
575 |  | On-Semi | 2023-04-20 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 10mm | SiC Anti-warpage chuck | Isolated
576 |  | On-Semi_1 | 2023-04-20 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Long | Dual MOD 8 units | 10mm | SiC Anti-warpage chuck | Isolated
577 |  | Sincere International Trading Company Limited | 2023-04-19 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
578 |  | University of Arizona | 2023-04-19 | Research | NX7 | Single 50µm | Standard | Minus-K
579 |  | Wuhan YMTC Fab2 #2 | 2023-04-18 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
580 |  | Ecole centrale de Lyon (LTDS) | 2023-04-17 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
581 |  | Wuhan YMTC Fab2 #3 | 2023-04-13 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 7 units | 50mm | SiC Anti-warpage chuck | Double Walled
582 |  | Yachay Tech University | 2023-04-13 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
583 |  | Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Research | NX10 | Single 100µm | Long | Mini450F |  |  | AE203
584 |  | 전남대학교 신소재공학부 | 2023-04-10 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
585 |  | MSSCORPS | 2023-04-07 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
586 |  | PSE DEMO_NX10 | 2023-04-06 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
587 |  | Leica Biosystems | 2023-04-06 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE101
588 |  | National Institutes of Technology (NITs) | 2023-04-05 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
589 |  | Fudan University(Department of Physics) | 2023-04-03 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
590 |  | 경북대 하이브리드부품연구원 | 2023-03-31 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
591 |  | 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE101
592 |  | Zhonghuan semiconductor | 2023-03-29 | Research | NX20 | Dual 100µm | Standard | Single MOD 2 units |  | Vacuum Sample Chuck | AE202
593 |  | OCI 주식회사 | 2023-03-28 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE202
594 |  | QMOT | 2023-03-27 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
595 |  | 한국기계연구원 | 2023-03-27 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
596 |  | Park Systems Europe GmbH | 2023-03-23 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
597 |  | PSI | 2023-03-23 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
598 |  | Park Systems UK | 2023-03-22 | Research | NX20 | Dual 100µm | Standard | Dual MOD 4 units |  | Vacuum Sample Chuck | AE401
599 |  | Zhejiang University Department of Chemistry | 2023-03-20 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
600 |  | National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
601 |  | TSMC F12 EBO | 2023-03-16 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
602 |  | INOMAR_NX10 | 2023-03-16 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE101
603 |  | Tokyo institute of technology | 2023-03-13 | Research | NX7 | Single 50µm | Standard | Accurion i4 |  |  | AE203
604 |  | University of Warwick | 2023-03-10 | Research | NX10 | Single 50µm | Standard | Mini450F |  |  | AE203
605 |  | Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Research | NX10 | Single 100µm | Standard | Mini450F |  |  | AE203
606 |  | Nanoplexus Limited | 2023-02-27 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
607 |  | 울산대학교 | 2023-02-15 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
608 |  | University of Reims | 2023-02-02 | Research | NX12 | Single 100µm | Standard | Single MOD 2 units |  |  | AE203
609 |  | 서울시립대 공동실습실험관 | 2023-02-01 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
610 |  | Samsung Mask | 2022-12-27 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
611 |  | Fujibo | 2022-12-01 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
612 |  | Canon(Hiratsuka #1) | 2022-11-29 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
613 |  | Canon(Hiratsuka #2) | 2022-11-29 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
614 |  | 경상국립대학교(화학과) | 2022-11-28 | Research | NX12 | Single 100µm | Standard | Mini450F |  |  | AE204
615 |  | Hoya | 2022-11-18 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
616 |  | Laboratory of Solid State Physics | 2022-11-03 | Research | FX40 | Single 100µm | FX Standard | Mini450F |  |  | FX40AE
617 |  | LG Display P9(H1) | 2022-09-01 | Industrial | NX-TSH |  | TSH 100µm
618 |  | LG Display P9(H2) | 2022-09-01 | Industrial | NX-TSH |  | TSH 100µm
619 |  | LG Display P10(H1) | 2022-08-26 | Industrial | NX-TSH |  | TSH 100µm
620 |  | LG Display P10(H2) | 2022-08-26 | Industrial | NX-TSH |  | TSH 100µm
621 |  | TSMC F12 P4 | 2022-08-25 | Industrial | NX-Wafer | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Fork type chuck | Isolated
622 |  | SMIC Shanghai | 2022-08-19 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 10 units | 50mm | SiC Anti-warpage chuck | Isolated
623 |  | NRMTC | 2022-08-17 | Industrial | NX-Mask |  | TSH 50µm | Dual MOD 8 units |  | Mask | Isolated
624 |  | Samsung SRD | 2022-04-15 | Industrial | NX-Hybrid WLI | Dual 100µm(300mm) | Auto Align Standard | Dual MOD 8 units | 50mm | SiC Anti-warpage chuck | Isolated
625 |  | Riken | 2022-03-14 | Industrial | NX-TSH |  | TSH 50µm | Dual MOD 8 units |  |  | AE402
626 |  | ASML (ST Instruments) | 2022-02-26 | Research | NX-Hivac | Single 100µm | Long | Single MOD 2 units |  |  | Chamber
627 |  | Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Research | NX10 | Single 50µm | Long | Mini450F |  |  | AE203
628 |  | Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Research | NX20 | Dual 100µm(300mm) | Long | Dual MOD 6 units |  | Vacuum Sample Chuck | AE402

## Measurements
SID | 장비명 | 종료일 | Check Items | Value
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Linearity by Z scanner position | 0.49
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Z Detector during Z Scan sweep | 13.81
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height calibration error | 0.080225
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drive calibration error | 0.25867
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Linearity by 4 different sample height | 1.0044
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Linearity by 4 different sample height | 1.0044
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Z Height during Z Scan sweep | 16.98
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.5
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4373
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.6311
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.371
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height noise | 0.024
| BYD Semiconductor(Jinan) | 2025-11-30 | Wafer placement(Loading) error on the load port1 (X offset) | 42.9
| BYD Semiconductor(Jinan) | 2025-11-30 | Wafer placement(Loading) error on the load port1 (Y offset) | 71.39
| BYD Semiconductor(Jinan) | 2025-11-30 | Wafer placement(Loading) error on the load port2 (X offset) | 46.99
| BYD Semiconductor(Jinan) | 2025-11-30 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.46
| BYD Semiconductor(Jinan) | 2025-11-30 | OpticalAxisShiftPerFocusX | -0.0068
| BYD Semiconductor(Jinan) | 2025-11-30 | OpticalAxisShiftPerFocusY | 0.0022
| BYD Semiconductor(Jinan) | 2025-11-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00024
| BYD Semiconductor(Jinan) | 2025-11-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00012
| BYD Semiconductor(Jinan) | 2025-11-30 | Lift pins synchronization offset check (X offset) | 14.77
| BYD Semiconductor(Jinan) | 2025-11-30 | Lift pins synchronization offset check (Y offset) | 37.84
| BYD Semiconductor(Jinan) | 2025-11-30 | Profiler noise @200um | 0.135
| BYD Semiconductor(Jinan) | 2025-11-30 | Roll motion to X+ full scan | 4.2
| BYD Semiconductor(Jinan) | 2025-11-30 | Position repeatability(10repeats, Position Based) | 0.46
| BYD Semiconductor(Jinan) | 2025-11-30 | OPM Repeatability @5mm | 2.496
| BYD Semiconductor(Jinan) | 2025-11-30 | OPM Repeatability @1mm | 1.488
| BYD Semiconductor(Jinan) | 2025-11-30 | Max OPM @5mm | 19.246
| BYD Semiconductor(Jinan) | 2025-11-30 | Max OPM @1mm | 10.778
| BYD Semiconductor(Jinan) | 2025-11-30 | Torque Profile (10mm, 200um/s) | 0.732
| BYD Semiconductor(Jinan) | 2025-11-30 | Torque Profile (10mm, 20um/s) | 0.632
| BYD Semiconductor(Jinan) | 2025-11-30 | Lateral(C-D) sensitivity | 76.4
| BYD Semiconductor(Jinan) | 2025-11-30 | Lateral(C-D) noise | 0.175
| BYD Semiconductor(Jinan) | 2025-11-30 | Min. Vibration Transmissbility (~100Hz) | 0.011
| BYD Semiconductor(Jinan) | 2025-11-30 | RMS of ZeroScan | 0.03
| BYD Semiconductor(Jinan) | 2025-11-30 | RMS of ZeroScan | 0.04
| BYD Semiconductor(Jinan) | 2025-11-30 | LightStrengthGain | 0.5
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of X Detector during X Scan sweep | 181.4632
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of X1 Detector during X Scan sweep | 179.58
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of X2 Detector during X Scan sweep | 183.33
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Y Detector during Y Scan sweep | 178.7022
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Y1 Detector during Y Scan sweep | 182.26
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Y2 Detector during Y Scan sweep | 175.13
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of X1 Detector Raw during X Scan sweep | 14.38
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of X2 Detector Raw during X Scan sweep | 14.04
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.17
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.73
| BYD Semiconductor(Jinan) | 2025-11-30 | X Detector stroke calibration error (nominal scale error X) | 0.481
| BYD Semiconductor(Jinan) | 2025-11-30 | Y Detector stroke calibration error (nominal scale error Y) | 0.216
| BYD Semiconductor(Jinan) | 2025-11-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 243.4561
| BYD Semiconductor(Jinan) | 2025-11-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 255.4453
| BYD Semiconductor(Jinan) | 2025-11-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 249.0308
| BYD Semiconductor(Jinan) | 2025-11-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 233.7315
| BYD Semiconductor(Jinan) | 2025-11-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 154.982
| BYD Semiconductor(Jinan) | 2025-11-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 150.951
| BYD Semiconductor(Jinan) | 2025-11-30 | X Non-linearity error | 2.04
| BYD Semiconductor(Jinan) | 2025-11-30 | Y Non-linearity error | 1.8
| BYD Semiconductor(Jinan) | 2025-11-30 | X axis OPM | 0.58
| BYD Semiconductor(Jinan) | 2025-11-30 | Y axis OPM | 1.515
| BYD Semiconductor(Jinan) | 2025-11-30 | X Ringing(Rpv) | 0.155
| BYD Semiconductor(Jinan) | 2025-11-30 | Y Ringing(Rpv) | 0.1
| BYD Semiconductor(Jinan) | 2025-11-30 | X Detector noise (XY Servo on) | 0.318
| BYD Semiconductor(Jinan) | 2025-11-30 | Y Detector noise (XY Servo on) | 0.3
| BYD Semiconductor(Jinan) | 2025-11-30 | X1 Detector noise (XY Servo on) | 0.554
| BYD Semiconductor(Jinan) | 2025-11-30 | Y1 Detector noise (XY Servo on) | 0.559
| BYD Semiconductor(Jinan) | 2025-11-30 | X2 Detector noise (XY Servo on) | 0.546
| BYD Semiconductor(Jinan) | 2025-11-30 | Y2 Detector noise (XY Servo on) | 0.522
| BYD Semiconductor(Jinan) | 2025-11-30 | RMS error after Stage Calibration | 2.45
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from -X to +X (X scan) | 90.0879
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from -Y to +Y (Y scan) | 71.777
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from -X to +X (Y scan) | 32.23
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from -Y to +Y (X scan) | 9.52
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from +X to -X (X scan) | 8.789
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from +Y to -Y (Y scan) | 39.55
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from +X to -X (Y scan) | 108.4
| BYD Semiconductor(Jinan) | 2025-11-30 | Image offset after moving from +Y to -Y (X scan) | 5.86
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X offset (Vision Pattern Matching) | 0.391
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.328
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.085
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.335
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X offset (In-Die Align) | 0.781
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y offset (In-Die Align) | 1.797
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X 1σ (In-Die Align) | 0.168
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y 1σ (In-Die Align) | 0.372
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X offset (LLC Translation) | 0.781
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y offset (LLC Translation) | 1.953
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X 1σ (LLC Translation) | 0.14
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y 1σ (LLC Translation) | 0.15
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X offset (Global Align) | 2.031
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y offset (Global Align) | 4.141
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning X 1σ (Global Align) | 0.501
| BYD Semiconductor(Jinan) | 2025-11-30 | XY Stage positioning Y 1σ (Global Align) | 1.004
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drift | 10.459
| BYD Semiconductor(Jinan) | 2025-11-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| BYD Semiconductor(Jinan) | 2025-11-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| BYD Semiconductor(Jinan) | 2025-11-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.62
| BYD Semiconductor(Jinan) | 2025-11-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| BYD Semiconductor(Jinan) | 2025-11-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| BYD Semiconductor(Jinan) | 2025-11-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| BYD Semiconductor(Jinan) | 2025-11-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.64
| BYD Semiconductor(Jinan) | 2025-11-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| BYD Semiconductor(Jinan) | 2025-11-30 | Cycle time for 15-point measurement | 1811.847
| BYD Semiconductor(Jinan) | 2025-11-30 | Short-term stability using Step Height Sample (3σ) | 0.345
| BYD Semiconductor(Jinan) | 2025-11-30 | Roughness Repeatability 1x1um(RSD%) | 1.395
| BYD Semiconductor(Jinan) | 2025-11-30 | Roughness Repeatability 1x1um(1σ) | 0.002
| BYD Semiconductor(Jinan) | 2025-11-30 | Roughness Repeatability 80x80um(RSD%) | 0.847
| BYD Semiconductor(Jinan) | 2025-11-30 | Roughness Repeatability 80x80um(1σ) | 0.002
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Linearity by Z scanner position | 0.42
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height calibrtaion error | 0.10494
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drive calibrtaion error | 0.21032
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Linearity by 4 different sample height | 0.99448
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Linearity by 3 different sample height | 0.99448
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range in voltage during Z Scan sweep | 13.97
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range in um during Z Scan sweep | 43.6792
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 62.5284
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.6995
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 62.6556
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.7273
| BYD Semiconductor(Jinan) | 2025-11-30 | Z Height noise | 0.067
| BYD Semiconductor(Jinan) | 2025-11-30 | Full range in um during Z Scan sweep | 0.014
| LAPIS Semiconductor | 2025-10-29 | Lateral(C-D) sensitivity | 15.1
| LAPIS Semiconductor | 2025-10-29 | Lateral(C-D) noise | 0.139
| LAPIS Semiconductor | 2025-10-29 | A-B noise | 7.669
| LAPIS Semiconductor | 2025-10-29 | NCM Amplitude noise | 12.126
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.32
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.343
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.34
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.29
| LAPIS Semiconductor | 2025-10-29 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| LAPIS Semiconductor | 2025-10-29 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.053
| LAPIS Semiconductor | 2025-10-29 | Roughness Repeatability 1x1um(RSD%) | 4.743
| LAPIS Semiconductor | 2025-10-29 | Roughness Repeatability 1x1um(1σ) | 0.007
| ZJEAGLES | 2025-10-28 | Z Linearity by Z scanner position | 0.7
| ZJEAGLES | 2025-10-28 | Z Linearity by 4 different sample height | 1.0043
| ZJEAGLES | 2025-10-28 | Full range of Z Detector during Z Scan sweep | 12.8
| ZJEAGLES | 2025-10-28 | Z Height calibration error | 0.15568
| ZJEAGLES | 2025-10-28 | Z Drive calibration error | 0.45533
| ZJEAGLES | 2025-10-28 | Full range of Z Height during Z Scan sweep | 17.7
| ZJEAGLES | 2025-10-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.7188
| ZJEAGLES | 2025-10-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.451
| ZJEAGLES | 2025-10-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.7377
| ZJEAGLES | 2025-10-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.562
| ZJEAGLES | 2025-10-28 | Z Height noise | 0.027
| ZJEAGLES | 2025-10-28 | OpticalAxisShiftPerFocusX | 0.0044
| ZJEAGLES | 2025-10-28 | OpticalAxisShiftPerFocusY | 0.001
| ZJEAGLES | 2025-10-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.0001
| ZJEAGLES | 2025-10-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| ZJEAGLES | 2025-10-28 | RMS of ZeroScan | 0.038
| ZJEAGLES | 2025-10-28 | LightStrengthGain | 0.84
| ZJEAGLES | 2025-10-28 | Full range of X Detector during X Scan sweep | 190.5
| ZJEAGLES | 2025-10-28 | Full range of X1 Detector during X Scan sweep | 183.8
| ZJEAGLES | 2025-10-28 | Full range of X2 Detector during X Scan sweep | 197.2
| ZJEAGLES | 2025-10-28 | Full range of Y Detector during Y Scan sweep | 186
| ZJEAGLES | 2025-10-28 | Full range of Y1 Detector during Y Scan sweep | 190
| ZJEAGLES | 2025-10-28 | Full range of Y2 Detector during Y Scan sweep | 182
| ZJEAGLES | 2025-10-28 | Full range of X1 Detector Raw during X Scan sweep | 14.77
| ZJEAGLES | 2025-10-28 | Full range of X2 Detector Raw during X Scan sweep | 15.24
| ZJEAGLES | 2025-10-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.58
| ZJEAGLES | 2025-10-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.61
| ZJEAGLES | 2025-10-28 | X Detector stroke calibration error (nominal scale error X) | 0.11
| ZJEAGLES | 2025-10-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.36
| ZJEAGLES | 2025-10-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 242.6473
| ZJEAGLES | 2025-10-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 232.9342
| ZJEAGLES | 2025-10-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 240.2389
| ZJEAGLES | 2025-10-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 238.4261
| ZJEAGLES | 2025-10-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 158.458
| ZJEAGLES | 2025-10-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 158.279
| ZJEAGLES | 2025-10-28 | XY Orthogonality | 90.27
| ZJEAGLES | 2025-10-28 | X Non-linearity error | 2.46
| ZJEAGLES | 2025-10-28 | Y Non-linearity error | 2.69
| ZJEAGLES | 2025-10-28 | X axis OPM | 1.13
| ZJEAGLES | 2025-10-28 | Y axis OPM | 1.05
| ZJEAGLES | 2025-10-28 | X Ringing(Rpv) | 0.55
| ZJEAGLES | 2025-10-28 | Y Ringing(Rpv) | 0.48
| ZJEAGLES | 2025-10-28 | X Detector noise (XY Servo on) | 0.311
| ZJEAGLES | 2025-10-28 | Y Detector noise (XY Servo on) | 0.302
| ZJEAGLES | 2025-10-28 | X1 Detector noise (XY Servo on) | 0.398
| ZJEAGLES | 2025-10-28 | Y1 Detector noise (XY Servo on) | 0.425
| ZJEAGLES | 2025-10-28 | X2 Detector noise (XY Servo on) | 0.447
| ZJEAGLES | 2025-10-28 | Y2 Detector noise (XY Servo on) | 0.419
| ZJEAGLES | 2025-10-28 | RMS error after Stage Calibration | 3.4
| ZJEAGLES | 2025-10-28 | Z Drift | 7.6
| ZJEAGLES | 2025-10-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| ZJEAGLES | 2025-10-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| ZJEAGLES | 2025-10-28 | Short-term stability using Step Height Sample (3σ) | 0.25
| ZJEAGLES | 2025-10-28 | Roughness Repeatability 1x1um(RSD%) | 1
| ZJEAGLES | 2025-10-28 | Roughness Repeatability 1x1um(1σ) | 0.001
| ZJEAGLES | 2025-10-28 | Roughness Repeatability 80x80um(RSD%) | 1
| ZJEAGLES | 2025-10-28 | Roughness Repeatability 80x80um(1σ) | 0.002
| Brazilian Center for Research in Physics | 2025-10-27 | Lateral(C-D) sensitivity | 43.2
| Brazilian Center for Research in Physics | 2025-10-27 | Lateral(C-D) noise | 0.063
| Brazilian Center for Research in Physics | 2025-10-27 | A-B noise | 9.921
| Brazilian Center for Research in Physics | 2025-10-27 | NCM Amplitude noise | 14.647
| Brazilian Center for Research in Physics | 2025-10-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Brazilian Center for Research in Physics | 2025-10-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| Brazilian Center for Research in Physics | 2025-10-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Brazilian Center for Research in Physics | 2025-10-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.121
| Brazilian Center for Research in Physics | 2025-10-27 | Roughness Repeatability 1x1um(RSD%) | 1.27
| Brazilian Center for Research in Physics | 2025-10-27 | Roughness Repeatability 1x1um(1σ) | 0.004
| Brazilian Center for Research in Physics | 2025-10-27 | Full range of Z Height during Z Scan sweep | 18.2
| Brazilian Center for Research in Physics | 2025-10-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.666
| Brazilian Center for Research in Physics | 2025-10-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.938
| Brazilian Center for Research in Physics | 2025-10-27 | Full range of Z Detector during Z Scan sweep | 13.2
| Brazilian Center for Research in Physics | 2025-10-27 | Z Drive calibration error | 0.16762
| Brazilian Center for Research in Physics | 2025-10-27 | Z Height calibration error | 0.41858
| Brazilian Center for Research in Physics | 2025-10-27 | Z linearity by sample height | 0.99078
| Brazilian Center for Research in Physics | 2025-10-27 | 1) Z Scanner position at +3μm | 0.004183
| Brazilian Center for Research in Physics | 2025-10-27 | 3) Z Scanner position at -3μm | 0.004335
| Brazilian Center for Research in Physics | 2025-10-27 | Z Height noise | 0.028
| Brazilian Center for Research in Physics | 2025-10-27 | Full range of X Detector during X Scan sweep | 66.5
| Brazilian Center for Research in Physics | 2025-10-27 | Full range of Y Detector during Y Scan sweep | 68.9
| Brazilian Center for Research in Physics | 2025-10-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.4
| Brazilian Center for Research in Physics | 2025-10-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.5
| Brazilian Center for Research in Physics | 2025-10-27 | X Detector stroke calibration error (nominal scale error X) | 0.162
| Brazilian Center for Research in Physics | 2025-10-27 | Y Detector stroke calibration error (nominal scale error Y) | 0.252
| Brazilian Center for Research in Physics | 2025-10-27 | X Scan Stroke | 55
| Brazilian Center for Research in Physics | 2025-10-27 | Y Scan Stroke | 55
| Brazilian Center for Research in Physics | 2025-10-27 | Full range of X Detector Raw during X Scan sweep | 13.5
| Brazilian Center for Research in Physics | 2025-10-27 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| Brazilian Center for Research in Physics | 2025-10-27 | X Detector noise (XY Servo on) | 0.247
| Brazilian Center for Research in Physics | 2025-10-27 | Y Detector noise (XY Servo on) | 0.254
| Brazilian Center for Research in Physics | 2025-10-27 | X Detector noise (XY Servo off) | 0.239
| Brazilian Center for Research in Physics | 2025-10-27 | Y Detector noise (XY Servo off) | 0.244
| Brazilian Center for Research in Physics | 2025-10-27 | X axis OPM | 0.443
| Brazilian Center for Research in Physics | 2025-10-27 | Y axis OPM | 0.491
| Brazilian Center for Research in Physics | 2025-10-27 | X Non-linearity | 0.046
| Brazilian Center for Research in Physics | 2025-10-27 | Y Non-linearity | 0.081
| OEINO | 2025-10-27 | Z Linearity by Z scanner position | 0.9
| OEINO | 2025-10-27 | Full range of Z Detector during Z Scan sweep | 12.76
| OEINO | 2025-10-27 | Z Height calibration error | 0.041197
| OEINO | 2025-10-27 | Z Drive calibration error | 0.057892
| OEINO | 2025-10-27 | Z Linearity by 4 different sample height | 1.0049
| OEINO | 2025-10-27 | Full range of Z Height during Z Scan sweep | 17.1628
| OEINO | 2025-10-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.6783
| OEINO | 2025-10-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.212
| OEINO | 2025-10-27 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.6987
| OEINO | 2025-10-27 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6987
| OEINO | 2025-10-27 | Z Height noise | 0.028
| OEINO | 2025-10-27 | Wafer placement(Loading) error on the load port1 (X offset) | 29.05
| OEINO | 2025-10-27 | Wafer placement(Loading) error on the load port1 (Y offset) | 17.21
| OEINO | 2025-10-27 | Wafer placement(Loading) error on the load port2 (X offset) | 34.01
| OEINO | 2025-10-27 | Wafer placement(Loading) error on the load port2 (Y offset) | 53.17
| OEINO | 2025-10-27 | OpticalAxisShiftPerFocusX | -0.0031
| OEINO | 2025-10-27 | OpticalAxisShiftPerFocusY | -0.0013
| OEINO | 2025-10-27 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00043
| OEINO | 2025-10-27 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| OEINO | 2025-10-27 | Lift pins synchronization offset check (X offset) | 24.16
| OEINO | 2025-10-27 | Lift pins synchronization offset check (Y offset) | 8.1
| OEINO | 2025-10-27 | Profiler noise @200um | 0.155
| OEINO | 2025-10-27 | Roll motion to X+ full scan | 11.1
| OEINO | 2025-10-27 | Position repeatability(10repeats, Position Based) | 0.26
| OEINO | 2025-10-27 | OPM Repeatability @25mm | 1.682
| OEINO | 2025-10-27 | OPM Repeatability @10mm | 2.034
| OEINO | 2025-10-27 | OPM Repeatability @5mm | 0.777
| OEINO | 2025-10-27 | OPM Repeatability @1mm | 0.935
| OEINO | 2025-10-27 | Max OPM @25mm | 51.474
| OEINO | 2025-10-27 | Max OPM @10mm | 28.636
| OEINO | 2025-10-27 | Max OPM @5mm | 24.168
| OEINO | 2025-10-27 | Max OPM @1mm | 8.574
| OEINO | 2025-10-27 | Torque Profile (50mm, 200um/s) | 0.47
| OEINO | 2025-10-27 | Torque Profile (50mm, 20um/s) | 0.617
| OEINO | 2025-10-27 | Lateral(C-D) sensitivity | 51.9
| OEINO | 2025-10-27 | Lateral(C-D) noise | 0.18
| OEINO | 2025-10-27 | Min. Vibration Transmissbility (~100Hz) | 0.057
| OEINO | 2025-10-27 | RMS of ZeroScan | 0.087
| OEINO | 2025-10-27 | LightStrengthGain | 0.43
| OEINO | 2025-10-27 | Full range of X Detector during X Scan sweep | 185.2552
| OEINO | 2025-10-27 | Full range of X1 Detector during X Scan sweep | 190.3383
| OEINO | 2025-10-27 | Full range of X2 Detector during X Scan sweep | 180.172
| OEINO | 2025-10-27 | Full range of Y Detector during Y Scan sweep | 185.4462
| OEINO | 2025-10-27 | Full range of Y1 Detector during Y Scan sweep | 184.1824
| OEINO | 2025-10-27 | Full range of Y2 Detector during Y Scan sweep | 186.71
| OEINO | 2025-10-27 | Full range of X1 Detector Raw during X Scan sweep | 14.18
| OEINO | 2025-10-27 | Full range of X2 Detector Raw during X Scan sweep | 14.18
| OEINO | 2025-10-27 | Full range of Y1 Detector Raw during Y Scan sweep | 13.97
| OEINO | 2025-10-27 | Full range of Y2 Detector Raw during Y Scan sweep | 14.36
| OEINO | 2025-10-27 | X Detector stroke calibration error (nominal scale error X) | 0.002
| OEINO | 2025-10-27 | Y Detector stroke calibration error (nominal scale error Y) | 0.049
| OEINO | 2025-10-27 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 245.547
| OEINO | 2025-10-27 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 243.596
| OEINO | 2025-10-27 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 254.268
| OEINO | 2025-10-27 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 247.659
| OEINO | 2025-10-27 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 175.691
| OEINO | 2025-10-27 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 169.83
| OEINO | 2025-10-27 | X Non-linearity error | 1.39
| OEINO | 2025-10-27 | Y Non-linearity error | 1.19
| OEINO | 2025-10-27 | X axis OPM | 0.583
| OEINO | 2025-10-27 | Y axis OPM | 0.323
| OEINO | 2025-10-27 | X Ringing(Rpv) | 0.131
| OEINO | 2025-10-27 | Y Ringing(Rpv) | 0.257
| OEINO | 2025-10-27 | X Detector noise (XY Servo on) | 0.313
| OEINO | 2025-10-27 | Y Detector noise (XY Servo on) | 0.296
| OEINO | 2025-10-27 | X1 Detector noise (XY Servo on) | 0.468
| OEINO | 2025-10-27 | Y1 Detector noise (XY Servo on) | 0.398
| OEINO | 2025-10-27 | X2 Detector noise (XY Servo on) | 0.426
| OEINO | 2025-10-27 | Y2 Detector noise (XY Servo on) | 0.451
| OEINO | 2025-10-27 | RMS error after Stage Calibration | 3.2
| OEINO | 2025-10-27 | Image offset after moving from -X to +X (X scan) | 77.6367
| OEINO | 2025-10-27 | Image offset after moving from -Y to +Y (Y scan) | 73.2422
| OEINO | 2025-10-27 | Image offset after moving from -X to +X (Y scan) | 47.6074
| OEINO | 2025-10-27 | Image offset after moving from -Y to +Y (X scan) | 22.7051
| OEINO | 2025-10-27 | Image offset after moving from +X to -X (X scan) | 65.918
| OEINO | 2025-10-27 | Image offset after moving from +Y to -Y (Y scan) | 25.6348
| OEINO | 2025-10-27 | Image offset after moving from +X to -X (Y scan) | 8.0566
| OEINO | 2025-10-27 | Image offset after moving from +Y to -Y (X scan) | 28.5645
| OEINO | 2025-10-27 | XY Stage positioning X offset (Vision Pattern Matching) | 3.0908
| OEINO | 2025-10-27 | XY Stage positioning Y offset (Vision Pattern Matching) | 3.1396
| OEINO | 2025-10-27 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.6789
| OEINO | 2025-10-27 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.6405
| OEINO | 2025-10-27 | XY Stage positioning X offset (In-Die Align) | 3.2471
| OEINO | 2025-10-27 | XY Stage positioning Y offset (In-Die Align) | 3.2861
| OEINO | 2025-10-27 | XY Stage positioning X 1σ (In-Die Align) | 0.6661
| OEINO | 2025-10-27 | XY Stage positioning Y 1σ (In-Die Align) | 0.7994
| OEINO | 2025-10-27 | XY Stage positioning X offset (LLC Translation) | 1.2354
| OEINO | 2025-10-27 | XY Stage positioning Y offset (LLC Translation) | 2.1533
| OEINO | 2025-10-27 | XY Stage positioning X 1σ (LLC Translation) | 0.2763
| OEINO | 2025-10-27 | XY Stage positioning Y 1σ (LLC Translation) | 0.4674
| OEINO | 2025-10-27 | XY Stage positioning X offset (Global Align) | 3.7207
| OEINO | 2025-10-27 | XY Stage positioning Y offset (Global Align) | 5.8008
| OEINO | 2025-10-27 | XY Stage positioning X 1σ (Global Align) | 0.6885
| OEINO | 2025-10-27 | XY Stage positioning Y 1σ (Global Align) | 1.3804
| OEINO | 2025-10-27 | Z Drift | 5.014
| OEINO | 2025-10-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.63
| OEINO | 2025-10-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| OEINO | 2025-10-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.52
| OEINO | 2025-10-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| OEINO | 2025-10-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| OEINO | 2025-10-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| OEINO | 2025-10-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.75
| OEINO | 2025-10-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| OEINO | 2025-10-27 | Cycle time for 15-point measurement | 1684.586
| OEINO | 2025-10-27 | Short-term stability using Step Height Sample (3σ) | 0.40133
| OEINO | 2025-10-27 | Roughness Repeatability 1x1um(RSD%) | 3.259
| OEINO | 2025-10-27 | Roughness Repeatability 1x1um(1σ) | 0.004
| OEINO | 2025-10-27 | Roughness Repeatability 80x80um(RSD%) | 2.332
| OEINO | 2025-10-27 | Roughness Repeatability 80x80um(1σ) | 0.002
| OEINO | 2025-10-27 | Z Linearity by Z scanner position | 0.34928
| OEINO | 2025-10-27 | Z Height calibrtaion error | 0.045967
| OEINO | 2025-10-27 | Z Drive calibrtaion error | 0.49501
| OEINO | 2025-10-27 | Z Linearity by 4 different sample height | 0.99974
| OEINO | 2025-10-27 | Full range in voltage during Z Scan sweep | 14.92
| OEINO | 2025-10-27 | Full range in um during Z Scan sweep | 44.4992
| OEINO | 2025-10-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 59.2786
| OEINO | 2025-10-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.278
| OEINO | 2025-10-27 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 59.3999
| OEINO | 2025-10-27 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.2927
| OEINO | 2025-10-27 | Z Height noise | 0.062
| 램테크놀러지㈜ | 2025-10-21 | SLD current limit check | 104
| 램테크놀러지㈜ | 2025-10-21 | SLD beam power at the end of the lens | 0.29
| 램테크놀러지㈜ | 2025-10-21 | Lateral(C-D) sensitivity | 37
| 램테크놀러지㈜ | 2025-10-21 | Lateral(C-D) noise | 0.115
| 램테크놀러지㈜ | 2025-10-21 | A-B noise | 15.696
| 램테크놀러지㈜ | 2025-10-21 | NCM Amplitude noise | 5.969
| 램테크놀러지㈜ | 2025-10-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| 램테크놀러지㈜ | 2025-10-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 램테크놀러지㈜ | 2025-10-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 램테크놀러지㈜ | 2025-10-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 램테크놀러지㈜ | 2025-10-21 | Roughness Repeatability 1x1um(RSD%) | 4.337
| 램테크놀러지㈜ | 2025-10-21 | Roughness Repeatability 1x1um(1σ) | 0.005
| 램테크놀러지㈜ | 2025-10-21 | Full range of Z Height during Z Scan sweep | 17.3
| 램테크놀러지㈜ | 2025-10-21 | Full range of Z Detector during Z Scan sweep | 14.4
| 램테크놀러지㈜ | 2025-10-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.161
| 램테크놀러지㈜ | 2025-10-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.568
| 램테크놀러지㈜ | 2025-10-21 | Z Drive calibrtaion error | 0.47989
| 램테크놀러지㈜ | 2025-10-21 | Z Height calibrtaion error | 0.26628
| 램테크놀러지㈜ | 2025-10-21 | 2) of 20nm | 8.824
| 램테크놀러지㈜ | 2025-10-21 | Z linearity by sample height | 1.0106
| 램테크놀러지㈜ | 2025-10-21 | 1) Z Scanner position at +3μm | 0.006557
| 램테크놀러지㈜ | 2025-10-21 | 3) Z Scanner position at -3μm | 0.001303
| 램테크놀러지㈜ | 2025-10-21 | Z Height noise | 0.022
| 램테크놀러지㈜ | 2025-10-21 | Full range of X Detector during X Scan sweep | 105.9
| 램테크놀러지㈜ | 2025-10-21 | Full range of Y Detector during Y Scan sweep | 104.9
| 램테크놀러지㈜ | 2025-10-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 150
| 램테크놀러지㈜ | 2025-10-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 154.2
| 램테크놀러지㈜ | 2025-10-21 | X Detector stroke calibration error (nominal scale error X) | 0.19
| 램테크놀러지㈜ | 2025-10-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.055
| 램테크놀러지㈜ | 2025-10-21 | X Scan Stroke | 105.9
| 램테크놀러지㈜ | 2025-10-21 | Y Scan Stroke | 104.9
| 램테크놀러지㈜ | 2025-10-21 | Full range of X Detector Raw during X Scan sweep | 14.1
| 램테크놀러지㈜ | 2025-10-21 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| 램테크놀러지㈜ | 2025-10-21 | X Detector noise (XYServo off) | 0.247
| 램테크놀러지㈜ | 2025-10-21 | Y Detector noise (XYServo off) | 0.248
| 램테크놀러지㈜ | 2025-10-21 | X Detector noise (XY Servo on) | 0.259
| 램테크놀러지㈜ | 2025-10-21 | Y Detector noise (XY Servo on) | 0.262
| 램테크놀러지㈜ | 2025-10-21 | X axis OPM | 0.889
| 램테크놀러지㈜ | 2025-10-21 | Y axis OPM | 1.71
| 램테크놀러지㈜ | 2025-10-21 | X Non-linearity | 0.166
| 램테크놀러지㈜ | 2025-10-21 | Y Non-linearity | 0.114
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Linearity by Z scanner position | 0.5
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Linearity by 4 different sample height | 1.0027
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Z Detector during Z Scan sweep | 12.74
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Height calibration error | 0.003036
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Drive calibration error | 0.11492
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Z Height during Z Scan sweep | 17.6372
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.0852
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.055
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.2014
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6789
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Height noise | 0.027
| WD Storage Technologies Ltd #3 | 2025-10-21 | OpticalAxisShiftPerFocusX | 0.0033
| WD Storage Technologies Ltd #3 | 2025-10-21 | OpticalAxisShiftPerFocusY | -0.0005
| WD Storage Technologies Ltd #3 | 2025-10-21 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000207
| WD Storage Technologies Ltd #3 | 2025-10-21 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4.83e-05
| WD Storage Technologies Ltd #3 | 2025-10-21 | RMS of ZeroScan | 0.035
| WD Storage Technologies Ltd #3 | 2025-10-21 | LightStrengthGain | 0.59
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of X Detector during X Scan sweep | 181.4632
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of X1 Detector during X Scan sweep | 179.5873
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of X2 Detector during X Scan sweep | 183.3391
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Y Detector during Y Scan sweep | 178.7022
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Y1 Detector during Y Scan sweep | 182.2659
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Y2 Detector during Y Scan sweep | 175.1386
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of X1 Detector Raw during X Scan sweep | 14.38
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of X2 Detector Raw during X Scan sweep | 14.04
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Y1 Detector Raw during Y Scan sweep | 14.17
| WD Storage Technologies Ltd #3 | 2025-10-21 | Full range of Y2 Detector Raw during Y Scan sweep | 14.73
| WD Storage Technologies Ltd #3 | 2025-10-21 | X Detector stroke calibration error (nominal scale error X) | 0.35
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.277
| WD Storage Technologies Ltd #3 | 2025-10-21 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 239.922
| WD Storage Technologies Ltd #3 | 2025-10-21 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 245.127
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 245.819
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 252.352
| WD Storage Technologies Ltd #3 | 2025-10-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 150.035
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 158.816
| WD Storage Technologies Ltd #3 | 2025-10-21 | XY Orthogonality | 89.869
| WD Storage Technologies Ltd #3 | 2025-10-21 | X axis OPM | 0.873
| WD Storage Technologies Ltd #3 | 2025-10-21 | X axis OPM raw Rpv | 0.586
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y axis OPM | 0.787
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y axis OPM raw Rpv | 1.942
| WD Storage Technologies Ltd #3 | 2025-10-21 | X Ringing(Rpv) | 0.532
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y Ringing(Rpv) | 0.431
| WD Storage Technologies Ltd #3 | 2025-10-21 | X Detector noise (XY Servo on) | 0.291
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y Detector noise (XY Servo on) | 0.307
| WD Storage Technologies Ltd #3 | 2025-10-21 | X1 Detector noise (XY Servo on) | 0.387
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y1 Detector noise (XY Servo on) | 0.406
| WD Storage Technologies Ltd #3 | 2025-10-21 | X2 Detector noise (XY Servo on) | 0.422
| WD Storage Technologies Ltd #3 | 2025-10-21 | Y2 Detector noise (XY Servo on) | 0.436
| WD Storage Technologies Ltd #3 | 2025-10-21 | StageSafetyXMax | -30969.9
| WD Storage Technologies Ltd #3 | 2025-10-21 | StageSafetyXMin | -88160.1
| WD Storage Technologies Ltd #3 | 2025-10-21 | StageSafetyY | 57177.41
| WD Storage Technologies Ltd #3 | 2025-10-21 | StageSafetyYMax | 37945.56
| WD Storage Technologies Ltd #3 | 2025-10-21 | StageSafetyYMin | -19231.9
| WD Storage Technologies Ltd #3 | 2025-10-21 | Check X axis upper limit sensor function | Yes
| WD Storage Technologies Ltd #3 | 2025-10-21 | Check X axis lower limit sensor function | Yes
| WD Storage Technologies Ltd #3 | 2025-10-21 | Check Y axis upper limit sensor function | Yes
| WD Storage Technologies Ltd #3 | 2025-10-21 | Check Y axis lower limit sensor function | Yes
| WD Storage Technologies Ltd #3 | 2025-10-21 | X axis Angle Range | 4.868
| WD Storage Technologies Ltd #3 | 2025-10-21 | Z Drift | 1.042
| WD Storage Technologies Ltd #3 | 2025-10-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.95
| WD Storage Technologies Ltd #3 | 2025-10-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.17
| WD Storage Technologies Ltd #3 | 2025-10-21 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.577
| WD Storage Technologies Ltd #3 | 2025-10-21 | Roughness Repeatability 1x1um(RSD%) | 1.947
| WD Storage Technologies Ltd #3 | 2025-10-21 | Roughness Repeatability 1x1um(1σ) | 0.007
| WD Storage Technologies Ltd #3 | 2025-10-21 | Roughness Repeatability 80x80um(RSD%) | 2.341
| WD Storage Technologies Ltd #3 | 2025-10-21 | Roughness Repeatability 80x80um(1σ) | 0.013
| LG Display #3 | 2025-10-21 | OpticalAxisShiftPerFocusX | 0.0001
| LG Display #3 | 2025-10-21 | OpticalAxisShiftPerFocusY | 0.0015
| LG Display #3 | 2025-10-21 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| LG Display #3 | 2025-10-21 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00011
| LG Display #3 | 2025-10-21 | Lateral(C-D) noise | 1.29
| LG Display #3 | 2025-10-21 | Min. Vibration Transmissbility (~100Hz) | 0.093
| LG Display #3 | 2025-10-21 | RMS of ZeroScan | 1.9
| LG Display #3 | 2025-10-21 | Z Linearity by Z scanner position | 1.3
| LG Display #3 | 2025-10-21 | Z Linearity by 4 different sample height | 1.006
| LG Display #3 | 2025-10-21 | Z Height noise | 0.026
| LG Display #3 | 2025-10-21 | Z Height calibrtaion error | 0.15178
| LG Display #3 | 2025-10-21 | Z Drive calibrtaion error | 0.21683
| LG Display #3 | 2025-10-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.5
| LG Display #3 | 2025-10-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.7
| LG Display #3 | 2025-10-21 | XY Orthogonality | 90.1
| LG Display #3 | 2025-10-21 | X detector noise (XY Servo on) | 0.26
| LG Display #3 | 2025-10-21 | Y detector noise (XY Servo on) | 0.26
| LG Display #3 | 2025-10-21 | X Ringing(Rpv) | 0.21
| LG Display #3 | 2025-10-21 | Y Ringing(Rpv) | 0.18
| LG Display #3 | 2025-10-21 | Full range of Z Detector during Z Scan sweep | 14.89
| LG Display #3 | 2025-10-21 | Full range of Z Height during Z Scan sweep | 18.7
| LG Display #3 | 2025-10-21 | Full range of X Detector Raw during X Scan sweep | 14
| LG Display #3 | 2025-10-21 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| LG Display #3 | 2025-10-21 | X Detector stroke calibration error (nominal scale error X) | 0.39
| LG Display #3 | 2025-10-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.23
| LG Display #3 | 2025-10-21 | Full range in um during X Scan sweep (XY Servo off) | 113.2
| LG Display #3 | 2025-10-21 | Full range in um during Y Scan sweep (XY Servo off) | 111.9
| LG Display #3 | 2025-10-21 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 162
| LG Display #3 | 2025-10-21 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 155.7
| LG Display #3 | 2025-10-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 109.5
| LG Display #3 | 2025-10-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 111.3
| LG Display #3 | 2025-10-21 | X axis OPM | 2.17
| LG Display #3 | 2025-10-21 | Y axis OPM | 1.9
| LG Display #3 | 2025-10-21 | Z Linearity by 4 different sample height | 1.002
| LG Display #3 | 2025-10-21 | Z Height noise | 0.029
| LG Display #3 | 2025-10-21 | Z Height calibration error | 0.08673
| LG Display #3 | 2025-10-21 | Z Drive calibration error | 0.17346
| LG Display #3 | 2025-10-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.5
| LG Display #3 | 2025-10-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.5
| LG Display #3 | 2025-10-21 | Full range of Z Detector during Z Scan sweep | 13.82
| LG Display #3 | 2025-10-21 | Full range of Z Height during Z Scan sweep | 18.8
| LG Display #3 | 2025-10-21 | Full range of X Detector Raw during X Scan sweep | 15
| LG Display #3 | 2025-10-21 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| LG Display #3 | 2025-10-21 | X Detector stroke calibration error (nominal scale error X) | 0.12
| LG Display #3 | 2025-10-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.1
| LG Display #3 | 2025-10-21 | Full range in um during X Scan sweep (XY Servo off) | 109.5
| LG Display #3 | 2025-10-21 | Full range in um during Y Scan sweep (XY Servo off) | 115
| LG Display #3 | 2025-10-21 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 145.612
| LG Display #3 | 2025-10-21 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 155.344
| LG Display #3 | 2025-10-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 105.7
| LG Display #3 | 2025-10-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 105.1
| LG Display #3 | 2025-10-21 | X axis OPM | 4.82
| LG Display #3 | 2025-10-21 | Y axis OPM | 3.5
| LG Display #3 | 2025-10-21 | LightStrengthGain | 0.87
| LG Display #3 | 2025-10-21 | X axis, Repeatability @1st position | 0.06
| LG Display #3 | 2025-10-21 | X axis, Repeatability @ 2nd position | 0.13
| LG Display #3 | 2025-10-21 | X axis, Repeatability @ 3rd position | 0.03
| LG Display #3 | 2025-10-21 | X axis, Repeatability @ 4th position | 0.02
| LG Display #3 | 2025-10-21 | X axis, Accuracy @1st position | 0.36
| LG Display #3 | 2025-10-21 | X axis, Accuracy @ 2nd position | 0.65
| LG Display #3 | 2025-10-21 | X axis, Accuracy @ 3rd position | 0.12
| LG Display #3 | 2025-10-21 | X axis, Accuracy @ 4th position | 0.2
| LG Display #3 | 2025-10-21 | Y axis, Repeatability @1st position | 0.15
| LG Display #3 | 2025-10-21 | Y axis, Repeatability @ 2nd position | 0.21
| LG Display #3 | 2025-10-21 | Y axis, Repeatability @ 3rd position | 0.05
| LG Display #3 | 2025-10-21 | Y axis, Repeatability @ 4th position | 0.05
| LG Display #3 | 2025-10-21 | Y axis, Accuracy @1st position | 0.77
| LG Display #3 | 2025-10-21 | Y axis, Accuracy @ 2nd position | 0.89
| LG Display #3 | 2025-10-21 | Y axis, Accuracy @ 3rd position | 0.43
| LG Display #3 | 2025-10-21 | Y axis, Accuracy @ 4th position | 0.51
| LG Display #3 | 2025-10-21 | Image offset after moving from -X to +X (X scan) | 405.273
| LG Display #3 | 2025-10-21 | Image offset after moving from -Y to +Y (Y scan) | 253.906
| LG Display #3 | 2025-10-21 | Image offset after moving from -X to +X (Y scan) | 92.7734
| LG Display #3 | 2025-10-21 | Image offset after moving from -Y to +Y (X scan) | 170.898
| LG Display #3 | 2025-10-21 | Image offset after moving from +X to -X (X scan) | 458.984
| LG Display #3 | 2025-10-21 | Image offset after moving from +Y to -Y (Y scan) | 454.102
| LG Display #3 | 2025-10-21 | Image offset after moving from +X to -X (Y scan) | 87.8906
| LG Display #3 | 2025-10-21 | Image offset after moving from +Y to -Y (X scan) | 302.734
| LG Display #3 | 2025-10-21 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ) | 0.44
| LG Display #3 | 2025-10-21 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%) | 0.01
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (X direction-Left) | 54
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (X direction-Right) | 55.53
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (Y direction-Left) | 54
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (Y direction-Right) | 55.6
| LG Display #3 | 2025-10-21 | X-Scan Repeatability (10 repeats, RSD%) | 0.04
| LG Display #3 | 2025-10-21 | Y-Scan Repeatability (10 repeats, RSD%) | 0.03
| LG Display #3 | 2025-10-21 | Mechanical Align Accuracy | 168.3
| LG Display #3 | 2025-10-21 | Check the Multi Peattern Upload | 12
| LG Display #3 | 2025-10-21 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ) | 1.4
| LG Display #3 | 2025-10-21 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%) | 0.03
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (X direction-Left) | 54.52
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (X direction-Right) | 55.08
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (Y direction-Left) | 53.85
| LG Display #3 | 2025-10-21 | Z-XY Orthogonality (Y direction-Right) | 55.09
| LG Display #3 | 2025-10-21 | X-Scan Repeatability (10 repeats, RSD%) | 0.06
| LG Display #3 | 2025-10-21 | Y-Scan Repeatability (10 repeats, RSD%) | 0.03
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Linearity by Z scanner position | 0.87
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Linearity by 4 different sample height | 1.0004
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Z Detector during Z Scan sweep | 12.87
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Height calibration error | 0.040546
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Drive calibration error | 0.20794
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Z Height during Z Scan sweep | 17.8024
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.6709
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.1184
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.7251
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7508
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Height noise | 0.029
| WD Storage Technologies Ltd #4 | 2025-10-21 | OpticalAxisShiftPerFocusX | 0.0021
| WD Storage Technologies Ltd #4 | 2025-10-21 | OpticalAxisShiftPerFocusY | 0.001
| WD Storage Technologies Ltd #4 | 2025-10-21 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000251
| WD Storage Technologies Ltd #4 | 2025-10-21 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5.68e-05
| WD Storage Technologies Ltd #4 | 2025-10-21 | RMS of ZeroScan | 0.026
| WD Storage Technologies Ltd #4 | 2025-10-21 | LightStrengthGain | 0.5
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of X Detector during X Scan sweep | 179.7799
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of X1 Detector during X Scan sweep | 179.7799
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of X2 Detector during X Scan sweep | 184.2684
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Y Detector during Y Scan sweep | 179.267
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Y1 Detector during Y Scan sweep | 179.267
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Y2 Detector during Y Scan sweep | 174.5321
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of X1 Detector Raw during X Scan sweep | 14.62
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of X2 Detector Raw during X Scan sweep | 14.47
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Y1 Detector Raw during Y Scan sweep | 14.15
| WD Storage Technologies Ltd #4 | 2025-10-21 | Full range of Y2 Detector Raw during Y Scan sweep | 14.56
| WD Storage Technologies Ltd #4 | 2025-10-21 | X Detector stroke calibration error (nominal scale error X) | 0.091
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.094
| WD Storage Technologies Ltd #4 | 2025-10-21 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 246.179
| WD Storage Technologies Ltd #4 | 2025-10-21 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 256.155
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 242.337
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 247.012
| WD Storage Technologies Ltd #4 | 2025-10-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 156.28
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 153.965
| WD Storage Technologies Ltd #4 | 2025-10-21 | XY Orthogonality | 90.165
| WD Storage Technologies Ltd #4 | 2025-10-21 | X axis OPM | 0.475
| WD Storage Technologies Ltd #4 | 2025-10-21 | X axis OPM raw Rpv | 10.468
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y axis OPM | 0.82
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y axis OPM raw Rpv | 2.135
| WD Storage Technologies Ltd #4 | 2025-10-21 | X Ringing(Rpv) | 0.532
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y Ringing(Rpv) | 0.431
| WD Storage Technologies Ltd #4 | 2025-10-21 | X Detector noise (XY Servo on) | 0.266
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y Detector noise (XY Servo on) | 0.214
| WD Storage Technologies Ltd #4 | 2025-10-21 | X1 Detector noise (XY Servo on) | 0.353
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y1 Detector noise (XY Servo on) | 0.265
| WD Storage Technologies Ltd #4 | 2025-10-21 | X2 Detector noise (XY Servo on) | 0.381
| WD Storage Technologies Ltd #4 | 2025-10-21 | Y2 Detector noise (XY Servo on) | 0.324
| WD Storage Technologies Ltd #4 | 2025-10-21 | Z Drift | 2.262
| WD Storage Technologies Ltd #4 | 2025-10-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| WD Storage Technologies Ltd #4 | 2025-10-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.01
| WD Storage Technologies Ltd #4 | 2025-10-21 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.596
| WD Storage Technologies Ltd #4 | 2025-10-21 | Roughness Repeatability 1x1um(RSD%) | 3.663
| WD Storage Technologies Ltd #4 | 2025-10-21 | Roughness Repeatability 1x1um(1σ) | 0.285
| WD Storage Technologies Ltd #4 | 2025-10-21 | Roughness Repeatability 80x80um(RSD%) | 1.4
| WD Storage Technologies Ltd #4 | 2025-10-21 | Roughness Repeatability 80x80um(1σ) | 0.007
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Linearity by Z scanner position | 0.78
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Linearity by 4 different sample height | 1.004
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Z Detector during Z Scan sweep | 12.93
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Height calibration error | 0.12511
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Drive calibration error | 0.28361
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Z Height during Z Scan sweep | 17.517
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.1014
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2478
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.1514
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7634
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Height noise | 0.026
| WD Storage Technologies Ltd #5 | 2025-10-21 | OpticalAxisShiftPerFocusX | 0.0026
| WD Storage Technologies Ltd #5 | 2025-10-21 | OpticalAxisShiftPerFocusY | 0.0002
| WD Storage Technologies Ltd #5 | 2025-10-21 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00016
| WD Storage Technologies Ltd #5 | 2025-10-21 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| WD Storage Technologies Ltd #5 | 2025-10-21 | RMS of ZeroScan | 0.026
| WD Storage Technologies Ltd #5 | 2025-10-21 | LightStrengthGain | 0.5
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of X Detector during X Scan sweep | 177.9139
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of X1 Detector during X Scan sweep | 185.5331
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of X2 Detector during X Scan sweep | 177.9139
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Y Detector during Y Scan sweep | 184.2607
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Y1 Detector during Y Scan sweep | 184.2607
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Y2 Detector during Y Scan sweep | 182.3194
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of X2 Detector Raw during X Scan sweep | 14.15
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Y1 Detector Raw during Y Scan sweep | 14.11
| WD Storage Technologies Ltd #5 | 2025-10-21 | Full range of Y2 Detector Raw during Y Scan sweep | 14.61
| WD Storage Technologies Ltd #5 | 2025-10-21 | X Detector stroke calibration error (nominal scale error X) | 0.076
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.192
| WD Storage Technologies Ltd #5 | 2025-10-21 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 257.527
| WD Storage Technologies Ltd #5 | 2025-10-21 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 251.664
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 261.582
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 249.79
| WD Storage Technologies Ltd #5 | 2025-10-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 160.983
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 154.599
| WD Storage Technologies Ltd #5 | 2025-10-21 | XY Orthogonality | 90.142
| WD Storage Technologies Ltd #5 | 2025-10-21 | X axis OPM | 0.65
| WD Storage Technologies Ltd #5 | 2025-10-21 | X axis OPM raw Rpv | 9.423
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y axis OPM | 0.5
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y axis OPM raw Rpv | 2.408
| WD Storage Technologies Ltd #5 | 2025-10-21 | X Ringing(Rpv) | 0.684
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y Ringing(Rpv) | 0.531
| WD Storage Technologies Ltd #5 | 2025-10-21 | X Detector noise (XY Servo on) | 0.297
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y Detector noise (XY Servo on) | 0.31
| WD Storage Technologies Ltd #5 | 2025-10-21 | X1 Detector noise (XY Servo on) | 0.43
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y1 Detector noise (XY Servo on) | 0.44
| WD Storage Technologies Ltd #5 | 2025-10-21 | X2 Detector noise (XY Servo on) | 0.402
| WD Storage Technologies Ltd #5 | 2025-10-21 | Y2 Detector noise (XY Servo on) | 0.424
| WD Storage Technologies Ltd #5 | 2025-10-21 | Z Drift | 3.353
| WD Storage Technologies Ltd #5 | 2025-10-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| WD Storage Technologies Ltd #5 | 2025-10-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| WD Storage Technologies Ltd #5 | 2025-10-21 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.527
| WD Storage Technologies Ltd #5 | 2025-10-21 | Roughness Repeatability 1x1um(RSD%) | 1.5
| WD Storage Technologies Ltd #5 | 2025-10-21 | Roughness Repeatability 1x1um(1σ) | 0.014
| WD Storage Technologies Ltd #5 | 2025-10-21 | Roughness Repeatability 80x80um(RSD%) | 3.531
| WD Storage Technologies Ltd #5 | 2025-10-21 | Roughness Repeatability 80x80um(1σ) | 0.023
| PSI EAST OFFICE | 2025-10-16 | Lateral(C-D) sensitivity | 26.9
| PSI EAST OFFICE | 2025-10-16 | Lateral(C-D) noise | 0.085
| PSI EAST OFFICE | 2025-10-16 | A-B noise | 8.491
| PSI EAST OFFICE | 2025-10-16 | NCM Amplitude noise | 8.89
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.19
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| PSI EAST OFFICE | 2025-10-16 | Roughness Repeatability 1x1um(RSD%) | 3.314
| PSI EAST OFFICE | 2025-10-16 | Roughness Repeatability 1x1um(1σ) | 0.011
| PSI EAST OFFICE | 2025-10-16 | Full range of Z Height during Z Scan sweep | 18
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.971
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.125
| PSI EAST OFFICE | 2025-10-16 | Full range of Z Detector during Z Scan sweep | 12.9
| PSI EAST OFFICE | 2025-10-16 | Z Drive calibration error | 0.31513
| PSI EAST OFFICE | 2025-10-16 | Z Height calibration error | 0.24234
| PSI EAST OFFICE | 2025-10-16 | Z linearity by sample height | 0.99489
| PSI EAST OFFICE | 2025-10-16 | 1) Z Scanner position at +3μm | 0.004486
| PSI EAST OFFICE | 2025-10-16 | 3) Z Scanner position at -3μm | 0.002893
| PSI EAST OFFICE | 2025-10-16 | Z Height noise | 0.029
| PSI EAST OFFICE | 2025-10-16 | Full range of X Detector during X Scan sweep | 114.2
| PSI EAST OFFICE | 2025-10-16 | Full range of Y Detector during Y Scan sweep | 113
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.3
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.8
| PSI EAST OFFICE | 2025-10-16 | X Detector stroke calibration error (nominal scale error X) | 0.228
| PSI EAST OFFICE | 2025-10-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.297
| PSI EAST OFFICE | 2025-10-16 | X Scan Stroke | 114.2
| PSI EAST OFFICE | 2025-10-16 | Y Scan Stroke | 113
| PSI EAST OFFICE | 2025-10-16 | Full range of X Detector Raw during X Scan sweep | 14
| PSI EAST OFFICE | 2025-10-16 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| PSI EAST OFFICE | 2025-10-16 | X Detector noise (XY Servo on) | 0.274
| PSI EAST OFFICE | 2025-10-16 | Y Detector noise (XY Servo on) | 0.268
| PSI EAST OFFICE | 2025-10-16 | X Detector noise (XY servo off) | 0.267
| PSI EAST OFFICE | 2025-10-16 | Y Detector noise (XY servo off) | 0.265
| PSI EAST OFFICE | 2025-10-16 | X axis OPM | 0.905
| PSI EAST OFFICE | 2025-10-16 | Y axis OPM | 0.651
| PSI EAST OFFICE | 2025-10-16 | X Non-linearity | 0.125
| PSI EAST OFFICE | 2025-10-16 | Y Non-linearity | 0.251
| PSI EAST OFFICE | 2025-10-16 | Lateral(C-D) sensitivity | 26.9
| PSI EAST OFFICE | 2025-10-16 | Lateral(C-D) noise | 0.085
| PSI EAST OFFICE | 2025-10-16 | A-B noise | 8.491
| PSI EAST OFFICE | 2025-10-16 | NCM Amplitude noise | 8.89
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.19
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| PSI EAST OFFICE | 2025-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| PSI EAST OFFICE | 2025-10-16 | Roughness Repeatability 1x1um(RSD%) | 3.314
| PSI EAST OFFICE | 2025-10-16 | Roughness Repeatability 1x1um(1σ) | 0.011
| PSI EAST OFFICE | 2025-10-16 | Full range of Z Height during Z Scan sweep | 18
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.971
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.125
| PSI EAST OFFICE | 2025-10-16 | Full range of Z Detector during Z Scan sweep | 12.9
| PSI EAST OFFICE | 2025-10-16 | Z Drive calibration error | 0.31513
| PSI EAST OFFICE | 2025-10-16 | Z Height calibration error | 0.24234
| PSI EAST OFFICE | 2025-10-16 | Z linearity by sample height | 0.99489
| PSI EAST OFFICE | 2025-10-16 | 1) Z Scanner position at +3μm | 0.004486
| PSI EAST OFFICE | 2025-10-16 | 3) Z Scanner position at -3μm | 0.002893
| PSI EAST OFFICE | 2025-10-16 | Z Height noise | 0.029
| PSI EAST OFFICE | 2025-10-16 | Full range of X Detector during X Scan sweep | 114.2
| PSI EAST OFFICE | 2025-10-16 | Full range of Y Detector during Y Scan sweep | 113
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.3
| PSI EAST OFFICE | 2025-10-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.8
| PSI EAST OFFICE | 2025-10-16 | X Detector stroke calibration error (nominal scale error X) | 0.228
| PSI EAST OFFICE | 2025-10-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.297
| PSI EAST OFFICE | 2025-10-16 | X Scan Stroke | 114.2
| PSI EAST OFFICE | 2025-10-16 | Y Scan Stroke | 113
| PSI EAST OFFICE | 2025-10-16 | Full range of X Detector Raw during X Scan sweep | 14
| PSI EAST OFFICE | 2025-10-16 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| PSI EAST OFFICE | 2025-10-16 | X Detector noise (XY Servo on) | 0.274
| PSI EAST OFFICE | 2025-10-16 | Y Detector noise (XY Servo on) | 0.268
| PSI EAST OFFICE | 2025-10-16 | X Detector noise (XY servo off) | 0.267
| PSI EAST OFFICE | 2025-10-16 | Y Detector noise (XY servo off) | 0.265
| PSI EAST OFFICE | 2025-10-16 | X axis OPM | 0.905
| PSI EAST OFFICE | 2025-10-16 | Y axis OPM | 0.651
| PSI EAST OFFICE | 2025-10-16 | X Non-linearity | 0.125
| PSI EAST OFFICE | 2025-10-16 | Y Non-linearity | 0.251
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | Lateral(C-D) sensitivity | 15.2
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | Lateral(C-D) noise | 0.044
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | A-B noise | 5.458
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | NCM Amplitude noise | 9.44
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.34
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.35
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.28
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.36
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | Roughness Repeatability 1x1um(RSD%) | 0.883
| Kitagawa Gresstech Co., Ltd | 2025-10-16 | Roughness Repeatability 1x1um(1σ) | 0.003
| Zhejiang SDS | 2025-10-14 | Z Linearity by Z scanner position | 0.32
| Zhejiang SDS | 2025-10-14 | Full range of Z Detector during Z Scan sweep | 12.29
| Zhejiang SDS | 2025-10-14 | Z Height calibration error | 0.008673
| Zhejiang SDS | 2025-10-14 | Z Drive calibration error | 0.18474
| Zhejiang SDS | 2025-10-14 | Z Linearity by 4 different sample height | 0.99917
| Zhejiang SDS | 2025-10-14 | Full range of Z Height during Z Scan sweep | 16.7745
| Zhejiang SDS | 2025-10-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.4811
| Zhejiang SDS | 2025-10-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.0983
| Zhejiang SDS | 2025-10-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5648
| Zhejiang SDS | 2025-10-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1417
| Zhejiang SDS | 2025-10-14 | Z Height noise | 0.03
| Zhejiang SDS | 2025-10-14 | Wafer placement(Loading) error on the load port1 (X offset) | 18.4
| Zhejiang SDS | 2025-10-14 | Wafer placement(Loading) error on the load port1 (Y offset) | 14.26
| Zhejiang SDS | 2025-10-14 | Wafer placement(Loading) error on the load port2 (X offset) | 35.6
| Zhejiang SDS | 2025-10-14 | Wafer placement(Loading) error on the load port2 (Y offset) | 41.02
| Zhejiang SDS | 2025-10-14 | OpticalAxisShiftPerFocusX | -0.0003
| Zhejiang SDS | 2025-10-14 | OpticalAxisShiftPerFocusY | -0.0006
| Zhejiang SDS | 2025-10-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7.89e-05
| Zhejiang SDS | 2025-10-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8.76e-05
| Zhejiang SDS | 2025-10-14 | Lift pins synchronization offset check (X offset) | 24.14
| Zhejiang SDS | 2025-10-14 | Lift pins synchronization offset check (Y offset) | 13.22
| Zhejiang SDS | 2025-10-14 | Profiler noise @200um | 0.114
| Zhejiang SDS | 2025-10-14 | Roll motion to X+ full scan | 7.8
| Zhejiang SDS | 2025-10-14 | Position repeatability(10repeats, Position Based) | 0.34
| Zhejiang SDS | 2025-10-14 | OPM Repeatability @25mm | 1.214
| Zhejiang SDS | 2025-10-14 | OPM Repeatability @10mm | 0.966
| Zhejiang SDS | 2025-10-14 | OPM Repeatability @5mm | 0.884
| Zhejiang SDS | 2025-10-14 | OPM Repeatability @1mm | 0.584
| Zhejiang SDS | 2025-10-14 | Max OPM @25mm | 90.088
| Zhejiang SDS | 2025-10-14 | Max OPM @10mm | 48.563
| Zhejiang SDS | 2025-10-14 | Max OPM @5mm | 22.74
| Zhejiang SDS | 2025-10-14 | Max OPM @1mm | 6.751
| Zhejiang SDS | 2025-10-14 | Torque Profile (50mm, 200um/s) | 0.53771
| Zhejiang SDS | 2025-10-14 | Torque Profile (50mm, 20um/s) | 0.68683
| Zhejiang SDS | 2025-10-14 | Lateral(C-D) sensitivity | 15.6
| Zhejiang SDS | 2025-10-14 | Lateral(C-D) noise | 0.138
| Zhejiang SDS | 2025-10-14 | Min. Vibration Transmissbility (~100Hz) | 0.041
| Zhejiang SDS | 2025-10-14 | RMS of ZeroScan | 0.049
| Zhejiang SDS | 2025-10-14 | LightStrengthGain | 0.45
| Zhejiang SDS | 2025-10-14 | Full range of X Detector during X Scan sweep | 212.7937
| Zhejiang SDS | 2025-10-14 | Full range of X1 Detector during X Scan sweep | 208.2337
| Zhejiang SDS | 2025-10-14 | Full range of X2 Detector during X Scan sweep | 212.7937
| Zhejiang SDS | 2025-10-14 | Full range of Y Detector during Y Scan sweep | 205.4233
| Zhejiang SDS | 2025-10-14 | Full range of Y1 Detector during Y Scan sweep | 205.4233
| Zhejiang SDS | 2025-10-14 | Full range of Y2 Detector during Y Scan sweep | 210.0524
| Zhejiang SDS | 2025-10-14 | Full range of X1 Detector Raw during X Scan sweep | 14.43
| Zhejiang SDS | 2025-10-14 | Full range of X2 Detector Raw during X Scan sweep | 14.34
| Zhejiang SDS | 2025-10-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.47
| Zhejiang SDS | 2025-10-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.36
| Zhejiang SDS | 2025-10-14 | X Detector stroke calibration error (nominal scale error X) | 0.188
| Zhejiang SDS | 2025-10-14 | Y Detector stroke calibration error (nominal scale error Y) | 0.407
| Zhejiang SDS | 2025-10-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 276.5022
| Zhejiang SDS | 2025-10-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 290.7311
| Zhejiang SDS | 2025-10-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 278.0589
| Zhejiang SDS | 2025-10-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 290.0378
| Zhejiang SDS | 2025-10-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 184.417
| Zhejiang SDS | 2025-10-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 175.718
| Zhejiang SDS | 2025-10-14 | X Non-linearity error | 3.904
| Zhejiang SDS | 2025-10-14 | Y Non-linearity error | 3.3794
| Zhejiang SDS | 2025-10-14 | X axis OPM | 0.901
| Zhejiang SDS | 2025-10-14 | Y axis OPM | 1.125
| Zhejiang SDS | 2025-10-14 | X Ringing(Rpv) | 0.101
| Zhejiang SDS | 2025-10-14 | Y Ringing(Rpv) | 0.092
| Zhejiang SDS | 2025-10-14 | X Detector noise (XY Servo on) | 0.372
| Zhejiang SDS | 2025-10-14 | Y Detector noise (XY Servo on) | 0.355
| Zhejiang SDS | 2025-10-14 | X1 Detector noise (XY Servo on) | 0.485
| Zhejiang SDS | 2025-10-14 | Y1 Detector noise (XY Servo on) | 0.471
| Zhejiang SDS | 2025-10-14 | X2 Detector noise (XY Servo on) | 0.496
| Zhejiang SDS | 2025-10-14 | Y2 Detector noise (XY Servo on) | 0.484
| Zhejiang SDS | 2025-10-14 | RMS error after Stage Calibration | 2.67
| Zhejiang SDS | 2025-10-14 | Image offset after moving from -X to +X (X scan) | 23.4375
| Zhejiang SDS | 2025-10-14 | Image offset after moving from -Y to +Y (Y scan) | 114.9902
| Zhejiang SDS | 2025-10-14 | Image offset after moving from -X to +X (Y scan) | 16.1133
| Zhejiang SDS | 2025-10-14 | Image offset after moving from -Y to +Y (X scan) | 25.6348
| Zhejiang SDS | 2025-10-14 | Image offset after moving from +X to -X (X scan) | 8.0566
| Zhejiang SDS | 2025-10-14 | Image offset after moving from +Y to -Y (Y scan) | 11.7188
| Zhejiang SDS | 2025-10-14 | Image offset after moving from +X to -X (Y scan) | 32.2266
| Zhejiang SDS | 2025-10-14 | Image offset after moving from +Y to -Y (X scan) | 27.832
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X offset (Vision Pattern Matching) | 0.49316
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.4551
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.09377
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.32664
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X offset (In-Die Align) | 1.3574
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y offset (In-Die Align) | 2.0361
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X 1σ (In-Die Align) | 0.28988
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y 1σ (In-Die Align) | 0.41163
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X offset (LLC Translation) | 2.2461
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y offset (LLC Translation) | 2.8467
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X 1σ (LLC Translation) | 0.56699
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y 1σ (LLC Translation) | 0.52676
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X offset (Global Align) | 5.4248
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y offset (Global Align) | 6.1182
| Zhejiang SDS | 2025-10-14 | XY Stage positioning X 1σ (Global Align) | 1.0651
| Zhejiang SDS | 2025-10-14 | XY Stage positioning Y 1σ (Global Align) | 1.4375
| Zhejiang SDS | 2025-10-14 | Z Drift | 16.2796
| Zhejiang SDS | 2025-10-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27105
| Zhejiang SDS | 2025-10-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.049121
| Zhejiang SDS | 2025-10-14 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.693
| Zhejiang SDS | 2025-10-14 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1265
| Zhejiang SDS | 2025-10-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37541
| Zhejiang SDS | 2025-10-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.068096
| Zhejiang SDS | 2025-10-14 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.80989
| Zhejiang SDS | 2025-10-14 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14776
| Zhejiang SDS | 2025-10-14 | Cycle time for 15-point measurement | 1768.829
| Zhejiang SDS | 2025-10-14 | Short-term stability using Step Height Sample (3σ) | 0.87392
| Zhejiang SDS | 2025-10-14 | Roughness Repeatability 1x1um(RSD%) | 3.4811
| Zhejiang SDS | 2025-10-14 | Roughness Repeatability 1x1um(1σ) | 0.002983
| Zhejiang SDS | 2025-10-14 | Roughness Repeatability 80x80um(RSD%) | 0.28865
| Zhejiang SDS | 2025-10-14 | Roughness Repeatability 80x80um(1σ) | 0.001033
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Linearity by Z scanner position | 1.2
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Linearity by 4 different sample height | 1.0048
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Z Detector during Z Scan sweep | 12.53
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Height calibration error | 0.02732
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Drive calibration error | 0.433
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Z Height during Z Scan sweep | 17.1033
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.2863
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.407
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.3308
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4036
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Height noise | 0.027
| WD Storage Technologies Ltd #1 | 2025-10-14 | OpticalAxisShiftPerFocusX | 0.0033
| WD Storage Technologies Ltd #1 | 2025-10-14 | OpticalAxisShiftPerFocusY | -0.0005
| WD Storage Technologies Ltd #1 | 2025-10-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| WD Storage Technologies Ltd #1 | 2025-10-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.0001
| WD Storage Technologies Ltd #1 | 2025-10-14 | RMS of ZeroScan | 0.026
| WD Storage Technologies Ltd #1 | 2025-10-14 | LightStrengthGain | 0.5
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of X Detector during X Scan sweep | 182.9189
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of X1 Detector during X Scan sweep | 182.9189
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of X2 Detector during X Scan sweep | 184.6356
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Y Detector during Y Scan sweep | 188.4989
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Y1 Detector during Y Scan sweep | 188.5422
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Y2 Detector during Y Scan sweep | 188.4989
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of X1 Detector Raw during X Scan sweep | 14.5
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of X2 Detector Raw during X Scan sweep | 14.47
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.54
| WD Storage Technologies Ltd #1 | 2025-10-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.55
| WD Storage Technologies Ltd #1 | 2025-10-14 | X Detector stroke calibration error (nominal scale error X) | 0.131
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y Detector stroke calibration error (nominal scale error Y) | 0.307
| WD Storage Technologies Ltd #1 | 2025-10-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 252.22
| WD Storage Technologies Ltd #1 | 2025-10-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 256.647
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 256.804
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 261.022
| WD Storage Technologies Ltd #1 | 2025-10-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 156.513
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 166.383
| WD Storage Technologies Ltd #1 | 2025-10-14 | XY Orthogonality | 89.817
| WD Storage Technologies Ltd #1 | 2025-10-14 | X axis OPM | 0.664
| WD Storage Technologies Ltd #1 | 2025-10-14 | X axis OPM raw Rpv | 2.129
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y axis OPM | 0.44
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y axis OPM raw Rpv | 1.755
| WD Storage Technologies Ltd #1 | 2025-10-14 | X Ringing(Rpv) | 0.389
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y Ringing(Rpv) | 0.408
| WD Storage Technologies Ltd #1 | 2025-10-14 | X Detector noise (XY Servo on) | 0.28
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y Detector noise (XY Servo on) | 0.314
| WD Storage Technologies Ltd #1 | 2025-10-14 | X1 Detector noise (XY Servo on) | 0.389
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y1 Detector noise (XY Servo on) | 0.438
| WD Storage Technologies Ltd #1 | 2025-10-14 | X2 Detector noise (XY Servo on) | 0.389
| WD Storage Technologies Ltd #1 | 2025-10-14 | Y2 Detector noise (XY Servo on) | 0.437
| WD Storage Technologies Ltd #1 | 2025-10-14 | Z Drift | 1.183
| WD Storage Technologies Ltd #1 | 2025-10-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.63
| WD Storage Technologies Ltd #1 | 2025-10-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| WD Storage Technologies Ltd #1 | 2025-10-14 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.301
| WD Storage Technologies Ltd #1 | 2025-10-14 | Roughness Repeatability 1x1um(RSD%) | 2.157
| WD Storage Technologies Ltd #1 | 2025-10-14 | Roughness Repeatability 1x1um(1σ) | 0.003
| WD Storage Technologies Ltd #1 | 2025-10-14 | Roughness Repeatability 80x80um(RSD%) | 4.181
| WD Storage Technologies Ltd #1 | 2025-10-14 | Roughness Repeatability 80x80um(1σ) | 0.019
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Linearity by Z scanner position | 0.23
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Linearity by 4 different sample height | 1.0008
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Z Detector during Z Scan sweep | 13.06
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Height calibration error | 0.11167
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Drive calibration error | 0.043365
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Z Height during Z Scan sweep | 17.8353
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.1718
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2373
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.4552
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8794
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Height noise | 0.026
| WD Storage Technologies Ltd #2 | 2025-10-14 | OpticalAxisShiftPerFocusX | 0.0023
| WD Storage Technologies Ltd #2 | 2025-10-14 | OpticalAxisShiftPerFocusY | 0.0013
| WD Storage Technologies Ltd #2 | 2025-10-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 6.75e-05
| WD Storage Technologies Ltd #2 | 2025-10-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3.16e-05
| WD Storage Technologies Ltd #2 | 2025-10-14 | RMS of ZeroScan | 0.033
| WD Storage Technologies Ltd #2 | 2025-10-14 | LightStrengthGain | 0.55
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of X Detector during X Scan sweep | 179.8796
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of X1 Detector during X Scan sweep | 179.8796
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of X2 Detector during X Scan sweep | 183.2523
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Y Detector during Y Scan sweep | 174.5509
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Y1 Detector during Y Scan sweep | 176.4089
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Y2 Detector during Y Scan sweep | 174.5509
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of X1 Detector Raw during X Scan sweep | 14.9
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of X2 Detector Raw during X Scan sweep | 14.56
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.78
| WD Storage Technologies Ltd #2 | 2025-10-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.91
| WD Storage Technologies Ltd #2 | 2025-10-14 | X Detector stroke calibration error (nominal scale error X) | 0.006
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y Detector stroke calibration error (nominal scale error Y) | 0.373
| WD Storage Technologies Ltd #2 | 2025-10-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.795
| WD Storage Technologies Ltd #2 | 2025-10-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 254.758
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 245.881
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 239.628
| WD Storage Technologies Ltd #2 | 2025-10-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.058
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 158.889
| WD Storage Technologies Ltd #2 | 2025-10-14 | XY Orthogonality | 90.112
| WD Storage Technologies Ltd #2 | 2025-10-14 | X axis OPM | 0.837
| WD Storage Technologies Ltd #2 | 2025-10-14 | X axis OPM raw Rpv | 3.766
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y axis OPM | 0.586
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y axis OPM raw Rpv | 4.273
| WD Storage Technologies Ltd #2 | 2025-10-14 | X Ringing(Rpv) | 0.406
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y Ringing(Rpv) | 0.387
| WD Storage Technologies Ltd #2 | 2025-10-14 | X Detector noise (XY Servo on) | 0.244
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y Detector noise (XY Servo on) | 0.195
| WD Storage Technologies Ltd #2 | 2025-10-14 | X1 Detector noise (XY Servo on) | 0.369
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y1 Detector noise (XY Servo on) | 0.317
| WD Storage Technologies Ltd #2 | 2025-10-14 | X2 Detector noise (XY Servo on) | 0.304
| WD Storage Technologies Ltd #2 | 2025-10-14 | Y2 Detector noise (XY Servo on) | 0.212
| WD Storage Technologies Ltd #2 | 2025-10-14 | Z Drift | 0.978
| WD Storage Technologies Ltd #2 | 2025-10-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.8
| WD Storage Technologies Ltd #2 | 2025-10-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| WD Storage Technologies Ltd #2 | 2025-10-14 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.943
| WD Storage Technologies Ltd #2 | 2025-10-14 | Roughness Repeatability 1x1um(RSD%) | 2.157
| WD Storage Technologies Ltd #2 | 2025-10-14 | Roughness Repeatability 1x1um(1σ) | 0.003
| WD Storage Technologies Ltd #2 | 2025-10-14 | Roughness Repeatability 80x80um(RSD%) | 0.981
| WD Storage Technologies Ltd #2 | 2025-10-14 | Roughness Repeatability 80x80um(1σ) | 0.005
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Linearity by Z scanner position | 0.15
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Linearity by 4 different sample height | 0.99724
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Z Detector during Z Scan sweep | 13.15
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Height calibration error | 0.035776
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Drive calibration error | 0.24848
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Z Height during Z Scan sweep | 18.2091
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.8938
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.8591
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.9153
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12.0249
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Height noise | 0.027
| Fujian Jingan Optoelectronics | 2025-10-13 | OpticalAxisShiftPerFocusX | 0.0038
| Fujian Jingan Optoelectronics | 2025-10-13 | OpticalAxisShiftPerFocusY | -0.0011
| Fujian Jingan Optoelectronics | 2025-10-13 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00037
| Fujian Jingan Optoelectronics | 2025-10-13 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0
| Fujian Jingan Optoelectronics | 2025-10-13 | Lateral(C-D) sensitivity | 11
| Fujian Jingan Optoelectronics | 2025-10-13 | Lateral(C-D) noise | 0.11258
| Fujian Jingan Optoelectronics | 2025-10-13 | RMS of ZeroScan | 0.029
| Fujian Jingan Optoelectronics | 2025-10-13 | LightStrengthGain | 0.55
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of X Detector during X Scan sweep | 183.4093
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of X1 Detector during X Scan sweep | 183.4093
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of X2 Detector during X Scan sweep | 183.3394
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Y Detector during Y Scan sweep | 174.3323
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Y1 Detector during Y Scan sweep | 179.9715
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Y2 Detector during Y Scan sweep | 174.3323
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of X1 Detector Raw during X Scan sweep | 14.73
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of X2 Detector Raw during X Scan sweep | 14.29
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.51
| Fujian Jingan Optoelectronics | 2025-10-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.62
| Fujian Jingan Optoelectronics | 2025-10-13 | X Detector stroke calibration error (nominal scale error X) | 0.053
| Fujian Jingan Optoelectronics | 2025-10-13 | Y Detector stroke calibration error (nominal scale error Y) | 0.037
| Fujian Jingan Optoelectronics | 2025-10-13 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 233.296
| Fujian Jingan Optoelectronics | 2025-10-13 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 237.287
| Fujian Jingan Optoelectronics | 2025-10-13 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 243.718
| Fujian Jingan Optoelectronics | 2025-10-13 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 239.527
| Fujian Jingan Optoelectronics | 2025-10-13 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 158.832
| Fujian Jingan Optoelectronics | 2025-10-13 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 156.179
| Fujian Jingan Optoelectronics | 2025-10-13 | XY Orthogonality | 90.015
| Fujian Jingan Optoelectronics | 2025-10-13 | X axis OPM | 0.167
| Fujian Jingan Optoelectronics | 2025-10-13 | Y axis OPM | 0.497
| Fujian Jingan Optoelectronics | 2025-10-13 | X Ringing(Rpv) | 0.042
| Fujian Jingan Optoelectronics | 2025-10-13 | Y Ringing(Rpv) | 0.053
| Fujian Jingan Optoelectronics | 2025-10-13 | X Detector noise (XY Servo on) | 0.299
| Fujian Jingan Optoelectronics | 2025-10-13 | Y Detector noise (XY Servo on) | 0.297
| Fujian Jingan Optoelectronics | 2025-10-13 | X1 Detector noise (XY Servo on) | 0.407
| Fujian Jingan Optoelectronics | 2025-10-13 | Y1 Detector noise (XY Servo on) | 0.42
| Fujian Jingan Optoelectronics | 2025-10-13 | X2 Detector noise (XY Servo on) | 0.422
| Fujian Jingan Optoelectronics | 2025-10-13 | Y2 Detector noise (XY Servo on) | 0.414
| Fujian Jingan Optoelectronics | 2025-10-13 | Z Drift | 12.7504
| Fujian Jingan Optoelectronics | 2025-10-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.10144
| Fujian Jingan Optoelectronics | 2025-10-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.019203
| Fujian Jingan Optoelectronics | 2025-10-13 | Short-term stability using Step Height Sample (3σ) | 0.1545
| Fujian Jingan Optoelectronics | 2025-10-13 | Roughness Repeatability 1x1um(RSD%) | 1.3095
| Fujian Jingan Optoelectronics | 2025-10-13 | Roughness Repeatability 1x1um(1σ) | 0.000471
| Fujian Jingan Optoelectronics | 2025-10-13 | Roughness Repeatability 80x80um(RSD%) | 0.63002
| Fujian Jingan Optoelectronics | 2025-10-13 | Roughness Repeatability 80x80um(1σ) | 0.000568
| Micron F6 | 2025-10-02 | Z Linearity by Z scanner position | 0.23
| Micron F6 | 2025-10-02 | Full range of Z Detector during Z Scan sweep | 12.98
| Micron F6 | 2025-10-02 | Z Height calibration error | 0.010841
| Micron F6 | 2025-10-02 | Z Drive calibration error | 0.30572
| Micron F6 | 2025-10-02 | Z Linearity by 4 different sample height | 1.0003
| Micron F6 | 2025-10-02 | Full range of Z Height during Z Scan sweep | 18.5098
| Micron F6 | 2025-10-02 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.3279
| Micron F6 | 2025-10-02 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4517
| Micron F6 | 2025-10-02 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.3617
| Micron F6 | 2025-10-02 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5087
| Micron F6 | 2025-10-02 | Z Height noise | 0.03
| Micron F6 | 2025-10-02 | Wafer placement(Loading) error on the load port1 (X offset) | 20.1
| Micron F6 | 2025-10-02 | Wafer placement(Loading) error on the load port1 (Y offset) | 37.06
| Micron F6 | 2025-10-02 | Wafer placement(Loading) error on the load port2 (X offset) | 68.86
| Micron F6 | 2025-10-02 | Wafer placement(Loading) error on the load port2 (Y offset) | 19.1
| Micron F6 | 2025-10-02 | OpticalAxisShiftPerFocusX | -0.0015
| Micron F6 | 2025-10-02 | OpticalAxisShiftPerFocusY | -0.0002
| Micron F6 | 2025-10-02 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| Micron F6 | 2025-10-02 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0
| Micron F6 | 2025-10-02 | Lift pins synchronization offset check (X offset) | 45.18
| Micron F6 | 2025-10-02 | Lift pins synchronization offset check (Y offset) | 16.2
| Micron F6 | 2025-10-02 | Profiler noise @200um | 0.118
| Micron F6 | 2025-10-02 | Roll motion to X+ full scan | 3.3
| Micron F6 | 2025-10-02 | Position repeatability(10repeats, Position Based) | 0.52
| Micron F6 | 2025-10-02 | OPM Repeatability @25mm | 1.695
| Micron F6 | 2025-10-02 | OPM Repeatability @10mm | 0.99
| Micron F6 | 2025-10-02 | OPM Repeatability @5mm | 0.842
| Micron F6 | 2025-10-02 | OPM Repeatability @1mm | 0.559
| Micron F6 | 2025-10-02 | Max OPM @25mm | 54.915
| Micron F6 | 2025-10-02 | Max OPM @10mm | 19.089
| Micron F6 | 2025-10-02 | Max OPM @5mm | 8.036
| Micron F6 | 2025-10-02 | Max OPM @1mm | 3.777
| Micron F6 | 2025-10-02 | Max OPM @25mm | 0.588
| Micron F6 | 2025-10-02 | Max OPM @10mm | 0.7
| Micron F6 | 2025-10-02 | Max OPM @5mm | 50
| Micron F6 | 2025-10-02 | Max OPM @1mm | 25.5
| Micron F6 | 2025-10-02 | Torque Profile (10mm, 200um/s) | -25.5
| Micron F6 | 2025-10-02 | Torque Profile (10mm, 20um/s) | 55
| Micron F6 | 2025-10-02 | Torque Profile (50mm, 200um/s) | 0.588
| Micron F6 | 2025-10-02 | Torque Profile (50mm, 20um/s) | 0.7
| Micron F6 | 2025-10-02 | Lateral(C-D) sensitivity | 15.6
| Micron F6 | 2025-10-02 | Lateral(C-D) noise | 0.212
| Micron F6 | 2025-10-02 | Min. Vibration Transmissbility (~100Hz) | 0.041
| Micron F6 | 2025-10-02 | RMS of ZeroScan | 0.069
| Micron F6 | 2025-10-02 | LightStrengthGain | 0.46
| Micron F6 | 2025-10-02 | Full range of X Detector during X Scan sweep | 197.192
| Micron F6 | 2025-10-02 | Full range of X1 Detector during X Scan sweep | 191.5401
| Micron F6 | 2025-10-02 | Full range of X2 Detector during X Scan sweep | 202.844
| Micron F6 | 2025-10-02 | Full range of Y Detector during Y Scan sweep | 203.1054
| Micron F6 | 2025-10-02 | Full range of Y1 Detector during Y Scan sweep | 196.7035
| Micron F6 | 2025-10-02 | Full range of Y2 Detector during Y Scan sweep | 202.7692
| Micron F6 | 2025-10-02 | Full range of X1 Detector Raw during X Scan sweep | 14.32
| Micron F6 | 2025-10-02 | Full range of X2 Detector Raw during X Scan sweep | 14.61
| Micron F6 | 2025-10-02 | Full range of Y1 Detector Raw during Y Scan sweep | 14.56
| Micron F6 | 2025-10-02 | Full range of Y2 Detector Raw during Y Scan sweep | 14.51
| Micron F6 | 2025-10-02 | X Detector stroke calibration error (nominal scale error X) | 0.015
| Micron F6 | 2025-10-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.447
| Micron F6 | 2025-10-02 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 297.6419
| Micron F6 | 2025-10-02 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 299.4266
| Micron F6 | 2025-10-02 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 297.458
| Micron F6 | 2025-10-02 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 297.0163
| Micron F6 | 2025-10-02 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.542
| Micron F6 | 2025-10-02 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 183.482
| Micron F6 | 2025-10-02 | X Non-linearity error | 3.37
| Micron F6 | 2025-10-02 | Y Non-linearity error | 3.05
| Micron F6 | 2025-10-02 | X axis OPM | 0.204
| Micron F6 | 2025-10-02 | Y axis OPM | 0.39
| Micron F6 | 2025-10-02 | X Ringing(Rpv) | 0.067
| Micron F6 | 2025-10-02 | Y Ringing(Rpv) | 0.063
| Micron F6 | 2025-10-02 | X Detector noise (XY Servo on) | 0.416
| Micron F6 | 2025-10-02 | Y Detector noise (XY Servo on) | 0.513
| Micron F6 | 2025-10-02 | X1 Detector noise (XY Servo on) | 0.552
| Micron F6 | 2025-10-02 | Y1 Detector noise (XY Servo on) | 0.613
| Micron F6 | 2025-10-02 | X2 Detector noise (XY Servo on) | 0.558
| Micron F6 | 2025-10-02 | Y2 Detector noise (XY Servo on) | 0.627
| Micron F6 | 2025-10-02 | X Detector noise (XY Servo on) | 160
| Micron F6 | 2025-10-02 | Y Detector noise (XY Servo on) | 50
| Micron F6 | 2025-10-02 | X1 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | Y1 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | X2 Detector noise (XY Servo on) | 2.36
| Micron F6 | 2025-10-02 | Y2 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | X Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | Y Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | X1 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | Y1 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | X2 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | Y2 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | X Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | Y Detector noise (XY Servo on) | Enabled
| Micron F6 | 2025-10-02 | X1 Detector noise (XY Servo on) | Yes
| Micron F6 | 2025-10-02 | Y1 Detector noise (XY Servo on) | V2BCH
| Micron F6 | 2025-10-02 | X2 Detector noise (XY Servo on) | HCMNEW
| Micron F6 | 2025-10-02 | Y2 Detector noise (XY Servo on) | Blue
| Micron F6 | 2025-10-02 | RMS error after Stage Calibration | 2.36
| Micron F6 | 2025-10-02 | Image offset after moving from -X to +X (X scan) | 12.4512
| Micron F6 | 2025-10-02 | Image offset after moving from -Y to +Y (Y scan) | 68.1152
| Micron F6 | 2025-10-02 | Image offset after moving from -X to +X (Y scan) | 32.2266
| Micron F6 | 2025-10-02 | Image offset after moving from -Y to +Y (X scan) | 16.1133
| Micron F6 | 2025-10-02 | Image offset after moving from +X to -X (X scan) | 17.5781
| Micron F6 | 2025-10-02 | Image offset after moving from +Y to -Y (Y scan) | 66.6504
| Micron F6 | 2025-10-02 | Image offset after moving from +X to -X (Y scan) | 42.4805
| Micron F6 | 2025-10-02 | Image offset after moving from +Y to -Y (X scan) | 6.5918
| Micron F6 | 2025-10-02 | XY Stage positioning X offset (Vision Pattern Matching) | 0.86914
| Micron F6 | 2025-10-02 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.3486
| Micron F6 | 2025-10-02 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.17611
| Micron F6 | 2025-10-02 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.39969
| Micron F6 | 2025-10-02 | XY Stage positioning X offset (In-Die Align) | 1.2988
| Micron F6 | 2025-10-02 | XY Stage positioning Y offset (In-Die Align) | 2.0068
| Micron F6 | 2025-10-02 | XY Stage positioning X 1σ (In-Die Align) | 0.2716
| Micron F6 | 2025-10-02 | XY Stage positioning Y 1σ (In-Die Align) | 0.47713
| Micron F6 | 2025-10-02 | XY Stage positioning X offset (LLC Translation) | 1.8799
| Micron F6 | 2025-10-02 | XY Stage positioning Y offset (LLC Translation) | 2.4609
| Micron F6 | 2025-10-02 | XY Stage positioning X 1σ (LLC Translation) | 0.35688
| Micron F6 | 2025-10-02 | XY Stage positioning Y 1σ (LLC Translation) | 0.51836
| Micron F6 | 2025-10-02 | XY Stage positioning X offset (Global Align) | 4.0137
| Micron F6 | 2025-10-02 | XY Stage positioning Y offset (Global Align) | 6.2402
| Micron F6 | 2025-10-02 | XY Stage positioning X 1σ (Global Align) | 0.90669
| Micron F6 | 2025-10-02 | XY Stage positioning Y 1σ (Global Align) | 1.9735
| Micron F6 | 2025-10-02 | Z Drift | 7.5238
| Micron F6 | 2025-10-02 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27186
| Micron F6 | 2025-10-02 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.049301
| Micron F6 | 2025-10-02 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.53993
| Micron F6 | 2025-10-02 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.098025
| Micron F6 | 2025-10-02 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25074
| Micron F6 | 2025-10-02 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.045474
| Micron F6 | 2025-10-02 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.77971
| Micron F6 | 2025-10-02 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14105
| Micron F6 | 2025-10-02 | Cycle time for 15-point measurement | 1639.904
| Micron F6 | 2025-10-02 | Short-term stability using Step Height Sample (3σ) | 0.39
| Micron F6 | 2025-10-02 | Roughness Repeatability 1x1um(RSD%) | 1.057
| Micron F6 | 2025-10-02 | Roughness Repeatability 1x1um(1σ) | 0.001
| Micron F6 | 2025-10-02 | Roughness Repeatability 80x80um(RSD%) | 1.035
| Micron F6 | 2025-10-02 | Roughness Repeatability 80x80um(1σ) | 0.001
| 포항공과대학교_화학공학과 | 2025-10-01 | Lateral(C-D) sensitivity | 50.9
| 포항공과대학교_화학공학과 | 2025-10-01 | Lateral(C-D) noise | 0.133
| 포항공과대학교_화학공학과 | 2025-10-01 | A-B noise | 6.43
| 포항공과대학교_화학공학과 | 2025-10-01 | NCM Amplitude noise | 10.303
| 포항공과대학교_화학공학과 | 2025-10-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.53
| 포항공과대학교_화학공학과 | 2025-10-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| 포항공과대학교_화학공학과 | 2025-10-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.48
| 포항공과대학교_화학공학과 | 2025-10-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.43
| 포항공과대학교_화학공학과 | 2025-10-01 | Roughness Repeatability 1x1um(RSD%) | 1.581
| 포항공과대학교_화학공학과 | 2025-10-01 | Roughness Repeatability 1x1um(1σ) | 0.004
| 포항공과대학교_화학공학과 | 2025-10-01 | Full range of Z Height during Z Scan sweep | 18.2
| 포항공과대학교_화학공학과 | 2025-10-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.279
| 포항공과대학교_화학공학과 | 2025-10-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.982
| 포항공과대학교_화학공학과 | 2025-10-01 | Full range of Z Detector during Z Scan sweep | 12.4
| 포항공과대학교_화학공학과 | 2025-10-01 | Z Drive calibration error | 0.11707
| 포항공과대학교_화학공학과 | 2025-10-01 | Z Height calibration error | 0.22146
| 포항공과대학교_화학공학과 | 2025-10-01 | Z linearity by sample height | 1.018
| 포항공과대학교_화학공학과 | 2025-10-01 | 1) Z Scanner position at +3μm | 0.004771
| 포항공과대학교_화학공학과 | 2025-10-01 | 3) Z Scanner position at -3μm | 0.001503
| 포항공과대학교_화학공학과 | 2025-10-01 | Z Height noise | 0.03
| 포항공과대학교_화학공학과 | 2025-10-01 | Full range of X Detector during X Scan sweep | 191.8
| 포항공과대학교_화학공학과 | 2025-10-01 | Full range of Y Detector during Y Scan sweep | 186.9
| 포항공과대학교_화학공학과 | 2025-10-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 270.2
| 포항공과대학교_화학공학과 | 2025-10-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 266.6
| 포항공과대학교_화학공학과 | 2025-10-01 | X Detector stroke calibration error (nominal scale error X) | 0.039
| 포항공과대학교_화학공학과 | 2025-10-01 | Y Detector stroke calibration error (nominal scale error Y) | 0.05
| 포항공과대학교_화학공학과 | 2025-10-01 | X Scan Stroke | 191.8
| 포항공과대학교_화학공학과 | 2025-10-01 | Y Scan Stroke | 186.9
| 포항공과대학교_화학공학과 | 2025-10-01 | Full range of X Detector Raw during X Scan sweep | 14.2
| 포항공과대학교_화학공학과 | 2025-10-01 | Full range of Y Detector Raw during Y Scan sweep | 14
| 포항공과대학교_화학공학과 | 2025-10-01 | X Detector noise (XY Servo on) | 0.315
| 포항공과대학교_화학공학과 | 2025-10-01 | Y Detector noise (XY Servo on) | 0.383
| 포항공과대학교_화학공학과 | 2025-10-01 | X Detector noise (XY servo off) | 0.312
| 포항공과대학교_화학공학과 | 2025-10-01 | Y Detector noise (XY servo off) | 0.382
| 포항공과대학교_화학공학과 | 2025-10-01 | X axis OPM | 0.886
| 포항공과대학교_화학공학과 | 2025-10-01 | Y axis OPM | 1.151
| 포항공과대학교_화학공학과 | 2025-10-01 | X Non-linearity | 0.042
| 포항공과대학교_화학공학과 | 2025-10-01 | Y Non-linearity | 0.152
| Dalian Institute of Chemical Physics | 2025-10-01 | Lateral(C-D) sensitivity | 12
| Dalian Institute of Chemical Physics | 2025-10-01 | Lateral(C-D) noise | 0.119
| Dalian Institute of Chemical Physics | 2025-10-01 | A-B noise | 13.498
| Dalian Institute of Chemical Physics | 2025-10-01 | NCM Amplitude noise | 11.479
| Dalian Institute of Chemical Physics | 2025-10-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.03
| Dalian Institute of Chemical Physics | 2025-10-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.022
| Dalian Institute of Chemical Physics | 2025-10-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.027
| Dalian Institute of Chemical Physics | 2025-10-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.018
| Dalian Institute of Chemical Physics | 2025-10-01 | Roughness Repeatability 1x1um(RSD%) | 0.776
| Dalian Institute of Chemical Physics | 2025-10-01 | Roughness Repeatability 1x1um(1σ) | 0.004
| CNR-IFN | 2025-09-29 | Lateral(C-D) sensitivity | 10.2
| CNR-IFN | 2025-09-29 | Lateral(C-D) noise | 0.106
| CNR-IFN | 2025-09-29 | A-B noise | 5.556
| CNR-IFN | 2025-09-29 | NCM Amplitude noise | 9.994
| CNR-IFN | 2025-09-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| CNR-IFN | 2025-09-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| CNR-IFN | 2025-09-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| CNR-IFN | 2025-09-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| CNR-IFN | 2025-09-29 | Roughness Repeatability 1x1um(RSD%) | 1.179
| CNR-IFN | 2025-09-29 | Roughness Repeatability 1x1um(1σ) | 0.004
| CNR-IFN | 2025-09-29 | Full range of Z Height during Z Scan sweep | 18.3
| CNR-IFN | 2025-09-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.844
| CNR-IFN | 2025-09-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.242
| CNR-IFN | 2025-09-29 | Full range of Z Detector during Z Scan sweep | 13.1
| CNR-IFN | 2025-09-29 | Z Drive calibration error | 0.42912
| CNR-IFN | 2025-09-29 | Z Height calibration error | 0.085249
| CNR-IFN | 2025-09-29 | Z linearity by sample height | 1.0048
| CNR-IFN | 2025-09-29 | 1) Z Scanner position at +3μm | 0.007208
| CNR-IFN | 2025-09-29 | 3) Z Scanner position at -3μm | 0.005301
| CNR-IFN | 2025-09-29 | Z Height noise | 0.027
| CNR-IFN | 2025-09-29 | Full range of X Detector during X Scan sweep | 108.9
| CNR-IFN | 2025-09-29 | Full range of Y Detector during Y Scan sweep | 110
| CNR-IFN | 2025-09-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 151.7
| CNR-IFN | 2025-09-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 152.5
| CNR-IFN | 2025-09-29 | X Detector stroke calibration error (nominal scale error X) | 0.026
| CNR-IFN | 2025-09-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.451
| CNR-IFN | 2025-09-29 | X Scan Stroke | 108.9
| CNR-IFN | 2025-09-29 | Y Scan Stroke | 110
| CNR-IFN | 2025-09-29 | Full range of X Detector Raw during X Scan sweep | 14.4
| CNR-IFN | 2025-09-29 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| CNR-IFN | 2025-09-29 | X Detector noise (XY Servo on) | 0.259
| CNR-IFN | 2025-09-29 | Y Detector noise (XY Servo on) | 0.261
| CNR-IFN | 2025-09-29 | X Detector noise (XY servo off) | 0.255
| CNR-IFN | 2025-09-29 | Y Detector noise (XY servo off) | 0.256
| CNR-IFN | 2025-09-29 | X axis OPM | 1.587
| CNR-IFN | 2025-09-29 | Y axis OPM | 1.791
| CNR-IFN | 2025-09-29 | X Non-linearity | 0.169
| CNR-IFN | 2025-09-29 | Y Non-linearity | 0.103
| CNR-IFN | 2025-09-29 | Full range of X Detector during X Scan sweep | 20
| CNR-IFN | 2025-09-29 | Full range of Y Detector during Y Scan sweep | 20.1
| CNR-IFN | 2025-09-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 30.5
| CNR-IFN | 2025-09-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 29.9
| CNR-IFN | 2025-09-29 | X Detector stroke calibration error (nominal scale error X) | 0.075
| CNR-IFN | 2025-09-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.477
| CNR-IFN | 2025-09-29 | X Scan Stroke | 20
| CNR-IFN | 2025-09-29 | Y Scan Stroke | 20.1
| CNR-IFN | 2025-09-29 | Full range of X Detector Raw during X Scan sweep | 13.1
| CNR-IFN | 2025-09-29 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| CNR-IFN | 2025-09-29 | X Detector noise (XY Servo on) | 0.022
| CNR-IFN | 2025-09-29 | Y Detector noise (XY Servo on) | 0.023
| CNR-IFN | 2025-09-29 | X Detector noise (XY servo off) | 0.024
| CNR-IFN | 2025-09-29 | Y Detector noise (XY servo off) | 0.025
| CNR-IFN | 2025-09-29 | X axis OPM | 0.438
| CNR-IFN | 2025-09-29 | Y axis OPM | 0.881
| CNR-IFN | 2025-09-29 | X Non-linearity | 0.026
| CNR-IFN | 2025-09-29 | Y Non-linearity | 0.038
| Intel Hillsboro #5 | 2025-09-26 | Z Linearity by Z scanner position | 0.6
| Intel Hillsboro #5 | 2025-09-26 | Z Linearity by 4 different sample height | 0.99547
| Intel Hillsboro #5 | 2025-09-26 | Full range of Z Detector during Z Scan sweep | 14.1
| Intel Hillsboro #5 | 2025-09-26 | Z Height calibration error | 0.039502
| Intel Hillsboro #5 | 2025-09-26 | Z Drive calibration error | 0.025753
| Intel Hillsboro #5 | 2025-09-26 | Full range of Z Height during Z Scan sweep | 41
| Intel Hillsboro #5 | 2025-09-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 58.7
| Intel Hillsboro #5 | 2025-09-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 26.7
| Intel Hillsboro #5 | 2025-09-26 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 66.6
| Intel Hillsboro #5 | 2025-09-26 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 26.9
| Intel Hillsboro #5 | 2025-09-26 | Z Height noise | 0.06
| Intel Hillsboro #5 | 2025-09-26 | Full range of Z Height during Z Scan sweep | 0.251
| Intel Hillsboro #5 | 2025-09-26 | Wafer placement(Loading) error on the load port1 (X offset) | 21.7
| Intel Hillsboro #5 | 2025-09-26 | Wafer placement(Loading) error on the load port1 (Y offset) | 24.1
| Intel Hillsboro #5 | 2025-09-26 | Wafer placement(Loading) error on the load port2 (X offset) | 50.1
| Intel Hillsboro #5 | 2025-09-26 | Wafer placement(Loading) error on the load port2 (Y offset) | 6.6
| Intel Hillsboro #5 | 2025-09-26 | OpticalAxisShiftPerFocusX | -0.007
| Intel Hillsboro #5 | 2025-09-26 | OpticalAxisShiftPerFocusY | 0.0021
| Intel Hillsboro #5 | 2025-09-26 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 9.14e-19
| Intel Hillsboro #5 | 2025-09-26 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4.83e-05
| Intel Hillsboro #5 | 2025-09-26 | Lift pins synchronization offset check (X offset) | 7.1
| Intel Hillsboro #5 | 2025-09-26 | Lift pins synchronization offset check (Y offset) | 8.3
| Intel Hillsboro #5 | 2025-09-26 | Max OPM @25mm | 97.896
| Intel Hillsboro #5 | 2025-09-26 | Max OPM @10mm | 37.031
| Intel Hillsboro #5 | 2025-09-26 | Max OPM @5mm | 22.265
| Intel Hillsboro #5 | 2025-09-26 | Max OPM @1mm | 7.138
| Intel Hillsboro #5 | 2025-09-26 | OPM Repeatability @25mm | 4.97
| Intel Hillsboro #5 | 2025-09-26 | OPM Repeatability @10mm | 1.733
| Intel Hillsboro #5 | 2025-09-26 | OPM Repeatability @5mm | 1.073
| Intel Hillsboro #5 | 2025-09-26 | OPM Repeatability @1mm | 0.7
| Intel Hillsboro #5 | 2025-09-26 | Torque Profile (50mm, 200um/s) | 0.33
| Intel Hillsboro #5 | 2025-09-26 | Torque Profile (50mm, 20um/s) | 0.745
| Intel Hillsboro #5 | 2025-09-26 | Profiler noise @200um | 0.163
| Intel Hillsboro #5 | 2025-09-26 | Roll motion to X+ full scan | 20
| Intel Hillsboro #5 | 2025-09-26 | Position repeatability(10repeats, Position Based) | 0.46
| Intel Hillsboro #5 | 2025-09-26 | Lateral(C-D) sensitivity | 22.7
| Intel Hillsboro #5 | 2025-09-26 | Lateral(C-D) noise | 0.134
| Intel Hillsboro #5 | 2025-09-26 | Min. Vibration Transmissbility (~100Hz) | 0.05
| Intel Hillsboro #5 | 2025-09-26 | RMS of ZeroScan | 0.062
| Intel Hillsboro #5 | 2025-09-26 | LightStrengthGain | 0.15
| Intel Hillsboro #5 | 2025-09-26 | Full range of X Detector during X Scan sweep | 224.1
| Intel Hillsboro #5 | 2025-09-26 | Full range of X1 Detector during X Scan sweep | 221.9
| Intel Hillsboro #5 | 2025-09-26 | Full range of X2 Detector during X Scan sweep | 226.3
| Intel Hillsboro #5 | 2025-09-26 | Full range of Y Detector during Y Scan sweep | 229.4
| Intel Hillsboro #5 | 2025-09-26 | Full range of Y1 Detector during Y Scan sweep | 232.8
| Intel Hillsboro #5 | 2025-09-26 | Full range of Y2 Detector during Y Scan sweep | 226
| Intel Hillsboro #5 | 2025-09-26 | Full range of X1 Detector Raw during X Scan sweep | 14.68
| Intel Hillsboro #5 | 2025-09-26 | Full range of X2 Detector Raw during X Scan sweep | 15.04
| Intel Hillsboro #5 | 2025-09-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.61
| Intel Hillsboro #5 | 2025-09-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| Intel Hillsboro #5 | 2025-09-26 | X Detector stroke calibration error (nominal scale error X) | 0.47
| Intel Hillsboro #5 | 2025-09-26 | Y Detector stroke calibration error (nominal scale error Y) | 0.2
| Intel Hillsboro #5 | 2025-09-26 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 272.95
| Intel Hillsboro #5 | 2025-09-26 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 267.0826
| Intel Hillsboro #5 | 2025-09-26 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 283.7472
| Intel Hillsboro #5 | 2025-09-26 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 289.4104
| Intel Hillsboro #5 | 2025-09-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 176.463
| Intel Hillsboro #5 | 2025-09-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 181.724
| Intel Hillsboro #5 | 2025-09-26 | XY Orthogonality | 90.13
| Intel Hillsboro #5 | 2025-09-26 | X Non-linearity error | 3.39
| Intel Hillsboro #5 | 2025-09-26 | Y Non-linearity error | 2.65
| Intel Hillsboro #5 | 2025-09-26 | X axis OPM | 0.257
| Intel Hillsboro #5 | 2025-09-26 | Y axis OPM | 1.664
| Intel Hillsboro #5 | 2025-09-26 | X Ringing(Rpv) | 0.136
| Intel Hillsboro #5 | 2025-09-26 | Y Ringing(Rpv) | 0.116
| Intel Hillsboro #5 | 2025-09-26 | X Detector noise (XY Servo on) | 0.341
| Intel Hillsboro #5 | 2025-09-26 | Y Detector noise (XY Servo on) | 0.358
| Intel Hillsboro #5 | 2025-09-26 | X1 Detector noise (XY Servo on) | 0.466
| Intel Hillsboro #5 | 2025-09-26 | Y1 Detector noise (XY Servo on) | 0.509
| Intel Hillsboro #5 | 2025-09-26 | X2 Detector noise (XY Servo on) | 0.469
| Intel Hillsboro #5 | 2025-09-26 | Y2 Detector noise (XY Servo on) | 0.467
| Intel Hillsboro #5 | 2025-09-26 | RMS error after Stage Calibration | 2.16
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from -X to +X (X scan) | 95.9473
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from -Y to +Y (Y scan) | 146.484
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from -X to +X (Y scan) | 45.4102
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from -Y to +Y (X scan) | 65.1855
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from +X to -X (X scan) | 78.3691
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from +Y to -Y (Y scan) | 76.1719
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from +X to -X (Y scan) | 12.4512
| Intel Hillsboro #5 | 2025-09-26 | Image offset after moving from +Y to -Y (X scan) | 43.9453
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X offset (Vision Pattern Matching) | 0.391
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.5244
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.083
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.464
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X offset (In-Die Align) | 1.016
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y offset (In-Die Align) | 3.174
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X 1σ (In-Die Align) | 0.195
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y 1σ (In-Die Align) | 0.653
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X offset (LLC Translation) | 1.226
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y offset (LLC Translation) | 2.363
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X 1σ (LLC Translation) | 0.23
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y 1σ (LLC Translation) | 0.524
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X offset (Global Align) | 2.822
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y offset (Global Align) | 6.421
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning X 1σ (Global Align) | 0.6
| Intel Hillsboro #5 | 2025-09-26 | XY Stage positioning Y 1σ (Global Align) | 1.988
| Intel Hillsboro #5 | 2025-09-26 | Z Drift | 18.5
| Intel Hillsboro #5 | 2025-09-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.21
| Intel Hillsboro #5 | 2025-09-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Intel Hillsboro #5 | 2025-09-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.38
| Intel Hillsboro #5 | 2025-09-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Intel Hillsboro #5 | 2025-09-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.5
| Intel Hillsboro #5 | 2025-09-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Intel Hillsboro #5 | 2025-09-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.39
| Intel Hillsboro #5 | 2025-09-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Intel Hillsboro #5 | 2025-09-26 | Cycle time for 15-point measurement | 1338.529
| Intel Hillsboro #5 | 2025-09-26 | Short-term stability using Step Height Sample (3σ) | 0.44
| Intel Hillsboro #5 | 2025-09-26 | Roughness Repeatability 1x1um(RSD%) | 2.1
| Intel Hillsboro #5 | 2025-09-26 | Roughness Repeatability 1x1um(1σ) | 0.002
| Intel Hillsboro #5 | 2025-09-26 | Roughness Repeatability 80x80um(RSD%) | 1.911
| Intel Hillsboro #5 | 2025-09-26 | Roughness Repeatability 80x80um(1σ) | 0.01
| Intel Hillsboro #5 | 2025-09-26 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer) | 1.87
| Intel Hillsboro #5 | 2025-09-26 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer) | 0.34
| Intel Hillsboro #5 | 2025-09-26 | Z Detector Max | 0.4
| Intel Hillsboro #5 | 2025-09-26 | Z Detector Min | -0.32
| Intel Hillsboro #5 | 2025-09-26 | WLI_Scale(Z Calibration) | 34.4
| PSI EAST OFFICE | 2025-09-25 | SLD current limit check | 112
| PSI EAST OFFICE | 2025-09-25 | SLD beam power at the end of the lens | 0.28
| PSI EAST OFFICE | 2025-09-25 | Lateral(C-D) sensitivity | 37
| PSI EAST OFFICE | 2025-09-25 | Lateral(C-D) noise | 0.067
| PSI EAST OFFICE | 2025-09-25 | A-B noise | 9.922
| PSI EAST OFFICE | 2025-09-25 | NCM Amplitude noise | 6.518
| PSI EAST OFFICE | 2025-09-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| PSI EAST OFFICE | 2025-09-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| PSI EAST OFFICE | 2025-09-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| PSI EAST OFFICE | 2025-09-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| PSI EAST OFFICE | 2025-09-25 | Roughness Repeatability 1x1um(RSD%) | 0.827
| PSI EAST OFFICE | 2025-09-25 | Roughness Repeatability 1x1um(1σ) | 0.002
| LMU Munich Nanophotonics Group | 2025-09-24 | Lateral(C-D) sensitivity | 13
| LMU Munich Nanophotonics Group | 2025-09-24 | Lateral(C-D) noise | 0.077
| LMU Munich Nanophotonics Group | 2025-09-24 | A-B noise | 10.13
| LMU Munich Nanophotonics Group | 2025-09-24 | NCM Amplitude noise | 13.95
| LMU Munich Nanophotonics Group | 2025-09-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| LMU Munich Nanophotonics Group | 2025-09-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| LMU Munich Nanophotonics Group | 2025-09-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| LMU Munich Nanophotonics Group | 2025-09-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| LMU Munich Nanophotonics Group | 2025-09-24 | Roughness Repeatability 1x1um(RSD%) | 0.928
| LMU Munich Nanophotonics Group | 2025-09-24 | Roughness Repeatability 1x1um(1σ) | 0.002
| LMU Munich Nanophotonics Group | 2025-09-24 | Full range of Z Height during Z Scan sweep | 16.7
| LMU Munich Nanophotonics Group | 2025-09-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.125
| LMU Munich Nanophotonics Group | 2025-09-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.305
| LMU Munich Nanophotonics Group | 2025-09-24 | Full range of Z Detector during Z Scan sweep | 12.7
| LMU Munich Nanophotonics Group | 2025-09-24 | Z Drive calibration error | 0.29215
| LMU Munich Nanophotonics Group | 2025-09-24 | Z Height calibration error | 0.25862
| LMU Munich Nanophotonics Group | 2025-09-24 | Z linearity by sample height | 0.99154
| LMU Munich Nanophotonics Group | 2025-09-24 | 1) Z Scanner position at +3μm | 0.004163
| LMU Munich Nanophotonics Group | 2025-09-24 | 3) Z Scanner position at -3μm | 0.002339
| LMU Munich Nanophotonics Group | 2025-09-24 | Z Height noise | 0.031
| LMU Munich Nanophotonics Group | 2025-09-24 | Full range of X Detector during X Scan sweep | 74.7
| LMU Munich Nanophotonics Group | 2025-09-24 | Full range of Y Detector during Y Scan sweep | 75.3
| LMU Munich Nanophotonics Group | 2025-09-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 109.4
| LMU Munich Nanophotonics Group | 2025-09-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 110.4
| LMU Munich Nanophotonics Group | 2025-09-24 | X Detector stroke calibration error (nominal scale error X) | 0.12
| LMU Munich Nanophotonics Group | 2025-09-24 | Y Detector stroke calibration error (nominal scale error Y) | 0.4
| LMU Munich Nanophotonics Group | 2025-09-24 | X Scan Stroke | 55
| LMU Munich Nanophotonics Group | 2025-09-24 | Y Scan Stroke | 55
| LMU Munich Nanophotonics Group | 2025-09-24 | Full range of X Detector Raw during X Scan sweep | 13.6
| LMU Munich Nanophotonics Group | 2025-09-24 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| LMU Munich Nanophotonics Group | 2025-09-24 | X Detector noise (XY Servo on) | 0.25
| LMU Munich Nanophotonics Group | 2025-09-24 | Y Detector noise (XY Servo on) | 0.25
| LMU Munich Nanophotonics Group | 2025-09-24 | X Detector noise (XY Servo off) | 0.244
| LMU Munich Nanophotonics Group | 2025-09-24 | Y Detector noise (XY Servo off) | 0.245
| LMU Munich Nanophotonics Group | 2025-09-24 | X axis OPM | 0.9974
| LMU Munich Nanophotonics Group | 2025-09-24 | Y axis OPM | 1.23
| LMU Munich Nanophotonics Group | 2025-09-24 | X Non-linearity | 0.063
| LMU Munich Nanophotonics Group | 2025-09-24 | Y Non-linearity | 0.088
| University of Electronic Science and Technology of China | 2025-09-23 | A-B noise | 6.745
| University of Electronic Science and Technology of China | 2025-09-23 | NCM Amplitude noise | 13.121
| University of Electronic Science and Technology of China | 2025-09-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| University of Electronic Science and Technology of China | 2025-09-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| University of Electronic Science and Technology of China | 2025-09-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| University of Electronic Science and Technology of China | 2025-09-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| University of Electronic Science and Technology of China | 2025-09-23 | Full range of Z Height during Z Scan sweep | 17.7
| University of Electronic Science and Technology of China | 2025-09-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.749
| University of Electronic Science and Technology of China | 2025-09-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.998
| University of Electronic Science and Technology of China | 2025-09-23 | Full range of Z Detector during Z Scan sweep | 12.8
| University of Electronic Science and Technology of China | 2025-09-23 | Z Drive calibration error | 0.10345
| University of Electronic Science and Technology of China | 2025-09-23 | Z Height calibration error | 0.019157
| University of Electronic Science and Technology of China | 2025-09-23 | Z linearity by sample height | 0.99794
| University of Electronic Science and Technology of China | 2025-09-23 | 1) Z Scanner position at +3μm | 0.010165
| University of Electronic Science and Technology of China | 2025-09-23 | 3) Z Scanner position at -3μm | 0.003446
| University of Electronic Science and Technology of China | 2025-09-23 | Z Height noise | 0.019
| University of Electronic Science and Technology of China | 2025-09-23 | Full range of X Detector during X Scan sweep | 61.8
| University of Electronic Science and Technology of China | 2025-09-23 | Full range of Y Detector during Y Scan sweep | 68.2
| University of Electronic Science and Technology of China | 2025-09-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 90.4
| University of Electronic Science and Technology of China | 2025-09-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 91.4
| University of Electronic Science and Technology of China | 2025-09-23 | X Detector stroke calibration error (nominal scale error X) | 0.023
| University of Electronic Science and Technology of China | 2025-09-23 | Y Detector stroke calibration error (nominal scale error Y) | 0.376
| University of Electronic Science and Technology of China | 2025-09-23 | X Scan Stroke | 55
| University of Electronic Science and Technology of China | 2025-09-23 | Y Scan Stroke | 55
| University of Electronic Science and Technology of China | 2025-09-23 | Full range of X Detector Raw during X Scan sweep | 13.7
| University of Electronic Science and Technology of China | 2025-09-23 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| University of Electronic Science and Technology of China | 2025-09-23 | X Detector noise (XY Servo on) | 0.405
| University of Electronic Science and Technology of China | 2025-09-23 | Y Detector noise (XY Servo on) | 0.404
| University of Electronic Science and Technology of China | 2025-09-23 | X Detector noise (XY Servo off) | 0.42
| University of Electronic Science and Technology of China | 2025-09-23 | Y Detector noise (XY Servo off) | 0.424
| University of Electronic Science and Technology of China | 2025-09-23 | X axis OPM | 0.89
| University of Electronic Science and Technology of China | 2025-09-23 | Y axis OPM | 0.939
| University of Electronic Science and Technology of China | 2025-09-23 | X Non-linearity | 0.039
| University of Electronic Science and Technology of China | 2025-09-23 | Y Non-linearity | 0.076
| Analog Devices GEN, Trias INC. | 2025-09-22 | SLD current limit check | 100
| Analog Devices GEN, Trias INC. | 2025-09-22 | SLD beam power at the end of the lens | 0.25
| Analog Devices GEN, Trias INC. | 2025-09-22 | Lateral(C-D) sensitivity | 8.8
| Analog Devices GEN, Trias INC. | 2025-09-22 | Lateral(C-D) noise | 0.094
| Analog Devices GEN, Trias INC. | 2025-09-22 | A-B noise | 10.773
| Analog Devices GEN, Trias INC. | 2025-09-22 | NCM Amplitude noise | 5.485
| Analog Devices GEN, Trias INC. | 2025-09-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Analog Devices GEN, Trias INC. | 2025-09-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Analog Devices GEN, Trias INC. | 2025-09-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Analog Devices GEN, Trias INC. | 2025-09-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Analog Devices GEN, Trias INC. | 2025-09-22 | Roughness Repeatability 1x1um(RSD%) | 0.632
| Analog Devices GEN, Trias INC. | 2025-09-22 | Roughness Repeatability 1x1um(1σ) | 0.001
| 한국에너지기술연구원 | 2025-09-19 | Lateral(C-D) sensitivity | 63.3
| 한국에너지기술연구원 | 2025-09-19 | Lateral(C-D) noise | 0.116
| 한국에너지기술연구원 | 2025-09-19 | A-B noise | 5.993
| 한국에너지기술연구원 | 2025-09-19 | NCM Amplitude noise | 8.652
| 한국에너지기술연구원 | 2025-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| 한국에너지기술연구원 | 2025-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 한국에너지기술연구원 | 2025-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한국에너지기술연구원 | 2025-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 한국에너지기술연구원 | 2025-09-19 | Roughness Repeatability 1x1um(RSD%) | 2.862
| 한국에너지기술연구원 | 2025-09-19 | Roughness Repeatability 1x1um(1σ) | 0.005
| 한국에너지기술연구원 | 2025-09-19 | Full range of Z Height during Z Scan sweep | 17.8
| 한국에너지기술연구원 | 2025-09-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.12
| 한국에너지기술연구원 | 2025-09-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.046
| 한국에너지기술연구원 | 2025-09-19 | Full range of Z Detector during Z Scan sweep | 12.7
| 한국에너지기술연구원 | 2025-09-19 | Z Drive calibration error | 0.088123
| 한국에너지기술연구원 | 2025-09-19 | Z Height calibration error | 0.29981
| 한국에너지기술연구원 | 2025-09-19 | Z linearity by sample height | 1.0014
| 한국에너지기술연구원 | 2025-09-19 | 1) Z Scanner position at +3μm | 0.007191
| 한국에너지기술연구원 | 2025-09-19 | 3) Z Scanner position at -3μm | 0.002872
| 한국에너지기술연구원 | 2025-09-19 | Z Height noise | 0.029
| 한국에너지기술연구원 | 2025-09-19 | Full range of X Detector during X Scan sweep | 118.8
| 한국에너지기술연구원 | 2025-09-19 | Full range of Y Detector during Y Scan sweep | 112.3
| 한국에너지기술연구원 | 2025-09-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167
| 한국에너지기술연구원 | 2025-09-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 172.1
| 한국에너지기술연구원 | 2025-09-19 | X Detector stroke calibration error (nominal scale error X) | 0.015
| 한국에너지기술연구원 | 2025-09-19 | Y Detector stroke calibration error (nominal scale error Y) | 0.109
| 한국에너지기술연구원 | 2025-09-19 | X Scan Stroke | 118.8
| 한국에너지기술연구원 | 2025-09-19 | Y Scan Stroke | 112.3
| 한국에너지기술연구원 | 2025-09-19 | Full range of X Detector Raw during X Scan sweep | 14.2
| 한국에너지기술연구원 | 2025-09-19 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| 한국에너지기술연구원 | 2025-09-19 | X Detector noise (XY Servo on) | 0.282
| 한국에너지기술연구원 | 2025-09-19 | Y Detector noise (XY Servo on) | 0.267
| 한국에너지기술연구원 | 2025-09-19 | X Detector noise (XY servo off) | 0.277
| 한국에너지기술연구원 | 2025-09-19 | Y Detector noise (XY servo off) | 0.265
| 한국에너지기술연구원 | 2025-09-19 | X axis OPM | 1.236
| 한국에너지기술연구원 | 2025-09-19 | Y axis OPM | 0.741
| 한국에너지기술연구원 | 2025-09-19 | X Non-linearity | 0.145
| 한국에너지기술연구원 | 2025-09-19 | Y Non-linearity | 0.219
| UNIST_신소재공학과 | 2025-09-18 | A-B noise | 8.438
| UNIST_신소재공학과 | 2025-09-18 | NCM Amplitude noise | 14.151
| UNIST_신소재공학과 | 2025-09-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| UNIST_신소재공학과 | 2025-09-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| UNIST_신소재공학과 | 2025-09-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| UNIST_신소재공학과 | 2025-09-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| UNIST_신소재공학과 | 2025-09-18 | Full range of Z Height during Z Scan sweep | 18.1
| UNIST_신소재공학과 | 2025-09-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.216
| UNIST_신소재공학과 | 2025-09-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.959
| UNIST_신소재공학과 | 2025-09-18 | Full range of Z Detector during Z Scan sweep | 12.8
| UNIST_신소재공학과 | 2025-09-18 | Z Drive calibration error | 0.049808
| UNIST_신소재공학과 | 2025-09-18 | Z Height calibration error | 0.22126
| UNIST_신소재공학과 | 2025-09-18 | Z linearity by sample height | 0.9979
| UNIST_신소재공학과 | 2025-09-18 | 1) Z Scanner position at +3μm | 0.003282
| UNIST_신소재공학과 | 2025-09-18 | 3) Z Scanner position at -3μm | 0.000775
| UNIST_신소재공학과 | 2025-09-18 | Z Height noise | 0.029
| UNIST_신소재공학과 | 2025-09-18 | Full range of X Detector during X Scan sweep | 66
| UNIST_신소재공학과 | 2025-09-18 | Full range of Y Detector during Y Scan sweep | 67.3
| UNIST_신소재공학과 | 2025-09-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 92.6
| UNIST_신소재공학과 | 2025-09-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 95.5
| UNIST_신소재공학과 | 2025-09-18 | X Detector stroke calibration error (nominal scale error X) | 0.108
| UNIST_신소재공학과 | 2025-09-18 | Y Detector stroke calibration error (nominal scale error Y) | 0.202
| UNIST_신소재공학과 | 2025-09-18 | X Scan Stroke | 55
| UNIST_신소재공학과 | 2025-09-18 | Y Scan Stroke | 55
| UNIST_신소재공학과 | 2025-09-18 | Full range of X Detector Raw during X Scan sweep | 14.3
| UNIST_신소재공학과 | 2025-09-18 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| UNIST_신소재공학과 | 2025-09-18 | X Detector noise (XY Servo on) | 0.34
| UNIST_신소재공학과 | 2025-09-18 | Y Detector noise (XY Servo on) | 0.348
| UNIST_신소재공학과 | 2025-09-18 | X Detector noise (XY Servo off) | 0.37
| UNIST_신소재공학과 | 2025-09-18 | Y Detector noise (XY Servo off) | 0.376
| UNIST_신소재공학과 | 2025-09-18 | X axis OPM | 0.982
| UNIST_신소재공학과 | 2025-09-18 | Y axis OPM | 1.392
| UNIST_신소재공학과 | 2025-09-18 | X Non-linearity | 0.023
| UNIST_신소재공학과 | 2025-09-18 | Y Non-linearity | 0.09
| Intel Hillsboro#6 | 2025-09-17 | Z Linearity by Z scanner position | 0.15
| Intel Hillsboro#6 | 2025-09-17 | Z Linearity by 4 different sample height | 1.002
| Intel Hillsboro#6 | 2025-09-17 | Full range of Z Detector during Z Scan sweep | 13.93
| Intel Hillsboro#6 | 2025-09-17 | Z Height calibration error | 0.001964
| Intel Hillsboro#6 | 2025-09-17 | Z Drive calibration error | 0.080751
| Intel Hillsboro#6 | 2025-09-17 | Full range of Z Height during Z Scan sweep | 43.8201
| Intel Hillsboro#6 | 2025-09-17 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 60.1385
| Intel Hillsboro#6 | 2025-09-17 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.9948
| Intel Hillsboro#6 | 2025-09-17 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 54.8648
| Intel Hillsboro#6 | 2025-09-17 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.8238
| Intel Hillsboro#6 | 2025-09-17 | Z Height noise | 0.052
| Intel Hillsboro#6 | 2025-09-17 | Wafer placement(Loading) error on the load port1 (X offset) | 19.05
| Intel Hillsboro#6 | 2025-09-17 | Wafer placement(Loading) error on the load port1 (Y offset) | 18.46
| Intel Hillsboro#6 | 2025-09-17 | Wafer placement(Loading) error on the load port2 (X offset) | 24.3
| Intel Hillsboro#6 | 2025-09-17 | Wafer placement(Loading) error on the load port2 (Y offset) | 17.91
| Intel Hillsboro#6 | 2025-09-17 | OpticalAxisShiftPerFocusX | 0.0006
| Intel Hillsboro#6 | 2025-09-17 | OpticalAxisShiftPerFocusY | 0.0006
| Intel Hillsboro#6 | 2025-09-17 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00022
| Intel Hillsboro#6 | 2025-09-17 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| Intel Hillsboro#6 | 2025-09-17 | Lift pins synchronization offset check (X offset) | 4.78
| Intel Hillsboro#6 | 2025-09-17 | Lift pins synchronization offset check (Y offset) | 40.36
| Intel Hillsboro#6 | 2025-09-17 | Max OPM @25mm | 135.773
| Intel Hillsboro#6 | 2025-09-17 | Max OPM @10mm | 41.137
| Intel Hillsboro#6 | 2025-09-17 | Max OPM @5mm | 32.829
| Intel Hillsboro#6 | 2025-09-17 | Max OPM @1mm | 9.893
| Intel Hillsboro#6 | 2025-09-17 | OPM Repeatability @25mm | 2.756
| Intel Hillsboro#6 | 2025-09-17 | OPM Repeatability @10mm | 1.611
| Intel Hillsboro#6 | 2025-09-17 | OPM Repeatability @5mm | 1.043
| Intel Hillsboro#6 | 2025-09-17 | OPM Repeatability @1mm | 0.711
| Intel Hillsboro#6 | 2025-09-17 | Torque Profile (50mm, 200um/s) | 0.357
| Intel Hillsboro#6 | 2025-09-17 | Torque Profile (50mm, 20um/s) | 0.432
| Intel Hillsboro#6 | 2025-09-17 | Profiler noise @200um | 0.129
| Intel Hillsboro#6 | 2025-09-17 | Roll motion to X+ full scan | 25.5
| Intel Hillsboro#6 | 2025-09-17 | Position repeatability(10repeats, Position Based) | 0.52
| Intel Hillsboro#6 | 2025-09-17 | Lateral(C-D) sensitivity | 15
| Intel Hillsboro#6 | 2025-09-17 | Lateral(C-D) noise | 0.238
| Intel Hillsboro#6 | 2025-09-17 | Min. Vibration Transmissbility (~100Hz) | 0.076
| Intel Hillsboro#6 | 2025-09-17 | RMS of ZeroScan | 0.047
| Intel Hillsboro#6 | 2025-09-17 | LightStrengthGain | 0.19
| Intel Hillsboro#6 | 2025-09-17 | Full range of X Detector during X Scan sweep | 213.5062
| Intel Hillsboro#6 | 2025-09-17 | Full range of X1 Detector during X Scan sweep | 212.956
| Intel Hillsboro#6 | 2025-09-17 | Full range of X2 Detector during X Scan sweep | 214.05
| Intel Hillsboro#6 | 2025-09-17 | Full range of Y Detector during Y Scan sweep | 217.033
| Intel Hillsboro#6 | 2025-09-17 | Full range of Y1 Detector during Y Scan sweep | 224.4
| Intel Hillsboro#6 | 2025-09-17 | Full range of Y2 Detector during Y Scan sweep | 209.6671
| Intel Hillsboro#6 | 2025-09-17 | Full range of X1 Detector Raw during X Scan sweep | 14.24
| Intel Hillsboro#6 | 2025-09-17 | Full range of X2 Detector Raw during X Scan sweep | 14.38
| Intel Hillsboro#6 | 2025-09-17 | Full range of Y1 Detector Raw during Y Scan sweep | 14.54
| Intel Hillsboro#6 | 2025-09-17 | Full range of Y2 Detector Raw during Y Scan sweep | 14.32
| Intel Hillsboro#6 | 2025-09-17 | X Detector stroke calibration error (nominal scale error X) | 0.22
| Intel Hillsboro#6 | 2025-09-17 | Y Detector stroke calibration error (nominal scale error Y) | 0.385
| Intel Hillsboro#6 | 2025-09-17 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 277.7995
| Intel Hillsboro#6 | 2025-09-17 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 282.8794
| Intel Hillsboro#6 | 2025-09-17 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 284.1796
| Intel Hillsboro#6 | 2025-09-17 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 271.3101
| Intel Hillsboro#6 | 2025-09-17 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 178.526
| Intel Hillsboro#6 | 2025-09-17 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 187.638
| Intel Hillsboro#6 | 2025-09-17 | XY Orthogonality | 90.1
| Intel Hillsboro#6 | 2025-09-17 | X Non-linearity error | 3.09
| Intel Hillsboro#6 | 2025-09-17 | Y Non-linearity error | 3.05
| Intel Hillsboro#6 | 2025-09-17 | X axis OPM | 0.651
| Intel Hillsboro#6 | 2025-09-17 | Y axis OPM | 0.619
| Intel Hillsboro#6 | 2025-09-17 | X Ringing(Rpv) | 0.133
| Intel Hillsboro#6 | 2025-09-17 | Y Ringing(Rpv) | 0.082
| Intel Hillsboro#6 | 2025-09-17 | X Detector noise (XY Servo on) | 0.299
| Intel Hillsboro#6 | 2025-09-17 | Y Detector noise (XY Servo on) | 0.362
| Intel Hillsboro#6 | 2025-09-17 | X1 Detector noise (XY Servo on) | 0.303
| Intel Hillsboro#6 | 2025-09-17 | Y1 Detector noise (XY Servo on) | 0.5
| Intel Hillsboro#6 | 2025-09-17 | X2 Detector noise (XY Servo on) | 0.482
| Intel Hillsboro#6 | 2025-09-17 | Y2 Detector noise (XY Servo on) | 0.472
| Intel Hillsboro#6 | 2025-09-17 | RMS error after Stage Calibration | 2.26
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from -X to +X (X scan) | 38.8183
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from -Y to +Y (Y scan) | 57.861
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from -X to +X (Y scan) | 29.296
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from -Y to +Y (X scan) | 6.5917
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from +X to -X (X scan) | 18.3105
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from +Y to -Y (Y scan) | 46.1425
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from +X to -X (Y scan) | 13.916
| Intel Hillsboro#6 | 2025-09-17 | Image offset after moving from +Y to -Y (X scan) | 13.1835
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X offset (Vision Pattern Matching) | 1.4697
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.1182
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.2587
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.1803
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X offset (In-Die Align) | 1.1865
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y offset (In-Die Align) | 1.3477
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X 1σ (In-Die Align) | 0.289
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y 1σ (In-Die Align) | 0.2729
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X offset (LLC Translation) | 1.5381
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y offset (LLC Translation) | 1.1523
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X 1σ (LLC Translation) | 0.3769
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y 1σ (LLC Translation) | 0.2331
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X offset (Global Align) | 1.6895
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y offset (Global Align) | 2
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning X 1σ (Global Align) | 0.3731
| Intel Hillsboro#6 | 2025-09-17 | XY Stage positioning Y 1σ (Global Align) | 0.5254
| Intel Hillsboro#6 | 2025-09-17 | Z Drift | 7.871
| Intel Hillsboro#6 | 2025-09-17 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.28
| Intel Hillsboro#6 | 2025-09-17 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Intel Hillsboro#6 | 2025-09-17 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.48
| Intel Hillsboro#6 | 2025-09-17 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Intel Hillsboro#6 | 2025-09-17 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.16
| Intel Hillsboro#6 | 2025-09-17 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| Intel Hillsboro#6 | 2025-09-17 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Intel Hillsboro#6 | 2025-09-17 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Intel Hillsboro#6 | 2025-09-17 | Cycle time for 15-point measurement | 2266.663
| Intel Hillsboro#6 | 2025-09-17 | Short-term stability using Step Height Sample (3σ) | 0.18
| Intel Hillsboro#6 | 2025-09-17 | Roughness Repeatability 1x1um(RSD%) | 2.196
| Intel Hillsboro#6 | 2025-09-17 | Roughness Repeatability 1x1um(1σ) | 0.001
| Intel Hillsboro#6 | 2025-09-17 | Roughness Repeatability 80x80um(RSD%) | 3.927
| Intel Hillsboro#6 | 2025-09-17 | Roughness Repeatability 80x80um(1σ) | 0.004
| Intel Hillsboro#6 | 2025-09-17 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer) | 1.14
| Intel Hillsboro#6 | 2025-09-17 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer) | 0.21
| Intel Hillsboro#6 | 2025-09-17 | Z Detector Max | 0.28249
| Intel Hillsboro#6 | 2025-09-17 | Z Detector Min | -0.36627
| Intel Hillsboro#6 | 2025-09-17 | WLI_Scale(Z Calibration) | 36.6
| PSS Demo | 2025-09-12 | SLD current limit check | 92
| PSS Demo | 2025-09-12 | SLD beam power at the end of the lens | 0.28
| PSS Demo | 2025-09-12 | Lateral(C-D) sensitivity | 5.9
| PSS Demo | 2025-09-12 | Lateral(C-D) noise | 0.113
| PSS Demo | 2025-09-12 | A-B noise | 10.531
| PSS Demo | 2025-09-12 | NCM Amplitude noise | 6.643
| PSS Demo | 2025-09-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| PSS Demo | 2025-09-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| PSS Demo | 2025-09-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| PSS Demo | 2025-09-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| PSS Demo | 2025-09-12 | Roughness Repeatability 1x1um(RSD%) | 0.89
| PSS Demo | 2025-09-12 | Roughness Repeatability 1x1um(1σ) | 0.002
| PSS Demo | 2025-09-12 | Full range of Z Height during Z Scan sweep | 17.3
| PSS Demo | 2025-09-12 | Full range of Z Detector during Z Scan sweep | 13.3
| PSS Demo | 2025-09-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.969
| PSS Demo | 2025-09-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.589
| PSS Demo | 2025-09-12 | Z Drive calibrtaion error | 0.40613
| PSS Demo | 2025-09-12 | Z Height calibrtaion error | 0.20019
| PSS Demo | 2025-09-12 | 2) of 20nm | 9.644
| PSS Demo | 2025-09-12 | Z linearity by sample height | 1.0151
| PSS Demo | 2025-09-12 | 1) Z Scanner position at +3μm | 0.000353
| PSS Demo | 2025-09-12 | 3) Z Scanner position at -3μm | 0.000264
| PSS Demo | 2025-09-12 | Z Height noise | 0.017
| PSS Demo | 2025-09-12 | Full range of X Detector during X Scan sweep | 104.9
| PSS Demo | 2025-09-12 | Full range of Y Detector during Y Scan sweep | 107
| PSS Demo | 2025-09-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.5
| PSS Demo | 2025-09-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 152
| PSS Demo | 2025-09-12 | X Detector stroke calibration error (nominal scale error X) | 0.356
| PSS Demo | 2025-09-12 | Y Detector stroke calibration error (nominal scale error Y) | 0.465
| PSS Demo | 2025-09-12 | X Scan Stroke | 104.9
| PSS Demo | 2025-09-12 | Y Scan Stroke | 107
| PSS Demo | 2025-09-12 | Full range of X Detector Raw during X Scan sweep | 13.6
| PSS Demo | 2025-09-12 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| PSS Demo | 2025-09-12 | X Detector noise (XYServo off) | 0.242
| PSS Demo | 2025-09-12 | Y Detector noise (XYServo off) | 0.237
| PSS Demo | 2025-09-12 | X Detector noise (XY Servo on) | 0.25
| PSS Demo | 2025-09-12 | Y Detector noise (XY Servo on) | 0.245
| PSS Demo | 2025-09-12 | X axis OPM | 1.373
| PSS Demo | 2025-09-12 | Y axis OPM | 0.707
| PSS Demo | 2025-09-12 | X Non-linearity | 0.222
| PSS Demo | 2025-09-12 | Y Non-linearity | 0.17
| Clemson University | 2025-09-12 | Lateral(C-D) sensitivity | 15.4
| Clemson University | 2025-09-12 | Lateral(C-D) noise | 0.045
| Clemson University | 2025-09-12 | A-B noise | 9.609
| Clemson University | 2025-09-12 | NCM Amplitude noise | 5.58
| Clemson University | 2025-09-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Clemson University | 2025-09-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| Clemson University | 2025-09-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Clemson University | 2025-09-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Clemson University | 2025-09-12 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0
| Clemson University | 2025-09-12 | Roughness Repeatability 1x1um(RSD%) | 2.361
| Clemson University | 2025-09-12 | Roughness Repeatability 1x1um(1σ) | 0.004
| Clemson University | 2025-09-12 | Full range of Z Height during Z Scan sweep | 17.1
| Clemson University | 2025-09-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.398
| Clemson University | 2025-09-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.579
| Clemson University | 2025-09-12 | Full range of Z Detector during Z Scan sweep | 13
| Clemson University | 2025-09-12 | Z Height calibration error | 0.015326
| Clemson University | 2025-09-12 | Z linearity by sample height | 1.0148
| Clemson University | 2025-09-12 | 1) Z Scanner position at +3μm | 0.005133
| Clemson University | 2025-09-12 | 3) Z Scanner position at -3μm | 0.007093
| Clemson University | 2025-09-12 | Z Height noise | 0.027
| Clemson University | 2025-09-12 | Full range of X Detector during X Scan sweep | 115.6
| Clemson University | 2025-09-12 | Full range of Y Detector during Y Scan sweep | 116.3
| Clemson University | 2025-09-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 165.5
| Clemson University | 2025-09-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 169.7
| Clemson University | 2025-09-12 | X Detector stroke calibration error (nominal scale error X) | 0.082
| Clemson University | 2025-09-12 | Y Detector stroke calibration error (nominal scale error Y) | 0.018
| Clemson University | 2025-09-12 | X Scan Stroke | 115.85
| Clemson University | 2025-09-12 | Y Scan Stroke | 116.2445
| Clemson University | 2025-09-12 | Full range of X Detector Raw during X Scan sweep | 14
| Clemson University | 2025-09-12 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Clemson University | 2025-09-12 | X Detector noise (XY Servo on) | 0.256
| Clemson University | 2025-09-12 | Y Detector noise (XY Servo on) | 0.26
| Clemson University | 2025-09-12 | X Detector noise (XY servo off) | 0.257
| Clemson University | 2025-09-12 | Y Detector noise (XY servo off) | 0.255
| Clemson University | 2025-09-12 | X axis OPM | 0.599
| Clemson University | 2025-09-12 | Y axis OPM | 1.345
| Clemson University | 2025-09-12 | X Non-linearity | 0.165
| Clemson University | 2025-09-12 | Y Non-linearity | 0.264
| Lumentum | 2025-09-12 | Z Linearity by Z scanner position | 0.2
| Lumentum | 2025-09-12 | Z Linearity by 4 different sample height | 1.005
| Lumentum | 2025-09-12 | Full range of Z Detector during Z Scan sweep | 12.35
| Lumentum | 2025-09-12 | Z Height calibration error | 0.19468
| Lumentum | 2025-09-12 | Z Drive calibration error | 0.12979
| Lumentum | 2025-09-12 | Full range of Z Height during Z Scan sweep | 17.415
| Lumentum | 2025-09-12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.2
| Lumentum | 2025-09-12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6
| Lumentum | 2025-09-12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.2
| Lumentum | 2025-09-12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6
| Lumentum | 2025-09-12 | Z Height noise | 0.027
| Lumentum | 2025-09-12 | OpticalAxisShiftPerFocusX | 0.0041
| Lumentum | 2025-09-12 | OpticalAxisShiftPerFocusY | -0.0011
| Lumentum | 2025-09-12 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| Lumentum | 2025-09-12 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Lumentum | 2025-09-12 | Lateral(C-D) sensitivity | 9.1
| Lumentum | 2025-09-12 | Lateral(C-D) noise | 0.216
| Lumentum | 2025-09-12 | RMS of ZeroScan | 0.032
| Lumentum | 2025-09-12 | LightStrengthGain | 1
| Lumentum | 2025-09-12 | Full range of X Detector during X Scan sweep | 169.2
| Lumentum | 2025-09-12 | Full range of X1 Detector during X Scan sweep | 166.9
| Lumentum | 2025-09-12 | Full range of X2 Detector during X Scan sweep | 169.2
| Lumentum | 2025-09-12 | Full range of Y Detector during Y Scan sweep | 184.9
| Lumentum | 2025-09-12 | Full range of Y1 Detector during Y Scan sweep | 184.9
| Lumentum | 2025-09-12 | Full range of Y2 Detector during Y Scan sweep | 178.9
| Lumentum | 2025-09-12 | Full range of X1 Detector Raw during X Scan sweep | 14.45
| Lumentum | 2025-09-12 | Full range of X2 Detector Raw during X Scan sweep | 14.47
| Lumentum | 2025-09-12 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| Lumentum | 2025-09-12 | Full range of Y2 Detector Raw during Y Scan sweep | 14.74
| Lumentum | 2025-09-12 | X Detector stroke calibration error (nominal scale error X) | 0.31
| Lumentum | 2025-09-12 | Y Detector stroke calibration error (nominal scale error Y) | 0.03
| Lumentum | 2025-09-12 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 244.0999
| Lumentum | 2025-09-12 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 239.1329
| Lumentum | 2025-09-12 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 237.7459
| Lumentum | 2025-09-12 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 235.3299
| Lumentum | 2025-09-12 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 165.017
| Lumentum | 2025-09-12 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 159.364
| Lumentum | 2025-09-12 | XY Orthogonality | 90.117
| Lumentum | 2025-09-12 | X Non-linearity error | 2.47
| Lumentum | 2025-09-12 | Y Non-linearity error | 2.81
| Lumentum | 2025-09-12 | X axis OPM | 0.368
| Lumentum | 2025-09-12 | Y axis OPM | 1.02
| Lumentum | 2025-09-12 | X Ringing(Rpv) | 0.293
| Lumentum | 2025-09-12 | Y Ringing(Rpv) | 0.299
| Lumentum | 2025-09-12 | X Detector noise (XY Servo on) | 0.358
| Lumentum | 2025-09-12 | Y Detector noise (XY Servo on) | 0.319
| Lumentum | 2025-09-12 | X1 Detector noise (XY Servo on) | 0.446
| Lumentum | 2025-09-12 | Y1 Detector noise (XY Servo on) | 0.444
| Lumentum | 2025-09-12 | X2 Detector noise (XY Servo on) | 0.49
| Lumentum | 2025-09-12 | Y2 Detector noise (XY Servo on) | 0.446
| Lumentum | 2025-09-12 | Z Drift | 13.8
| Lumentum | 2025-09-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.92
| Lumentum | 2025-09-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.17
| Lumentum | 2025-09-12 | Short-term stability using Step Height Sample (3σ) | 0.12
| Lumentum | 2025-09-12 | Roughness Repeatability 1x1um(RSD%) | 2.5
| Lumentum | 2025-09-12 | Roughness Repeatability 1x1um(1σ) | 0.005
| Lumentum | 2025-09-12 | Roughness Repeatability 80x80um(RSD%) | 1.639
| Lumentum | 2025-09-12 | Roughness Repeatability 80x80um(1σ) | 0.009
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Lateral(C-D) sensitivity | 27.5
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Lateral(C-D) noise | 0.059
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | A-B noise | 6.278
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | NCM Amplitude noise | 11.817
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Roughness Repeatability 1x1um(RSD%) | 1.94
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Roughness Repeatability 1x1um(1σ) | 0.005
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Full range of Z Height during Z Scan sweep | 18.5
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.589
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.548
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Full range of Z Detector during Z Scan sweep | 13.4
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Z Drive calibration error | 0.24138
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Z Height calibration error | 0.19923
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Z linearity by sample height | 1.0008
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | 1) Z Scanner position at +3μm | 0.003078
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | 3) Z Scanner position at -3μm | 0.001649
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Z Height noise | 0.029
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Full range of X Detector during X Scan sweep | 69.5
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Full range of Y Detector during Y Scan sweep | 67.7
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.2
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.8
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | X Detector stroke calibration error (nominal scale error X) | 0.244
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Y Detector stroke calibration error (nominal scale error Y) | 0.234
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | X Scan Stroke | 55
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Y Scan Stroke | 55
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Full range of X Detector Raw during X Scan sweep | 13.8
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | X Detector noise (XY Servo on) | 0.262
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Y Detector noise (XY Servo on) | 0.258
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | X Detector noise (XY Servo off) | 0.254
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Y Detector noise (XY Servo off) | 0.25
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | X axis OPM | 0.905
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Y axis OPM | 0.589
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | X Non-linearity | 0.065
| Uzkimyosanoat Joint Stock Company | 2025-09-10 | Y Non-linearity | 0.064
| Momentive Performance Materials (India) | 2025-09-08 | Lateral(C-D) sensitivity | 16.2
| Momentive Performance Materials (India) | 2025-09-08 | Lateral(C-D) noise | 0.106
| Momentive Performance Materials (India) | 2025-09-08 | A-B noise | 7.455
| Momentive Performance Materials (India) | 2025-09-08 | NCM Amplitude noise | 15.357
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.4
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.45
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.42
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.46
| Momentive Performance Materials (India) | 2025-09-08 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.41
| Momentive Performance Materials (India) | 2025-09-08 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Momentive Performance Materials (India) | 2025-09-08 | Roughness Repeatability 1x1um(RSD%) | 2.203
| Momentive Performance Materials (India) | 2025-09-08 | Roughness Repeatability 1x1um(1σ) | 0.003
| 서울대학교_에너지자원공학과 | 2025-09-03 | Lateral(C-D) sensitivity | 10.9
| 서울대학교_에너지자원공학과 | 2025-09-03 | Lateral(C-D) noise | 0.095
| 서울대학교_에너지자원공학과 | 2025-09-03 | A-B noise | 7.94
| 서울대학교_에너지자원공학과 | 2025-09-03 | NCM Amplitude noise | 8.153
| 서울대학교_에너지자원공학과 | 2025-09-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 서울대학교_에너지자원공학과 | 2025-09-03 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.18
| 서울대학교_에너지자원공학과 | 2025-09-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 서울대학교_에너지자원공학과 | 2025-09-03 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 서울대학교_에너지자원공학과 | 2025-09-03 | Roughness Repeatability 1x1um(RSD%) | 0.654
| 서울대학교_에너지자원공학과 | 2025-09-03 | Roughness Repeatability 1x1um(1σ) | 0.001
| 서울대학교_에너지자원공학과 | 2025-09-03 | Full range of Z Height during Z Scan sweep | 17.5
| 서울대학교_에너지자원공학과 | 2025-09-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.062
| 서울대학교_에너지자원공학과 | 2025-09-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.687
| 서울대학교_에너지자원공학과 | 2025-09-03 | Full range of Z Detector during Z Scan sweep | 12.9
| 서울대학교_에너지자원공학과 | 2025-09-03 | Z Drive calibration error | 0.45211
| 서울대학교_에너지자원공학과 | 2025-09-03 | Z Height calibration error | 0.16571
| 서울대학교_에너지자원공학과 | 2025-09-03 | Z linearity by sample height | 0.99409
| 서울대학교_에너지자원공학과 | 2025-09-03 | 1) Z Scanner position at +3μm | 0.00714
| 서울대학교_에너지자원공학과 | 2025-09-03 | 3) Z Scanner position at -3μm | 0.003032
| 서울대학교_에너지자원공학과 | 2025-09-03 | Z Height noise | 0.026
| 서울대학교_에너지자원공학과 | 2025-09-03 | Full range of X Detector during X Scan sweep | 63.6
| 서울대학교_에너지자원공학과 | 2025-09-03 | Full range of Y Detector during Y Scan sweep | 66.7
| 서울대학교_에너지자원공학과 | 2025-09-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97
| 서울대학교_에너지자원공학과 | 2025-09-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.3
| 서울대학교_에너지자원공학과 | 2025-09-03 | X Detector stroke calibration error (nominal scale error X) | 0.201
| 서울대학교_에너지자원공학과 | 2025-09-03 | Y Detector stroke calibration error (nominal scale error Y) | 0.041
| 서울대학교_에너지자원공학과 | 2025-09-03 | X Scan Stroke | 55
| 서울대학교_에너지자원공학과 | 2025-09-03 | Y Scan Stroke | 55
| 서울대학교_에너지자원공학과 | 2025-09-03 | Full range of X Detector Raw during X Scan sweep | 13.1
| 서울대학교_에너지자원공학과 | 2025-09-03 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| 서울대학교_에너지자원공학과 | 2025-09-03 | X Detector noise (XY Servo on) | 0.261
| 서울대학교_에너지자원공학과 | 2025-09-03 | Y Detector noise (XY Servo on) | 0.256
| 서울대학교_에너지자원공학과 | 2025-09-03 | X Detector noise (XY Servo off) | 0.273
| 서울대학교_에너지자원공학과 | 2025-09-03 | Y Detector noise (XY Servo off) | 0.285
| 서울대학교_에너지자원공학과 | 2025-09-03 | X axis OPM | 1.306
| 서울대학교_에너지자원공학과 | 2025-09-03 | Y axis OPM | 1.074
| 서울대학교_에너지자원공학과 | 2025-09-03 | X Non-linearity | 0.072
| 서울대학교_에너지자원공학과 | 2025-09-03 | Y Non-linearity | 0.066
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | SLD current limit check | 108
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | SLD beam power at the end of the lens | 0.29
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Lateral(C-D) sensitivity | 61.5
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Lateral(C-D) noise | 0.085
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | A-B noise | 10.572
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | NCM Amplitude noise | 5.859
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Roughness Repeatability 1x1um(RSD%) | 2.375
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Roughness Repeatability 1x1um(1σ) | 0.007
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Full range of Z Height during Z Scan sweep | 16.7
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Full range of Z Detector during Z Scan sweep | 14.4
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.224
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.68
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Z Drive calibrtaion error | 0.068008
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Z Height calibrtaion error | 0.15421
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | 2) of 20nm | 0.971
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Z linearity by sample height | 1.0017
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | 1) Z Scanner position at +3μm | 0.002064
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | 3) Z Scanner position at -3μm | 0.002451
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Z Height noise | 0.018
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Full range of X Detector during X Scan sweep | 109.8
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Full range of Y Detector during Y Scan sweep | 113.2
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 156.5
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | X Detector stroke calibration error (nominal scale error X) | 0.112
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.32
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | X Scan Stroke | 106.1
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Y Scan Stroke | 108.7
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Full range of X Detector Raw during X Scan sweep | 13.7
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | X Detector noise (XYServo off) | 0.261
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Y Detector noise (XYServo off) | 0.258
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | X Detector noise (XY Servo on) | 0.27
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Y Detector noise (XY Servo on) | 0.264
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | X axis OPM | 0.871
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Y axis OPM | 0.569
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | X Non-linearity | 0.089
| 전북대학교_중재적메카노바이오센터 | 2025-09-02 | Y Non-linearity | 0.223
| 한국나노마이스터고등학고 | 2025-08-29 | A-B noise | 6.919
| 한국나노마이스터고등학고 | 2025-08-29 | NCM Amplitude noise | 15.394
| 한국나노마이스터고등학고 | 2025-08-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| 한국나노마이스터고등학고 | 2025-08-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| 한국나노마이스터고등학고 | 2025-08-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| 한국나노마이스터고등학고 | 2025-08-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 한국나노마이스터고등학고 | 2025-08-29 | Full range of Z Height during Z Scan sweep | 18.1
| 한국나노마이스터고등학고 | 2025-08-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.853
| 한국나노마이스터고등학고 | 2025-08-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.939
| 한국나노마이스터고등학고 | 2025-08-29 | Full range of Z Detector during Z Scan sweep | 13.5
| 한국나노마이스터고등학고 | 2025-08-29 | Z Drive calibration error | 0.24234
| 한국나노마이스터고등학고 | 2025-08-29 | Z Height calibration error | 0.20785
| 한국나노마이스터고등학고 | 2025-08-29 | Z linearity by sample height | 1.0015
| 한국나노마이스터고등학고 | 2025-08-29 | 1) Z Scanner position at +3μm | 0.008007
| 한국나노마이스터고등학고 | 2025-08-29 | 3) Z Scanner position at -3μm | 0.004185
| 한국나노마이스터고등학고 | 2025-08-29 | Z Height noise | 0.027
| 한국나노마이스터고등학고 | 2025-08-29 | Full range of X Detector during X Scan sweep | 69.1
| 한국나노마이스터고등학고 | 2025-08-29 | Full range of Y Detector during Y Scan sweep | 72.3
| 한국나노마이스터고등학고 | 2025-08-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 89.2
| 한국나노마이스터고등학고 | 2025-08-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 93.3
| 한국나노마이스터고등학고 | 2025-08-29 | X Detector stroke calibration error (nominal scale error X) | 0.144
| 한국나노마이스터고등학고 | 2025-08-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.332
| 한국나노마이스터고등학고 | 2025-08-29 | X Scan Stroke | 55
| 한국나노마이스터고등학고 | 2025-08-29 | Y Scan Stroke | 55
| 한국나노마이스터고등학고 | 2025-08-29 | Full range of X Detector Raw during X Scan sweep | 14.4
| 한국나노마이스터고등학고 | 2025-08-29 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| 한국나노마이스터고등학고 | 2025-08-29 | X Detector noise (XY Servo on) | 0.35
| 한국나노마이스터고등학고 | 2025-08-29 | Y Detector noise (XY Servo on) | 0.371
| 한국나노마이스터고등학고 | 2025-08-29 | X Detector noise (XY Servo off) | 0.373
| 한국나노마이스터고등학고 | 2025-08-29 | Y Detector noise (XY Servo off) | 0.382
| 한국나노마이스터고등학고 | 2025-08-29 | X axis OPM | 0.973
| 한국나노마이스터고등학고 | 2025-08-29 | Y axis OPM | 0.524
| 한국나노마이스터고등학고 | 2025-08-29 | X Non-linearity | 0.043
| 한국나노마이스터고등학고 | 2025-08-29 | Y Non-linearity | 0.091
| IOE CAS (Chengdu) | 2025-08-28 | Z Linearity by Z scanner position | 0.53
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Z Detector during Z Scan sweep | 12.27
| IOE CAS (Chengdu) | 2025-08-28 | Z Height calibration error | 0.028154
| IOE CAS (Chengdu) | 2025-08-28 | Z Drive calibration error | 0.14251
| IOE CAS (Chengdu) | 2025-08-28 | Z Linearity by 4 different sample height | 0.99563
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Z Height during Z Scan sweep | 17.5117
| IOE CAS (Chengdu) | 2025-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.5418
| IOE CAS (Chengdu) | 2025-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6131
| IOE CAS (Chengdu) | 2025-08-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.5482
| IOE CAS (Chengdu) | 2025-08-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6306
| IOE CAS (Chengdu) | 2025-08-28 | Z Height noise | 0.032
| IOE CAS (Chengdu) | 2025-08-28 | Wafer placement(Loading) error on the load port1 (X offset) | 7.45
| IOE CAS (Chengdu) | 2025-08-28 | Wafer placement(Loading) error on the load port1 (Y offset) | 29.96
| IOE CAS (Chengdu) | 2025-08-28 | Wafer placement(Loading) error on the load port2 (X offset) | 19.35
| IOE CAS (Chengdu) | 2025-08-28 | Wafer placement(Loading) error on the load port2 (Y offset) | 29.02
| IOE CAS (Chengdu) | 2025-08-28 | OpticalAxisShiftPerFocusX | -0.002
| IOE CAS (Chengdu) | 2025-08-28 | OpticalAxisShiftPerFocusY | 0.0005
| IOE CAS (Chengdu) | 2025-08-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000242
| IOE CAS (Chengdu) | 2025-08-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6.75e-05
| IOE CAS (Chengdu) | 2025-08-28 | Profiler noise @200um | 0.144
| IOE CAS (Chengdu) | 2025-08-28 | Roll motion to X+ full scan | 26.1
| IOE CAS (Chengdu) | 2025-08-28 | Position repeatability(10repeats, Position Based) | 0.52
| IOE CAS (Chengdu) | 2025-08-28 | OPM Repeatability @25mm | 0.983
| IOE CAS (Chengdu) | 2025-08-28 | OPM Repeatability @10mm | 1.162
| IOE CAS (Chengdu) | 2025-08-28 | OPM Repeatability @5mm | 0.878
| IOE CAS (Chengdu) | 2025-08-28 | OPM Repeatability @1mm | 0.935
| IOE CAS (Chengdu) | 2025-08-28 | Max OPM @25mm | 20.918
| IOE CAS (Chengdu) | 2025-08-28 | Max OPM @10mm | 9.005
| IOE CAS (Chengdu) | 2025-08-28 | Max OPM @5mm | 9.552
| IOE CAS (Chengdu) | 2025-08-28 | Max OPM @1mm | 9.581
| IOE CAS (Chengdu) | 2025-08-28 | Torque Profile (50mm, 200um/s) | 0.357
| IOE CAS (Chengdu) | 2025-08-28 | Torque Profile (50mm, 20um/s) | 0.432
| IOE CAS (Chengdu) | 2025-08-28 | Lateral(C-D) sensitivity | 45.9
| IOE CAS (Chengdu) | 2025-08-28 | Lateral(C-D) noise | 0.15821
| IOE CAS (Chengdu) | 2025-08-28 | Min. Vibration Transmissbility (~100Hz) | 0.032362
| IOE CAS (Chengdu) | 2025-08-28 | RMS of ZeroScan | 0.043
| IOE CAS (Chengdu) | 2025-08-28 | LightStrengthGain | 0.56
| IOE CAS (Chengdu) | 2025-08-28 | Full range of X Detector during X Scan sweep | 219.895
| IOE CAS (Chengdu) | 2025-08-28 | Full range of X1 Detector during X Scan sweep | 219.927
| IOE CAS (Chengdu) | 2025-08-28 | Full range of X2 Detector during X Scan sweep | 219.862
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Y Detector during Y Scan sweep | 208.806
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Y1 Detector during Y Scan sweep | 198.019
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Y2 Detector during Y Scan sweep | 219.592
| IOE CAS (Chengdu) | 2025-08-28 | Full range of X1 Detector Raw during X Scan sweep | 13.94
| IOE CAS (Chengdu) | 2025-08-28 | Full range of X2 Detector Raw during X Scan sweep | 14.21
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.31
| IOE CAS (Chengdu) | 2025-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.17
| IOE CAS (Chengdu) | 2025-08-28 | X Detector stroke calibration error (nominal scale error X) | 0.499
| IOE CAS (Chengdu) | 2025-08-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.241
| IOE CAS (Chengdu) | 2025-08-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 298.3204
| IOE CAS (Chengdu) | 2025-08-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 292.3854
| IOE CAS (Chengdu) | 2025-08-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 266.5094
| IOE CAS (Chengdu) | 2025-08-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 272.6814
| IOE CAS (Chengdu) | 2025-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 180.088
| IOE CAS (Chengdu) | 2025-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 170.465
| IOE CAS (Chengdu) | 2025-08-28 | X Non-linearity error | 2.86
| IOE CAS (Chengdu) | 2025-08-28 | Y Non-linearity error | 3.04
| IOE CAS (Chengdu) | 2025-08-28 | X axis OPM | 1.055
| IOE CAS (Chengdu) | 2025-08-28 | Y axis OPM | 0.881
| IOE CAS (Chengdu) | 2025-08-28 | X Ringing(Rpv) | 0.092
| IOE CAS (Chengdu) | 2025-08-28 | Y Ringing(Rpv) | 0.131
| IOE CAS (Chengdu) | 2025-08-28 | X Detector noise (XY Servo on) | 0.352
| IOE CAS (Chengdu) | 2025-08-28 | Y Detector noise (XY Servo on) | 0.347
| IOE CAS (Chengdu) | 2025-08-28 | X1 Detector noise (XY Servo on) | 0.48
| IOE CAS (Chengdu) | 2025-08-28 | Y1 Detector noise (XY Servo on) | 0.479
| IOE CAS (Chengdu) | 2025-08-28 | X2 Detector noise (XY Servo on) | 0.479
| IOE CAS (Chengdu) | 2025-08-28 | Y2 Detector noise (XY Servo on) | 0.462
| IOE CAS (Chengdu) | 2025-08-28 | RMS error after Stage Calibration | 2.86
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from -X to +X (X scan) | 43.2129
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from -Y to +Y (Y scan) | 106.9336
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from -X to +X (Y scan) | 28.5645
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from -Y to +Y (X scan) | 35.1562
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from +X to -X (X scan) | 43.2129
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from +Y to -Y (Y scan) | 48.3398
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from +X to -X (Y scan) | 9.5215
| IOE CAS (Chengdu) | 2025-08-28 | Image offset after moving from +Y to -Y (X scan) | 13.916
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X offset (Vision Pattern Matching) | 1.3086
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y offset (Vision Pattern Matching) | 0.9131
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.3032
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.1726
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X offset (In-Die Align) | 1.9385
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y offset (In-Die Align) | 1.499
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X 1σ (In-Die Align) | 0.4212
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y 1σ (In-Die Align) | 0.3139
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X offset (LLC Translation) | 2.0361
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y offset (LLC Translation) | 2.002
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X 1σ (LLC Translation) | 0.5086
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y 1σ (LLC Translation) | 0.4166
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X offset (Global Align) | 2.6611
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y offset (Global Align) | 3.6182
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning X 1σ (Global Align) | 0.6095
| IOE CAS (Chengdu) | 2025-08-28 | XY Stage positioning Y 1σ (Global Align) | 1.1007
| IOE CAS (Chengdu) | 2025-08-28 | Z Drift | 8.2057
| IOE CAS (Chengdu) | 2025-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.19101
| IOE CAS (Chengdu) | 2025-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.035767
| IOE CAS (Chengdu) | 2025-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24831
| IOE CAS (Chengdu) | 2025-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.046708
| IOE CAS (Chengdu) | 2025-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17223
| IOE CAS (Chengdu) | 2025-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.032243
| IOE CAS (Chengdu) | 2025-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1997
| IOE CAS (Chengdu) | 2025-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.037553
| IOE CAS (Chengdu) | 2025-08-28 | Cycle time for 15-point measurement | 1653.069
| IOE CAS (Chengdu) | 2025-08-28 | Short-term stability w/ 180nm Step Height Sample (3σ) | 0.17636
| IOE CAS (Chengdu) | 2025-08-28 | Roughness Repeatability 1x1um(RSD%) | 1.3317
| IOE CAS (Chengdu) | 2025-08-28 | Roughness Repeatability 1x1um(1σ) | 0.001197
| IOE CAS (Chengdu) | 2025-08-28 | Roughness Repeatability 80x80um(RSD%) | 0.40796
| IOE CAS (Chengdu) | 2025-08-28 | Roughness Repeatability 80x80um(1σ) | 0.00137
| IOE CAS (Chengdu) | 2025-08-28 | Laser intensity with AC160TS cantilever | 2.83
| IOE CAS (Chengdu) | 2025-08-28 | Check if the driving % changes proportionally as the NCM amplitude changes | Yes
| IOE CAS (Chengdu) | 2025-08-28 | NCM driving % at an amplitude of 1 nm | 0.04
| IOE CAS (Chengdu) | 2025-08-28 | NCM driving % at an amplitude of 20 nm | 0.86
| IOE CAS (Chengdu) | 2025-08-28 | NCM driving % at an amplitude of 100 nm | 4.26
| IOE CAS (Chengdu) | 2025-08-28 | Z Linearity by Z scanner position | 0.6
| IOE CAS (Chengdu) | 2025-08-28 | Z Linearity by 4 different sample height | 0.994
| IOE CAS (Chengdu) | 2025-08-28 | - 940nm Step Height Sample (10 Repeats) | 943.33
| IOE CAS (Chengdu) | 2025-08-28 | - 200~500nm Step Height Sample (10 Repeats) | 458.19
| IOE CAS (Chengdu) | 2025-08-28 | - 50~200nm Step Height Sample (10 Repeats) | 184.24
| IOE CAS (Chengdu) | 2025-08-28 | - Under 50nm Step Height Sample (10 Repeats) | 16.22
| IOE CAS (Chengdu) | 2025-08-28 | Full range in voltage during Z Scan sweep | 13.2
| IOE CAS (Chengdu) | 2025-08-28 | Full range in um during Z Scan sweep | 42.4
| IOE CAS (Chengdu) | 2025-08-28 | Z Height noise | 0.07
| IOE CAS (Chengdu) | 2025-08-28 | A-B after Auto PSPD Alignment | 0.04
| Samsung(SEC) P2D | 2025-08-25 | Z Linearity by Z scanner position | 0.73
| Samsung(SEC) P2D | 2025-08-25 | Full range of Z Detector during Z Scan sweep | 12.74
| Samsung(SEC) P2D | 2025-08-25 | Z Height calibration error | 0.044016
| Samsung(SEC) P2D | 2025-08-25 | Z Drive calibration error | 0.20273
| Samsung(SEC) P2D | 2025-08-25 | Z Linearity by 4 different sample height | 0.99475
| Samsung(SEC) P2D | 2025-08-25 | Full range of Z Height during Z Scan sweep | 16.354
| Samsung(SEC) P2D | 2025-08-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.5443
| Samsung(SEC) P2D | 2025-08-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3353
| Samsung(SEC) P2D | 2025-08-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5742
| Samsung(SEC) P2D | 2025-08-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4216
| Samsung(SEC) P2D | 2025-08-25 | Z Height noise | 0.027
| Samsung(SEC) P2D | 2025-08-25 | Wafer placement(Loading) error on the load port1 (X offset) | 47.7
| Samsung(SEC) P2D | 2025-08-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 21.1
| Samsung(SEC) P2D | 2025-08-25 | Wafer placement(Loading) error on the load port2 (X offset) | 13.3
| Samsung(SEC) P2D | 2025-08-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 22.65
| Samsung(SEC) P2D | 2025-08-25 | OpticalAxisShiftPerFocusX | 0.0018
| Samsung(SEC) P2D | 2025-08-25 | OpticalAxisShiftPerFocusY | -0.0049
| Samsung(SEC) P2D | 2025-08-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00094
| Samsung(SEC) P2D | 2025-08-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00038
| Samsung(SEC) P2D | 2025-08-25 | Lift pins synchronization offset check (X offset) | 46.67
| Samsung(SEC) P2D | 2025-08-25 | Lift pins synchronization offset check (Y offset) | 19.02
| Samsung(SEC) P2D | 2025-08-25 | Profiler noise @200um | 0.11
| Samsung(SEC) P2D | 2025-08-25 | Roll motion to X+ full scan | 69.7
| Samsung(SEC) P2D | 2025-08-25 | Position repeatability(10repeats, Position Based) | 0.4
| Samsung(SEC) P2D | 2025-08-25 | OPM Repeatability @25mm | 2.77
| Samsung(SEC) P2D | 2025-08-25 | OPM Repeatability @10mm | 2.85
| Samsung(SEC) P2D | 2025-08-25 | OPM Repeatability @5mm | 0.963
| Samsung(SEC) P2D | 2025-08-25 | OPM Repeatability @1mm | 0.545
| Samsung(SEC) P2D | 2025-08-25 | Max OPM @25mm | 164.992
| Samsung(SEC) P2D | 2025-08-25 | Max OPM @10mm | 48.682
| Samsung(SEC) P2D | 2025-08-25 | Max OPM @5mm | 34.324
| Samsung(SEC) P2D | 2025-08-25 | Max OPM @1mm | 8.188
| Samsung(SEC) P2D | 2025-08-25 | Torque Profile (50mm, 200um/s) | 0.73
| Samsung(SEC) P2D | 2025-08-25 | Torque Profile (50mm, 20um/s) | 0.739
| Samsung(SEC) P2D | 2025-08-25 | Lateral(C-D) sensitivity | 13.2
| Samsung(SEC) P2D | 2025-08-25 | Lateral(C-D) noise | 0.232
| Samsung(SEC) P2D | 2025-08-25 | Min. Vibration Transmissbility (~100Hz) | 0.02
| Samsung(SEC) P2D | 2025-08-25 | RMS of ZeroScan | 0.042
| Samsung(SEC) P2D | 2025-08-25 | LightStrengthGain | 0.53
| Samsung(SEC) P2D | 2025-08-25 | Full range of X Detector during X Scan sweep | 197.1112
| Samsung(SEC) P2D | 2025-08-25 | Full range of X1 Detector during X Scan sweep | 211.3975
| Samsung(SEC) P2D | 2025-08-25 | Full range of X2 Detector during X Scan sweep | 197.1112
| Samsung(SEC) P2D | 2025-08-25 | Full range of Y Detector during Y Scan sweep | 193.7353
| Samsung(SEC) P2D | 2025-08-25 | Full range of Y1 Detector during Y Scan sweep | 193.7353
| Samsung(SEC) P2D | 2025-08-25 | Full range of Y2 Detector during Y Scan sweep | 200.6991
| Samsung(SEC) P2D | 2025-08-25 | Full range of X1 Detector Raw during X Scan sweep | 14.63
| Samsung(SEC) P2D | 2025-08-25 | Full range of X2 Detector Raw during X Scan sweep | 14.56
| Samsung(SEC) P2D | 2025-08-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.61
| Samsung(SEC) P2D | 2025-08-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.67
| Samsung(SEC) P2D | 2025-08-25 | X Detector stroke calibration error (nominal scale error X) | 0.121
| Samsung(SEC) P2D | 2025-08-25 | Y Detector stroke calibration error (nominal scale error Y) | 0.443
| Samsung(SEC) P2D | 2025-08-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 264.151
| Samsung(SEC) P2D | 2025-08-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 267.7272
| Samsung(SEC) P2D | 2025-08-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 253.811
| Samsung(SEC) P2D | 2025-08-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 265.4847
| Samsung(SEC) P2D | 2025-08-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 175.374
| Samsung(SEC) P2D | 2025-08-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 172.457
| Samsung(SEC) P2D | 2025-08-25 | X Non-linearity error | 3.65
| Samsung(SEC) P2D | 2025-08-25 | Y Non-linearity error | 3.09
| Samsung(SEC) P2D | 2025-08-25 | X axis OPM | 3.591
| Samsung(SEC) P2D | 2025-08-25 | Y axis OPM | 0.798
| Samsung(SEC) P2D | 2025-08-25 | X Ringing(Rpv) | 0.124
| Samsung(SEC) P2D | 2025-08-25 | Y Ringing(Rpv) | 0.069
| Samsung(SEC) P2D | 2025-08-25 | X Detector noise (XY Servo on) | 0.357
| Samsung(SEC) P2D | 2025-08-25 | Y Detector noise (XY Servo on) | 0.398
| Samsung(SEC) P2D | 2025-08-25 | X1 Detector noise (XY Servo on) | 0.482
| Samsung(SEC) P2D | 2025-08-25 | Y1 Detector noise (XY Servo on) | 0.512
| Samsung(SEC) P2D | 2025-08-25 | X2 Detector noise (XY Servo on) | 0.482
| Samsung(SEC) P2D | 2025-08-25 | Y2 Detector noise (XY Servo on) | 0.513
| Samsung(SEC) P2D | 2025-08-25 | RMS error after Stage Calibration | 1.98
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from -X to +X (X scan) | 12.4512
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from -Y to +Y (Y scan) | 67.3828
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from -X to +X (Y scan) | 79.834
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from -Y to +Y (X scan) | 38.8184
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from +X to -X (X scan) | 51.2695
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from +Y to -Y (Y scan) | 27.0996
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from +X to -X (Y scan) | 16.8457
| Samsung(SEC) P2D | 2025-08-25 | Image offset after moving from +Y to -Y (X scan) | 62.2559
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X offset (Vision Pattern Matching) | 0.98633
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.7578
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.1906
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3778
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X offset (In-Die Align) | 1.2061
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y offset (In-Die Align) | 2.3975
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X 1σ (In-Die Align) | 0.256
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y 1σ (In-Die Align) | 0.4532
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X offset (LLC Translation) | 1.5723
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y offset (LLC Translation) | 2.6221
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X 1σ (LLC Translation) | 0.3392
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y 1σ (LLC Translation) | 0.5155
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X offset (Global Align) | 5.9668
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y offset (Global Align) | 5.6641
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning X 1σ (Global Align) | 1.3483
| Samsung(SEC) P2D | 2025-08-25 | XY Stage positioning Y 1σ (Global Align) | 1.565
| Samsung(SEC) P2D | 2025-08-25 | Z Drift | 10.277
| Samsung(SEC) P2D | 2025-08-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.41495
| Samsung(SEC) P2D | 2025-08-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.074798
| Samsung(SEC) P2D | 2025-08-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.44546
| Samsung(SEC) P2D | 2025-08-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.080667
| Samsung(SEC) P2D | 2025-08-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13524
| Samsung(SEC) P2D | 2025-08-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.024389
| Samsung(SEC) P2D | 2025-08-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.16301
| Samsung(SEC) P2D | 2025-08-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.029527
| Samsung(SEC) P2D | 2025-08-25 | Cycle time for 15-point measurement | 1737.061
| Samsung(SEC) P2D | 2025-08-25 | Short-term stability using Step Height Sample (3σ) | 0.19
| Samsung(SEC) P2D | 2025-08-25 | Roughness Repeatability 1x1um(RSD%) | 0.7
| Sotacarbo Spa | 2025-08-20 | Lateral(C-D) sensitivity | 13.7
| Sotacarbo Spa | 2025-08-20 | Lateral(C-D) noise | 0.085
| Sotacarbo Spa | 2025-08-20 | A-B noise | 8.163
| Sotacarbo Spa | 2025-08-20 | NCM Amplitude noise | 14.696
| Sotacarbo Spa | 2025-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| Sotacarbo Spa | 2025-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| Sotacarbo Spa | 2025-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Sotacarbo Spa | 2025-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Sotacarbo Spa | 2025-08-20 | Roughness Repeatability 1x1um(RSD%) | 0.514
| Sotacarbo Spa | 2025-08-20 | Roughness Repeatability 1x1um(1σ) | 0.002
| Sotacarbo Spa | 2025-08-20 | Full range of Z Height during Z Scan sweep | 17.6
| Sotacarbo Spa | 2025-08-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.749
| Sotacarbo Spa | 2025-08-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.873
| Sotacarbo Spa | 2025-08-20 | Full range of Z Detector during Z Scan sweep | 12.7
| Sotacarbo Spa | 2025-08-20 | Z Drive calibration error | 0.26149
| Sotacarbo Spa | 2025-08-20 | Z Height calibration error | 0.37356
| Sotacarbo Spa | 2025-08-20 | Z linearity by sample height | 1.0036
| Sotacarbo Spa | 2025-08-20 | 1) Z Scanner position at +3μm | 0.007162
| Sotacarbo Spa | 2025-08-20 | 3) Z Scanner position at -3μm | 0.001101
| Sotacarbo Spa | 2025-08-20 | Z Height noise | 0.028
| Sotacarbo Spa | 2025-08-20 | Full range of X Detector during X Scan sweep | 116.9
| Sotacarbo Spa | 2025-08-20 | Full range of Y Detector during Y Scan sweep | 114.9
| Sotacarbo Spa | 2025-08-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 173.5
| Sotacarbo Spa | 2025-08-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 177.5
| Sotacarbo Spa | 2025-08-20 | X Detector stroke calibration error (nominal scale error X) | 0.178
| Sotacarbo Spa | 2025-08-20 | Y Detector stroke calibration error (nominal scale error Y) | 0.071
| Sotacarbo Spa | 2025-08-20 | X Scan Stroke | 116.9
| Sotacarbo Spa | 2025-08-20 | Y Scan Stroke | 114.9
| Sotacarbo Spa | 2025-08-20 | Full range of X Detector Raw during X Scan sweep | 13.5
| Sotacarbo Spa | 2025-08-20 | Full range of Y Detector Raw during Y Scan sweep | 13
| Sotacarbo Spa | 2025-08-20 | X Detector noise (XY Servo on) | 0.275
| Sotacarbo Spa | 2025-08-20 | Y Detector noise (XY Servo on) | 0.296
| Sotacarbo Spa | 2025-08-20 | X Detector noise (XY servo off) | 0.271
| Sotacarbo Spa | 2025-08-20 | Y Detector noise (XY servo off) | 0.288
| Sotacarbo Spa | 2025-08-20 | X axis OPM | 1.861
| Sotacarbo Spa | 2025-08-20 | Y axis OPM | 0.966
| Sotacarbo Spa | 2025-08-20 | X Non-linearity | 0.176
| Sotacarbo Spa | 2025-08-20 | Y Non-linearity | 0.114
| 고려대학교_융합대학원 | 2025-08-19 | Lateral(C-D) sensitivity | 13
| 고려대학교_융합대학원 | 2025-08-19 | Lateral(C-D) noise | 0.063
| 고려대학교_융합대학원 | 2025-08-19 | A-B noise | 11.887
| 고려대학교_융합대학원 | 2025-08-19 | NCM Amplitude noise | 12.002
| 고려대학교_융합대학원 | 2025-08-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 고려대학교_융합대학원 | 2025-08-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 고려대학교_융합대학원 | 2025-08-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 고려대학교_융합대학원 | 2025-08-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| 고려대학교_융합대학원 | 2025-08-19 | Roughness Repeatability 1x1um(RSD%) | 2.359
| 고려대학교_융합대학원 | 2025-08-19 | Roughness Repeatability 1x1um(1σ) | 0.005
| 고려대학교_융합대학원 | 2025-08-19 | Full range of Z Height during Z Scan sweep | 17.5
| 고려대학교_융합대학원 | 2025-08-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.631
| 고려대학교_융합대학원 | 2025-08-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.581
| 고려대학교_융합대학원 | 2025-08-19 | Full range of Z Detector during Z Scan sweep | 12.7
| 고려대학교_융합대학원 | 2025-08-19 | Z Drive calibration error | 0.19444
| 고려대학교_융합대학원 | 2025-08-19 | Z Height calibration error | 0.050766
| 고려대학교_융합대학원 | 2025-08-19 | Z linearity by sample height | 1.0121
| 고려대학교_융합대학원 | 2025-08-19 | 1) Z Scanner position at +3μm | 0.006622
| 고려대학교_융합대학원 | 2025-08-19 | 3) Z Scanner position at -3μm | 0.001153
| 고려대학교_융합대학원 | 2025-08-19 | Z Height noise | 0.028
| 고려대학교_융합대학원 | 2025-08-19 | Full range of X Detector during X Scan sweep | 66.4
| 고려대학교_융합대학원 | 2025-08-19 | Full range of Y Detector during Y Scan sweep | 66.4
| 고려대학교_융합대학원 | 2025-08-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.7
| 고려대학교_융합대학원 | 2025-08-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.9
| 고려대학교_융합대학원 | 2025-08-19 | X Detector stroke calibration error (nominal scale error X) | 0.031
| 고려대학교_융합대학원 | 2025-08-19 | Y Detector stroke calibration error (nominal scale error Y) | 0.189
| 고려대학교_융합대학원 | 2025-08-19 | X Scan Stroke | 55
| 고려대학교_융합대학원 | 2025-08-19 | Y Scan Stroke | 55
| 고려대학교_융합대학원 | 2025-08-19 | Full range of X Detector Raw during X Scan sweep | 13.6
| 고려대학교_융합대학원 | 2025-08-19 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| 고려대학교_융합대학원 | 2025-08-19 | X Detector noise (XY Servo on) | 0.262
| 고려대학교_융합대학원 | 2025-08-19 | Y Detector noise (XY Servo on) | 0.261
| 고려대학교_융합대학원 | 2025-08-19 | X Detector noise (XY Servo off) | 0.255
| 고려대학교_융합대학원 | 2025-08-19 | Y Detector noise (XY Servo off) | 0.254
| 고려대학교_융합대학원 | 2025-08-19 | X axis OPM | 1.077
| 고려대학교_융합대학원 | 2025-08-19 | Y axis OPM | 0.693
| 고려대학교_융합대학원 | 2025-08-19 | X Non-linearity | 0.065
| 고려대학교_융합대학원 | 2025-08-19 | Y Non-linearity | 0.038
| 공주대학교 공동실험실습관 | 2025-08-11 | SLD current limit check | 148
| 공주대학교 공동실험실습관 | 2025-08-11 | SLD beam power at the end of the lens | 0.3
| 공주대학교 공동실험실습관 | 2025-08-11 | Lateral(C-D) sensitivity | 29.2
| 공주대학교 공동실험실습관 | 2025-08-11 | Lateral(C-D) noise | 0.093
| 공주대학교 공동실험실습관 | 2025-08-11 | A-B noise | 12.367
| 공주대학교 공동실험실습관 | 2025-08-11 | NCM Amplitude noise | 5.335
| 공주대학교 공동실험실습관 | 2025-08-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 공주대학교 공동실험실습관 | 2025-08-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| 공주대학교 공동실험실습관 | 2025-08-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 공주대학교 공동실험실습관 | 2025-08-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 공주대학교 공동실험실습관 | 2025-08-11 | Roughness Repeatability 1x1um(RSD%) | 3.338
| 공주대학교 공동실험실습관 | 2025-08-11 | Roughness Repeatability 1x1um(1σ) | 0.004
| 공주대학교 공동실험실습관 | 2025-08-11 | Full range of Z Height during Z Scan sweep | 16.9
| 공주대학교 공동실험실습관 | 2025-08-11 | Full range of Z Detector during Z Scan sweep | 14.2
| 공주대학교 공동실험실습관 | 2025-08-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.88
| 공주대학교 공동실험실습관 | 2025-08-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.683
| 공주대학교 공동실험실습관 | 2025-08-11 | Z Drive calibrtaion error | 0.25479
| 공주대학교 공동실험실습관 | 2025-08-11 | Z Height calibrtaion error | 0.06705
| 공주대학교 공동실험실습관 | 2025-08-11 | NCM driving% with AC160TS cantilever | 2.634
| 공주대학교 공동실험실습관 | 2025-08-11 | Z linearity by sample height | 1.011
| 공주대학교 공동실험실습관 | 2025-08-11 | 1) Z Scanner position at +3μm | 0.006539
| 공주대학교 공동실험실습관 | 2025-08-11 | 3) Z Scanner position at -3μm | 0.003234
| 공주대학교 공동실험실습관 | 2025-08-11 | Z Height noise | 0.021
| 공주대학교 공동실험실습관 | 2025-08-11 | Full range of X Detector during X Scan sweep | 104.3
| 공주대학교 공동실험실습관 | 2025-08-11 | Full range of Y Detector during Y Scan sweep | 108.2
| 공주대학교 공동실험실습관 | 2025-08-11 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 152.6
| 공주대학교 공동실험실습관 | 2025-08-11 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 156
| 공주대학교 공동실험실습관 | 2025-08-11 | X Detector stroke calibration error (nominal scale error X) | 0.368
| 공주대학교 공동실험실습관 | 2025-08-11 | Y Detector stroke calibration error (nominal scale error Y) | 0.397
| 공주대학교 공동실험실습관 | 2025-08-11 | X Scan Stroke | 104.3
| 공주대학교 공동실험실습관 | 2025-08-11 | Y Scan Stroke | 108.2
| 공주대학교 공동실험실습관 | 2025-08-11 | Full range of X Detector Raw during X Scan sweep | 13.7
| 공주대학교 공동실험실습관 | 2025-08-11 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 공주대학교 공동실험실습관 | 2025-08-11 | X Detector noise (XYServo off) | 0.242
| 공주대학교 공동실험실습관 | 2025-08-11 | Y Detector noise (XYServo off) | 0.251
| 공주대학교 공동실험실습관 | 2025-08-11 | X Detector noise (XY Servo on) | 0.25
| 공주대학교 공동실험실습관 | 2025-08-11 | Y Detector noise (XY Servo on) | 0.258
| 공주대학교 공동실험실습관 | 2025-08-11 | X axis OPM | 0.4598
| 공주대학교 공동실험실습관 | 2025-08-11 | Y axis OPM | 0.8068
| 공주대학교 공동실험실습관 | 2025-08-11 | X Non-linearity | 0.189
| 공주대학교 공동실험실습관 | 2025-08-11 | Y Non-linearity | 0.123
| 한국생산기술연구소 | 2025-08-11 | A-B noise | 7.042
| 한국생산기술연구소 | 2025-08-11 | NCM Amplitude noise | 13.934
| 한국생산기술연구소 | 2025-08-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| 한국생산기술연구소 | 2025-08-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| 한국생산기술연구소 | 2025-08-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 한국생산기술연구소 | 2025-08-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 한국생산기술연구소 | 2025-08-11 | Full range of Z Height during Z Scan sweep | 18
| 한국생산기술연구소 | 2025-08-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.402
| 한국생산기술연구소 | 2025-08-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.935
| 한국생산기술연구소 | 2025-08-11 | Full range of Z Detector during Z Scan sweep | 13.1
| 한국생산기술연구소 | 2025-08-11 | Z Drive calibration error | 0.16092
| 한국생산기술연구소 | 2025-08-11 | Z Height calibration error | 0.047893
| 한국생산기술연구소 | 2025-08-11 | Z linearity by sample height | 1.0048
| 한국생산기술연구소 | 2025-08-11 | 1) Z Scanner position at +3μm | 0.001587
| 한국생산기술연구소 | 2025-08-11 | 3) Z Scanner position at -3μm | 0.000249
| 한국생산기술연구소 | 2025-08-11 | Z Height noise | 0.034
| 한국생산기술연구소 | 2025-08-11 | Full range of X Detector during X Scan sweep | 66.2
| 한국생산기술연구소 | 2025-08-11 | Full range of Y Detector during Y Scan sweep | 64.8
| 한국생산기술연구소 | 2025-08-11 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 93.4
| 한국생산기술연구소 | 2025-08-11 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 92.7
| 한국생산기술연구소 | 2025-08-11 | X Detector stroke calibration error (nominal scale error X) | 0.007
| 한국생산기술연구소 | 2025-08-11 | Y Detector stroke calibration error (nominal scale error Y) | 0.454
| 한국생산기술연구소 | 2025-08-11 | X Scan Stroke | 55
| 한국생산기술연구소 | 2025-08-11 | Y Scan Stroke | 55
| 한국생산기술연구소 | 2025-08-11 | Full range of X Detector Raw during X Scan sweep | 14.2
| 한국생산기술연구소 | 2025-08-11 | Full range of Y Detector Raw during Y Scan sweep | 14
| 한국생산기술연구소 | 2025-08-11 | X Detector noise (XY Servo on) | 0.368
| 한국생산기술연구소 | 2025-08-11 | Y Detector noise (XY Servo on) | 0.367
| 한국생산기술연구소 | 2025-08-11 | X Detector noise (XY Servo off) | 0.383
| 한국생산기술연구소 | 2025-08-11 | Y Detector noise (XY Servo off) | 0.3841
| 한국생산기술연구소 | 2025-08-11 | X axis OPM | 0.812
| 한국생산기술연구소 | 2025-08-11 | Y axis OPM | 0.957
| 한국생산기술연구소 | 2025-08-11 | X Non-linearity | 0.024
| 한국생산기술연구소 | 2025-08-11 | Y Non-linearity | 0.018
| Galaxycore(2) | 2025-08-11 | Z Linearity by Z scanner position | 0.04
| Galaxycore(2) | 2025-08-11 | Full range of Z Detector during Z Scan sweep | 12.97
| Galaxycore(2) | 2025-08-11 | Z Height calibration error | 0.035356
| Galaxycore(2) | 2025-08-11 | Z Drive calibration error | 0.29921
| Galaxycore(2) | 2025-08-11 | Z Linearity by 4 different sample height | 0.99496
| Galaxycore(2) | 2025-08-11 | Full range of Z Height during Z Scan sweep | 43.2
| Galaxycore(2) | 2025-08-11 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 66.2726
| Galaxycore(2) | 2025-08-11 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 26.4034
| Galaxycore(2) | 2025-08-11 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 66.5562
| Galaxycore(2) | 2025-08-11 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.5853
| Galaxycore(2) | 2025-08-11 | Z Height noise | 0.07
| Galaxycore(2) | 2025-08-11 | Full range of Z Height during Z Scan sweep | 16.729
| Galaxycore(2) | 2025-08-11 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.252
| Galaxycore(2) | 2025-08-11 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.0835
| Galaxycore(2) | 2025-08-11 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.2349
| Galaxycore(2) | 2025-08-11 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1362
| Galaxycore(2) | 2025-08-11 | Z Height noise | 0.025
| Galaxycore(2) | 2025-08-11 | Wafer placement(Loading) error on the load port1 (X offset) | 37.75
| Galaxycore(2) | 2025-08-11 | Wafer placement(Loading) error on the load port1 (Y offset) | 23.36
| Galaxycore(2) | 2025-08-11 | Wafer placement(Loading) error on the load port2 (X offset) | 74.46
| Galaxycore(2) | 2025-08-11 | Wafer placement(Loading) error on the load port2 (Y offset) | 96.91
| Galaxycore(2) | 2025-08-11 | OpticalAxisShiftPerFocusX | -0.0051
| Galaxycore(2) | 2025-08-11 | OpticalAxisShiftPerFocusY | -0.0018
| Galaxycore(2) | 2025-08-11 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00021
| Galaxycore(2) | 2025-08-11 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00013
| Galaxycore(2) | 2025-08-11 | Lift pins synchronization offset check (X offset) | 20.9
| Galaxycore(2) | 2025-08-11 | Lift pins synchronization offset check (Y offset) | 6.06
| Galaxycore(2) | 2025-08-11 | Profiler noise @200um | 0.127
| Galaxycore(2) | 2025-08-11 | Roll motion to X+ full scan | 33.4
| Galaxycore(2) | 2025-08-11 | Position repeatability(10repeats, Position Based) | 0.56
| Galaxycore(2) | 2025-08-11 | OPM Repeatability @25mm | 4.242
| Galaxycore(2) | 2025-08-11 | OPM Repeatability @10mm | 1.484
| Galaxycore(2) | 2025-08-11 | OPM Repeatability @5mm | 1.0001
| Galaxycore(2) | 2025-08-11 | OPM Repeatability @1mm | 0.981
| Galaxycore(2) | 2025-08-11 | Max OPM @25mm | 176.317
| Galaxycore(2) | 2025-08-11 | Max OPM @10mm | 27.596
| Galaxycore(2) | 2025-08-11 | Max OPM @5mm | 17.892
| Galaxycore(2) | 2025-08-11 | Max OPM @1mm | 6.97
| Galaxycore(2) | 2025-08-11 | Torque Profile (50mm, 200um/s) | 0.509
| Galaxycore(2) | 2025-08-11 | Torque Profile (50mm, 20um/s) | 0.705
| Galaxycore(2) | 2025-08-11 | Lateral(C-D) sensitivity | 26.7
| Galaxycore(2) | 2025-08-11 | Lateral(C-D) noise | 0.234
| Galaxycore(2) | 2025-08-11 | Min. Vibration Transmissbility (~100Hz) | 0.022
| Galaxycore(2) | 2025-08-11 | RMS of ZeroScan | 0.055
| Galaxycore(2) | 2025-08-11 | LightStrengthGain | 0.41
| Galaxycore(2) | 2025-08-11 | Full range of X Detector during X Scan sweep | 217.3
| Galaxycore(2) | 2025-08-11 | Full range of X1 Detector during X Scan sweep | 217.6
| Galaxycore(2) | 2025-08-11 | Full range of X2 Detector during X Scan sweep | 217
| Galaxycore(2) | 2025-08-11 | Full range of Y Detector during Y Scan sweep | 211.3
| Galaxycore(2) | 2025-08-11 | Full range of Y1 Detector during Y Scan sweep | 211.2
| Galaxycore(2) | 2025-08-11 | Full range of Y2 Detector during Y Scan sweep | 211.4
| Galaxycore(2) | 2025-08-11 | Full range of X1 Detector Raw during X Scan sweep | 13.99
| Galaxycore(2) | 2025-08-11 | Full range of X2 Detector Raw during X Scan sweep | 14.08
| Galaxycore(2) | 2025-08-11 | Full range of Y1 Detector Raw during Y Scan sweep | 14.11
| Galaxycore(2) | 2025-08-11 | Full range of Y2 Detector Raw during Y Scan sweep | 13.92
| Galaxycore(2) | 2025-08-11 | X Detector stroke calibration error (nominal scale error X) | 0.283
| Galaxycore(2) | 2025-08-11 | Y Detector stroke calibration error (nominal scale error Y) | 0.213
| Galaxycore(2) | 2025-08-11 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 302.295
| Galaxycore(2) | 2025-08-11 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 305.528
| Galaxycore(2) | 2025-08-11 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 296.8681
| Galaxycore(2) | 2025-08-11 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 300.9057
| Galaxycore(2) | 2025-08-11 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 181.962
| Galaxycore(2) | 2025-08-11 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 181.51
| Galaxycore(2) | 2025-08-11 | X Non-linearity error | 1.56
| Galaxycore(2) | 2025-08-11 | Y Non-linearity error | 1.54
| Galaxycore(2) | 2025-08-11 | X axis OPM | 1.027
| Galaxycore(2) | 2025-08-11 | Y axis OPM | 0.364
| Galaxycore(2) | 2025-08-11 | X Ringing(Rpv) | 0.208
| Galaxycore(2) | 2025-08-11 | Y Ringing(Rpv) | 0.108
| Galaxycore(2) | 2025-08-11 | X Detector noise (XY Servo on) | 0.38
| Galaxycore(2) | 2025-08-11 | Y Detector noise (XY Servo on) | 0.358
| Galaxycore(2) | 2025-08-11 | X1 Detector noise (XY Servo on) | 0.511
| Galaxycore(2) | 2025-08-11 | Y1 Detector noise (XY Servo on) | 0.502
| Galaxycore(2) | 2025-08-11 | X2 Detector noise (XY Servo on) | 0.522
| Galaxycore(2) | 2025-08-11 | Y2 Detector noise (XY Servo on) | 0.496
| Galaxycore(2) | 2025-08-11 | RMS error after Stage Calibration | 3.06
| Galaxycore(2) | 2025-08-11 | Image offset after moving from -X to +X (X scan) | 22.7051
| Galaxycore(2) | 2025-08-11 | Image offset after moving from -Y to +Y (Y scan) | 112.061
| Galaxycore(2) | 2025-08-11 | Image offset after moving from -X to +X (Y scan) | 15.3809
| Galaxycore(2) | 2025-08-11 | Image offset after moving from -Y to +Y (X scan) | 13.1836
| Galaxycore(2) | 2025-08-11 | Image offset after moving from +X to -X (X scan) | 12.4512
| Galaxycore(2) | 2025-08-11 | Image offset after moving from +Y to -Y (Y scan) | 89.3555
| Galaxycore(2) | 2025-08-11 | Image offset after moving from +X to -X (Y scan) | 21.9727
| Galaxycore(2) | 2025-08-11 | Image offset after moving from +Y to -Y (X scan) | 9.5215
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X offset (Vision Pattern Matching) | 0.669
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.021
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.122
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.21
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X offset (In-Die Align) | 1.309
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y offset (In-Die Align) | 1.479
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X 1σ (In-Die Align) | 0.249
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y 1σ (In-Die Align) | 0.272
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X offset (LLC Translation) | 1.597
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y offset (LLC Translation) | 1.846
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X 1σ (LLC Translation) | 0.294
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y 1σ (LLC Translation) | 0.463
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X offset (Global Align) | 2.607
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y offset (Global Align) | 4.644
| Galaxycore(2) | 2025-08-11 | XY Stage positioning X 1σ (Global Align) | 0.577
| Galaxycore(2) | 2025-08-11 | XY Stage positioning Y 1σ (Global Align) | 1.282
| Galaxycore(2) | 2025-08-11 | Z Drift | 17.675
| Galaxycore(2) | 2025-08-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| Galaxycore(2) | 2025-08-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Galaxycore(2) | 2025-08-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.76
| Galaxycore(2) | 2025-08-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| Galaxycore(2) | 2025-08-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| Galaxycore(2) | 2025-08-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.01
| Galaxycore(2) | 2025-08-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| Galaxycore(2) | 2025-08-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Galaxycore(2) | 2025-08-11 | Cycle time for 15-point measurement | 1825.22
| Galaxycore(2) | 2025-08-11 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.23
| Galaxycore(2) | 2025-08-11 | Roughness Repeatability 1x1um(RSD%) | 3.19
| Galaxycore(2) | 2025-08-11 | Roughness Repeatability 1x1um(1σ) | 0.002
| Galaxycore(2) | 2025-08-11 | Roughness Repeatability 80x80um(RSD%) | 0.587
| Galaxycore(2) | 2025-08-11 | Roughness Repeatability 80x80um(1σ) | 0.001
| GRANDIT#2 | 2025-08-08 | Lateral(C-D) sensitivity | 19.9
| GRANDIT#2 | 2025-08-08 | Lateral(C-D) noise | 0.06
| GRANDIT#2 | 2025-08-08 | A-B noise | 8.467
| GRANDIT#2 | 2025-08-08 | NCM Amplitude noise | 10.216
| GRANDIT#2 | 2025-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| GRANDIT#2 | 2025-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| GRANDIT#2 | 2025-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| GRANDIT#2 | 2025-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| GRANDIT#2 | 2025-08-08 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0
| GRANDIT#2 | 2025-08-08 | Roughness Repeatability 1x1um(RSD%) | 0.533
| GRANDIT#2 | 2025-08-08 | Roughness Repeatability 1x1um(1σ) | 0.002
| GRANDIT#2 | 2025-08-08 | Full range of Z Height during Z Scan sweep | 17
| GRANDIT#2 | 2025-08-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.22
| GRANDIT#2 | 2025-08-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.533
| GRANDIT#2 | 2025-08-08 | Full range of Z Detector during Z Scan sweep | 13
| GRANDIT#2 | 2025-08-08 | Z Drive calibration error | 0.26724
| GRANDIT#2 | 2025-08-08 | Z Height calibration error | 0.18678
| GRANDIT#2 | 2025-08-08 | Z linearity by sample height | 0.99166
| GRANDIT#2 | 2025-08-08 | 1) Z Scanner position at +3μm | 0.007886
| GRANDIT#2 | 2025-08-08 | 3) Z Scanner position at -3μm | 0.014998
| GRANDIT#2 | 2025-08-08 | Z Height noise | 0.027
| GRANDIT#2 | 2025-08-08 | Full range of X Detector during X Scan sweep | 108.1
| GRANDIT#2 | 2025-08-08 | Full range of Y Detector during Y Scan sweep | 109.3
| GRANDIT#2 | 2025-08-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.6
| GRANDIT#2 | 2025-08-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.1
| GRANDIT#2 | 2025-08-08 | X Detector stroke calibration error (nominal scale error X) | 0.116
| GRANDIT#2 | 2025-08-08 | Y Detector stroke calibration error (nominal scale error Y) | 0.035
| GRANDIT#2 | 2025-08-08 | X Scan Stroke | 108.6
| GRANDIT#2 | 2025-08-08 | Y Scan Stroke | 109.7
| GRANDIT#2 | 2025-08-08 | Full range of X Detector Raw during X Scan sweep | 13.6
| GRANDIT#2 | 2025-08-08 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| GRANDIT#2 | 2025-08-08 | X Detector noise (XY Servo on) | 0.321
| GRANDIT#2 | 2025-08-08 | Y Detector noise (XY Servo on) | 0.277
| GRANDIT#2 | 2025-08-08 | X Detector noise (XY servo off) | 0.269
| GRANDIT#2 | 2025-08-08 | Y Detector noise (XY servo off) | 0.267
| GRANDIT#2 | 2025-08-08 | X axis OPM | 1.2864
| GRANDIT#2 | 2025-08-08 | Y axis OPM | 1.2698
| GRANDIT#2 | 2025-08-08 | X Non-linearity | 0.137
| GRANDIT#2 | 2025-08-08 | Y Non-linearity | 0.1
| 강원대학교 전기전자공학과 | 2025-08-06 | A-B noise | 11.566
| 강원대학교 전기전자공학과 | 2025-08-06 | NCM Amplitude noise | 13.528
| 강원대학교 전기전자공학과 | 2025-08-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| 강원대학교 전기전자공학과 | 2025-08-06 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| 강원대학교 전기전자공학과 | 2025-08-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 강원대학교 전기전자공학과 | 2025-08-06 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| 강원대학교 전기전자공학과 | 2025-08-06 | Full range of Z Height during Z Scan sweep | 17.9
| 강원대학교 전기전자공학과 | 2025-08-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.34
| 강원대학교 전기전자공학과 | 2025-08-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.826
| 강원대학교 전기전자공학과 | 2025-08-06 | Full range of Z Detector during Z Scan sweep | 13.1
| 강원대학교 전기전자공학과 | 2025-08-06 | Z Drive calibration error | 0.23946
| 강원대학교 전기전자공학과 | 2025-08-06 | Z Height calibration error | 0.2251
| 강원대학교 전기전자공학과 | 2025-08-06 | NCM driving% with AC160TS cantilever | 0.361
| 강원대학교 전기전자공학과 | 2025-08-06 | Z linearity by sample height | 0.99836
| 강원대학교 전기전자공학과 | 2025-08-06 | 1) Z Scanner position at +3μm | 0.014195
| 강원대학교 전기전자공학과 | 2025-08-06 | 3) Z Scanner position at -3μm | 0.008699
| 강원대학교 전기전자공학과 | 2025-08-06 | Z Height noise | 0.028
| 강원대학교 전기전자공학과 | 2025-08-06 | Full range of X Detector during X Scan sweep | 64
| 강원대학교 전기전자공학과 | 2025-08-06 | Full range of Y Detector during Y Scan sweep | 60
| 강원대학교 전기전자공학과 | 2025-08-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 87.2
| 강원대학교 전기전자공학과 | 2025-08-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 88.8
| 강원대학교 전기전자공학과 | 2025-08-06 | X Detector stroke calibration error (nominal scale error X) | 0.167
| 강원대학교 전기전자공학과 | 2025-08-06 | Y Detector stroke calibration error (nominal scale error Y) | 0.283
| 강원대학교 전기전자공학과 | 2025-08-06 | X Scan Stroke | 55
| 강원대학교 전기전자공학과 | 2025-08-06 | Y Scan Stroke | 55
| 강원대학교 전기전자공학과 | 2025-08-06 | Full range of X Detector Raw during X Scan sweep | 14.7
| 강원대학교 전기전자공학과 | 2025-08-06 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| 강원대학교 전기전자공학과 | 2025-08-06 | X Detector noise (XY Servo on) | 0.391
| 강원대학교 전기전자공학과 | 2025-08-06 | Y Detector noise (XY Servo on) | 0.395
| 강원대학교 전기전자공학과 | 2025-08-06 | X Detector noise (XY Servo off) | 0.439
| 강원대학교 전기전자공학과 | 2025-08-06 | Y Detector noise (XY Servo off) | 0.445
| 강원대학교 전기전자공학과 | 2025-08-06 | X axis OPM | 1.014
| 강원대학교 전기전자공학과 | 2025-08-06 | Y axis OPM | 0.5
| 강원대학교 전기전자공학과 | 2025-08-06 | X Non-linearity | 0.062
| 강원대학교 전기전자공학과 | 2025-08-06 | Y Non-linearity | 0.032
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Lateral(C-D) sensitivity | 80.3
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Lateral(C-D) noise | 0.088
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | A-B noise | 5.603
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | NCM Amplitude noise | 9.398
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.26
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.32
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.28
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Roughness Repeatability 1x1um(RSD%) | 1.098
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Roughness Repeatability 1x1um(1σ) | 0.002
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Full range of Z Height during Z Scan sweep | 18.2
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.986
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.88
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Full range of Z Detector during Z Scan sweep | 13.5
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Z Drive calibration error | 0.39707
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Z Height calibration error | 0.25268
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Z linearity by sample height | 0.99558
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | 1) Z Scanner position at +3μm | 0.011088
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | 3) Z Scanner position at -3μm | 0.004614
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Z Height noise | 0.027
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Full range of X Detector during X Scan sweep | 134.2
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Full range of Y Detector during Y Scan sweep | 135
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 190
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 201.6
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | X Detector stroke calibration error (nominal scale error X) | 0.269
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Y Detector stroke calibration error (nominal scale error Y) | 0.1
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Resonance Frequency in DB | 600
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Servo Cut-off Frequency in DB | 180
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Full range of X Detector Raw during X Scan sweep | 14.1
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | X Detector noise (XY Servo on) | 0.278
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Y Detector noise (XY Servo on) | 0.303
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | X Detector noise (XY servo off) | 0.268
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Y Detector noise (XY servo off) | 0.286
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | X axis OPM | 1.4268
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Y axis OPM | 0.9296
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | X Non-linearity | 0.236
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Y Non-linearity | 0.158
| Hon Hai Technology Group (FOXCONN) | 2025-08-06 | Check chuck surface no damage | Checked
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Lateral(C-D) sensitivity | 10.1
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Lateral(C-D) noise | 0.112
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | A-B noise | 5.588
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | NCM Amplitude noise | 9.729
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.33
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.39
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.35
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.39
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.37
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Roughness Repeatability 1x1um(RSD%) | 3.172
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Roughness Repeatability 1x1um(1σ) | 0.004
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Z Height during Z Scan sweep | 18.1
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.32
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.006
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Z Detector during Z Scan sweep | 12.4
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Z Drive calibration error | 0.25575
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Z Height calibration error | 0.062261
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | NCM driving% with AC160TS cantilever | 0.596
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Z linearity by sample height | 1.0064
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | 1) Z Scanner position at +3μm | 0.001404
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | 3) Z Scanner position at -3μm | 0.00101
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Z Height noise | 0.03
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of X Detector during X Scan sweep | 177.1
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of X1 Detector during X Scan sweep | 181.8
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of X2 Detector during X Scan sweep | 172.3
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Y Detector during Y Scan sweep | 179
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Y1 Detector during Y Scan sweep | 177.1
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Y2 Detector during Y Scan sweep | 180.9
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 250.8
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 235.5
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 240
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 244
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X Detector stroke calibration error (nominal scale error X) | 0.151
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y Detector stroke calibration error (nominal scale error Y) | 0.274
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X Scan Stroke | 176.2
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y Scan Stroke | 178.5
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of X1 Detector Raw during X Scan sweep | 14.5
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of X2 Detector Raw during X Scan sweep | 14.6
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Y1 Detector Raw during Y Scan sweep | 14.7
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Full range of Y2 Detector Raw during Y Scan sweep | 14.8
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X Detector noise (XY servo off) | 0.224
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X1 Detector noise (XY servo off) | 0.374
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X2 Detector noise (XY servo off) | 0.235
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y Detector noise (XY servo off) | 0.238
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y1 Detector noise (XY servo off) | 0.326
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y2 Detector noise (XY servo off) | 0.337
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X Detector noise (XY servo on) | 0.24
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X1 Detector noise (XY servo on) | 0.386
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X2 Detector noise (XY servo on) | 0.26
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y Detector noise (XY servo on) | 0.261
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y1 Detector noise (XY servo on) | 0.354
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y2 Detector noise (XY servo on) | 0.357
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X axis OPM | 0.8848
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y axis OPM | 1.173
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | X Non-linearity | 0.292
| National Research Council Institute for Microelectronics and Microsystems | 2025-08-05 | Y Non-linearity | 0.234
| RATC Demo | 2025-08-05 | Lateral(C-D) sensitivity | 15.2
| RATC Demo | 2025-08-05 | Lateral(C-D) noise | 0.054
| RATC Demo | 2025-08-05 | A-B noise | 7.076
| RATC Demo | 2025-08-05 | NCM Amplitude noise | 12.379
| RATC Demo | 2025-08-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| RATC Demo | 2025-08-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| RATC Demo | 2025-08-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| RATC Demo | 2025-08-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| RATC Demo | 2025-08-05 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0
| RATC Demo | 2025-08-05 | Roughness Repeatability 1x1um(RSD%) | 0.371
| RATC Demo | 2025-08-05 | Roughness Repeatability 1x1um(1σ) | 0.001
| RATC Demo | 2025-08-05 | Full range of Z Height during Z Scan sweep | 18.8
| RATC Demo | 2025-08-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.376
| RATC Demo | 2025-08-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.417
| RATC Demo | 2025-08-05 | Full range of Z Detector during Z Scan sweep | 13.2
| RATC Demo | 2025-08-05 | Z Height calibration error | 0.15517
| RATC Demo | 2025-08-05 | Z linearity by sample height | 1.0113
| RATC Demo | 2025-08-05 | 1) Z Scanner position at +3μm | 0.002916
| RATC Demo | 2025-08-05 | 3) Z Scanner position at -3μm | 0.006863
| RATC Demo | 2025-08-05 | Z Height noise | 0.034
| RATC Demo | 2025-08-05 | Full range of X Detector during X Scan sweep | 69.3
| RATC Demo | 2025-08-05 | Full range of Y Detector during Y Scan sweep | 70.8
| RATC Demo | 2025-08-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.8
| RATC Demo | 2025-08-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.6
| RATC Demo | 2025-08-05 | X Detector stroke calibration error (nominal scale error X) | 0.136
| RATC Demo | 2025-08-05 | Y Detector stroke calibration error (nominal scale error Y) | 0.225
| RATC Demo | 2025-08-05 | X Scan Stroke | 55
| RATC Demo | 2025-08-05 | Y Scan Stroke | 55
| RATC Demo | 2025-08-05 | Full range of X Detector Raw during X Scan sweep | 14
| RATC Demo | 2025-08-05 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| RATC Demo | 2025-08-05 | X Detector noise (XY Servo on) | 0.249
| RATC Demo | 2025-08-05 | Y Detector noise (XY Servo on) | 0.254
| RATC Demo | 2025-08-05 | X Detector noise (XY Servo off) | 0.242
| RATC Demo | 2025-08-05 | Y Detector noise (XY Servo off) | 0.246
| RATC Demo | 2025-08-05 | X axis OPM | 0.6436
| RATC Demo | 2025-08-05 | Y axis OPM | 0.5688
| RATC Demo | 2025-08-05 | X Non-linearity | 0.06
| RATC Demo | 2025-08-05 | Y Non-linearity | 0.055
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Linearity by Z scanner position | 0.69
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Z Detector during Z Scan sweep | 12.42
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Height calibration error | 0.010624
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Drive calibration error | 0.19688
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Linearity by 4 different sample height | 0.99535
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Z Height during Z Scan sweep | 17.03
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.4822
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.0056
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.515
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.236
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Height noise | 0.027
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Wafer placement(Loading) error on the load port1 (X offset) | 52.75
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Wafer placement(Loading) error on the load port1 (Y offset) | 21.6
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Wafer placement(Loading) error on the load port2 (X offset) | 18.45
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Wafer placement(Loading) error on the load port2 (Y offset) | 34.1
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OpticalAxisShiftPerFocusX | -0.0087
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OpticalAxisShiftPerFocusY | 0.0037
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00022
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00023
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Lift pins synchronization offset check (X offset) | 33.8
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Lift pins synchronization offset check (Y offset) | 37.16
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Profiler noise @200um | 0.109
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Roll motion to X+ full scan | 9.7
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Position repeatability(10repeats, Position Based) | 0.62
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OPM Repeatability @25mm | 2.408
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OPM Repeatability @10mm | 1.663
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OPM Repeatability @5mm | 0.916
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | OPM Repeatability @1mm | 0.533
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Max OPM @25mm | 51.091
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Max OPM @10mm | 34.426
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Max OPM @5mm | 18.829
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Max OPM @1mm | 7.094
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Torque Profile (50mm, 200um/s) | 0.629
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Torque Profile (50mm, 20um/s) | 0.689
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Lateral(C-D) sensitivity | 18.2
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Lateral(C-D) noise | 0.189
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Min. Vibration Transmissbility (~100Hz) | 0.018
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | RMS of ZeroScan | 0.041
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | LightStrengthGain | 0.5
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of X Detector during X Scan sweep | 197.1804
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of X1 Detector during X Scan sweep | 197.1804
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of X2 Detector during X Scan sweep | 201.2361
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Y Detector during Y Scan sweep | 201.505
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Y1 Detector during Y Scan sweep | 202.8088
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Y2 Detector during Y Scan sweep | 201.505
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of X1 Detector Raw during X Scan sweep | 14.55
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of X2 Detector Raw during X Scan sweep | 14.05
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.57
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Full range of Y2 Detector Raw during Y Scan sweep | 14.46
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X Detector stroke calibration error (nominal scale error X) | 0.2
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y Detector stroke calibration error (nominal scale error Y) | 0.026
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 274.83
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 289.933
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 278.629
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 281.727
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 178.782
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 181.45
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X Non-linearity error | 2.99
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y Non-linearity error | 2.79
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X axis OPM | 0.611
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y axis OPM | 1.56
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X Ringing(Rpv) | 0.191
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y Ringing(Rpv) | 0.226
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X Detector noise (XY Servo on) | 0.345
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y Detector noise (XY Servo on) | 0.321
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X1 Detector noise (XY Servo on) | 0.439
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y1 Detector noise (XY Servo on) | 0.501
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | X2 Detector noise (XY Servo on) | 0.468
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Y2 Detector noise (XY Servo on) | 0.343
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | RMS error after Stage Calibration | 1.47
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from -X to +X (X scan) | 48.3398
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from -Y to +Y (Y scan) | 27.0996
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from -X to +X (Y scan) | 19.043
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from -Y to +Y (X scan) | 19.7754
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from +X to -X (X scan) | 59.3262
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from +Y to -Y (Y scan) | 25.6348
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from +X to -X (Y scan) | 46.1426
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Image offset after moving from +Y to -Y (X scan) | 35.8887
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X offset (Vision Pattern Matching) | 1.748
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.5186
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.42866
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3088
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X offset (In-Die Align) | 2.4219
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y offset (In-Die Align) | 3.0469
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X 1σ (In-Die Align) | 0.51888
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y 1σ (In-Die Align) | 0.71406
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X offset (LLC Translation) | 2.417
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y offset (LLC Translation) | 1.8848
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X 1σ (LLC Translation) | 0.48489
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y 1σ (LLC Translation) | 0.38112
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X offset (Global Align) | 2.4365
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y offset (Global Align) | 3.4277
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning X 1σ (Global Align) | 0.59313
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | XY Stage positioning Y 1σ (Global Align) | 1.0076
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Z Drift | 19.885
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13148
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.023703
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.69332
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12488
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.73388
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13201
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.81073
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14596
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Cycle time for 15-point measurement | 1577.877
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.331
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Roughness Repeatability 1x1um(RSD%) | 1.527
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Roughness Repeatability 1x1um(1σ) | 0.002
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Roughness Repeatability 80x80um(RSD%) | 1.2
| Sony Semiconductor Manufacturing Corporation Nagasaki | 2025-08-04 | Roughness Repeatability 80x80um(1σ) | 0.002
| University of Padova | 2025-08-01 | A-B noise | 7.13
| University of Padova | 2025-08-01 | NCM Amplitude noise | 14.96
| University of Padova | 2025-08-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| University of Padova | 2025-08-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| University of Padova | 2025-08-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| University of Padova | 2025-08-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| University of Padova | 2025-08-01 | Full range of Z Height during Z Scan sweep | 17.8
| University of Padova | 2025-08-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.593
| University of Padova | 2025-08-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.051
| University of Padova | 2025-08-01 | Full range of Z Detector during Z Scan sweep | 12.9
| University of Padova | 2025-08-01 | Z Drive calibration error | 0.061303
| University of Padova | 2025-08-01 | Z Height calibration error | 0.21456
| University of Padova | 2025-08-01 | Z linearity by sample height | 1.0124
| University of Padova | 2025-08-01 | 1) Z Scanner position at +3μm | 0.006029
| University of Padova | 2025-08-01 | 3) Z Scanner position at -3μm | 0.001284
| University of Padova | 2025-08-01 | Z Height noise | 0.03
| University of Padova | 2025-08-01 | Full range of X Detector during X Scan sweep | 68.1
| University of Padova | 2025-08-01 | Full range of Y Detector during Y Scan sweep | 66.9
| University of Padova | 2025-08-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 96.2
| University of Padova | 2025-08-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.2
| University of Padova | 2025-08-01 | X Detector stroke calibration error (nominal scale error X) | 0.192
| University of Padova | 2025-08-01 | Y Detector stroke calibration error (nominal scale error Y) | 0.243
| University of Padova | 2025-08-01 | X Scan Stroke | 55
| University of Padova | 2025-08-01 | Y Scan Stroke | 55
| University of Padova | 2025-08-01 | Full range of X Detector Raw during X Scan sweep | 14.2
| University of Padova | 2025-08-01 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| University of Padova | 2025-08-01 | X Detector noise (XY Servo on) | 0.368
| University of Padova | 2025-08-01 | Y Detector noise (XY Servo on) | 0.377
| University of Padova | 2025-08-01 | X Detector noise (XY Servo off) | 0.397
| University of Padova | 2025-08-01 | Y Detector noise (XY Servo off) | 0.408
| University of Padova | 2025-08-01 | X axis OPM | 0.969
| University of Padova | 2025-08-01 | Y axis OPM | 0.7746
| University of Padova | 2025-08-01 | X Non-linearity | 0.033
| University of Padova | 2025-08-01 | Y Non-linearity | 0.041
| 강원대학교_공동실험실습관 | 2025-07-30 | SLD current limit check | 104
| 강원대학교_공동실험실습관 | 2025-07-30 | SLD beam power at the end of the lens | 0.27
| 강원대학교_공동실험실습관 | 2025-07-30 | Lateral(C-D) sensitivity | 11.9
| 강원대학교_공동실험실습관 | 2025-07-30 | Lateral(C-D) noise | 0.112
| 강원대학교_공동실험실습관 | 2025-07-30 | A-B noise | 14.33
| 강원대학교_공동실험실습관 | 2025-07-30 | NCM Amplitude noise | 10.55
| 강원대학교_공동실험실습관 | 2025-07-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 강원대학교_공동실험실습관 | 2025-07-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 강원대학교_공동실험실습관 | 2025-07-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 강원대학교_공동실험실습관 | 2025-07-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 강원대학교_공동실험실습관 | 2025-07-30 | Roughness Repeatability 1x1um(RSD%) | 0.82
| 강원대학교_공동실험실습관 | 2025-07-30 | Roughness Repeatability 1x1um(1σ) | 0.002
| 강원대학교_공동실험실습관 | 2025-07-30 | Full range of Z Height during Z Scan sweep | 17.2
| 강원대학교_공동실험실습관 | 2025-07-30 | Full range of Z Detector during Z Scan sweep | 14.5
| 강원대학교_공동실험실습관 | 2025-07-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.692
| 강원대학교_공동실험실습관 | 2025-07-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.204
| 강원대학교_공동실험실습관 | 2025-07-30 | Z Drive calibrtaion error | 0.076628
| 강원대학교_공동실험실습관 | 2025-07-30 | Z Height calibrtaion error | 0.047893
| 강원대학교_공동실험실습관 | 2025-07-30 | 2) of 20nm | 2.963
| 강원대학교_공동실험실습관 | 2025-07-30 | Z linearity by sample height | 1.0011
| 강원대학교_공동실험실습관 | 2025-07-30 | 1) Z Scanner position at +3μm | 0.008272
| 강원대학교_공동실험실습관 | 2025-07-30 | 3) Z Scanner position at -3μm | 0.004973
| 강원대학교_공동실험실습관 | 2025-07-30 | Z Height noise | 0.028
| 강원대학교_공동실험실습관 | 2025-07-30 | Full range of X Detector during X Scan sweep | 104.7
| 강원대학교_공동실험실습관 | 2025-07-30 | Full range of Y Detector during Y Scan sweep | 103.6
| 강원대학교_공동실험실습관 | 2025-07-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 146.2
| 강원대학교_공동실험실습관 | 2025-07-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 149.5
| 강원대학교_공동실험실습관 | 2025-07-30 | X Detector stroke calibration error (nominal scale error X) | 0.106
| 강원대학교_공동실험실습관 | 2025-07-30 | Y Detector stroke calibration error (nominal scale error Y) | 0.457
| 강원대학교_공동실험실습관 | 2025-07-30 | X Scan Stroke | 104.7
| 강원대학교_공동실험실습관 | 2025-07-30 | Y Scan Stroke | 103.6
| 강원대학교_공동실험실습관 | 2025-07-30 | Full range of X Detector Raw during X Scan sweep | 14.3
| 강원대학교_공동실험실습관 | 2025-07-30 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 강원대학교_공동실험실습관 | 2025-07-30 | X Detector noise (XYServo off) | 0.24
| 강원대학교_공동실험실습관 | 2025-07-30 | Y Detector noise (XYServo off) | 0.24
| 강원대학교_공동실험실습관 | 2025-07-30 | X Detector noise (XY Servo on) | 0.25
| 강원대학교_공동실험실습관 | 2025-07-30 | Y Detector noise (XY Servo on) | 0.24
| 강원대학교_공동실험실습관 | 2025-07-30 | X axis OPM | 1.42
| 강원대학교_공동실험실습관 | 2025-07-30 | Y axis OPM | 0.92
| 강원대학교_공동실험실습관 | 2025-07-30 | X Non-linearity | 0.13
| 강원대학교_공동실험실습관 | 2025-07-30 | Y Non-linearity | 0.15
| Intel Hillsboro#4 | 2025-07-29 | Z Linearity by Z scanner position | 0.6
| Intel Hillsboro#4 | 2025-07-29 | Z Linearity by 4 different sample height | 0.99729
| Intel Hillsboro#4 | 2025-07-29 | Full range of Z Detector during Z Scan sweep | 12.78
| Intel Hillsboro#4 | 2025-07-29 | Z Height calibration error | 0.001951
| Intel Hillsboro#4 | 2025-07-29 | Z Drive calibration error | 0.20078
| Intel Hillsboro#4 | 2025-07-29 | Full range of Z Height during Z Scan sweep | 42.7
| Intel Hillsboro#4 | 2025-07-29 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 62.7835
| Intel Hillsboro#4 | 2025-07-29 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.7453
| Intel Hillsboro#4 | 2025-07-29 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 67.33
| Intel Hillsboro#4 | 2025-07-29 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.9783
| Intel Hillsboro#4 | 2025-07-29 | Z Height noise | 0.082
| Intel Hillsboro#4 | 2025-07-29 | Wafer placement(Loading) error on the load port1 (X offset) | 32.3
| Intel Hillsboro#4 | 2025-07-29 | Wafer placement(Loading) error on the load port1 (Y offset) | 20
| Intel Hillsboro#4 | 2025-07-29 | Wafer placement(Loading) error on the load port2 (X offset) | 52.8
| Intel Hillsboro#4 | 2025-07-29 | Wafer placement(Loading) error on the load port2 (Y offset) | 14.8
| Intel Hillsboro#4 | 2025-07-29 | OpticalAxisShiftPerFocusX | -0.0053
| Intel Hillsboro#4 | 2025-07-29 | OpticalAxisShiftPerFocusY | -0.0017
| Intel Hillsboro#4 | 2025-07-29 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 3e-05
| Intel Hillsboro#4 | 2025-07-29 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| Intel Hillsboro#4 | 2025-07-29 | Lift pins synchronization offset check (X offset) | 0.37
| Intel Hillsboro#4 | 2025-07-29 | Lift pins synchronization offset check (Y offset) | 42.16
| Intel Hillsboro#4 | 2025-07-29 | Max OPM @25mm | 117.38
| Intel Hillsboro#4 | 2025-07-29 | Max OPM @10mm | 27.425
| Intel Hillsboro#4 | 2025-07-29 | Max OPM @5mm | 22.587
| Intel Hillsboro#4 | 2025-07-29 | Max OPM @1mm | 7.359
| Intel Hillsboro#4 | 2025-07-29 | OPM Repeatability @25mm | 2.011
| Intel Hillsboro#4 | 2025-07-29 | OPM Repeatability @10mm | 1.229
| Intel Hillsboro#4 | 2025-07-29 | OPM Repeatability @5mm | 1.041
| Intel Hillsboro#4 | 2025-07-29 | OPM Repeatability @1mm | 0.633
| Intel Hillsboro#4 | 2025-07-29 | Torque Profile (50mm, 200um/s) | 0.569
| Intel Hillsboro#4 | 2025-07-29 | Torque Profile (50mm, 20um/s) | 0.902
| Intel Hillsboro#4 | 2025-07-29 | Profiler noise @200um | 0.162
| Intel Hillsboro#4 | 2025-07-29 | Roll motion to X+ full scan | 6
| Intel Hillsboro#4 | 2025-07-29 | Position repeatability(10repeats, Position Based) | 0.3
| Intel Hillsboro#4 | 2025-07-29 | Lateral(C-D) sensitivity | 76.6
| Intel Hillsboro#4 | 2025-07-29 | Lateral(C-D) noise | 0.172
| Intel Hillsboro#4 | 2025-07-29 | Min. Vibration Transmissbility (~100Hz) | 0.025
| Intel Hillsboro#4 | 2025-07-29 | RMS of ZeroScan | 0.052
| Intel Hillsboro#4 | 2025-07-29 | Full range of X Detector during X Scan sweep | 194.02
| Intel Hillsboro#4 | 2025-07-29 | Full range of X1 Detector during X Scan sweep | 191.48
| Intel Hillsboro#4 | 2025-07-29 | Full range of X2 Detector during X Scan sweep | 194.02
| Intel Hillsboro#4 | 2025-07-29 | Full range of Y Detector during Y Scan sweep | 179.04
| Intel Hillsboro#4 | 2025-07-29 | Full range of Y1 Detector during Y Scan sweep | 179.04
| Intel Hillsboro#4 | 2025-07-29 | Full range of Y2 Detector during Y Scan sweep | 176.95
| Intel Hillsboro#4 | 2025-07-29 | Full range of X1 Detector Raw during X Scan sweep | 14.21
| Intel Hillsboro#4 | 2025-07-29 | Full range of X2 Detector Raw during X Scan sweep | 14.55
| Intel Hillsboro#4 | 2025-07-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.18
| Intel Hillsboro#4 | 2025-07-29 | Full range of Y2 Detector Raw during Y Scan sweep | 14.13
| Intel Hillsboro#4 | 2025-07-29 | X Detector stroke calibration error (nominal scale error X) | 0.3
| Intel Hillsboro#4 | 2025-07-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.447
| Intel Hillsboro#4 | 2025-07-29 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 267.84
| Intel Hillsboro#4 | 2025-07-29 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 273.85
| Intel Hillsboro#4 | 2025-07-29 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 265.16
| Intel Hillsboro#4 | 2025-07-29 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 274.48
| Intel Hillsboro#4 | 2025-07-29 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 175.971
| Intel Hillsboro#4 | 2025-07-29 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 177.983
| Intel Hillsboro#4 | 2025-07-29 | XY Orthogonality | 90.293
| Intel Hillsboro#4 | 2025-07-29 | X Non-linearity error | 3.07
| Intel Hillsboro#4 | 2025-07-29 | Y Non-linearity error | 2.74
| Intel Hillsboro#4 | 2025-07-29 | X axis OPM | 0.904
| Intel Hillsboro#4 | 2025-07-29 | Y axis OPM | 0.557
| Intel Hillsboro#4 | 2025-07-29 | X Ringing(Rpv) | 0.144
| Intel Hillsboro#4 | 2025-07-29 | Y Ringing(Rpv) | 0.204
| Intel Hillsboro#4 | 2025-07-29 | X Detector noise (XY Servo on) | 0.306
| Intel Hillsboro#4 | 2025-07-29 | Y Detector noise (XY Servo on) | 0.435
| Intel Hillsboro#4 | 2025-07-29 | X1 Detector noise (XY Servo on) | 0.359
| Intel Hillsboro#4 | 2025-07-29 | Y1 Detector noise (XY Servo on) | 0.466
| Intel Hillsboro#4 | 2025-07-29 | X2 Detector noise (XY Servo on) | 0.401
| Intel Hillsboro#4 | 2025-07-29 | Y2 Detector noise (XY Servo on) | 0.49
| Intel Hillsboro#4 | 2025-07-29 | RMS error after Stage Calibration | 1.78
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from -X to +X (X scan) | 13.916
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from -Y to +Y (Y scan) | 76.9043
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from -X to +X (Y scan) | 144.287
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from -Y to +Y (X scan) | 18.31
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from +X to -X (X scan) | 40.2832
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from +Y to -Y (Y scan) | 145.0195
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from +X to -X (Y scan) | 65.917
| Intel Hillsboro#4 | 2025-07-29 | Image offset after moving from +Y to -Y (X scan) | 15.38
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X offset (Vision Pattern Matching) | 1.1719
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.0352
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.2674
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.1635
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X offset (In-Die Align) | 1.8408
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y offset (In-Die Align) | 1.8311
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X 1σ (In-Die Align) | 0.3609
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y 1σ (In-Die Align) | 0.4224
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X offset (LLC Translation) | 1.7676
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y offset (LLC Translation) | 1.5967
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X 1σ (LLC Translation) | 0.3857
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y 1σ (LLC Translation) | 0.3004
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X offset (Global Align) | 2.4902
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y offset (Global Align) | 4.4678
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning X 1σ (Global Align) | 0.6277
| Intel Hillsboro#4 | 2025-07-29 | XY Stage positioning Y 1σ (Global Align) | 1.3999
| Intel Hillsboro#4 | 2025-07-29 | Z Drift | 3.512
| Intel Hillsboro#4 | 2025-07-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| Intel Hillsboro#4 | 2025-07-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| Intel Hillsboro#4 | 2025-07-29 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| Intel Hillsboro#4 | 2025-07-29 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Intel Hillsboro#4 | 2025-07-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.2
| Intel Hillsboro#4 | 2025-07-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Intel Hillsboro#4 | 2025-07-29 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.64
| Intel Hillsboro#4 | 2025-07-29 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| Intel Hillsboro#4 | 2025-07-29 | Cycle time for 15-point measurement | 1652.389
| Intel Hillsboro#4 | 2025-07-29 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.368
| Intel Hillsboro#4 | 2025-07-29 | Roughness Repeatability 1x1um(RSD%) | 0.882
| Intel Hillsboro#4 | 2025-07-29 | Roughness Repeatability 1x1um(1σ) | 0.001
| Intel Hillsboro#4 | 2025-07-29 | Roughness Repeatability 80x80um(RSD%) | 1.102
| Intel Hillsboro#4 | 2025-07-29 | Roughness Repeatability 80x80um(1σ) | 0.002
| Intel Hillsboro#4 | 2025-07-29 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer) | 0.88
| Intel Hillsboro#4 | 2025-07-29 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer) | 0.16
| Intel Hillsboro#4 | 2025-07-29 | Z Detector Max | 0.42108
| Intel Hillsboro#4 | 2025-07-29 | Z Detector Min | -0.31257
| Intel Hillsboro#4 | 2025-07-29 | WLI_Scale(Z Calibration) | 27.7
| University of North Carolina | 2025-07-25 | Lateral(C-D) sensitivity | 28.1
| University of North Carolina | 2025-07-25 | Lateral(C-D) noise | 0.086
| University of North Carolina | 2025-07-25 | A-B noise | 7.661
| University of North Carolina | 2025-07-25 | NCM Amplitude noise | 10.218
| University of North Carolina | 2025-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| University of North Carolina | 2025-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| University of North Carolina | 2025-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| University of North Carolina | 2025-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of North Carolina | 2025-07-25 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0
| University of North Carolina | 2025-07-25 | Roughness Repeatability 1x1um(RSD%) | 2.816
| University of North Carolina | 2025-07-25 | Roughness Repeatability 1x1um(1σ) | 0.006
| University of North Carolina | 2025-07-25 | Full range of X Detector during X Scan sweep | 72.3
| University of North Carolina | 2025-07-25 | Full range of Y Detector during Y Scan sweep | 73.7
| University of North Carolina | 2025-07-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.8
| University of North Carolina | 2025-07-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 104
| University of North Carolina | 2025-07-25 | X Detector stroke calibration error (nominal scale error X) | 0.019
| University of North Carolina | 2025-07-25 | Y Detector stroke calibration error (nominal scale error Y) | 0.179
| University of North Carolina | 2025-07-25 | X Scan Stroke | 55
| University of North Carolina | 2025-07-25 | Y Scan Stroke | 55
| University of North Carolina | 2025-07-25 | Full range of X Detector Raw during X Scan sweep | 13.8
| University of North Carolina | 2025-07-25 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| University of North Carolina | 2025-07-25 | X Detector noise (XY Servo on) | 0.256
| University of North Carolina | 2025-07-25 | Y Detector noise (XY Servo on) | 0.254
| University of North Carolina | 2025-07-25 | X Detector noise (XY Servo off) | 0.247
| University of North Carolina | 2025-07-25 | Y Detector noise (XY Servo off) | 0.244
| University of North Carolina | 2025-07-25 | X axis OPM | 1.2276
| University of North Carolina | 2025-07-25 | Y axis OPM | 0.779
| University of North Carolina | 2025-07-25 | X Non-linearity | 0.031
| University of North Carolina | 2025-07-25 | Y Non-linearity | 0.033
| University of California at Berkeley | 2025-07-25 | Lateral(C-D) sensitivity | 14.1
| University of California at Berkeley | 2025-07-25 | Lateral(C-D) noise | 0.084
| University of California at Berkeley | 2025-07-25 | A-B noise | 9.105
| University of California at Berkeley | 2025-07-25 | NCM Amplitude noise | 16.037
| University of California at Berkeley | 2025-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| University of California at Berkeley | 2025-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| University of California at Berkeley | 2025-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| University of California at Berkeley | 2025-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| University of California at Berkeley | 2025-07-25 | Roughness Repeatability 1x1um(RSD%) | 1.555
| University of California at Berkeley | 2025-07-25 | Roughness Repeatability 1x1um(1σ) | 0.002
| University of California at Berkeley | 2025-07-25 | Full range of Z Height during Z Scan sweep | 17.3
| University of California at Berkeley | 2025-07-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.185
| University of California at Berkeley | 2025-07-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.993
| University of California at Berkeley | 2025-07-25 | Full range of Z Detector during Z Scan sweep | 12.3
| University of California at Berkeley | 2025-07-25 | Z Drive calibration error | 0.38027
| University of California at Berkeley | 2025-07-25 | Z Height calibration error | 0.42816
| University of California at Berkeley | 2025-07-25 | Z linearity by sample height | 1.0185
| University of California at Berkeley | 2025-07-25 | 1) Z Scanner position at +3μm | 0.004085
| University of California at Berkeley | 2025-07-25 | 3) Z Scanner position at -3μm | 0.000758
| University of California at Berkeley | 2025-07-25 | Z Height noise | 0.028
| University of California at Berkeley | 2025-07-25 | Full range of X Detector during X Scan sweep | 66.1
| University of California at Berkeley | 2025-07-25 | Full range of Y Detector during Y Scan sweep | 69.1
| University of California at Berkeley | 2025-07-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99
| University of California at Berkeley | 2025-07-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 97.2
| University of California at Berkeley | 2025-07-25 | X Detector stroke calibration error (nominal scale error X) | 0.107
| University of California at Berkeley | 2025-07-25 | Y Detector stroke calibration error (nominal scale error Y) | 0.274
| University of California at Berkeley | 2025-07-25 | X Scan Stroke | 55
| University of California at Berkeley | 2025-07-25 | Y Scan Stroke | 55
| University of California at Berkeley | 2025-07-25 | Full range of X Detector Raw during X Scan sweep | 13.3
| University of California at Berkeley | 2025-07-25 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| University of California at Berkeley | 2025-07-25 | X Detector noise (XY Servo on) | 0.245
| University of California at Berkeley | 2025-07-25 | Y Detector noise (XY Servo on) | 0.245
| University of California at Berkeley | 2025-07-25 | X Detector noise (XY Servo off) | 0.238
| University of California at Berkeley | 2025-07-25 | Y Detector noise (XY Servo off) | 0.238
| University of California at Berkeley | 2025-07-25 | X axis OPM | 0.9662
| University of California at Berkeley | 2025-07-25 | Y axis OPM | 0.3594
| University of California at Berkeley | 2025-07-25 | X Non-linearity | 0.089
| University of California at Berkeley | 2025-07-25 | Y Non-linearity | 0.141
| Analog Devices Inc | 2025-07-24 | SLD current limit check | 100
| Analog Devices Inc | 2025-07-24 | SLD beam power at the end of the lens | 0.24
| Analog Devices Inc | 2025-07-24 | Lateral(C-D) sensitivity | 23.7
| Analog Devices Inc | 2025-07-24 | Lateral(C-D) noise | 0.117
| Analog Devices Inc | 2025-07-24 | A-B noise | 6.034
| Analog Devices Inc | 2025-07-24 | NCM Amplitude noise | 12.32
| Analog Devices Inc | 2025-07-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Analog Devices Inc | 2025-07-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Analog Devices Inc | 2025-07-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Analog Devices Inc | 2025-07-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Analog Devices Inc | 2025-07-24 | Roughness Repeatability 1x1um(RSD%) | 1.972
| Analog Devices Inc | 2025-07-24 | Roughness Repeatability 1x1um(1σ) | 0.002
| Analog Devices Inc | 2025-07-24 | Full range of Z Height during Z Scan sweep | 17.6
| Analog Devices Inc | 2025-07-24 | Full range of Z Detector during Z Scan sweep | 14.6
| Analog Devices Inc | 2025-07-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.089
| Analog Devices Inc | 2025-07-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.487
| Analog Devices Inc | 2025-07-24 | Z Drive calibrtaion error | 0.20115
| Analog Devices Inc | 2025-07-24 | Z Height calibrtaion error | 0.2864
| Analog Devices Inc | 2025-07-24 | 2) of 20nm | 0.517
| Analog Devices Inc | 2025-07-24 | Z linearity by sample height | 1.0112
| Analog Devices Inc | 2025-07-24 | 1) Z Scanner position at +3μm | 0.0006
| Analog Devices Inc | 2025-07-24 | 3) Z Scanner position at -3μm | 0.00167
| Analog Devices Inc | 2025-07-24 | Z Height noise | 0.017
| Analog Devices Inc | 2025-07-24 | Full range of X Detector during X Scan sweep | 101.2
| Analog Devices Inc | 2025-07-24 | Full range of Y Detector during Y Scan sweep | 101.6
| Analog Devices Inc | 2025-07-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 142.8
| Analog Devices Inc | 2025-07-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 147.9
| Analog Devices Inc | 2025-07-24 | X Detector stroke calibration error (nominal scale error X) | 0.226
| Analog Devices Inc | 2025-07-24 | Y Detector stroke calibration error (nominal scale error Y) | 0.132
| Analog Devices Inc | 2025-07-24 | X Scan Stroke | 101.2
| Analog Devices Inc | 2025-07-24 | Y Scan Stroke | 101.6
| Analog Devices Inc | 2025-07-24 | Full range of X Detector Raw during X Scan sweep | 14.1
| Analog Devices Inc | 2025-07-24 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Analog Devices Inc | 2025-07-24 | X Detector noise (XYServo off) | 0.234
| Analog Devices Inc | 2025-07-24 | Y Detector noise (XYServo off) | 0.238
| Analog Devices Inc | 2025-07-24 | X Detector noise (XY Servo on) | 0.242
| Analog Devices Inc | 2025-07-24 | Y Detector noise (XY Servo on) | 0.246
| Analog Devices Inc | 2025-07-24 | X axis OPM | 1.0998
| Analog Devices Inc | 2025-07-24 | Y axis OPM | 0.685
| Analog Devices Inc | 2025-07-24 | X Non-linearity | 0.154
| Analog Devices Inc | 2025-07-24 | Y Non-linearity | 0.155
| HAESUNG DS | 2025-07-23 | Lateral(C-D) sensitivity | 27.5
| HAESUNG DS | 2025-07-23 | Lateral(C-D) noise | 0.073
| HAESUNG DS | 2025-07-23 | A-B noise | 11.09
| HAESUNG DS | 2025-07-23 | NCM Amplitude noise | 16.625
| HAESUNG DS | 2025-07-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| HAESUNG DS | 2025-07-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.16
| HAESUNG DS | 2025-07-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| HAESUNG DS | 2025-07-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| HAESUNG DS | 2025-07-23 | Roughness Repeatability 1x1um(RSD%) | 1.09
| HAESUNG DS | 2025-07-23 | Roughness Repeatability 1x1um(1σ) | 0.002
| HAESUNG DS | 2025-07-23 | Full range of Z Height during Z Scan sweep | 18.4
| HAESUNG DS | 2025-07-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.438
| HAESUNG DS | 2025-07-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.024
| HAESUNG DS | 2025-07-23 | Full range of Z Detector during Z Scan sweep | 12.7
| HAESUNG DS | 2025-07-23 | Z Drive calibration error | 0.22989
| HAESUNG DS | 2025-07-23 | Z Height calibration error | 0.24904
| HAESUNG DS | 2025-07-23 | Z linearity by sample height | 0.99161
| HAESUNG DS | 2025-07-23 | 1) Z Scanner position at +3μm | 0.004459
| HAESUNG DS | 2025-07-23 | 3) Z Scanner position at -3μm | 0.003085
| HAESUNG DS | 2025-07-23 | Z Height noise | 0.03
| HAESUNG DS | 2025-07-23 | Full range of X Detector during X Scan sweep | 67.5
| HAESUNG DS | 2025-07-23 | Full range of Y Detector during Y Scan sweep | 67.9
| HAESUNG DS | 2025-07-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99
| HAESUNG DS | 2025-07-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.6
| HAESUNG DS | 2025-07-23 | X Detector stroke calibration error (nominal scale error X) | 0.019
| HAESUNG DS | 2025-07-23 | Y Detector stroke calibration error (nominal scale error Y) | 0.12
| HAESUNG DS | 2025-07-23 | X Scan Stroke | 55
| HAESUNG DS | 2025-07-23 | Y Scan Stroke | 55
| HAESUNG DS | 2025-07-23 | Full range of X Detector Raw during X Scan sweep | 13.7
| HAESUNG DS | 2025-07-23 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| HAESUNG DS | 2025-07-23 | X Detector noise (XY Servo on) | 0.246
| HAESUNG DS | 2025-07-23 | Y Detector noise (XY Servo on) | 0.244
| HAESUNG DS | 2025-07-23 | X Detector noise (XY Servo off) | 0.239
| HAESUNG DS | 2025-07-23 | Y Detector noise (XY Servo off) | 0.236
| HAESUNG DS | 2025-07-23 | X axis OPM | 1.4758
| HAESUNG DS | 2025-07-23 | Y axis OPM | 0.9836
| HAESUNG DS | 2025-07-23 | X Non-linearity | 0.063
| HAESUNG DS | 2025-07-23 | Y Non-linearity | 0.058
| Samsung NX-TSH 400TF | 2025-07-23 | Wafer placement(Loading) error on the load port1 (X offset) | 18.3
| Samsung NX-TSH 400TF | 2025-07-23 | Wafer placement(Loading) error on the load port1 (Y offset) | 51.4
| Samsung NX-TSH 400TF | 2025-07-23 | Wafer placement(Loading) error on the load port2 (X offset) | 42.4
| Samsung NX-TSH 400TF | 2025-07-23 | Wafer placement(Loading) error on the load port2 (Y offset) | 25.5
| Samsung NX-TSH 400TF | 2025-07-23 | OpticalAxisShiftPerFocusX | -0.0004
| Samsung NX-TSH 400TF | 2025-07-23 | OpticalAxisShiftPerFocusY | -0.0001
| Samsung NX-TSH 400TF | 2025-07-23 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00046
| Samsung NX-TSH 400TF | 2025-07-23 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00024
| Samsung NX-TSH 400TF | 2025-07-23 | Profiler noise @200um | 0.12
| Samsung NX-TSH 400TF | 2025-07-23 | Position repeatability(10repeats, Position Based) | 0.74
| Samsung NX-TSH 400TF | 2025-07-23 | OPM Repeatability @25mm | 3.843
| Samsung NX-TSH 400TF | 2025-07-23 | OPM Repeatability @10mm | 1.597
| Samsung NX-TSH 400TF | 2025-07-23 | OPM Repeatability @5mm | 1.518
| Samsung NX-TSH 400TF | 2025-07-23 | OPM Repeatability @1mm | 1.254
| Samsung NX-TSH 400TF | 2025-07-23 | Max OPM @25mm | 52.46
| Samsung NX-TSH 400TF | 2025-07-23 | Max OPM @10mm | 22.002
| Samsung NX-TSH 400TF | 2025-07-23 | Max OPM @5mm | 17.329
| Samsung NX-TSH 400TF | 2025-07-23 | Max OPM @1mm | 11.71
| Samsung NX-TSH 400TF | 2025-07-23 | Torque Profile (50mm, 200um/s) | 0.422
| Samsung NX-TSH 400TF | 2025-07-23 | Torque Profile (50mm, 20um/s) | 0.56
| Samsung NX-TSH 400TF | 2025-07-23 | Min. Vibration Transmissbility (~100Hz) | 0.02
| Samsung NX-TSH 400TF | 2025-07-23 | RMS of ZeroScan | 0.063
| Samsung NX-TSH 400TF | 2025-07-23 | Z Detector offset | 38.6
| Samsung NX-TSH 400TF | 2025-07-23 | Z Linearity by Z scanner position | 0.5
| Samsung NX-TSH 400TF | 2025-07-23 | Z Linearity by 4 different sample height | 1.0059
| Samsung NX-TSH 400TF | 2025-07-23 | Z Height noise | 0.033
| Samsung NX-TSH 400TF | 2025-07-23 | Z Height calibration error | 0.065048
| Samsung NX-TSH 400TF | 2025-07-23 | Z Drive calibration error | 0.065048
| Samsung NX-TSH 400TF | 2025-07-23 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.1
| Samsung NX-TSH 400TF | 2025-07-23 | A-B after Auto PSPD Alignment | 0.06
| Samsung NX-TSH 400TF | 2025-07-23 | Y detector noise (XY Servo on) | 0.09
| Samsung NX-TSH 400TF | 2025-07-23 | Check the full stroke image(servo on) | Yes
| Samsung NX-TSH 400TF | 2025-07-23 | X Ringing(Rpv) | 0.15
| Samsung NX-TSH 400TF | 2025-07-23 | Y Ringing(Rpv) | 0.09
| Samsung NX-TSH 400TF | 2025-07-23 | ResonanceFrequencyHz | 500
| Samsung NX-TSH 400TF | 2025-07-23 | ServoCutoffFrequencyHz | 200
| Samsung NX-TSH 400TF | 2025-07-23 | Full range of Z Detector during Z Scan sweep | 13.54
| Samsung NX-TSH 400TF | 2025-07-23 | Full range of Z Height during Z Scan sweep | 18.6
| Samsung NX-TSH 400TF | 2025-07-23 | Full range of X Detector Raw during X Scan sweep | 14.3
| Samsung NX-TSH 400TF | 2025-07-23 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Samsung NX-TSH 400TF | 2025-07-23 | X Detector stroke calibration error (nominal scale error X) | 0.02
| Samsung NX-TSH 400TF | 2025-07-23 | Y Detector stroke calibration error (nominal scale error Y) | 0.12
| Samsung NX-TSH 400TF | 2025-07-23 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 154.1
| Samsung NX-TSH 400TF | 2025-07-23 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 157.2
| Samsung NX-TSH 400TF | 2025-07-23 | FullPixelHeight | 1828
| Samsung NX-TSH 400TF | 2025-07-23 | RMS error after Stage Calibration | 4.99
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from -X to +X (X scan) | 30.0293
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from -Y to +Y (Y scan) | 21.9727
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from -X to +X (Y scan) | 24.1699
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from -Y to +Y (X scan) | 76.1719
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from +X to -X (X scan) | 21.9727
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from +Y to -Y (Y scan) | 65.918
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from +X to -X (Y scan) | 32.959
| Samsung NX-TSH 400TF | 2025-07-23 | Image offset after moving from +Y to -Y (X scan) | 21.2402
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X offset (Vision Pattern Matching) | 0.40527
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.3574
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.085509
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.23298
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X offset (In-Die Align) | 2.1143
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y offset (In-Die Align) | 2.0898
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X 1σ (In-Die Align) | 0.39568
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y 1σ (In-Die Align) | 0.3601
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X offset (LLC Translation) | 2.666
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y offset (LLC Translation) | 1.7285
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X 1σ (LLC Translation) | 0.52479
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y 1σ (LLC Translation) | 0.31171
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X offset (Global Align) | 2.8906
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y offset (Global Align) | 2.0215
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning X 1σ (Global Align) | 0.53028
| Samsung NX-TSH 400TF | 2025-07-23 | XY Stage positioning Y 1σ (Global Align) | 0.4358
| Samsung NX-TSH 400TF | 2025-07-23 | Z Drift | 36.795
| Samsung NX-TSH 400TF | 2025-07-23 | Z-XY Orthogonality (X direction-Left) | 54.15
| Samsung NX-TSH 400TF | 2025-07-23 | Z-XY Orthogonality (X direction-Right) | 54.49
| Samsung NX-TSH 400TF | 2025-07-23 | Z-XY Orthogonality (Y direction-Left) | 54.21
| Samsung NX-TSH 400TF | 2025-07-23 | Z-XY Orthogonality (Y direction-Right) | 55.01
| Samsung NX-TSH 400TF | 2025-07-23 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.28
| Samsung NX-TSH 400TF | 2025-07-23 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Samsung NX-TSH 400TF | 2025-07-23 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.77
| Samsung NX-TSH 400TF | 2025-07-23 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| Samsung NX-TSH 400TF | 2025-07-23 | Cycle time for 15-point measurement | 1619.7
| Samsung NX-TSH 400TF | 2025-07-23 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.5
| Samsung NX-TSH 400TF | 2025-07-23 | Roughness Repeatability 1x1um(RSD%) | 2.7
| Samsung NX-TSH 400TF | 2025-07-23 | Roughness Repeatability 1x1um(1σ) | 0.002
| Samsung NX-TSH 400TF | 2025-07-23 | Roughness Repeatability 80x80um(RSD%) | 0.9
| Samsung NX-TSH 400TF | 2025-07-23 | Roughness Repeatability 80x80um(1σ) | 0.006
| Samsung NX-TSH 400TF | 2025-07-23 | Serial Port# | 8
| Samsung NX-TSH 400TF | 2025-07-23 | Off-axis Camera X Position(ΔX) | -106003
| 아주대학교_기계공학과 | 2025-07-17 | A-B noise | 7.329
| 아주대학교_기계공학과 | 2025-07-17 | NCM Amplitude noise | 12.894
| 아주대학교_기계공학과 | 2025-07-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 아주대학교_기계공학과 | 2025-07-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| 아주대학교_기계공학과 | 2025-07-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 아주대학교_기계공학과 | 2025-07-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 아주대학교_기계공학과 | 2025-07-17 | Full range of Z Height during Z Scan sweep | 18
| 아주대학교_기계공학과 | 2025-07-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.257
| 아주대학교_기계공학과 | 2025-07-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.935
| 아주대학교_기계공학과 | 2025-07-17 | Full range of Z Detector during Z Scan sweep | 12.7
| 아주대학교_기계공학과 | 2025-07-17 | Z Drive calibration error | 0.13697
| 아주대학교_기계공학과 | 2025-07-17 | Z Height calibration error | 0.11782
| 아주대학교_기계공학과 | 2025-07-17 | Z linearity by sample height | 1.0033
| 아주대학교_기계공학과 | 2025-07-17 | 1) Z Scanner position at +3μm | 0.002225
| 아주대학교_기계공학과 | 2025-07-17 | 3) Z Scanner position at -3μm | 0.003577
| 아주대학교_기계공학과 | 2025-07-17 | Z Height noise | 0.03
| 아주대학교_기계공학과 | 2025-07-17 | Full range of X Detector during X Scan sweep | 65.9
| 아주대학교_기계공학과 | 2025-07-17 | Full range of Y Detector during Y Scan sweep | 66
| 아주대학교_기계공학과 | 2025-07-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 89.8
| 아주대학교_기계공학과 | 2025-07-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 91.1
| 아주대학교_기계공학과 | 2025-07-17 | X Detector stroke calibration error (nominal scale error X) | 0.181
| 아주대학교_기계공학과 | 2025-07-17 | Y Detector stroke calibration error (nominal scale error Y) | 0.435
| 아주대학교_기계공학과 | 2025-07-17 | X Scan Stroke | 55
| 아주대학교_기계공학과 | 2025-07-17 | Y Scan Stroke | 55
| 아주대학교_기계공학과 | 2025-07-17 | Full range of X Detector Raw during X Scan sweep | 14.7
| 아주대학교_기계공학과 | 2025-07-17 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 아주대학교_기계공학과 | 2025-07-17 | X Detector noise (XY Servo on) | 0.322
| 아주대학교_기계공학과 | 2025-07-17 | Y Detector noise (XY Servo on) | 0.322
| 아주대학교_기계공학과 | 2025-07-17 | X Detector noise (XY Servo off) | 0.334
| 아주대학교_기계공학과 | 2025-07-17 | Y Detector noise (XY Servo off) | 0.337
| 아주대학교_기계공학과 | 2025-07-17 | X axis OPM | 1.1298
| 아주대학교_기계공학과 | 2025-07-17 | Y axis OPM | 0.901
| 아주대학교_기계공학과 | 2025-07-17 | X Non-linearity | 0.034
| 아주대학교_기계공학과 | 2025-07-17 | Y Non-linearity | 0.059
| Vellore Institute of Technology | 2025-07-10 | Lateral(C-D) sensitivity | 12.3
| Vellore Institute of Technology | 2025-07-10 | Lateral(C-D) noise | 0.105
| Vellore Institute of Technology | 2025-07-10 | A-B noise | 6.717
| Vellore Institute of Technology | 2025-07-10 | NCM Amplitude noise | 11.403
| Vellore Institute of Technology | 2025-07-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Vellore Institute of Technology | 2025-07-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| Vellore Institute of Technology | 2025-07-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Vellore Institute of Technology | 2025-07-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| Vellore Institute of Technology | 2025-07-10 | Roughness Repeatability 1x1um(RSD%) | 2.764
| Vellore Institute of Technology | 2025-07-10 | Roughness Repeatability 1x1um(1σ) | 0.004
| Micron | 2025-07-10 | Z Linearity by Z scanner position | 0.78
| Micron | 2025-07-10 | Z Linearity by 4 different sample height | 0.99761
| Micron | 2025-07-10 | Full range of Z Detector during Z Scan sweep | 13.58
| Micron | 2025-07-10 | Full range of Z Height during Z Scan sweep | 17.651
| Micron | 2025-07-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.316
| Micron | 2025-07-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3288
| Micron | 2025-07-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.2965
| Micron | 2025-07-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.3607
| Micron | 2025-07-10 | Z Height noise | 0.028
| Micron | 2025-07-10 | Wafer placement(Loading) error on the load port1 (X offset) | 22.45
| Micron | 2025-07-10 | Wafer placement(Loading) error on the load port1 (Y offset) | 20.71
| Micron | 2025-07-10 | Wafer placement(Loading) error on the load port2 (X offset) | 29.1
| Micron | 2025-07-10 | Wafer placement(Loading) error on the load port2 (Y offset) | 20
| Micron | 2025-07-10 | OpticalAxisShiftPerFocusX | -0.0046
| Micron | 2025-07-10 | OpticalAxisShiftPerFocusY | 0.0024
| Micron | 2025-07-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 3e-05
| Micron | 2025-07-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Micron | 2025-07-10 | Lift pins synchronization offset check (X offset) | 29.08
| Micron | 2025-07-10 | Lift pins synchronization offset check (Y offset) | 63.09
| Micron | 2025-07-10 | Max OPM @25mm | 159.411
| Micron | 2025-07-10 | Max OPM @10mm | 51.436
| Micron | 2025-07-10 | Max OPM @5mm | 22.059
| Micron | 2025-07-10 | Max OPM @1mm | 12.017
| Micron | 2025-07-10 | OPM Repeatability @25mm | 4.867
| Micron | 2025-07-10 | OPM Repeatability @10mm | 2.342
| Micron | 2025-07-10 | OPM Repeatability @5mm | 1.602
| Micron | 2025-07-10 | OPM Repeatability @1mm | 1.356
| Micron | 2025-07-10 | Torque Profile (50mm, 200um/s) | 1.058
| Micron | 2025-07-10 | Torque Profile (50mm, 20um/s) | 1.591
| Micron | 2025-07-10 | Profiler noise @200um | 0.197
| Micron | 2025-07-10 | Roll motion to X+ full scan | 7.3
| Micron | 2025-07-10 | Position repeatability(10repeats, Position Based) | 0.36
| Micron | 2025-07-10 | Lateral(C-D) sensitivity | 86.2
| Micron | 2025-07-10 | Lateral(C-D) noise | 0.206
| Micron | 2025-07-10 | Min. Vibration Transmissbility (~100Hz) | 0.031
| Micron | 2025-07-10 | RMS of ZeroScan | 0.138
| Micron | 2025-07-10 | Full range of X Detector during X Scan sweep | 216.9231
| Micron | 2025-07-10 | Full range of X1 Detector during X Scan sweep | 205.3924
| Micron | 2025-07-10 | Full range of X2 Detector during X Scan sweep | 216.9231
| Micron | 2025-07-10 | Full range of Y Detector during Y Scan sweep | 196.4621
| Micron | 2025-07-10 | Full range of Y1 Detector during Y Scan sweep | 196.4621
| Micron | 2025-07-10 | Full range of Y2 Detector during Y Scan sweep | 215.1302
| Micron | 2025-07-10 | Full range of X1 Detector Raw during X Scan sweep | 14.19
| Micron | 2025-07-10 | Full range of X2 Detector Raw during X Scan sweep | 14.12
| Micron | 2025-07-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.14
| Micron | 2025-07-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14.04
| Micron | 2025-07-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 278.274
| Micron | 2025-07-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 291.5547
| Micron | 2025-07-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 280.2552
| Micron | 2025-07-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 287.0176
| Micron | 2025-07-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.191
| Micron | 2025-07-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 186.046
| Micron | 2025-07-10 | XY Orthogonality | 89.94
| Micron | 2025-07-10 | X Non-linearity error | 3.09
| Micron | 2025-07-10 | Y Non-linearity error | 3.59
| Micron | 2025-07-10 | X axis OPM | 0.243
| Micron | 2025-07-10 | Y axis OPM | 1.653
| Micron | 2025-07-10 | X Ringing(Rpv) | 0.133
| Micron | 2025-07-10 | Y Ringing(Rpv) | 0.109
| Micron | 2025-07-10 | X Detector noise (XY Servo on) | 0.286
| Micron | 2025-07-10 | Y Detector noise (XY Servo on) | 0.411
| Micron | 2025-07-10 | X1 Detector noise (XY Servo on) | 0.353
| Micron | 2025-07-10 | Y1 Detector noise (XY Servo on) | 0.476
| Micron | 2025-07-10 | X2 Detector noise (XY Servo on) | 0.376
| Micron | 2025-07-10 | Y2 Detector noise (XY Servo on) | 0.462
| Micron | 2025-07-10 | RMS error after Stage Calibration | 2.26
| Micron | 2025-07-10 | Image offset after moving from -X to +X (X scan) | 32.2266
| Micron | 2025-07-10 | Image offset after moving from -Y to +Y (Y scan) | 88.6231
| Micron | 2025-07-10 | Image offset after moving from -X to +X (Y scan) | 47.6074
| Micron | 2025-07-10 | Image offset after moving from -Y to +Y (X scan) | 16.1133
| Micron | 2025-07-10 | Image offset after moving from +X to -X (X scan) | 22.7051
| Micron | 2025-07-10 | Image offset after moving from +Y to -Y (Y scan) | 22.7051
| Micron | 2025-07-10 | Image offset after moving from +X to -X (Y scan) | 50.5371
| Micron | 2025-07-10 | Image offset after moving from +Y to -Y (X scan) | 12.4512
| Micron | 2025-07-10 | XY Stage positioning X offset (Vision Pattern Matching) | 0.8252
| Micron | 2025-07-10 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.0117
| Micron | 2025-07-10 | XY Stage positioning X offset (In-Die Align) | 1.1182
| Micron | 2025-07-10 | XY Stage positioning Y offset (In-Die Align) | 4.6143
| Micron | 2025-07-10 | XY Stage positioning X offset (LLC Translation) | 2.8955
| Micron | 2025-07-10 | XY Stage positioning Y offset (LLC Translation) | 1.7871
| Micron | 2025-07-10 | XY Stage positioning X 1σ (LLC Translation) | 0.465
| Micron | 2025-07-10 | XY Stage positioning Y 1σ (LLC Translation) | 0.3851
| Micron | 2025-07-10 | XY Stage positioning X offset (Global Align) | 3.125
| Micron | 2025-07-10 | XY Stage positioning Y offset (Global Align) | 7.1143
| Micron | 2025-07-10 | Z Drift | 12.8869
| Micron | 2025-07-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.068675
| Micron | 2025-07-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.012376
| Micron | 2025-07-10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.85
| Micron | 2025-07-10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| Micron | 2025-07-10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06595
| Micron | 2025-07-10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.011888
| Micron | 2025-07-10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.93
| Micron | 2025-07-10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.17
| Micron | 2025-07-10 | Cycle time for 15-point measurement | 1270.67
| Micron | 2025-07-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.309
| Micron | 2025-07-10 | Roughness Repeatability 1x1um(10Repeats, RSD%) | 2.29
| Micron | 2025-07-10 | Roughness Repeatability 80x80um(10Repeats, RSD%) | 2.11
| Neel Institute | 2025-07-08 | Lateral(C-D) sensitivity | 25.6
| Neel Institute | 2025-07-08 | Lateral(C-D) noise | 0.089
| Neel Institute | 2025-07-08 | A-B noise | 7.716
| Neel Institute | 2025-07-08 | NCM Amplitude noise | 14.804
| Neel Institute | 2025-07-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Neel Institute | 2025-07-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Neel Institute | 2025-07-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Neel Institute | 2025-07-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Neel Institute | 2025-07-08 | Roughness Repeatability 1x1um(RSD%) | 0.751
| Neel Institute | 2025-07-08 | Roughness Repeatability 1x1um(1σ) | 0.001
| Neel Institute | 2025-07-08 | Full range of Z Height during Z Scan sweep | 17.9
| Neel Institute | 2025-07-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.318
| Neel Institute | 2025-07-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.821
| Neel Institute | 2025-07-08 | Full range of Z Detector during Z Scan sweep | 13.1
| Neel Institute | 2025-07-08 | Z Drive calibration error | 0.14655
| Neel Institute | 2025-07-08 | Z Height calibration error | 0.40038
| Neel Institute | 2025-07-08 | NCM driving% with AC160TS cantilever | 0.412
| Neel Institute | 2025-07-08 | Z linearity by sample height | 0.9902
| Neel Institute | 2025-07-08 | 1) Z Scanner position at +3μm | 0.00436
| Neel Institute | 2025-07-08 | 3) Z Scanner position at -3μm | 0.000424
| Neel Institute | 2025-07-08 | Z Height noise | 0.029
| Neel Institute | 2025-07-08 | Full range of X Detector during X Scan sweep | 107.1
| Neel Institute | 2025-07-08 | Full range of Y Detector during Y Scan sweep | 107.7
| Neel Institute | 2025-07-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.8
| Neel Institute | 2025-07-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 161.9
| Neel Institute | 2025-07-08 | X Detector stroke calibration error (nominal scale error X) | 0.054
| Neel Institute | 2025-07-08 | Y Detector stroke calibration error (nominal scale error Y) | 0.24
| Neel Institute | 2025-07-08 | X Scan Stroke | 107.1
| Neel Institute | 2025-07-08 | Y Scan Stroke | 107.7
| Neel Institute | 2025-07-08 | Full range of X Detector Raw during X Scan sweep | 13.4
| Neel Institute | 2025-07-08 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| Neel Institute | 2025-07-08 | X Detector noise (XY Servo on) | 0.259
| Neel Institute | 2025-07-08 | Y Detector noise (XY Servo on) | 0.262
| Neel Institute | 2025-07-08 | X Detector noise (XY servo off) | 0.256
| Neel Institute | 2025-07-08 | Y Detector noise (XY servo off) | 0.258
| Neel Institute | 2025-07-08 | X axis OPM | 0.6704
| Neel Institute | 2025-07-08 | Y axis OPM | 1.1316
| Neel Institute | 2025-07-08 | X Non-linearity | 0.172
| Neel Institute | 2025-07-08 | Y Non-linearity | 0.19
| Tomoegawa Corporation | 2025-07-07 | Z Scanner position at -3μm | 0.000277
| University of Utah | 2025-07-07 | Z Scanner position at -3μm | 0.002869
| Tomoegawa Corporation | 2025-07-07 | Z Scanner position at +3μm | 0.000353
| University of Utah | 2025-07-07 | Z Scanner position at +3μm | 0.004755
| Tomoegawa Corporation | 2025-07-07 | Z linearity by sample height | 1.0101
| University of Utah | 2025-07-07 | Z linearity by sample height | 1.006
| Tomoegawa Corporation | 2025-07-07 | Z Height noise | 0.03
| University of Utah | 2025-07-07 | Z Height noise | 0.026
| Tomoegawa Corporation | 2025-07-07 | Z Height calibration error | 0.16858
| University of Utah | 2025-07-07 | Z Height calibration error | 0.11024
| Tomoegawa Corporation | 2025-07-07 | Z Drive calibration error | 0.22031
| University of Utah | 2025-07-07 | Z Drive calibration error | 0.34146
| Tomoegawa Corporation | 2025-07-07 | Y Scan Stroke | 55
| University of Utah | 2025-07-07 | Y Scan Stroke | 120.9
| Tomoegawa Corporation | 2025-07-07 | Y Non-linearity | 0.128
| University of Utah | 2025-07-07 | Y Non-linearity | 0.243
| Tomoegawa Corporation | 2025-07-07 | Y Detector stroke calibration error (nominal scale error Y) | 0.071
| University of Utah | 2025-07-07 | Y Detector stroke calibration error (nominal scale error Y) | 0.319
| Tomoegawa Corporation | 2025-07-07 | Y Detector noise (XY Servo on) | 0.254
| University of Utah | 2025-07-07 | Y Detector noise (XY Servo on) | 0.29
| Tomoegawa Corporation | 2025-07-07 | Y Detector noise (XY Servo off) | 0.249
| University of Utah | 2025-07-07 | Y Detector noise (XY servo off) | 0.291
| Tomoegawa Corporation | 2025-07-07 | Y axis OPM | 0.5014
| University of Utah | 2025-07-07 | Y axis OPM | 0.6558
| Tomoegawa Corporation | 2025-07-07 | X Scan Stroke | 55
| University of Utah | 2025-07-07 | X Scan Stroke | 113.1
| Tomoegawa Corporation | 2025-07-07 | X Non-linearity | 0.124
| University of Utah | 2025-07-07 | X Non-linearity | 0.272
| Tomoegawa Corporation | 2025-07-07 | X Detector stroke calibration error (nominal scale error X) | 0.23
| University of Utah | 2025-07-07 | X Detector stroke calibration error (nominal scale error X) | 0.039
| Tomoegawa Corporation | 2025-07-07 | X Detector noise (XY Servo on) | 0.254
| University of Utah | 2025-07-07 | X Detector noise (XY Servo on) | 0.289
| Tomoegawa Corporation | 2025-07-07 | X Detector noise (XY Servo off) | 0.249
| University of Utah | 2025-07-07 | X Detector noise (XY servo off) | 0.293
| Tomoegawa Corporation | 2025-07-07 | X axis OPM | 0.4004
| University of Utah | 2025-07-07 | X axis OPM | 0.7496
| Tomoegawa Corporation | 2025-07-07 | Roughness Repeatability 1x1um(RSD%) | 1.447
| University of Utah | 2025-07-07 | Roughness Repeatability 1x1um(RSD%) | 1.899
| Tomoegawa Corporation | 2025-07-07 | Roughness Repeatability 1x1um(1σ) | 0.001
| University of Utah | 2025-07-07 | Roughness Repeatability 1x1um(1σ) | 0.003
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| University of Utah | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| University of Utah | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of Utah | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| University of Utah | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Tomoegawa Corporation | 2025-07-07 | NCM driving% with AC160TS cantilever | 1.368
| University of Utah | 2025-07-07 | NCM driving% with AC160TS cantilever | 1.392
| Tomoegawa Corporation | 2025-07-07 | NCM Amplitude noise | 11.11
| University of Utah | 2025-07-07 | NCM Amplitude noise | 12.599
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.079
| University of Utah | 2025-07-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.828
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.112
| University of Utah | 2025-07-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.266
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 107.9
| University of Utah | 2025-07-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 166.1
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.6
| University of Utah | 2025-07-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.5
| Tomoegawa Corporation | 2025-07-07 | Lateral(C-D) sensitivity | 66.9
| University of Utah | 2025-07-07 | Lateral(C-D) sensitivity | 58.7
| Tomoegawa Corporation | 2025-07-07 | Lateral(C-D) noise | 0.095
| University of Utah | 2025-07-07 | Lateral(C-D) noise | 0.098
| Tomoegawa Corporation | 2025-07-07 | Full range of Z Height during Z Scan sweep | 18.1
| University of Utah | 2025-07-07 | Full range of Z Height during Z Scan sweep | 17.6
| Tomoegawa Corporation | 2025-07-07 | Full range of Z Detector during Z Scan sweep | 12.9
| University of Utah | 2025-07-07 | Full range of Z Detector during Z Scan sweep | 13.9
| Tomoegawa Corporation | 2025-07-07 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| University of Utah | 2025-07-07 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Tomoegawa Corporation | 2025-07-07 | Full range of Y Detector during Y Scan sweep | 72.9
| University of Utah | 2025-07-07 | Full range of Y Detector during Y Scan sweep | 120.9
| Tomoegawa Corporation | 2025-07-07 | Full range of X Detector Raw during X Scan sweep | 13.5
| University of Utah | 2025-07-07 | Full range of X Detector Raw during X Scan sweep | 13.8
| Tomoegawa Corporation | 2025-07-07 | Full range of X Detector during X Scan sweep | 71.9
| University of Utah | 2025-07-07 | Full range of X Detector during X Scan sweep | 113.1
| Tomoegawa Corporation | 2025-07-07 | A-B noise | 6.774
| University of Utah | 2025-07-07 | A-B noise | 6.707
| Tomoegawa Corporation | 2025-07-07 | Lateral(C-D) sensitivity | 66.9
| Tomoegawa Corporation | 2025-07-07 | Lateral(C-D) noise | 0.095
| Tomoegawa Corporation | 2025-07-07 | A-B noise | 6.774
| Tomoegawa Corporation | 2025-07-07 | NCM Amplitude noise | 11.11
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Tomoegawa Corporation | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Tomoegawa Corporation | 2025-07-07 | Roughness Repeatability 1x1um(RSD%) | 1.447
| Tomoegawa Corporation | 2025-07-07 | Roughness Repeatability 1x1um(1σ) | 0.001
| Tomoegawa Corporation | 2025-07-07 | Full range of Z Height during Z Scan sweep | 18.1
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.112
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.079
| Tomoegawa Corporation | 2025-07-07 | Full range of Z Detector during Z Scan sweep | 12.9
| Tomoegawa Corporation | 2025-07-07 | Z Drive calibration error | 0.22031
| Tomoegawa Corporation | 2025-07-07 | Z Height calibration error | 0.16858
| Tomoegawa Corporation | 2025-07-07 | NCM driving% with AC160TS cantilever | 1.368
| Tomoegawa Corporation | 2025-07-07 | Z linearity by sample height | 1.0101
| Tomoegawa Corporation | 2025-07-07 | 1) Z Scanner position at +3μm | 0.000353
| Tomoegawa Corporation | 2025-07-07 | 3) Z Scanner position at -3μm | 0.000277
| Tomoegawa Corporation | 2025-07-07 | Z Height noise | 0.03
| Tomoegawa Corporation | 2025-07-07 | Full range of X Detector during X Scan sweep | 71.9
| Tomoegawa Corporation | 2025-07-07 | Full range of Y Detector during Y Scan sweep | 72.9
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.6
| Tomoegawa Corporation | 2025-07-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 107.9
| Tomoegawa Corporation | 2025-07-07 | X Detector stroke calibration error (nominal scale error X) | 0.23
| Tomoegawa Corporation | 2025-07-07 | Y Detector stroke calibration error (nominal scale error Y) | 0.071
| Tomoegawa Corporation | 2025-07-07 | X Scan Stroke | 55
| Tomoegawa Corporation | 2025-07-07 | Y Scan Stroke | 55
| Tomoegawa Corporation | 2025-07-07 | Full range of X Detector Raw during X Scan sweep | 13.5
| Tomoegawa Corporation | 2025-07-07 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| Tomoegawa Corporation | 2025-07-07 | X Detector noise (XY Servo on) | 0.254
| Tomoegawa Corporation | 2025-07-07 | Y Detector noise (XY Servo on) | 0.254
| Tomoegawa Corporation | 2025-07-07 | X Detector noise (XY Servo off) | 0.249
| Tomoegawa Corporation | 2025-07-07 | Y Detector noise (XY Servo off) | 0.249
| Tomoegawa Corporation | 2025-07-07 | X axis OPM | 0.4004
| Tomoegawa Corporation | 2025-07-07 | Y axis OPM | 0.5014
| Tomoegawa Corporation | 2025-07-07 | X Non-linearity | 0.124
| Tomoegawa Corporation | 2025-07-07 | Y Non-linearity | 0.128
| University of Utah | 2025-07-07 | Lateral(C-D) sensitivity | 58.7
| University of Utah | 2025-07-07 | Lateral(C-D) noise | 0.098
| University of Utah | 2025-07-07 | A-B noise | 6.707
| University of Utah | 2025-07-07 | NCM Amplitude noise | 12.599
| University of Utah | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| University of Utah | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Utah | 2025-07-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| University of Utah | 2025-07-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| University of Utah | 2025-07-07 | Roughness Repeatability 1x1um(RSD%) | 1.899
| University of Utah | 2025-07-07 | Roughness Repeatability 1x1um(1σ) | 0.003
| University of Utah | 2025-07-07 | Full range of Z Height during Z Scan sweep | 17.6
| University of Utah | 2025-07-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.266
| University of Utah | 2025-07-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.828
| University of Utah | 2025-07-07 | Full range of Z Detector during Z Scan sweep | 13.9
| University of Utah | 2025-07-07 | Z Drive calibration error | 0.34146
| University of Utah | 2025-07-07 | Z Height calibration error | 0.11024
| University of Utah | 2025-07-07 | NCM driving% with AC160TS cantilever | 1.392
| University of Utah | 2025-07-07 | Z linearity by sample height | 1.006
| University of Utah | 2025-07-07 | 1) Z Scanner position at +3μm | 0.004755
| University of Utah | 2025-07-07 | 3) Z Scanner position at -3μm | 0.002869
| University of Utah | 2025-07-07 | Z Height noise | 0.026
| University of Utah | 2025-07-07 | Full range of X Detector during X Scan sweep | 113.1
| University of Utah | 2025-07-07 | Full range of Y Detector during Y Scan sweep | 120.9
| University of Utah | 2025-07-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.5
| University of Utah | 2025-07-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 166.1
| University of Utah | 2025-07-07 | X Detector stroke calibration error (nominal scale error X) | 0.039
| University of Utah | 2025-07-07 | Y Detector stroke calibration error (nominal scale error Y) | 0.319
| University of Utah | 2025-07-07 | X Scan Stroke | 113.1
| University of Utah | 2025-07-07 | Y Scan Stroke | 120.9
| University of Utah | 2025-07-07 | Full range of X Detector Raw during X Scan sweep | 13.8
| University of Utah | 2025-07-07 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| University of Utah | 2025-07-07 | X Detector noise (XY Servo on) | 0.289
| University of Utah | 2025-07-07 | Y Detector noise (XY Servo on) | 0.29
| University of Utah | 2025-07-07 | X Detector noise (XY servo off) | 0.293
| University of Utah | 2025-07-07 | Y Detector noise (XY servo off) | 0.291
| University of Utah | 2025-07-07 | X axis OPM | 0.7496
| University of Utah | 2025-07-07 | Y axis OPM | 0.6558
| University of Utah | 2025-07-07 | X Non-linearity | 0.272
| University of Utah | 2025-07-07 | Y Non-linearity | 0.243
| 한양대학교 융합전자공학과 | 2025-07-04 | A-B noise | 6.297
| 한양대학교 융합전자공학과 | 2025-07-04 | NCM Amplitude noise | 10.48
| 한양대학교 융합전자공학과 | 2025-07-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| 한양대학교 융합전자공학과 | 2025-07-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| 한양대학교 융합전자공학과 | 2025-07-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한양대학교 융합전자공학과 | 2025-07-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| 한양대학교 융합전자공학과 | 2025-07-04 | Full range of Z Height during Z Scan sweep | 17.8
| 한양대학교 융합전자공학과 | 2025-07-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.964
| 한양대학교 융합전자공학과 | 2025-07-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.977
| 한양대학교 융합전자공학과 | 2025-07-04 | Full range of Z Detector during Z Scan sweep | 12.7
| 한양대학교 융합전자공학과 | 2025-07-04 | Z Drive calibration error | 0.20785
| 한양대학교 융합전자공학과 | 2025-07-04 | Z Height calibration error | 0.082375
| 한양대학교 융합전자공학과 | 2025-07-04 | NCM driving% with AC160TS cantilever | 1.199
| 한양대학교 융합전자공학과 | 2025-07-04 | Z linearity by sample height | 1.0037
| 한양대학교 융합전자공학과 | 2025-07-04 | 1) Z Scanner position at +3μm | 0.00981
| 한양대학교 융합전자공학과 | 2025-07-04 | 3) Z Scanner position at -3μm | 0.005058
| 한양대학교 융합전자공학과 | 2025-07-04 | Z Height noise | 0.029
| 한양대학교 융합전자공학과 | 2025-07-04 | Full range of X Detector during X Scan sweep | 65
| 한양대학교 융합전자공학과 | 2025-07-04 | Full range of Y Detector during Y Scan sweep | 65.5
| 한양대학교 융합전자공학과 | 2025-07-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 87.7
| 한양대학교 융합전자공학과 | 2025-07-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 86.8
| 한양대학교 융합전자공학과 | 2025-07-04 | X Detector stroke calibration error (nominal scale error X) | 0.054
| 한양대학교 융합전자공학과 | 2025-07-04 | Y Detector stroke calibration error (nominal scale error Y) | 0.209
| 한양대학교 융합전자공학과 | 2025-07-04 | X Scan Stroke | 55
| 한양대학교 융합전자공학과 | 2025-07-04 | Y Scan Stroke | 55
| 한양대학교 융합전자공학과 | 2025-07-04 | Full range of X Detector Raw during X Scan sweep | 14.8
| 한양대학교 융합전자공학과 | 2025-07-04 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| 한양대학교 융합전자공학과 | 2025-07-04 | X Detector noise (XY Servo on) | 0.347
| 한양대학교 융합전자공학과 | 2025-07-04 | Y Detector noise (XY Servo on) | 0.344
| 한양대학교 융합전자공학과 | 2025-07-04 | X Detector noise (XY Servo off) | 0.372
| 한양대학교 융합전자공학과 | 2025-07-04 | Y Detector noise (XY Servo off) | 0.368
| 한양대학교 융합전자공학과 | 2025-07-04 | X axis OPM | 0.7124
| 한양대학교 융합전자공학과 | 2025-07-04 | Y axis OPM | 0.6184
| 한양대학교 융합전자공학과 | 2025-07-04 | X Non-linearity | 0.078
| 한양대학교 융합전자공학과 | 2025-07-04 | Y Non-linearity | 0.035
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Lateral(C-D) sensitivity | 26.4
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Lateral(C-D) noise | 0.084
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | A-B noise | 7.014
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | NCM Amplitude noise | 13.413
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Roughness Repeatability 1x1um(RSD%) | 1.33
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Roughness Repeatability 1x1um(1σ) | 0.002
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Full range of Z Height during Z Scan sweep | 16.7
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.599
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.418
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Full range of Z Detector during Z Scan sweep | 12.1
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Z Height calibration error | 0.055556
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | NCM driving% with AC160TS cantilever | 0.647
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Z linearity by sample height | 1.0092
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | 1) Z Scanner position at +3μm | 0.006537
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | 3) Z Scanner position at -3μm | 0.003102
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Z Height noise | 0.028
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Full range of X Detector during X Scan sweep | 116.8
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Full range of Y Detector during Y Scan sweep | 115.6
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 165.6
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 172.5
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | X Detector stroke calibration error (nominal scale error X) | 0.021
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Y Detector stroke calibration error (nominal scale error Y) | 0.149
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | X Scan Stroke | 116.748
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Y Scan Stroke | 115.575
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Full range of X Detector Raw during X Scan sweep | 14.1
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | X Detector noise (XY Servo on) | 0.282
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Y Detector noise (XY Servo on) | 0.265
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | X Detector noise (XY servo off) | 0.288
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Y Detector noise (XY servo off) | 0.261
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | X axis OPM | 1.0074
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Y axis OPM | 0.6494
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | X Non-linearity | 0.184
| Advanced Micro Devices (Singapore) Pte. Ltd. | 2025-07-01 | Y Non-linearity | 0.215
| University of Chinese Academy of Sciences | 2025-06-30 | SLD current limit check | 148
| University of Chinese Academy of Sciences | 2025-06-30 | SLD beam power at the end of the lens | 0.27
| University of Chinese Academy of Sciences | 2025-06-30 | Lateral(C-D) sensitivity | 11.7
| University of Chinese Academy of Sciences | 2025-06-30 | Lateral(C-D) noise | 0.114
| University of Chinese Academy of Sciences | 2025-06-30 | A-B noise | 5.152
| University of Chinese Academy of Sciences | 2025-06-30 | NCM Amplitude noise | 4.795
| University of Chinese Academy of Sciences | 2025-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| University of Chinese Academy of Sciences | 2025-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| University of Chinese Academy of Sciences | 2025-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| University of Chinese Academy of Sciences | 2025-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| University of Chinese Academy of Sciences | 2025-06-30 | Roughness Repeatability 1x1um(RSD%) | 1.286
| University of Chinese Academy of Sciences | 2025-06-30 | Roughness Repeatability 1x1um(1σ) | 0.001
| University of Chinese Academy of Sciences | 2025-06-30 | Full range of Z Height during Z Scan sweep | 16.4
| University of Chinese Academy of Sciences | 2025-06-30 | Full range of Z Detector during Z Scan sweep | 14.2
| University of Chinese Academy of Sciences | 2025-06-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.133
| University of Chinese Academy of Sciences | 2025-06-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.569
| University of Chinese Academy of Sciences | 2025-06-30 | Z Drive calibrtaion error | 0.17463
| University of Chinese Academy of Sciences | 2025-06-30 | Z Height calibrtaion error | 0.01561
| University of Chinese Academy of Sciences | 2025-06-30 | NCM driving% with AC160TS cantilever | 0.561
| University of Chinese Academy of Sciences | 2025-06-30 | Z linearity by sample height | 0.9877
| University of Chinese Academy of Sciences | 2025-06-30 | 1) Z Scanner position at +3μm | 0.018852
| University of Chinese Academy of Sciences | 2025-06-30 | 3) Z Scanner position at -3μm | 0.002713
| University of Chinese Academy of Sciences | 2025-06-30 | Z Height noise | 0.017
| University of Chinese Academy of Sciences | 2025-06-30 | Full range of X Detector during X Scan sweep | 101.7
| University of Chinese Academy of Sciences | 2025-06-30 | Full range of Y Detector during Y Scan sweep | 103.7
| University of Chinese Academy of Sciences | 2025-06-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 146.4
| University of Chinese Academy of Sciences | 2025-06-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.9
| University of Chinese Academy of Sciences | 2025-06-30 | X Detector stroke calibration error (nominal scale error X) | 0.036
| University of Chinese Academy of Sciences | 2025-06-30 | Y Detector stroke calibration error (nominal scale error Y) | 0.088
| University of Chinese Academy of Sciences | 2025-06-30 | X Scan Stroke | 102.5
| University of Chinese Academy of Sciences | 2025-06-30 | Y Scan Stroke | 104.7
| University of Chinese Academy of Sciences | 2025-06-30 | Full range of X Detector Raw during X Scan sweep | 13.9
| University of Chinese Academy of Sciences | 2025-06-30 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| University of Chinese Academy of Sciences | 2025-06-30 | X Detector noise (XYServo off) | 0.24
| University of Chinese Academy of Sciences | 2025-06-30 | Y Detector noise (XYServo off) | 0.235
| University of Chinese Academy of Sciences | 2025-06-30 | X Detector noise (XY Servo on) | 0.247
| University of Chinese Academy of Sciences | 2025-06-30 | Y Detector noise (XY Servo on) | 0.244
| University of Chinese Academy of Sciences | 2025-06-30 | X axis OPM | 1.063
| University of Chinese Academy of Sciences | 2025-06-30 | Y axis OPM | 0.695
| University of Chinese Academy of Sciences | 2025-06-30 | X Non-linearity | 0.195
| University of Chinese Academy of Sciences | 2025-06-30 | Y Non-linearity | 0.13
| Chung Yuan Christian University | 2025-06-27 | Z Scanner position at -3μm | 0.002843
| Chung Yuan Christian University | 2025-06-27 | Z Scanner position at +3μm | 0.00758
| Chung Yuan Christian University | 2025-06-27 | Z linearity by sample height | 1.0119
| Chung Yuan Christian University | 2025-06-27 | Z Height noise | 0.029
| Chung Yuan Christian University | 2025-06-27 | Z Height calibration error | 0.034483
| Chung Yuan Christian University | 2025-06-27 | Z Drive calibration error | 0.14464
| Chung Yuan Christian University | 2025-06-27 | Y Scan Stroke | 125.9
| Chung Yuan Christian University | 2025-06-27 | Y Non-linearity | 0.118
| Chung Yuan Christian University | 2025-06-27 | Y Detector stroke calibration error (nominal scale error Y) | 0.164
| Chung Yuan Christian University | 2025-06-27 | Y Detector noise (XY Servo on) | 0.525
| Chung Yuan Christian University | 2025-06-27 | Y Detector noise (XY Servo off) | 0.57
| Chung Yuan Christian University | 2025-06-27 | Y axis OPM | 1.6658
| Chung Yuan Christian University | 2025-06-27 | X Scan Stroke | 120.5
| Chung Yuan Christian University | 2025-06-27 | X Non-linearity | 0.064
| Chung Yuan Christian University | 2025-06-27 | X Detector stroke calibration error (nominal scale error X) | 0.013
| Chung Yuan Christian University | 2025-06-27 | X Detector noise (XY Servo on) | 0.532
| Chung Yuan Christian University | 2025-06-27 | X Detector noise (XY Servo off) | 0.567
| Chung Yuan Christian University | 2025-06-27 | X axis OPM | 1.432
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Chung Yuan Christian University | 2025-06-27 | NCM driving% with AC160TS cantilever | 0.85
| Chung Yuan Christian University | 2025-06-27 | NCM Amplitude noise | 18.377
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.341
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.277
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.5
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 165.8
| Chung Yuan Christian University | 2025-06-27 | Full range of Z Height during Z Scan sweep | 18.6
| Chung Yuan Christian University | 2025-06-27 | Full range of Z Detector during Z Scan sweep | 13.2
| Chung Yuan Christian University | 2025-06-27 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| Chung Yuan Christian University | 2025-06-27 | Full range of Y Detector during Y Scan sweep | 125.9
| Chung Yuan Christian University | 2025-06-27 | Full range of X Detector Raw during X Scan sweep | 14.5
| Chung Yuan Christian University | 2025-06-27 | Full range of X Detector during X Scan sweep | 120.5
| Chung Yuan Christian University | 2025-06-27 | A-B noise | 6.626
| Chung Yuan Christian University | 2025-06-27 | A-B noise | 6.626
| Chung Yuan Christian University | 2025-06-27 | NCM Amplitude noise | 18.377
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Chung Yuan Christian University | 2025-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Chung Yuan Christian University | 2025-06-27 | Full range of Z Height during Z Scan sweep | 18.6
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.277
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.341
| Chung Yuan Christian University | 2025-06-27 | Full range of Z Detector during Z Scan sweep | 13.2
| Chung Yuan Christian University | 2025-06-27 | Z Drive calibration error | 0.14464
| Chung Yuan Christian University | 2025-06-27 | Z Height calibration error | 0.034483
| Chung Yuan Christian University | 2025-06-27 | NCM driving% with AC160TS cantilever | 0.85
| Chung Yuan Christian University | 2025-06-27 | Z linearity by sample height | 1.0119
| Chung Yuan Christian University | 2025-06-27 | 1) Z Scanner position at +3μm | 0.00758
| Chung Yuan Christian University | 2025-06-27 | 3) Z Scanner position at -3μm | 0.002843
| Chung Yuan Christian University | 2025-06-27 | Z Height noise | 0.029
| Chung Yuan Christian University | 2025-06-27 | Full range of X Detector during X Scan sweep | 120.5
| Chung Yuan Christian University | 2025-06-27 | Full range of Y Detector during Y Scan sweep | 125.9
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 165.8
| Chung Yuan Christian University | 2025-06-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.5
| Chung Yuan Christian University | 2025-06-27 | X Detector stroke calibration error (nominal scale error X) | 0.013
| Chung Yuan Christian University | 2025-06-27 | Y Detector stroke calibration error (nominal scale error Y) | 0.164
| Chung Yuan Christian University | 2025-06-27 | X Scan Stroke | 120.5
| Chung Yuan Christian University | 2025-06-27 | Y Scan Stroke | 125.9
| Chung Yuan Christian University | 2025-06-27 | Full range of X Detector Raw during X Scan sweep | 14.5
| Chung Yuan Christian University | 2025-06-27 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| Chung Yuan Christian University | 2025-06-27 | X Detector noise (XY Servo on) | 0.532
| Chung Yuan Christian University | 2025-06-27 | Y Detector noise (XY Servo on) | 0.525
| Chung Yuan Christian University | 2025-06-27 | X Detector noise (XY Servo off) | 0.567
| Chung Yuan Christian University | 2025-06-27 | Y Detector noise (XY Servo off) | 0.57
| Chung Yuan Christian University | 2025-06-27 | X axis OPM | 1.432
| Chung Yuan Christian University | 2025-06-27 | Y axis OPM | 1.6658
| Chung Yuan Christian University | 2025-06-27 | X Non-linearity | 0.064
| Chung Yuan Christian University | 2025-06-27 | Y Non-linearity | 0.118
| Wuxi DIS Microelectronics | 2025-06-25 | OpticalAxisShiftPerFocusX | 0.0044
| Wuxi DIS Microelectronics | 2025-06-25 | OpticalAxisShiftPerFocusY | -0.001
| Wuxi DIS Microelectronics | 2025-06-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00028
| Wuxi DIS Microelectronics | 2025-06-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Wuxi DIS Microelectronics | 2025-06-25 | Min. Vibration Transmissbility (~100Hz) | 0.024
| Wuxi DIS Microelectronics | 2025-06-25 | RMS of ZeroScan | 0.044
| Wuxi DIS Microelectronics | 2025-06-25 | Z Linearity by Z scanner position | 1.2
| Wuxi DIS Microelectronics | 2025-06-25 | Z Linearity by 4 different sample height | 0.99463
| Wuxi DIS Microelectronics | 2025-06-25 | Z Height noise | 0.038
| Wuxi DIS Microelectronics | 2025-06-25 | Z Height calibration error | 0.10386
| Wuxi DIS Microelectronics | 2025-06-25 | Z Drive calibration error | 0.29011
| Wuxi DIS Microelectronics | 2025-06-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.0798
| Wuxi DIS Microelectronics | 2025-06-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.5123
| Wuxi DIS Microelectronics | 2025-06-25 | XY Orthogonality | 90.252
| Wuxi DIS Microelectronics | 2025-06-25 | X detector noise (XY Servo on) | 0.14
| Wuxi DIS Microelectronics | 2025-06-25 | Y detector noise (XY Servo on) | 0.163
| Wuxi DIS Microelectronics | 2025-06-25 | X Ringing(Rpv) | 0.401
| Wuxi DIS Microelectronics | 2025-06-25 | Y Ringing(Rpv) | 0.492
| Wuxi DIS Microelectronics | 2025-06-25 | Full range of Z Detector during Z Scan sweep | 13.81
| Wuxi DIS Microelectronics | 2025-06-25 | Full range of Z Height during Z Scan sweep | 18.809
| Wuxi DIS Microelectronics | 2025-06-25 | Full range of X Detector Raw during X Scan sweep | 14.58
| Wuxi DIS Microelectronics | 2025-06-25 | Full range of Y Detector Raw during Y Scan sweep | 14.99
| Wuxi DIS Microelectronics | 2025-06-25 | X Detector stroke calibration error (nominal scale error X) | 0.051
| Wuxi DIS Microelectronics | 2025-06-25 | Y Detector stroke calibration error (nominal scale error Y) | 0.051
| Wuxi DIS Microelectronics | 2025-06-25 | Full range of X Detector during X Scan sweep | 63.2241
| Wuxi DIS Microelectronics | 2025-06-25 | Full range of Y Detector during Y Scan sweep | 63.6377
| Wuxi DIS Microelectronics | 2025-06-25 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 82.8509
| Wuxi DIS Microelectronics | 2025-06-25 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 80.5737
| Wuxi DIS Microelectronics | 2025-06-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 62.8689
| Wuxi DIS Microelectronics | 2025-06-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 62.2318
| Wuxi DIS Microelectronics | 2025-06-25 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.5
| Wuxi DIS Microelectronics | 2025-06-25 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.5
| Wuxi DIS Microelectronics | 2025-06-25 | X axis OPM | 1.498
| Wuxi DIS Microelectronics | 2025-06-25 | Y axis OPM | 0.936
| Wuxi DIS Microelectronics | 2025-06-25 | RMS error after Stage Calibration | 1.6
| Wuxi DIS Microelectronics | 2025-06-25 | Scan offset accuracy (Precision) | 0.86516
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (X offset-Point Align1) | 0.39
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (Y offset-Point Align1) | 0.61
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (X offset-Point Align2) | 0.71
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (Y offset-Point Align2) | 0.93
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (X offset-Point Align3) | 0.7
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (Y offset-Point Align3) | 1.07
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (X offset) | -0.71
| Wuxi DIS Microelectronics | 2025-06-25 | Aligned stage accuracy (Y offset) | 0.68
| Wuxi DIS Microelectronics | 2025-06-25 | Imag offset after EFEM loading (X scan) | 39.0625
| Wuxi DIS Microelectronics | 2025-06-25 | Imag offset after EFEM loading (Y scan) | 87.8906
| Wuxi DIS Microelectronics | 2025-06-25 | Imag offset after Lock stage (X scan) | 53.7109
| Wuxi DIS Microelectronics | 2025-06-25 | Imag offset after Lock stage (Y scan) | 34.1797
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from -X to +X (X scan) | 39.0625
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from -Y to +Y (Y scan) | 53.7109
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from -X to +X (Y scan) | 73.2422
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from -Y to +Y (X scan) | 43.9453
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from +X to -X (X scan) | 34.1797
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from +Y to -Y (Y scan) | 14.6484
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from +X to -X (Y scan) | 58.5938
| Wuxi DIS Microelectronics | 2025-06-25 | Image offset after moving from +Y to -Y (X scan) | 34.1797
| Wuxi DIS Microelectronics | 2025-06-25 | Z Drift | 21.554
| Wuxi DIS Microelectronics | 2025-06-25 | Z-XY Orthogonality (X direction-Left) | 55.63
| Wuxi DIS Microelectronics | 2025-06-25 | Z-XY Orthogonality (X direction-Right) | 54.3
| Wuxi DIS Microelectronics | 2025-06-25 | Z-XY Orthogonality (Y direction-Left) | 53.87
| Wuxi DIS Microelectronics | 2025-06-25 | Z-XY Orthogonality (Y direction-Right) | 55.55
| Wuxi DIS Microelectronics | 2025-06-25 | Mask Roughness Repeatability (10Repeats, RSD%) | 0.0871
| Wuxi DIS Microelectronics | 2025-06-25 | Mask Roughness Repeatability (10Repeats, σ ) | 0.0431
| Wuxi DIS Microelectronics | 2025-06-25 | Mask exchange time | 4.11
| University of Iceland | 2025-06-24 | Lateral(C-D) sensitivity | 15.2
| University of Iceland | 2025-06-24 | Lateral(C-D) noise | 0.07
| University of Iceland | 2025-06-24 | A-B noise | 8.409
| University of Iceland | 2025-06-24 | NCM Amplitude noise | 12.114
| University of Iceland | 2025-06-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| University of Iceland | 2025-06-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| University of Iceland | 2025-06-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Iceland | 2025-06-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.09
| University of Iceland | 2025-06-24 | Roughness Repeatability 1x1um(RSD%) | 0.991
| University of Iceland | 2025-06-24 | Roughness Repeatability 1x1um(1σ) | 0.001
| University of Iceland | 2025-06-24 | Full range of Z Height during Z Scan sweep | 17.3
| University of Iceland | 2025-06-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.496
| University of Iceland | 2025-06-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.423
| University of Iceland | 2025-06-24 | Full range of Z Detector during Z Scan sweep | 12.6
| University of Iceland | 2025-06-24 | Z Drive calibration error | 0.23467
| University of Iceland | 2025-06-24 | Z Height calibration error | 0.025862
| University of Iceland | 2025-06-24 | NCM driving% with AC160TS cantilever | 3.686
| University of Iceland | 2025-06-24 | Z linearity by sample height | 0.98677
| University of Iceland | 2025-06-24 | 1) Z Scanner position at +3μm | 0.010669
| University of Iceland | 2025-06-24 | 3) Z Scanner position at -3μm | 0.005738
| University of Iceland | 2025-06-24 | Z Height noise | 0.029
| University of Iceland | 2025-06-24 | Full range of X Detector during X Scan sweep | 73.7
| University of Iceland | 2025-06-24 | Full range of Y Detector during Y Scan sweep | 73.8
| University of Iceland | 2025-06-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.4
| University of Iceland | 2025-06-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 112.9
| University of Iceland | 2025-06-24 | X Detector stroke calibration error (nominal scale error X) | 0.034
| University of Iceland | 2025-06-24 | Y Detector stroke calibration error (nominal scale error Y) | 0.409
| University of Iceland | 2025-06-24 | X Scan Stroke | 55
| University of Iceland | 2025-06-24 | Y Scan Stroke | 55
| University of Iceland | 2025-06-24 | Full range of X Detector Raw during X Scan sweep | 14.1
| University of Iceland | 2025-06-24 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| University of Iceland | 2025-06-24 | X Detector noise (XY Servo on) | 0.245
| University of Iceland | 2025-06-24 | Y Detector noise (XY Servo on) | 0.263
| University of Iceland | 2025-06-24 | X Detector noise (XY Servo off) | 0.237
| University of Iceland | 2025-06-24 | Y Detector noise (XY Servo off) | 0.254
| University of Iceland | 2025-06-24 | X axis OPM | 0.9906
| University of Iceland | 2025-06-24 | Y axis OPM | 0.8854
| University of Iceland | 2025-06-24 | X Non-linearity | 0.072
| University of Iceland | 2025-06-24 | Y Non-linearity | 0.112
| United Nova Technology | 2025-06-24 | Z Linearity by Z scanner position | 1.07
| United Nova Technology | 2025-06-24 | Z Linearity by 4 different sample height | 0.9955
| United Nova Technology | 2025-06-24 | Z Height noise | 0.028
| United Nova Technology | 2025-06-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.9928
| United Nova Technology | 2025-06-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.973
| United Nova Technology | 2025-06-24 | Z Height calibration error | 0.10971
| United Nova Technology | 2025-06-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2453
| United Nova Technology | 2025-06-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.1971
| United Nova Technology | 2025-06-24 | Z Drive calibration error | 0.086513
| United Nova Technology | 2025-06-24 | Z Drift | 6.645
| United Nova Technology | 2025-06-24 | Y2 Detector noise (XY Servo on) | 0.379
| United Nova Technology | 2025-06-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 273.4818
| United Nova Technology | 2025-06-24 | Y1 Detector noise (XY Servo on) | 0.284
| United Nova Technology | 2025-06-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 273.4818
| United Nova Technology | 2025-06-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 179.936
| United Nova Technology | 2025-06-24 | Y Ringing(Rpv) | 0.126
| United Nova Technology | 2025-06-24 | Y Non-linearity error | 2.66
| United Nova Technology | 2025-06-24 | Y Detector stroke calibration error (nominal scale error Y) | 0.487
| United Nova Technology | 2025-06-24 | Y Detector noise (XY Servo on) | 0.247
| United Nova Technology | 2025-06-24 | Y axis OPM | 0.578
| United Nova Technology | 2025-06-24 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.8701
| United Nova Technology | 2025-06-24 | XY Stage positioning Y offset (LLC Translation) | 2.5977
| United Nova Technology | 2025-06-24 | XY Stage positioning Y offset (In-Die Align) | 1.9873
| United Nova Technology | 2025-06-24 | XY Stage positioning Y offset (Global Align) | 6.7871
| United Nova Technology | 2025-06-24 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.2176
| United Nova Technology | 2025-06-24 | XY Stage positioning Y 1σ (LLC Translation) | 0.49634
| United Nova Technology | 2025-06-24 | XY Stage positioning Y 1σ (In-Die Align) | 0.36
| United Nova Technology | 2025-06-24 | XY Stage positioning Y 1σ (Global Align) | 1.0752
| United Nova Technology | 2025-06-24 | XY Stage positioning X offset (Vision Pattern Matching) | 1.9971
| United Nova Technology | 2025-06-24 | XY Stage positioning X offset (LLC Translation) | 1.9824
| United Nova Technology | 2025-06-24 | XY Stage positioning X offset (In-Die Align) | 1.5674
| United Nova Technology | 2025-06-24 | XY Stage positioning X offset (Global Align) | 1.9775
| United Nova Technology | 2025-06-24 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.4526
| United Nova Technology | 2025-06-24 | XY Stage positioning X 1σ (LLC Translation) | 0.45664
| United Nova Technology | 2025-06-24 | XY Stage positioning X 1σ (In-Die Align) | 0.3432
| United Nova Technology | 2025-06-24 | XY Stage positioning X 1σ (Global Align) | 0.44426
| United Nova Technology | 2025-06-24 | X2 Detector noise (XY Servo on) | 0.299
| United Nova Technology | 2025-06-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 285.6867
| United Nova Technology | 2025-06-24 | X1 Detector noise (XY Servo on) | 0.344
| United Nova Technology | 2025-06-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 292.3823
| United Nova Technology | 2025-06-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 181.026
| United Nova Technology | 2025-06-24 | X Ringing(Rpv) | 0.087
| United Nova Technology | 2025-06-24 | X Non-linearity error | 3.42
| United Nova Technology | 2025-06-24 | X Detector stroke calibration error (nominal scale error X) | 0.045
| United Nova Technology | 2025-06-24 | X Detector noise (XY Servo on) | 0.251
| United Nova Technology | 2025-06-24 | X axis OPM | 1.567
| United Nova Technology | 2025-06-24 | Wafer placement(Loading) error on the load port1 (Y offset) | 24.21
| United Nova Technology | 2025-06-24 | Wafer placement(Loading) error on the load port1 (X offset) | 16.09
| United Nova Technology | 2025-06-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.728
| United Nova Technology | 2025-06-24 | Roughness Repeatability 80x80um(RSD%) | 1.1572
| United Nova Technology | 2025-06-24 | Roughness Repeatability 80x80um(1σ) | 0.003433
| United Nova Technology | 2025-06-24 | Roughness Repeatability 1x1um(RSD%) | 4.5343
| United Nova Technology | 2025-06-24 | Roughness Repeatability 1x1um(1σ) | 0.00404
| United Nova Technology | 2025-06-24 | RMS of ZeroScan | 0.051
| United Nova Technology | 2025-06-24 | RMS error after Stage Calibration | 2.62
| United Nova Technology | 2025-06-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4e-05
| United Nova Technology | 2025-06-24 | OpticalAxisShiftPerFocusY | 0
| United Nova Technology | 2025-06-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00021
| United Nova Technology | 2025-06-24 | OpticalAxisShiftPerFocusX | -0.0021
| United Nova Technology | 2025-06-24 | Min. Vibration Transmissbility (~100Hz) | 0.023
| United Nova Technology | 2025-06-24 | Lateral(C-D) sensitivity | 8.6
| United Nova Technology | 2025-06-24 | Lateral(C-D) noise | 0.183
| United Nova Technology | 2025-06-24 | Image offset after moving from -Y to +Y (Y scan) | 10.9863
| United Nova Technology | 2025-06-24 | Image offset after moving from -Y to +Y (X scan) | 5.8594
| United Nova Technology | 2025-06-24 | Image offset after moving from -X to +X (Y scan) | 74.707
| United Nova Technology | 2025-06-24 | Image offset after moving from -X to +X (X scan) | 93.0176
| United Nova Technology | 2025-06-24 | Image offset after moving from +Y to -Y (Y scan) | 61.5234
| United Nova Technology | 2025-06-24 | Image offset after moving from +Y to -Y (X scan) | 5.1269
| United Nova Technology | 2025-06-24 | Image offset after moving from +X to -X (Y scan) | 17.5781
| United Nova Technology | 2025-06-24 | Image offset after moving from +X to -X (X scan) | 90.8203
| United Nova Technology | 2025-06-24 | Full range of Z Height during Z Scan sweep | 15.9316
| United Nova Technology | 2025-06-24 | Full range of Z Detector during Z Scan sweep | 10.99
| United Nova Technology | 2025-06-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.24
| United Nova Technology | 2025-06-24 | Full range of Y2 Detector during Y Scan sweep | 185.0321
| United Nova Technology | 2025-06-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.51
| United Nova Technology | 2025-06-24 | Full range of Y1 Detector during Y Scan sweep | 190.38
| United Nova Technology | 2025-06-24 | Full range of Y Detector during Y Scan sweep | 187.7061
| United Nova Technology | 2025-06-24 | Full range of X2 Detector Raw during X Scan sweep | 14.39
| United Nova Technology | 2025-06-24 | Full range of X2 Detector during X Scan sweep | 203.1832
| United Nova Technology | 2025-06-24 | Full range of X1 Detector Raw during X Scan sweep | 14.13
| United Nova Technology | 2025-06-24 | Full range of X1 Detector during X Scan sweep | 171.8248
| United Nova Technology | 2025-06-24 | Full range of X Detector during X Scan sweep | 187.504
| United Nova Technology | 2025-06-24 | Cycle time for 15-point measurement | 1743.104
| United Nova Technology | 2025-06-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| United Nova Technology | 2025-06-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| United Nova Technology | 2025-06-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| United Nova Technology | 2025-06-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.21
| Silicon Labs | 2025-06-23 | SLD current limit check | 148
| Silicon Labs | 2025-06-23 | SLD beam power at the end of the lens | 0.28
| Silicon Labs | 2025-06-23 | Lateral(C-D) sensitivity | 8
| Silicon Labs | 2025-06-23 | Lateral(C-D) noise | 0.105
| Silicon Labs | 2025-06-23 | A-B noise | 14.432
| Silicon Labs | 2025-06-23 | NCM Amplitude noise | 5.292
| Silicon Labs | 2025-06-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Silicon Labs | 2025-06-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Silicon Labs | 2025-06-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Silicon Labs | 2025-06-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Silicon Labs | 2025-06-23 | Roughness Repeatability 1x1um(RSD%) | 2.698
| Silicon Labs | 2025-06-23 | Roughness Repeatability 1x1um(1σ) | 0.006
| Silicon Labs | 2025-06-23 | Full range of Z Height during Z Scan sweep | 17.2
| Silicon Labs | 2025-06-23 | Full range of Z Detector during Z Scan sweep | 14
| Silicon Labs | 2025-06-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.5
| Silicon Labs | 2025-06-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.617
| Silicon Labs | 2025-06-23 | Z Drive calibrtaion error | 0.31034
| Silicon Labs | 2025-06-23 | Z Height calibrtaion error | 0.032567
| Silicon Labs | 2025-06-23 | NCM driving% with AC160TS cantilever | 0.72
| Silicon Labs | 2025-06-23 | Z linearity by sample height | 0.98976
| Silicon Labs | 2025-06-23 | 1) Z Scanner position at +3μm | 0.006306
| Silicon Labs | 2025-06-23 | 3) Z Scanner position at -3μm | 0.001432
| Silicon Labs | 2025-06-23 | Z Height noise | 0.017
| Silicon Labs | 2025-06-23 | Full range of X Detector during X Scan sweep | 107.3
| Silicon Labs | 2025-06-23 | Full range of Y Detector during Y Scan sweep | 105.9
| Silicon Labs | 2025-06-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.1
| Silicon Labs | 2025-06-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 158.6
| Silicon Labs | 2025-06-23 | X Detector stroke calibration error (nominal scale error X) | 0.02
| Silicon Labs | 2025-06-23 | Y Detector stroke calibration error (nominal scale error Y) | 0.327
| Silicon Labs | 2025-06-23 | X Scan Stroke | 107.3
| Silicon Labs | 2025-06-23 | Y Scan Stroke | 105.9
| Silicon Labs | 2025-06-23 | Full range of X Detector Raw during X Scan sweep | 13.9
| Silicon Labs | 2025-06-23 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| Silicon Labs | 2025-06-23 | X Detector noise (XYServo off) | 0.239
| Silicon Labs | 2025-06-23 | Y Detector noise (XYServo off) | 0.243
| Silicon Labs | 2025-06-23 | X Detector noise (XY Servo on) | 0.247
| Silicon Labs | 2025-06-23 | Y Detector noise (XY Servo on) | 0.251
| Silicon Labs | 2025-06-23 | X axis OPM | 1.1394
| Silicon Labs | 2025-06-23 | Y axis OPM | 1.209
| Silicon Labs | 2025-06-23 | X Non-linearity | 0.117
| Silicon Labs | 2025-06-23 | Y Non-linearity | 0.127
| Henkel Corporation, Ltd. | 2025-06-17 | Lateral(C-D) sensitivity | 18.4
| Henkel Corporation, Ltd. | 2025-06-17 | Lateral(C-D) noise | 0.112
| Henkel Corporation, Ltd. | 2025-06-17 | A-B noise | 7.299
| Henkel Corporation, Ltd. | 2025-06-17 | NCM Amplitude noise | 14.635
| Henkel Corporation, Ltd. | 2025-06-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Henkel Corporation, Ltd. | 2025-06-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Henkel Corporation, Ltd. | 2025-06-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Henkel Corporation, Ltd. | 2025-06-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Henkel Corporation, Ltd. | 2025-06-17 | Roughness Repeatability 1x1um(RSD%) | 0.881
| Henkel Corporation, Ltd. | 2025-06-17 | Roughness Repeatability 1x1um(1σ) | 0.002
| Henkel Corporation, Ltd. | 2025-06-17 | Full range of Z Height during Z Scan sweep | 17.4
| Henkel Corporation, Ltd. | 2025-06-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.549
| Henkel Corporation, Ltd. | 2025-06-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.605
| Henkel Corporation, Ltd. | 2025-06-17 | Full range of Z Detector during Z Scan sweep | 12.6
| Henkel Corporation, Ltd. | 2025-06-17 | Z Drive calibration error | 0.081418
| Henkel Corporation, Ltd. | 2025-06-17 | Z Height calibration error | 0.12835
| Henkel Corporation, Ltd. | 2025-06-17 | NCM driving% with AC160TS cantilever | 0.329
| Henkel Corporation, Ltd. | 2025-06-17 | Z linearity by sample height | 0.99746
| Henkel Corporation, Ltd. | 2025-06-17 | 1) Z Scanner position at +3μm | 0.005115
| Henkel Corporation, Ltd. | 2025-06-17 | 3) Z Scanner position at -3μm | 0.001248
| Henkel Corporation, Ltd. | 2025-06-17 | Z Height noise | 0.03
| Henkel Corporation, Ltd. | 2025-06-17 | Full range of X Detector during X Scan sweep | 114.7
| Henkel Corporation, Ltd. | 2025-06-17 | Full range of Y Detector during Y Scan sweep | 112.6
| Henkel Corporation, Ltd. | 2025-06-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 164.4
| Henkel Corporation, Ltd. | 2025-06-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 173.1
| Henkel Corporation, Ltd. | 2025-06-17 | X Detector stroke calibration error (nominal scale error X) | 0.003
| Henkel Corporation, Ltd. | 2025-06-17 | Y Detector stroke calibration error (nominal scale error Y) | 0.035
| Henkel Corporation, Ltd. | 2025-06-17 | X Scan Stroke | 113.3
| Henkel Corporation, Ltd. | 2025-06-17 | Y Scan Stroke | 113.2
| Henkel Corporation, Ltd. | 2025-06-17 | Full range of X Detector Raw during X Scan sweep | 14
| Henkel Corporation, Ltd. | 2025-06-17 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| Henkel Corporation, Ltd. | 2025-06-17 | X Detector noise (XY Servo on) | 0.253
| Henkel Corporation, Ltd. | 2025-06-17 | Y Detector noise (XY Servo on) | 0.266
| Henkel Corporation, Ltd. | 2025-06-17 | X Detector noise (XY servo off) | 0.254
| Henkel Corporation, Ltd. | 2025-06-17 | Y Detector noise (XY servo off) | 0.263
| Henkel Corporation, Ltd. | 2025-06-17 | X axis OPM | 1.52
| Henkel Corporation, Ltd. | 2025-06-17 | Y axis OPM | 1.71
| Henkel Corporation, Ltd. | 2025-06-17 | X Non-linearity | 0.22
| Henkel Corporation, Ltd. | 2025-06-17 | Y Non-linearity | 0.2
| 에이티아이 | 2025-06-17 | Lateral(C-D) sensitivity | 76.5
| 에이티아이 | 2025-06-17 | Lateral(C-D) noise | 0.144
| 에이티아이 | 2025-06-17 | A-B noise | 7.667
| 에이티아이 | 2025-06-17 | NCM Amplitude noise | 18.186
| 에이티아이 | 2025-06-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| 에이티아이 | 2025-06-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 에이티아이 | 2025-06-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| 에이티아이 | 2025-06-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 에이티아이 | 2025-06-17 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.25
| 에이티아이 | 2025-06-17 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.29
| 에이티아이 | 2025-06-17 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.27
| 에이티아이 | 2025-06-17 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.28
| 에이티아이 | 2025-06-17 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| 에이티아이 | 2025-06-17 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| 에이티아이 | 2025-06-17 | Roughness Repeatability 1x1um(RSD%) | 1.224
| 에이티아이 | 2025-06-17 | Roughness Repeatability 1x1um(1σ) | 0.003
| 에이티아이 | 2025-06-17 | Full range of Z Height during Z Scan sweep | 17.4
| 에이티아이 | 2025-06-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.068
| 에이티아이 | 2025-06-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.547
| 에이티아이 | 2025-06-17 | Full range of Z Detector during Z Scan sweep | 12.8
| 에이티아이 | 2025-06-17 | Z Drive calibration error | 0.049756
| 에이티아이 | 2025-06-17 | Z Height calibration error | 0.065366
| 에이티아이 | 2025-06-17 | NCM driving% with AC160TS cantilever | 2.306
| 에이티아이 | 2025-06-17 | Z linearity by sample height | 1.0013
| 에이티아이 | 2025-06-17 | 1) Z Scanner position at +3μm | 0.002671
| 에이티아이 | 2025-06-17 | 3) Z Scanner position at -3μm | 0.009613
| 에이티아이 | 2025-06-17 | Z Height noise | 0.028
| 에이티아이 | 2025-06-17 | Full range of X Detector during X Scan sweep | 212
| 에이티아이 | 2025-06-17 | Full range of X1 Detector during X Scan sweep | 211.6
| 에이티아이 | 2025-06-17 | Full range of X2 Detector during X Scan sweep | 212.4
| 에이티아이 | 2025-06-17 | Full range of Y Detector during Y Scan sweep | 210.8
| 에이티아이 | 2025-06-17 | Full range of Y1 Detector during Y Scan sweep | 206.4
| 에이티아이 | 2025-06-17 | Full range of Y2 Detector during Y Scan sweep | 215
| 에이티아이 | 2025-06-17 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 285.1
| 에이티아이 | 2025-06-17 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 285.5
| 에이티아이 | 2025-06-17 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 285.1
| 에이티아이 | 2025-06-17 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 296.2
| 에이티아이 | 2025-06-17 | X Detector stroke calibration error (nominal scale error X) | 0.367
| 에이티아이 | 2025-06-17 | Y Detector stroke calibration error (nominal scale error Y) | 0.255
| 에이티아이 | 2025-06-17 | X Scan Stroke | 211.8
| 에이티아이 | 2025-06-17 | Y Scan Stroke | 210.7
| 에이티아이 | 2025-06-17 | Full range of X1 Detector Raw during X Scan sweep | 14.8
| 에이티아이 | 2025-06-17 | Full range of X2 Detector Raw during X Scan sweep | 14.9
| 에이티아이 | 2025-06-17 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| 에이티아이 | 2025-06-17 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| 에이티아이 | 2025-06-17 | X Detector noise (XY servo off) | 0.29
| 에이티아이 | 2025-06-17 | X1 Detector noise (XY servo off) | 0.373
| 에이티아이 | 2025-06-17 | X2 Detector noise (XY servo off) | 0.418
| 에이티아이 | 2025-06-17 | Y Detector noise (XY servo off) | 0.291
| 에이티아이 | 2025-06-17 | Y1 Detector noise (XY servo off) | 0.389
| 에이티아이 | 2025-06-17 | Y2 Detector noise (XY servo off) | 0.413
| 에이티아이 | 2025-06-17 | X Detector noise (XY servo on) | 0.3
| 에이티아이 | 2025-06-17 | X1 Detector noise (XY servo on) | 0.383
| 에이티아이 | 2025-06-17 | X2 Detector noise (XY servo on) | 0.429
| 에이티아이 | 2025-06-17 | Y Detector noise (XY servo on) | 0.311
| 에이티아이 | 2025-06-17 | Y1 Detector noise (XY servo on) | 0.405
| 에이티아이 | 2025-06-17 | Y2 Detector noise (XY servo on) | 0.432
| 에이티아이 | 2025-06-17 | X axis OPM | 3.542
| 에이티아이 | 2025-06-17 | Y axis OPM | 2.281
| 에이티아이 | 2025-06-17 | X Non-linearity | 0.132
| 에이티아이 | 2025-06-17 | Y Non-linearity | 0.113
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Scanner position at -3μm | 0.00252
| 도레이첨담소재 | 2025-06-16 | Z Scanner position at -3μm | 0.00201
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Scanner position at +3μm | 0.004235
| 도레이첨담소재 | 2025-06-16 | Z Scanner position at +3μm | 0.001023
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z linearity by sample height | 1.0106
| 도레이첨담소재 | 2025-06-16 | Z linearity by sample height | 0.99159
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Height noise | 0.028
| 도레이첨담소재 | 2025-06-16 | Z Height noise | 0.029
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Height calibration error | 0.047893
| 도레이첨담소재 | 2025-06-16 | Z Height calibration error | 0.022989
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Drive calibration error | 0.35345
| 도레이첨담소재 | 2025-06-16 | Z Drive calibration error | 0.13218
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Scan Stroke | 130.4
| 도레이첨담소재 | 2025-06-16 | Y Scan Stroke | 55
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Non-linearity | 0.046
| 도레이첨담소재 | 2025-06-16 | Y Non-linearity | 0.057
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.095
| 도레이첨담소재 | 2025-06-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.292
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Detector noise (XY Servo on) | 0.421
| 도레이첨담소재 | 2025-06-16 | Y Detector noise (XY Servo on) | 0.249
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Detector noise (XY Servo off) | 0.476
| 도레이첨담소재 | 2025-06-16 | Y Detector noise (XY Servo off) | 0.241
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y axis OPM | 1.4742
| 도레이첨담소재 | 2025-06-16 | Y axis OPM | 0.9538
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Scan Stroke | 123.1
| 도레이첨담소재 | 2025-06-16 | X Scan Stroke | 55
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Non-linearity | 0.153
| 도레이첨담소재 | 2025-06-16 | X Non-linearity | 0.043
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Detector stroke calibration error (nominal scale error X) | 0.162
| 도레이첨담소재 | 2025-06-16 | X Detector stroke calibration error (nominal scale error X) | 0.398
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Detector noise (XY Servo on) | 0.404
| 도레이첨담소재 | 2025-06-16 | X Detector noise (XY Servo on) | 0.245
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Detector noise (XY Servo off) | 0.45
| 도레이첨담소재 | 2025-06-16 | X Detector noise (XY Servo off) | 0.239
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X axis OPM | 0.949
| 도레이첨담소재 | 2025-06-16 | X axis OPM | 0.9692
| 도레이첨담소재 | 2025-06-16 | Roughness Repeatability 1x1um(RSD%) | 1.317
| 도레이첨담소재 | 2025-06-16 | Roughness Repeatability 1x1um(1σ) | 0.004
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| 도레이첨담소재 | 2025-06-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.57
| 도레이첨담소재 | 2025-06-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 도레이첨담소재 | 2025-06-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| 도레이첨담소재 | 2025-06-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | NCM driving% with AC160TS cantilever | 0.458
| 도레이첨담소재 | 2025-06-16 | NCM driving% with AC160TS cantilever | 0.59
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | NCM Amplitude noise | 14.69
| 도레이첨담소재 | 2025-06-16 | NCM Amplitude noise | 10.737
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.919
| 도레이첨담소재 | 2025-06-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.568
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.513
| 도레이첨담소재 | 2025-06-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.589
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 183.1
| 도레이첨담소재 | 2025-06-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.8
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 178.3
| 도레이첨담소재 | 2025-06-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.8
| 도레이첨담소재 | 2025-06-16 | Lateral(C-D) sensitivity | 84.6
| 도레이첨담소재 | 2025-06-16 | Lateral(C-D) noise | 0.073
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Z Height during Z Scan sweep | 17.9
| 도레이첨담소재 | 2025-06-16 | Full range of Z Height during Z Scan sweep | 17.3
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Z Detector during Z Scan sweep | 13
| 도레이첨담소재 | 2025-06-16 | Full range of Z Detector during Z Scan sweep | 12.1
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| 도레이첨담소재 | 2025-06-16 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Y Detector during Y Scan sweep | 130.4
| 도레이첨담소재 | 2025-06-16 | Full range of Y Detector during Y Scan sweep | 67.6
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of X Detector Raw during X Scan sweep | 13.8
| 도레이첨담소재 | 2025-06-16 | Full range of X Detector Raw during X Scan sweep | 13.7
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of X Detector during X Scan sweep | 123.1
| 도레이첨담소재 | 2025-06-16 | Full range of X Detector during X Scan sweep | 67.6
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | A-B noise | 7.202
| 도레이첨담소재 | 2025-06-16 | A-B noise | 5.568
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | A-B noise | 7.202
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | NCM Amplitude noise | 14.69
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.57
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Z Height during Z Scan sweep | 17.9
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.513
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.919
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Z Detector during Z Scan sweep | 13
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Drive calibration error | 0.35345
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Height calibration error | 0.047893
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | NCM driving% with AC160TS cantilever | 0.458
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z linearity by sample height | 1.0106
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | 1) Z Scanner position at +3μm | 0.004235
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | 3) Z Scanner position at -3μm | 0.00252
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Z Height noise | 0.028
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of X Detector during X Scan sweep | 123.1
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Y Detector during Y Scan sweep | 130.4
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 178.3
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 183.1
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Detector stroke calibration error (nominal scale error X) | 0.162
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.095
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Scan Stroke | 123.1
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Scan Stroke | 130.4
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of X Detector Raw during X Scan sweep | 13.8
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Detector noise (XY Servo on) | 0.404
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Detector noise (XY Servo on) | 0.421
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Detector noise (XY Servo off) | 0.45
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Detector noise (XY Servo off) | 0.476
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X axis OPM | 0.949
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y axis OPM | 1.4742
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | X Non-linearity | 0.153
| Xidian University Guangzhou Institute of Technology | 2025-06-16 | Y Non-linearity | 0.046
| TSMC AP3(400TF) | 2025-06-15 | Z-XY Orthogonality (Y direction-Right) | 55.37
| TSMC AP3(400TF) | 2025-06-15 | Z-XY Orthogonality (Y direction-Left) | 53.77
| TSMC AP3(400TF) | 2025-06-15 | Z-XY Orthogonality (X direction-Right) | 54.75
| TSMC AP3(400TF) | 2025-06-15 | Z-XY Orthogonality (X direction-Left) | 54.19
| TSMC AP3(400TF) | 2025-06-15 | Z Linearity by Z scanner position | 0.99
| TSMC AP3(400TF) | 2025-06-15 | Z Linearity by 4 different sample height | 1.003
| TSMC AP3(400TF) | 2025-06-15 | Z Height noise | 0.027
| TSMC AP3(400TF) | 2025-06-15 | Z Height calibration error | 0.017954
| TSMC AP3(400TF) | 2025-06-15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.493
| TSMC AP3(400TF) | 2025-06-15 | Z Drive calibration error | 0.040883
| TSMC AP3(400TF) | 2025-06-15 | Z Drift | 13.808
| TSMC AP3(400TF) | 2025-06-15 | Z Detector offset | 57.3
| TSMC AP3(400TF) | 2025-06-15 | Y Ringing(Rpv) | 0.229
| TSMC AP3(400TF) | 2025-06-15 | Y Detector stroke calibration error (nominal scale error Y) | 0.254
| TSMC AP3(400TF) | 2025-06-15 | Y detector noise (XY Servo on) | 0.234
| TSMC AP3(400TF) | 2025-06-15 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 155.863
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Repeatability @1st position | 0.06
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Repeatability @ 4th position | 0.06
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Repeatability @ 3rd position | 0.07
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Repeatability @ 2nd position | 0.06
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Accuracy @1st position | 0.97
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Accuracy @ 4th position | 1.18
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Accuracy @ 3rd position | 0.92
| TSMC AP3(400TF) | 2025-06-15 | Y axis, Accuracy @ 2nd position | 0.63
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y offset (Vision Pattern Matching) | 0.918
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y offset (LLC Translation) | 1.6846
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y offset (In-Die Align) | 2.4121
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y offset (Global Align) | 3.999
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.1967
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y 1σ (LLC Translation) | 0.3799
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y 1σ (In-Die Align) | 0.4264
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning Y 1σ (Global Align) | 0.9671
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X offset (Vision Pattern Matching) | 1.2402
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X offset (LLC Translation) | 1.7041
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X offset (In-Die Align) | 2.5977
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X offset (Global Align) | 1.5527
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.2648
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X 1σ (LLC Translation) | 0.365
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X 1σ (In-Die Align) | 0.4923
| TSMC AP3(400TF) | 2025-06-15 | XY Stage positioning X 1σ (Global Align) | 0.3545
| TSMC AP3(400TF) | 2025-06-15 | X Ringing(Rpv) | 0.194
| TSMC AP3(400TF) | 2025-06-15 | X Detector stroke calibration error (nominal scale error X) | 0.284
| TSMC AP3(400TF) | 2025-06-15 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 159.078
| TSMC AP3(400TF) | 2025-06-15 | X axis, Repeatability @1st position | 0.06
| TSMC AP3(400TF) | 2025-06-15 | X axis, Repeatability @ 4th position | 0.07
| TSMC AP3(400TF) | 2025-06-15 | X axis, Repeatability @ 3rd position | 0.04
| TSMC AP3(400TF) | 2025-06-15 | X axis, Repeatability @ 2nd position | 0.05
| TSMC AP3(400TF) | 2025-06-15 | X axis, Accuracy @1st position | 0.47
| TSMC AP3(400TF) | 2025-06-15 | X axis, Accuracy @ 4th position | 0.4
| TSMC AP3(400TF) | 2025-06-15 | X axis, Accuracy @ 3rd position | 0.7
| TSMC AP3(400TF) | 2025-06-15 | X axis, Accuracy @ 2nd position | 0.47
| TSMC AP3(400TF) | 2025-06-15 | Wafer placement(Loading) error on the load port2 (Y offset) | 38.4
| TSMC AP3(400TF) | 2025-06-15 | Wafer placement(Loading) error on the load port2 (X offset) | 94.16
| TSMC AP3(400TF) | 2025-06-15 | Wafer placement(Loading) error on the load port1 (Y offset) | 96.22
| TSMC AP3(400TF) | 2025-06-15 | Wafer placement(Loading) error on the load port1 (X offset) | 27.9
| TSMC AP3(400TF) | 2025-06-15 | Torque Profile (50mm, 20um/s) | 0.947
| TSMC AP3(400TF) | 2025-06-15 | Torque Profile (50mm, 200um/s) | 0.446
| TSMC AP3(400TF) | 2025-06-15 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.256
| TSMC AP3(400TF) | 2025-06-15 | Roughness Repeatability 80x80um(RSD%) | 0.746
| TSMC AP3(400TF) | 2025-06-15 | Roughness Repeatability 80x80um(1σ) | 0.005
| TSMC AP3(400TF) | 2025-06-15 | Roughness Repeatability 1x1um(RSD%) | 3.805
| TSMC AP3(400TF) | 2025-06-15 | Roughness Repeatability 1x1um(1σ) | 0.005
| TSMC AP3(400TF) | 2025-06-15 | RMS of ZeroScan | 0.047
| TSMC AP3(400TF) | 2025-06-15 | RMS error after Stage Calibration | 4.73
| TSMC AP3(400TF) | 2025-06-15 | Profiler noise @200um | 0.092
| TSMC AP3(400TF) | 2025-06-15 | Position repeatability(10repeats, Position Based) | 0.42
| TSMC AP3(400TF) | 2025-06-15 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00013
| TSMC AP3(400TF) | 2025-06-15 | OpticalAxisShiftPerFocusY | -0.0009
| TSMC AP3(400TF) | 2025-06-15 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00016
| TSMC AP3(400TF) | 2025-06-15 | OpticalAxisShiftPerFocusX | -0.0042
| TSMC AP3(400TF) | 2025-06-15 | OPM Repeatability @5mm | 1.896
| TSMC AP3(400TF) | 2025-06-15 | OPM Repeatability @25mm | 5.501
| TSMC AP3(400TF) | 2025-06-15 | OPM Repeatability @1mm | 1.047
| TSMC AP3(400TF) | 2025-06-15 | OPM Repeatability @10mm | 4.493
| TSMC AP3(400TF) | 2025-06-15 | Off-axis Camera X Position(ΔX) | -103327
| TSMC AP3(400TF) | 2025-06-15 | Min. Vibration Transmissbility (~100Hz) | 0.036
| TSMC AP3(400TF) | 2025-06-15 | Max OPM @5mm | 20.732
| TSMC AP3(400TF) | 2025-06-15 | Max OPM @25mm | 36.485
| TSMC AP3(400TF) | 2025-06-15 | Max OPM @1mm | 6.875
| TSMC AP3(400TF) | 2025-06-15 | Max OPM @10mm | 23.118
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from -Y to +Y (Y scan) | 69.5801
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from -Y to +Y (X scan) | 22.7051
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from -X to +X (Y scan) | 16.8457
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from -X to +X (X scan) | 83.4909
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from +Y to -Y (Y scan) | 23.4375
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from +Y to -Y (X scan) | 19.7754
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from +X to -X (Y scan) | 21.9727
| TSMC AP3(400TF) | 2025-06-15 | Image offset after moving from +X to -X (X scan) | 52.7344
| TSMC AP3(400TF) | 2025-06-15 | FullPixelHeight | 1828
| TSMC AP3(400TF) | 2025-06-15 | Full range of Z Height during Z Scan sweep | 18.27
| TSMC AP3(400TF) | 2025-06-15 | Full range of Z Detector during Z Scan sweep | 13.5
| TSMC AP3(400TF) | 2025-06-15 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| TSMC AP3(400TF) | 2025-06-15 | Full range of X Detector Raw during X Scan sweep | 14.54
| TSMC AP3(400TF) | 2025-06-15 | Cycle time for 15-point measurement | 1519.478
| TSMC AP3(400TF) | 2025-06-15 | Check the full stroke image(servo on) | Yes
| TSMC AP3(400TF) | 2025-06-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| TSMC AP3(400TF) | 2025-06-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.75
| TSMC AP3(400TF) | 2025-06-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| TSMC AP3(400TF) | 2025-06-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| TSMC AP3(400TF) | 2025-06-15 | A-B after Auto PSPD Alignment | 0.009
| Intel Hillsboro#3 | 2025-06-13 | Z Linearity by Z scanner position | 0.84
| PSE demotool | 2025-06-13 | Z Linearity by Z scanner position | 0.3752
| Intel Hillsboro#3 | 2025-06-13 | Z Linearity by 4 different sample height | 1.0028
| PSE demotool | 2025-06-13 | Z Linearity by 4 different sample height | 0.99943
| Intel Hillsboro#3 | 2025-06-13 | Z Height noise | 0.062
| PSE demotool | 2025-06-13 | Z Height noise | 0.026
| Intel Hillsboro#3 | 2025-06-13 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 61.524
| PSE demotool | 2025-06-13 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.4936
| Intel Hillsboro#3 | 2025-06-13 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 61.5417
| PSE demotool | 2025-06-13 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.4864
| Intel Hillsboro#3 | 2025-06-13 | Z Height calibration error | 0.000434
| PSE demotool | 2025-06-13 | Z Height calibration error | 0.026019
| Intel Hillsboro#3 | 2025-06-13 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 28.312
| PSE demotool | 2025-06-13 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7315
| Intel Hillsboro#3 | 2025-06-13 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 28.3237
| PSE demotool | 2025-06-13 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6876
| Intel Hillsboro#3 | 2025-06-13 | Z Drive calibration error | 0.020598
| PSE demotool | 2025-06-13 | Z Drive calibration error | 0.018864
| Intel Hillsboro#3 | 2025-06-13 | Z Drift | 19.15
| PSE demotool | 2025-06-13 | Z Drift | 18.513
| Intel Hillsboro#3 | 2025-06-13 | Z Detector Min | -0.333
| Intel Hillsboro#3 | 2025-06-13 | Z Detector Max | 0.41765
| Intel Hillsboro#3 | 2025-06-13 | Y2 Detector noise (XY Servo on) | 0.455
| PSE demotool | 2025-06-13 | Y2 Detector noise (XY Servo on) | 0.35
| Intel Hillsboro#3 | 2025-06-13 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 271.877
| PSE demotool | 2025-06-13 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 273.08
| Intel Hillsboro#3 | 2025-06-13 | Y1 Detector noise (XY Servo on) | 0.442
| PSE demotool | 2025-06-13 | Y1 Detector noise (XY Servo on) | 0.513
| Intel Hillsboro#3 | 2025-06-13 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 262.855
| PSE demotool | 2025-06-13 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 282.099
| Intel Hillsboro#3 | 2025-06-13 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 171.751
| PSE demotool | 2025-06-13 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 175.129
| Intel Hillsboro#3 | 2025-06-13 | Y Ringing(Rpv) | 0.108
| PSE demotool | 2025-06-13 | Y Ringing(Rpv) | 0.085
| Intel Hillsboro#3 | 2025-06-13 | Y Non-linearity error | 2.38
| PSE demotool | 2025-06-13 | Y Non-linearity error | 3.72
| Intel Hillsboro#3 | 2025-06-13 | Y Detector stroke calibration error (nominal scale error Y) | 0.145
| PSE demotool | 2025-06-13 | Y Detector stroke calibration error (nominal scale error Y) | 0.162
| Intel Hillsboro#3 | 2025-06-13 | Y Detector noise (XY Servo on) | 0.333
| PSE demotool | 2025-06-13 | Y Detector noise (XY Servo on) | 0.34
| Intel Hillsboro#3 | 2025-06-13 | Y axis OPM | 0.941
| PSE demotool | 2025-06-13 | Y axis OPM | 3.163
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.2656
| PSE demotool | 2025-06-13 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.417
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning Y offset (LLC Translation) | 2.7588
| PSE demotool | 2025-06-13 | XY Stage positioning Y offset (LLC Translation) | 2.334
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning Y offset (In-Die Align) | 5.24
| PSE demotool | 2025-06-13 | XY Stage positioning Y offset (In-Die Align) | 2.3788
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning Y offset (Global Align) | 9.7363
| PSE demotool | 2025-06-13 | XY Stage positioning Y offset (Global Align) | 6.1572
| PSE demotool | 2025-06-13 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3836
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning Y 1σ (LLC Translation) | 0.56
| PSE demotool | 2025-06-13 | XY Stage positioning Y 1σ (LLC Translation) | 0.597
| PSE demotool | 2025-06-13 | XY Stage positioning Y 1σ (In-Die Align) | 0.5318
| PSE demotool | 2025-06-13 | XY Stage positioning Y 1σ (Global Align) | 1.4589
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning X offset (Vision Pattern Matching) | 1.8311
| PSE demotool | 2025-06-13 | XY Stage positioning X offset (Vision Pattern Matching) | 1.8115
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning X offset (LLC Translation) | 1.8262
| PSE demotool | 2025-06-13 | XY Stage positioning X offset (LLC Translation) | 2.1436
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning X offset (In-Die Align) | 2.7246
| PSE demotool | 2025-06-13 | XY Stage positioning X offset (In-Die Align) | 1.4453
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning X offset (Global Align) | 4.1992
| PSE demotool | 2025-06-13 | XY Stage positioning X offset (Global Align) | 3.1787
| PSE demotool | 2025-06-13 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.3338
| Intel Hillsboro#3 | 2025-06-13 | XY Stage positioning X 1σ (LLC Translation) | 0.3811
| PSE demotool | 2025-06-13 | XY Stage positioning X 1σ (LLC Translation) | 0.4857
| PSE demotool | 2025-06-13 | XY Stage positioning X 1σ (In-Die Align) | 0.2784
| PSE demotool | 2025-06-13 | XY Stage positioning X 1σ (Global Align) | 0.7746
| Intel Hillsboro#3 | 2025-06-13 | XY Orthogonality | 90.191
| PSE demotool | 2025-06-13 | XY Orthogonality | 90.019
| Intel Hillsboro#3 | 2025-06-13 | X2 Detector noise (XY Servo on) | 0.445
| PSE demotool | 2025-06-13 | X2 Detector noise (XY Servo on) | 0.49
| Intel Hillsboro#3 | 2025-06-13 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 256.915
| PSE demotool | 2025-06-13 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 270.039
| Intel Hillsboro#3 | 2025-06-13 | X1 Detector noise (XY Servo on) | 0.44
| PSE demotool | 2025-06-13 | X1 Detector noise (XY Servo on) | 0.315
| Intel Hillsboro#3 | 2025-06-13 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 263.223
| PSE demotool | 2025-06-13 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 290.225
| Intel Hillsboro#3 | 2025-06-13 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.291
| PSE demotool | 2025-06-13 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 177.314
| Intel Hillsboro#3 | 2025-06-13 | X Ringing(Rpv) | 0.127
| PSE demotool | 2025-06-13 | X Ringing(Rpv) | 0.04
| Intel Hillsboro#3 | 2025-06-13 | X Non-linearity error | 2.95
| PSE demotool | 2025-06-13 | X Non-linearity error | 3.08
| Intel Hillsboro#3 | 2025-06-13 | X Detector stroke calibration error (nominal scale error X) | 0.01
| PSE demotool | 2025-06-13 | X Detector stroke calibration error (nominal scale error X) | 0.361
| Intel Hillsboro#3 | 2025-06-13 | X Detector noise (XY Servo on) | 0.332
| PSE demotool | 2025-06-13 | X Detector noise (XY Servo on) | 0.307
| Intel Hillsboro#3 | 2025-06-13 | X axis OPM | 1.087
| PSE demotool | 2025-06-13 | X axis OPM | 1.433
| Intel Hillsboro#3 | 2025-06-13 | WLI_Scale(Z Calibration) | 29.0138
| PSE demotool | 2025-06-13 | WLI_Scale(C Calibration) | 33.705
| PSE demotool | 2025-06-13 | WLI_Scale(B Calibration) | 31.185
| PSE demotool | 2025-06-13 | WLI_Scale(A Calibration) | 33.667
| Intel Hillsboro#3 | 2025-06-13 | Wafer placement(Loading) error on the load port2 (Y offset) | 34.57
| PSE demotool | 2025-06-13 | Wafer placement(Loading) error on the load port2 (Y offset) | 11.85
| Intel Hillsboro#3 | 2025-06-13 | Wafer placement(Loading) error on the load port2 (X offset) | 31.2
| PSE demotool | 2025-06-13 | Wafer placement(Loading) error on the load port2 (X offset) | 13.6
| Intel Hillsboro#3 | 2025-06-13 | Wafer placement(Loading) error on the load port1 (Y offset) | 41.47
| PSE demotool | 2025-06-13 | Wafer placement(Loading) error on the load port1 (Y offset) | 26.3
| Intel Hillsboro#3 | 2025-06-13 | Wafer placement(Loading) error on the load port1 (X offset) | 18.75
| PSE demotool | 2025-06-13 | Wafer placement(Loading) error on the load port1 (X offset) | 31.4
| Intel Hillsboro#3 | 2025-06-13 | Torque Profile (50mm, 20um/s) | 1.602
| PSE demotool | 2025-06-13 | Torque Profile (50mm, 20um/s) | 0.584
| Intel Hillsboro#3 | 2025-06-13 | Torque Profile (50mm, 200um/s) | 0.821
| PSE demotool | 2025-06-13 | Torque Profile (50mm, 200um/s) | 0.438
| Intel Hillsboro#3 | 2025-06-13 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer) | 0.23
| Intel Hillsboro#3 | 2025-06-13 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer) | 1.27
| Intel Hillsboro#3 | 2025-06-13 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.68
| PSE demotool | 2025-06-13 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.36
| Intel Hillsboro#3 | 2025-06-13 | Roughness Repeatability 80x80um(RSD%) | 0.57927
| PSE demotool | 2025-06-13 | Roughness Repeatability 80x80um(RSD%) | 4.0697
| Intel Hillsboro#3 | 2025-06-13 | Roughness Repeatability 80x80um(1σ) | 0.002079
| PSE demotool | 2025-06-13 | Roughness Repeatability 80x80um(1σ) | 0.034544
| Intel Hillsboro#3 | 2025-06-13 | Roughness Repeatability 1x1um(RSD%) | 1.9095
| PSE demotool | 2025-06-13 | Roughness Repeatability 1x1um(RSD%) | 4.7801
| Intel Hillsboro#3 | 2025-06-13 | Roughness Repeatability 1x1um(1σ) | 0.003504
| PSE demotool | 2025-06-13 | Roughness Repeatability 1x1um(1σ) | 0.001874
| Intel Hillsboro#3 | 2025-06-13 | Roll motion to X+ full scan | 19.9
| PSE demotool | 2025-06-13 | Roll motion to X+ full scan | 20.4
| Intel Hillsboro#3 | 2025-06-13 | RMS of ZeroScan | 0.05
| PSE demotool | 2025-06-13 | RMS of ZeroScan | 0.039
| Intel Hillsboro#3 | 2025-06-13 | RMS error after Stage Calibration | 2.38
| PSE demotool | 2025-06-13 | RMS error after Stage Calibration | 3.41
| Intel Hillsboro#3 | 2025-06-13 | Profiler noise @200um | 0.168
| PSE demotool | 2025-06-13 | Profiler noise @200um | 0.089
| Intel Hillsboro#3 | 2025-06-13 | Position repeatability(10repeats, Position Based) | 0.84
| PSE demotool | 2025-06-13 | Position repeatability(10repeats, Position Based) | 0.96
| PSE demotool | 2025-06-13 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat | 0.5
| PSE demotool | 2025-06-13 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat | 0.57
| PSE demotool | 2025-06-13 | Port 1 lens(to OnAxis) Y Offset | 133.31
| PSE demotool | 2025-06-13 | Port 1 lens(to OnAxis) X Offset | -148413
| PSE demotool | 2025-06-13 | P Gain | 0.8
| Intel Hillsboro#3 | 2025-06-13 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| PSE demotool | 2025-06-13 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6.75e-05
| Intel Hillsboro#3 | 2025-06-13 | OpticalAxisShiftPerFocusY | 0.0063
| PSE demotool | 2025-06-13 | OpticalAxisShiftPerFocusY | 0.0001
| Intel Hillsboro#3 | 2025-06-13 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 9e-05
| PSE demotool | 2025-06-13 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 6.99e-05
| Intel Hillsboro#3 | 2025-06-13 | OpticalAxisShiftPerFocusX | -0.0085
| PSE demotool | 2025-06-13 | OpticalAxisShiftPerFocusX | -0.0038
| Intel Hillsboro#3 | 2025-06-13 | OPM Repeatability @5mm | 0.966
| PSE demotool | 2025-06-13 | OPM Repeatability @5mm | 1.241
| Intel Hillsboro#3 | 2025-06-13 | OPM Repeatability @25mm | 4.567
| PSE demotool | 2025-06-13 | OPM Repeatability @25mm | 6.805
| Intel Hillsboro#3 | 2025-06-13 | OPM Repeatability @1mm | 0.705
| PSE demotool | 2025-06-13 | OPM Repeatability @1mm | 0.746
| Intel Hillsboro#3 | 2025-06-13 | OPM Repeatability @10mm | 2.166
| PSE demotool | 2025-06-13 | OPM Repeatability @10mm | 2.23
| Intel Hillsboro#3 | 2025-06-13 | Min. Vibration Transmissbility (~100Hz) | 0.044
| PSE demotool | 2025-06-13 | Min. Vibration Transmissbility (~100Hz) | 0.028
| Intel Hillsboro#3 | 2025-06-13 | Max OPM @5mm | 26.058
| PSE demotool | 2025-06-13 | Max OPM @5mm | 33.167
| Intel Hillsboro#3 | 2025-06-13 | Max OPM @25mm | 188.734
| PSE demotool | 2025-06-13 | Max OPM @25mm | 124.496
| Intel Hillsboro#3 | 2025-06-13 | Max OPM @1mm | 7
| PSE demotool | 2025-06-13 | Max OPM @1mm | 7.926
| Intel Hillsboro#3 | 2025-06-13 | Max OPM @10mm | 54.058
| PSE demotool | 2025-06-13 | Max OPM @10mm | 44.309
| Intel Hillsboro#3 | 2025-06-13 | Lift pins synchronization offset check (Y offset) | 29.99
| PSE demotool | 2025-06-13 | Lift pins synchronization offset check (Y offset) | 38.66
| Intel Hillsboro#3 | 2025-06-13 | Lift pins synchronization offset check (X offset) | 2.02
| PSE demotool | 2025-06-13 | Lift pins synchronization offset check (X offset) | 38.02
| Intel Hillsboro#3 | 2025-06-13 | Lateral(C-D) sensitivity | 25.9
| PSE demotool | 2025-06-13 | Lateral(C-D) sensitivity | 22.9
| Intel Hillsboro#3 | 2025-06-13 | Lateral(C-D) noise | 0.155
| PSE demotool | 2025-06-13 | Lateral(C-D) noise | 0.169
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from -Y to +Y (Y scan) | 78.3691
| PSE demotool | 2025-06-13 | Image offset after moving from -Y to +Y (Y scan) | 102.5391
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from -Y to +Y (X scan) | 38.0859
| PSE demotool | 2025-06-13 | Image offset after moving from -Y to +Y (X scan) | 11.7188
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from -X to +X (Y scan) | 52.7344
| PSE demotool | 2025-06-13 | Image offset after moving from -X to +X (Y scan) | 61.5234
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from -X to +X (X scan) | 32.959
| PSE demotool | 2025-06-13 | Image offset after moving from -X to +X (X scan) | 13.1836
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from +Y to -Y (Y scan) | 145.752
| PSE demotool | 2025-06-13 | Image offset after moving from +Y to -Y (Y scan) | 27.0996
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from +Y to -Y (X scan) | 14.6484
| PSE demotool | 2025-06-13 | Image offset after moving from +Y to -Y (X scan) | 32.2266
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from +X to -X (Y scan) | 30.0293
| PSE demotool | 2025-06-13 | Image offset after moving from +X to -X (Y scan) | 17.5781
| Intel Hillsboro#3 | 2025-06-13 | Image offset after moving from +X to -X (X scan) | 15.3809
| PSE demotool | 2025-06-13 | Image offset after moving from +X to -X (X scan) | 65.1856
| PSE demotool | 2025-06-13 | I Gain | 0.8
| Intel Hillsboro#3 | 2025-06-13 | Full range of Z Height during Z Scan sweep | 43.2543
| PSE demotool | 2025-06-13 | Full range of Z Height during Z Scan sweep | 17.4
| Intel Hillsboro#3 | 2025-06-13 | Full range of Z Detector during Z Scan sweep | 14.06
| PSE demotool | 2025-06-13 | Full range of Z Detector during Z Scan sweep | 13.15
| Intel Hillsboro#3 | 2025-06-13 | Full range of Y2 Detector Raw during Y Scan sweep | 13.59
| PSE demotool | 2025-06-13 | Full range of Y2 Detector Raw during Y Scan sweep | 13.92
| Intel Hillsboro#3 | 2025-06-13 | Full range of Y2 Detector during Y Scan sweep | 195.4319
| PSE demotool | 2025-06-13 | Full range of Y2 Detector during Y Scan sweep | 195.25
| Intel Hillsboro#3 | 2025-06-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| PSE demotool | 2025-06-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.13
| Intel Hillsboro#3 | 2025-06-13 | Full range of Y1 Detector during Y Scan sweep | 197.0985
| PSE demotool | 2025-06-13 | Full range of Y1 Detector during Y Scan sweep | 200.6
| Intel Hillsboro#3 | 2025-06-13 | Full range of Y Detector during Y Scan sweep | 196.2652
| PSE demotool | 2025-06-13 | Full range of Y Detector during Y Scan sweep | 200.6
| Intel Hillsboro#3 | 2025-06-13 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| PSE demotool | 2025-06-13 | Full range of X2 Detector Raw during X Scan sweep | 14.45
| Intel Hillsboro#3 | 2025-06-13 | Full range of X2 Detector during X Scan sweep | 196.934
| PSE demotool | 2025-06-13 | Full range of X2 Detector during X Scan sweep | 198.16
| Intel Hillsboro#3 | 2025-06-13 | Full range of X1 Detector Raw during X Scan sweep | 14.06
| PSE demotool | 2025-06-13 | Full range of X1 Detector Raw during X Scan sweep | 14.02
| Intel Hillsboro#3 | 2025-06-13 | Full range of X1 Detector during X Scan sweep | 197.1125
| PSE demotool | 2025-06-13 | Full range of X1 Detector during X Scan sweep | 205.33
| Intel Hillsboro#3 | 2025-06-13 | Full range of X Detector during X Scan sweep | 197.0234
| PSE demotool | 2025-06-13 | Full range of X Detector during X Scan sweep | 205.33
| PSE demotool | 2025-06-13 | D Gain | 0.1
| Intel Hillsboro#3 | 2025-06-13 | Cycle time for 15-point measurement | 643.135
| PSE demotool | 2025-06-13 | Cycle time for 15-point measurement | 1461.389
| PSE demotool | 2025-06-13 | Check the Port 2 lens light strength | 33.7
| PSE demotool | 2025-06-13 | Check the Port 1 lens light strength | 13.7
| PSE demotool | 2025-06-13 | Check the focus deviation between the lenses(Δf) | 12.95
| PSE demotool | 2025-06-13 | Check the 25mm Moving – Y Axis Offset Check | 5.7
| PSE demotool | 2025-06-13 | C Scanner Detector Min | -0.25
| PSE demotool | 2025-06-13 | C Scanner Detector Max | 0.24
| PSE demotool | 2025-06-13 | B Scanner Detector Min | -0.22
| PSE demotool | 2025-06-13 | B Scanner Detector Max | 0.24
| PSE demotool | 2025-06-13 | Angle align between Port 2 lens and VLSI SHS: Y | -0.03
| PSE demotool | 2025-06-13 | Angle align between Port 2 lens and VLSI SHS: X | -0.017
| PSE demotool | 2025-06-13 | Angle align between Port 2 lens and Sample chuck: Y | -0.024
| PSE demotool | 2025-06-13 | Angle align between Port 2 lens and Sample chuck: X | -0.014
| PSE demotool | 2025-06-13 | Angle align between Port 1 lens and Sample chuck: Y | -0.006
| PSE demotool | 2025-06-13 | Angle align between Port 1 lens and Sample chuck: X | -0.002
| Intel Hillsboro#3 | 2025-06-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.18
| PSE demotool | 2025-06-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Intel Hillsboro#3 | 2025-06-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.94
| PSE demotool | 2025-06-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| Intel Hillsboro#3 | 2025-06-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.86
| PSE demotool | 2025-06-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.23
| Intel Hillsboro#3 | 2025-06-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.16
| PSE demotool | 2025-06-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Intel Hillsboro#3 | 2025-06-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| PSE demotool | 2025-06-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Intel Hillsboro#3 | 2025-06-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.16
| PSE demotool | 2025-06-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Intel Hillsboro#3 | 2025-06-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| PSE demotool | 2025-06-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Intel Hillsboro#3 | 2025-06-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.74
| PSE demotool | 2025-06-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.23
| PSE demotool | 2025-06-13 | A Scanner Detector Min | -0.26
| PSE demotool | 2025-06-13 | A Scanner Detector Max | 0.24
| PSF Demo Lab | 2025-06-12 | SLD current limit check | 148
| PSF Demo Lab | 2025-06-12 | SLD beam power at the end of the lens | 0.221
| PSF Demo Lab | 2025-06-12 | Lateral(C-D) sensitivity | 54.7
| PSF Demo Lab | 2025-06-12 | Lateral(C-D) noise | 0.107
| PSF Demo Lab | 2025-06-12 | A-B noise | 9.875
| PSF Demo Lab | 2025-06-12 | NCM Amplitude noise | 6.475
| PSF Demo Lab | 2025-06-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| PSF Demo Lab | 2025-06-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| PSF Demo Lab | 2025-06-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| PSF Demo Lab | 2025-06-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| PSF Demo Lab | 2025-06-12 | Roughness Repeatability 1x1um(RSD%) | 0.678
| PSF Demo Lab | 2025-06-12 | Roughness Repeatability 1x1um(1σ) | 0.001
| PSF Demo Lab | 2025-06-12 | Full range of Z Height during Z Scan sweep | 16.5
| PSF Demo Lab | 2025-06-12 | Full range of Z Detector during Z Scan sweep | 15.6
| PSF Demo Lab | 2025-06-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.187
| PSF Demo Lab | 2025-06-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.571
| PSF Demo Lab | 2025-06-12 | Z Drive calibrtaion error | 0.16475
| PSF Demo Lab | 2025-06-12 | Z Height calibrtaion error | 0.21648
| PSF Demo Lab | 2025-06-12 | NCM driving% with AC160TS cantilever | 1.032
| PSF Demo Lab | 2025-06-12 | Z linearity by sample height | 1.0174
| PSF Demo Lab | 2025-06-12 | 1) Z Scanner position at +3μm | 0.014195
| PSF Demo Lab | 2025-06-12 | 3) Z Scanner position at -3μm | 0.006264
| PSF Demo Lab | 2025-06-12 | Z Height noise | 0.015
| PSF Demo Lab | 2025-06-12 | Full range of X Detector during X Scan sweep | 106
| PSF Demo Lab | 2025-06-12 | Full range of Y Detector during Y Scan sweep | 108.4
| PSF Demo Lab | 2025-06-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.9
| PSF Demo Lab | 2025-06-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 155.8
| PSF Demo Lab | 2025-06-12 | X Detector stroke calibration error (nominal scale error X) | 0.012
| PSF Demo Lab | 2025-06-12 | Y Detector stroke calibration error (nominal scale error Y) | 0.001
| PSF Demo Lab | 2025-06-12 | X Scan Stroke | 106.1
| PSF Demo Lab | 2025-06-12 | Y Scan Stroke | 108.3
| PSF Demo Lab | 2025-06-12 | Full range of X Detector Raw during X Scan sweep | 13.7
| PSF Demo Lab | 2025-06-12 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| PSF Demo Lab | 2025-06-12 | X Detector noise (XYServo off) | 0.258
| PSF Demo Lab | 2025-06-12 | Y Detector noise (XYServo off) | 0.247
| PSF Demo Lab | 2025-06-12 | X Detector noise (XY Servo on) | 0.269
| PSF Demo Lab | 2025-06-12 | Y Detector noise (XY Servo on) | 0.261
| PSF Demo Lab | 2025-06-12 | X axis OPM | 0.7192
| PSF Demo Lab | 2025-06-12 | Y axis OPM | 1.0348
| PSF Demo Lab | 2025-06-12 | X Non-linearity | 0.182
| PSF Demo Lab | 2025-06-12 | Y Non-linearity | 0.088
| Dupont Taiwan | 2025-06-11 | Z Scanner position at -3μm | 0.011004
| Dupont Taiwan | 2025-06-11 | Z Scanner position at +3μm | 0.010092
| Dupont Taiwan | 2025-06-11 | Z linearity by sample height | 1.0174
| Dupont Taiwan | 2025-06-11 | Z Height noise | 0.027
| Dupont Taiwan | 2025-06-11 | Z Height calibration error | 0.055556
| Dupont Taiwan | 2025-06-11 | Y2 Detector noise (XY Servo on) | 0.247
| Dupont Taiwan | 2025-06-11 | Y2 Detector noise (XY Servo off) | 0.237
| Dupont Taiwan | 2025-06-11 | Y1 Detector noise (XY servo on) | 0.258
| Dupont Taiwan | 2025-06-11 | Y1 Detector noise (XY servo off) | 0.245
| Dupont Taiwan | 2025-06-11 | Y Scan Stroke | 195.09
| Dupont Taiwan | 2025-06-11 | Y Non-linearity | 0.11
| Dupont Taiwan | 2025-06-11 | Y Detector stroke calibration error (nominal scale error Y) | 0.27
| Dupont Taiwan | 2025-06-11 | Y Detector noise (XY servo on) | 0.189
| Dupont Taiwan | 2025-06-11 | Y Detector noise (XY servo off) | 0.179
| Dupont Taiwan | 2025-06-11 | Y axis OPM | 2.455
| Dupont Taiwan | 2025-06-11 | X2 Detector noise (XY Servo on) | 0.28
| Dupont Taiwan | 2025-06-11 | X2 Detector noise (XY Servo off) | 0.268
| Dupont Taiwan | 2025-06-11 | X1 Detector noise (XY servo on) | 0.25
| Dupont Taiwan | 2025-06-11 | X1 Detector noise (XY servo off) | 0.242
| Dupont Taiwan | 2025-06-11 | X Scan Stroke | 202.16
| Dupont Taiwan | 2025-06-11 | X Non-linearity | 0.063
| Dupont Taiwan | 2025-06-11 | X Detector stroke calibration error (nominal scale error X) | 0.036
| Dupont Taiwan | 2025-06-11 | X Detector noise (XY servo on) | 0.198
| Dupont Taiwan | 2025-06-11 | X Detector noise (XY servo off) | 0.186
| Dupont Taiwan | 2025-06-11 | X axis OPM | 1.546
| Dupont Taiwan | 2025-06-11 | Roughness Repeatability 1x1um(RSD%) | 0.743
| Dupont Taiwan | 2025-06-11 | Roughness Repeatability 1x1um(1σ) | 0.002
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.37
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.37
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.34
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.34
| Dupont Taiwan | 2025-06-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.35
| Dupont Taiwan | 2025-06-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Dupont Taiwan | 2025-06-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Dupont Taiwan | 2025-06-11 | NCM driving% with AC160TS cantilever | 1.814
| Dupont Taiwan | 2025-06-11 | NCM Amplitude noise | 13.338
| Dupont Taiwan | 2025-06-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.076
| Dupont Taiwan | 2025-06-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.203
| Dupont Taiwan | 2025-06-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 277.2
| Dupont Taiwan | 2025-06-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 280.2
| Dupont Taiwan | 2025-06-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 307.5
| Dupont Taiwan | 2025-06-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 287.1
| Dupont Taiwan | 2025-06-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Dupont Taiwan | 2025-06-11 | Lateral(C-D) sensitivity | 28.3
| Dupont Taiwan | 2025-06-11 | Lateral(C-D) noise | 0.104
| Dupont Taiwan | 2025-06-11 | Full range of Z Height during Z Scan sweep | 17.3
| Dupont Taiwan | 2025-06-11 | Full range of Z Detector during Z Scan sweep | 13.7
| Dupont Taiwan | 2025-06-11 | Full range of Y2 Detector Raw during Y Scan sweep | 13.9
| Dupont Taiwan | 2025-06-11 | Full range of Y2 Detector during Y Scan sweep | 192.2
| Dupont Taiwan | 2025-06-11 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| Dupont Taiwan | 2025-06-11 | Full range of Y1 Detector during Y Scan sweep | 197.8
| Dupont Taiwan | 2025-06-11 | Full range of Y Detector during Y Scan sweep | 195
| Dupont Taiwan | 2025-06-11 | Full range of X2 Detector Raw during X Scan sweep | 13.3
| Dupont Taiwan | 2025-06-11 | Full range of X2 Detector during X Scan sweep | 205.1
| Dupont Taiwan | 2025-06-11 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| Dupont Taiwan | 2025-06-11 | Full range of X1 Detector during X Scan sweep | 198.9
| Dupont Taiwan | 2025-06-11 | Full range of X Detector during X Scan sweep | 202
| Dupont Taiwan | 2025-06-11 | A-B noise | 6.621
| 울산과학기술원 물리학과 | 2025-06-10 | Z Scanner position at -3μm | 0.004429
| 울산과학기술원 물리학과 | 2025-06-10 | Z Scanner position at +3μm | 0.014341
| 울산과학기술원 물리학과 | 2025-06-10 | Z linearity by sample height | 0.99459
| 울산과학기술원 물리학과 | 2025-06-10 | Z Height noise | 0.03
| 울산과학기술원 물리학과 | 2025-06-10 | Z Height calibration error | 0.2251
| 울산과학기술원 물리학과 | 2025-06-10 | Z Drive calibration error | 0.091954
| 울산과학기술원 물리학과 | 2025-06-10 | Y Scan Stroke | 55
| 울산과학기술원 물리학과 | 2025-06-10 | Y Non-linearity | 0.048
| 울산과학기술원 물리학과 | 2025-06-10 | Y Detector stroke calibration error (nominal scale error Y) | 0.128
| 울산과학기술원 물리학과 | 2025-06-10 | Y Detector noise (XY Servo on) | 0.367
| 울산과학기술원 물리학과 | 2025-06-10 | Y Detector noise (XY Servo off) | 0.38
| 울산과학기술원 물리학과 | 2025-06-10 | Y axis OPM | 1.0578
| 울산과학기술원 물리학과 | 2025-06-10 | X Scan Stroke | 55
| 울산과학기술원 물리학과 | 2025-06-10 | X Non-linearity | 0.029
| 울산과학기술원 물리학과 | 2025-06-10 | X Detector stroke calibration error (nominal scale error X) | 0.043
| 울산과학기술원 물리학과 | 2025-06-10 | X Detector noise (XY Servo on) | 0.365
| 울산과학기술원 물리학과 | 2025-06-10 | X Detector noise (XY Servo off) | 0.384
| 울산과학기술원 물리학과 | 2025-06-10 | X axis OPM | 0.8002
| 울산과학기술원 물리학과 | 2025-06-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 울산과학기술원 물리학과 | 2025-06-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 울산과학기술원 물리학과 | 2025-06-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| 울산과학기술원 물리학과 | 2025-06-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| 울산과학기술원 물리학과 | 2025-06-10 | NCM driving% with AC160TS cantilever | 1.087
| 울산과학기술원 물리학과 | 2025-06-10 | NCM Amplitude noise | 12.933
| 울산과학기술원 물리학과 | 2025-06-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.02
| 울산과학기술원 물리학과 | 2025-06-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.494
| 울산과학기술원 물리학과 | 2025-06-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 90.1
| 울산과학기술원 물리학과 | 2025-06-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 90.7
| 울산과학기술원 물리학과 | 2025-06-10 | Full range of Z Height during Z Scan sweep | 18.4
| 울산과학기술원 물리학과 | 2025-06-10 | Full range of Z Detector during Z Scan sweep | 13.4
| 울산과학기술원 물리학과 | 2025-06-10 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 울산과학기술원 물리학과 | 2025-06-10 | Full range of Y Detector during Y Scan sweep | 65.4
| 울산과학기술원 물리학과 | 2025-06-10 | Full range of X Detector Raw during X Scan sweep | 14.4
| 울산과학기술원 물리학과 | 2025-06-10 | Full range of X Detector during X Scan sweep | 65.3
| 울산과학기술원 물리학과 | 2025-06-10 | A-B noise | 7.681
| PSJ DEMO | 2025-06-04 | Z Scanner position at -3μm | 0.007645
| National Metrology Institute | 2025-06-04 | Z Scanner position at -3μm | 0.005956
| PSJ DEMO | 2025-06-04 | Z Scanner position at +3μm | 0.014803
| National Metrology Institute | 2025-06-04 | Z Scanner position at +3μm | 0.014638
| PSJ DEMO | 2025-06-04 | Z linearity by sample height | 1.0046
| National Metrology Institute | 2025-06-04 | Z linearity by sample height | 1.0055
| PSJ DEMO | 2025-06-04 | Z Height noise | 0.021
| National Metrology Institute | 2025-06-04 | Z Height noise | 0.018
| PSJ DEMO | 2025-06-04 | Z Height calibration error | 0.018199
| National Metrology Institute | 2025-06-04 | Z Height calibration error | 0.1954
| PSJ DEMO | 2025-06-04 | Z Drive calibration error | 0.042146
| National Metrology Institute | 2025-06-04 | Z Drive calibration error | 0.011494
| National Metrology Institute | 2025-06-04 | Y2 Detector noise (XY Servo on) | 0.218
| National Metrology Institute | 2025-06-04 | Y2 Detector noise (XY Servo off) | 0.203
| National Metrology Institute | 2025-06-04 | Y1 Detector noise (XY servo on) | 0.238
| National Metrology Institute | 2025-06-04 | Y1 Detector noise (XY servo off) | 0.222
| PSJ DEMO | 2025-06-04 | Y Scan Stroke | 113.1
| National Metrology Institute | 2025-06-04 | Y Scan Stroke | 169.9
| PSJ DEMO | 2025-06-04 | Y Non-linearity | 0.293
| National Metrology Institute | 2025-06-04 | Y Non-linearity | 0.219
| PSJ DEMO | 2025-06-04 | Y Detector stroke calibration error (nominal scale error Y) | 0.313
| National Metrology Institute | 2025-06-04 | Y Detector stroke calibration error (nominal scale error Y) | 0.426
| PSJ DEMO | 2025-06-04 | Y Detector noise (XY Servo on) | 0.241
| National Metrology Institute | 2025-06-04 | Y Detector noise (XY servo on) | 0.163
| National Metrology Institute | 2025-06-04 | Y Detector noise (XY servo off) | 0.154
| PSJ DEMO | 2025-06-04 | Y Detector noise (XY Servo off) | 0.23
| National Metrology Institute | 2025-06-04 | Y axis OPM (-52000,-52000) | 0.495
| National Metrology Institute | 2025-06-04 | Y axis OPM (-52000,52000) | 1.228
| National Metrology Institute | 2025-06-04 | Y axis OPM (52000,-52000) | 2.191
| National Metrology Institute | 2025-06-04 | Y axis OPM (52000,52000) | 2.546
| National Metrology Institute | 2025-06-04 | Y axis OPM (0,0) | 1.2388
| PSJ DEMO | 2025-06-04 | Y axis OPM | 1.5268
| National Metrology Institute | 2025-06-04 | XY Orthogonality (-52000,-52000) | 90.24
| National Metrology Institute | 2025-06-04 | XY Orthogonality (-52000,52000) | 90.21
| National Metrology Institute | 2025-06-04 | XY Orthogonality (52000,-52000) | 90.226
| National Metrology Institute | 2025-06-04 | XY Orthogonality (52000,52000) | 89.844
| National Metrology Institute | 2025-06-04 | XY Orthogonality (0,0) | 90.026
| National Metrology Institute | 2025-06-04 | X2 Detector noise (XY Servo on) | 0.285
| National Metrology Institute | 2025-06-04 | X2 Detector noise (XY Servo off) | 0.255
| National Metrology Institute | 2025-06-04 | X1 Detector noise (XY servo on) | 0.395
| National Metrology Institute | 2025-06-04 | X1 Detector noise (XY servo off) | 0.372
| PSJ DEMO | 2025-06-04 | X Scan Stroke | 112.1
| National Metrology Institute | 2025-06-04 | X Scan Stroke | 171.1
| PSJ DEMO | 2025-06-04 | X Non-linearity | 0.198
| National Metrology Institute | 2025-06-04 | X Non-linearity | 0.236
| PSJ DEMO | 2025-06-04 | X Detector stroke calibration error (nominal scale error X) | 0.159
| National Metrology Institute | 2025-06-04 | X Detector stroke calibration error (nominal scale error X) | 0.166
| PSJ DEMO | 2025-06-04 | X Detector noise (XY Servo on) | 0.246
| National Metrology Institute | 2025-06-04 | X Detector noise (XY servo on) | 0.272
| National Metrology Institute | 2025-06-04 | X Detector noise (XY servo off) | 0.247
| PSJ DEMO | 2025-06-04 | X Detector noise (XY Servo off) | 0.237
| National Metrology Institute | 2025-06-04 | X axis OPM (-52000,-52000) | 1.285
| National Metrology Institute | 2025-06-04 | X axis OPM (-52000,52000) | 0.926
| National Metrology Institute | 2025-06-04 | X axis OPM (52000,-52000) | 2.127
| National Metrology Institute | 2025-06-04 | X axis OPM (52000,52000) | 1.995
| National Metrology Institute | 2025-06-04 | X axis OPM (0,0) | 1.0748
| PSJ DEMO | 2025-06-04 | X axis OPM | 1.7134
| PSJ DEMO | 2025-06-04 | SLD beam power at the end of the lens | 0.22
| National Metrology Institute | 2025-06-04 | SLD beam power at the end of the lens | 0.24
| PSJ DEMO | 2025-06-04 | Roughness Repeatability 1x1um(RSD%) | 1.599
| National Metrology Institute | 2025-06-04 | Roughness Repeatability 1x1um(RSD%) | 2.854
| PSJ DEMO | 2025-06-04 | Roughness Repeatability 1x1um(1σ) | 0.001
| National Metrology Institute | 2025-06-04 | Roughness Repeatability 1x1um(1σ) | 0.003
| PSJ DEMO | 2025-06-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| National Metrology Institute | 2025-06-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| PSJ DEMO | 2025-06-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| National Metrology Institute | 2025-06-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| National Metrology Institute | 2025-06-04 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.16
| National Metrology Institute | 2025-06-04 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.23
| National Metrology Institute | 2025-06-04 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.2
| National Metrology Institute | 2025-06-04 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.25
| National Metrology Institute | 2025-06-04 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.18
| PSJ DEMO | 2025-06-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| National Metrology Institute | 2025-06-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| PSJ DEMO | 2025-06-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| National Metrology Institute | 2025-06-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| PSJ DEMO | 2025-06-04 | NCM driving% with AC160TS cantilever | 3.536
| National Metrology Institute | 2025-06-04 | NCM driving% with AC160TS cantilever | 3.039
| PSJ DEMO | 2025-06-04 | NCM Amplitude noise | 9.735
| National Metrology Institute | 2025-06-04 | NCM Amplitude noise | 4.711
| PSJ DEMO | 2025-06-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.568
| National Metrology Institute | 2025-06-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.727
| PSJ DEMO | 2025-06-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.677
| National Metrology Institute | 2025-06-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.921
| National Metrology Institute | 2025-06-04 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 227.1
| National Metrology Institute | 2025-06-04 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 231.9
| PSJ DEMO | 2025-06-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.9
| National Metrology Institute | 2025-06-04 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 244.8
| National Metrology Institute | 2025-06-04 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 237.2
| PSJ DEMO | 2025-06-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.4
| National Metrology Institute | 2025-06-04 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.09
| PSJ DEMO | 2025-06-04 | Lateral(C-D) sensitivity | 8.6
| National Metrology Institute | 2025-06-04 | Lateral(C-D) sensitivity | 10.4
| PSJ DEMO | 2025-06-04 | Lateral(C-D) noise | 0.111
| National Metrology Institute | 2025-06-04 | Lateral(C-D) noise | 0.11
| PSJ DEMO | 2025-06-04 | Full range of Z Height during Z Scan sweep | 16.3
| National Metrology Institute | 2025-06-04 | Full range of Z Height during Z Scan sweep | 16.8
| PSJ DEMO | 2025-06-04 | Full range of Z Detector during Z Scan sweep | 13.2
| National Metrology Institute | 2025-06-04 | Full range of Z Detector during Z Scan sweep | 14
| National Metrology Institute | 2025-06-04 | Full range of Y2 Detector Raw during Y Scan sweep | 14.8
| National Metrology Institute | 2025-06-04 | Full range of Y2 Detector during Y Scan sweep | 168.2
| National Metrology Institute | 2025-06-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.8
| National Metrology Institute | 2025-06-04 | Full range of Y1 Detector during Y Scan sweep | 172
| PSJ DEMO | 2025-06-04 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| PSJ DEMO | 2025-06-04 | Full range of Y Detector during Y Scan sweep | 113.1
| National Metrology Institute | 2025-06-04 | Full range of Y Detector during Y Scan sweep | 170.1
| National Metrology Institute | 2025-06-04 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| National Metrology Institute | 2025-06-04 | Full range of X2 Detector during X Scan sweep | 167.6
| National Metrology Institute | 2025-06-04 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| National Metrology Institute | 2025-06-04 | Full range of X1 Detector during X Scan sweep | 168.7
| PSJ DEMO | 2025-06-04 | Full range of X Detector Raw during X Scan sweep | 13.8
| PSJ DEMO | 2025-06-04 | Full range of X Detector during X Scan sweep | 112.1
| National Metrology Institute | 2025-06-04 | Full range of X Detector during X Scan sweep | 170.9
| PSJ DEMO | 2025-06-04 | A-B noise | 4.867
| National Metrology Institute | 2025-06-04 | A-B noise | 6.068
| IME CAS(Beijing) | 2025-06-04 | Z Linearity by Z scanner position | 0.95
| IME CAS(Beijing) | 2025-06-04 | Z Linearity by Z scanner position | 0.31
| IME CAS(Beijing) | 2025-06-04 | Z Linearity by 4 different sample height | 1.0025
| IME CAS(Beijing) | 2025-06-04 | Z Linearity by 4 different sample height | 0.99491
| IME CAS(Beijing) | 2025-06-04 | Z Height noise | 0.03
| IME CAS(Beijing) | 2025-06-04 | Z Height noise | 0.067
| IME CAS(Beijing) | 2025-06-04 | Z Height calibration error | 0.01
| IME CAS(Beijing) | 2025-06-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 22.0567
| IME CAS(Beijing) | 2025-06-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 67.5018
| IME CAS(Beijing) | 2025-06-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 22.0507
| IME CAS(Beijing) | 2025-06-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 67.5018
| IME CAS(Beijing) | 2025-06-04 | Z Height calibration error | 0.19189
| IME CAS(Beijing) | 2025-06-04 | Z Drive calibration error | 0.08
| IME CAS(Beijing) | 2025-06-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5438
| IME CAS(Beijing) | 2025-06-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.7273
| IME CAS(Beijing) | 2025-06-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3359
| IME CAS(Beijing) | 2025-06-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.6995
| IME CAS(Beijing) | 2025-06-04 | Z Drive calibration error | 0.36839
| IME CAS(Beijing) | 2025-06-04 | Z Drift | 8.284
| IME CAS(Beijing) | 2025-06-04 | Y2 Detector noise (XY Servo on) | 0.393
| IME CAS(Beijing) | 2025-06-04 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.2021
| IME CAS(Beijing) | 2025-06-04 | Y1 Detector noise (XY Servo on) | 0.409
| IME CAS(Beijing) | 2025-06-04 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 275.0285
| IME CAS(Beijing) | 2025-06-04 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 185.804
| IME CAS(Beijing) | 2025-06-04 | Y Ringing(Rpv) | 0.684
| IME CAS(Beijing) | 2025-06-04 | Y Non-linearity error | 3.23
| IME CAS(Beijing) | 2025-06-04 | Y Detector stroke calibration error (nominal scale error Y) | 0.056
| IME CAS(Beijing) | 2025-06-04 | Y Detector noise (XY Servo on) | 0.35
| IME CAS(Beijing) | 2025-06-04 | Y axis OPM | 0.223
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.3145
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y offset (LLC Translation) | 1.9531
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y offset (In-Die Align) | 1.2939
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y offset (Global Align) | 4.2871
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3655
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y 1σ (LLC Translation) | 0.4871
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y 1σ (In-Die Align) | 0.3175
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning Y 1σ (Global Align) | 1.1432
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X offset (Vision Pattern Matching) | 3.3057
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X offset (LLC Translation) | 1.543
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X offset (In-Die Align) | 1.3477
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X offset (Global Align) | 2.5977
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.6097
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X 1σ (LLC Translation) | 0.3103
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X 1σ (In-Die Align) | 0.291
| IME CAS(Beijing) | 2025-06-04 | XY Stage positioning X 1σ (Global Align) | 0.6338
| IME CAS(Beijing) | 2025-06-04 | X2 Detector noise (XY Servo on) | 0.303
| IME CAS(Beijing) | 2025-06-04 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 280.1139
| IME CAS(Beijing) | 2025-06-04 | X1 Detector noise (XY Servo on) | 0.314
| IME CAS(Beijing) | 2025-06-04 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 274.6981
| IME CAS(Beijing) | 2025-06-04 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 178.593
| IME CAS(Beijing) | 2025-06-04 | X Ringing(Rpv) | 0.608
| IME CAS(Beijing) | 2025-06-04 | X Non-linearity error | 3.1
| IME CAS(Beijing) | 2025-06-04 | X Detector stroke calibration error (nominal scale error X) | 0.252
| IME CAS(Beijing) | 2025-06-04 | X Detector noise (XY Servo on) | 0.237
| IME CAS(Beijing) | 2025-06-04 | X axis OPM | 0.375
| IME CAS(Beijing) | 2025-06-04 | Torque Profile (50mm, 20um/s) | 0.345
| IME CAS(Beijing) | 2025-06-04 | Torque Profile (50mm, 200um/s) | 0.396
| IME CAS(Beijing) | 2025-06-04 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.906
| IME CAS(Beijing) | 2025-06-04 | Roughness Repeatability 80x80um(RSD%) | 0.196
| IME CAS(Beijing) | 2025-06-04 | Roughness Repeatability 80x80um(1σ) | 0.002
| IME CAS(Beijing) | 2025-06-04 | Roughness Repeatability 1x1um(RSD%) | 2.901
| IME CAS(Beijing) | 2025-06-04 | Roughness Repeatability 1x1um(1σ) | 0.001
| IME CAS(Beijing) | 2025-06-04 | RMS of ZeroScan | 0.047
| IME CAS(Beijing) | 2025-06-04 | RMS error after Stage Calibration | 2.26
| IME CAS(Beijing) | 2025-06-04 | Profiler noise @200um | 0.158
| IME CAS(Beijing) | 2025-06-04 | Position repeatability(10repeats, Position Based) | 0.66
| IME CAS(Beijing) | 2025-06-04 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0
| IME CAS(Beijing) | 2025-06-04 | OpticalAxisShiftPerFocusY | 0
| IME CAS(Beijing) | 2025-06-04 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| IME CAS(Beijing) | 2025-06-04 | OpticalAxisShiftPerFocusX | 0.0029
| IME CAS(Beijing) | 2025-06-04 | OPM Repeatability @5mm | 1.755
| IME CAS(Beijing) | 2025-06-04 | OPM Repeatability @25mm | 2.081
| IME CAS(Beijing) | 2025-06-04 | OPM Repeatability @1mm | 1.241
| IME CAS(Beijing) | 2025-06-04 | OPM Repeatability @10mm | 3.762
| IME CAS(Beijing) | 2025-06-04 | Min. Vibration Transmissbility (~100Hz) | 0.018
| IME CAS(Beijing) | 2025-06-04 | Max OPM @5mm | 15.355
| IME CAS(Beijing) | 2025-06-04 | Max OPM @25mm | 66.901
| IME CAS(Beijing) | 2025-06-04 | Max OPM @1mm | 11.41
| IME CAS(Beijing) | 2025-06-04 | Max OPM @10mm | 22.787
| IME CAS(Beijing) | 2025-06-04 | Lateral(C-D) sensitivity | 12
| IME CAS(Beijing) | 2025-06-04 | Lateral(C-D) noise | 0.136
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from -Y to +Y (Y scan) | 16.1133
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from -Y to +Y (X scan) | 18.311
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from -X to +X (Y scan) | 16.1133
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from -X to +X (X scan) | 120.1172
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from +Y to -Y (Y scan) | 46.875
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from +Y to -Y (X scan) | 49.0723
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from +X to -X (Y scan) | 24.9023
| IME CAS(Beijing) | 2025-06-04 | Image offset after moving from +X to -X (X scan) | 46.875
| IME CAS(Beijing) | 2025-06-04 | Full range of Z Height during Z Scan sweep | 16.95
| IME CAS(Beijing) | 2025-06-04 | Full range of Z Detector during Z Scan sweep | 15.49
| IME CAS(Beijing) | 2025-06-04 | Full range of Y2 Detector Raw during Y Scan sweep | 14.35
| IME CAS(Beijing) | 2025-06-04 | Full range of Y2 Detector during Y Scan sweep | 195.3074
| IME CAS(Beijing) | 2025-06-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.51
| IME CAS(Beijing) | 2025-06-04 | Full range of Y1 Detector during Y Scan sweep | 200.1169
| IME CAS(Beijing) | 2025-06-04 | Full range of Y Detector during Y Scan sweep | 195.3074
| IME CAS(Beijing) | 2025-06-04 | Full range of X2 Detector Raw during X Scan sweep | 14.52
| IME CAS(Beijing) | 2025-06-04 | Full range of X2 Detector during X Scan sweep | 201.7315
| IME CAS(Beijing) | 2025-06-04 | Full range of X1 Detector Raw during X Scan sweep | 14.37
| IME CAS(Beijing) | 2025-06-04 | Full range of X1 Detector during X Scan sweep | 203.2917
| IME CAS(Beijing) | 2025-06-04 | Full range of X Detector during X Scan sweep | 201.7315
| IME CAS(Beijing) | 2025-06-04 | Full range in voltage during Z Scan sweep | 13.07
| IME CAS(Beijing) | 2025-06-04 | Full range in um during Z Scan sweep | 44.1046
| IME CAS(Beijing) | 2025-06-04 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| IME CAS(Beijing) | 2025-06-04 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.33
| IME CAS(Beijing) | 2025-06-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| IME CAS(Beijing) | 2025-06-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| 단국대학교_융합반도체공학과 | 2025-06-02 | Z Scanner position at -3μm | 0.000605
| PS UK Demo Lab | 2025-06-02 | Z Scanner position at -3μm | 0.002413
| 단국대학교_융합반도체공학과 | 2025-06-02 | Z Scanner position at +3μm | 0.002583
| PS UK Demo Lab | 2025-06-02 | Z Scanner position at +3μm | 0.010452
| 단국대학교_융합반도체공학과 | 2025-06-02 | Z linearity by sample height | 0.98748
| PS UK Demo Lab | 2025-06-02 | Z linearity by sample height | 1.0063
| 단국대학교_융합반도체공학과 | 2025-06-02 | Z Height noise | 0.028
| PS UK Demo Lab | 2025-06-02 | Z Height noise | 0.017
| PS UK Demo Lab | 2025-06-02 | Z Height calibration error | 0.090996
| 단국대학교_융합반도체공학과 | 2025-06-02 | Z Height calibration error | 0.041188
| PS UK Demo Lab | 2025-06-02 | Z Drive calibration error | 0.007663
| 단국대학교_융합반도체공학과 | 2025-06-02 | Z Drive calibration error | 0.087165
| PS UK Demo Lab | 2025-06-02 | Y2 Detector noise (XY Servo on) | 0.224
| PS UK Demo Lab | 2025-06-02 | Y2 Detector noise (XY Servo off) | 0.214
| PS UK Demo Lab | 2025-06-02 | Y1 Detector noise (XY servo on) | 0.198
| PS UK Demo Lab | 2025-06-02 | Y1 Detector noise (XY servo off) | 0.192
| 단국대학교_융합반도체공학과 | 2025-06-02 | Y Scan Stroke | 55
| PS UK Demo Lab | 2025-06-02 | Y Scan Stroke | 166.5
| 단국대학교_융합반도체공학과 | 2025-06-02 | Y Non-linearity | 0.037
| PS UK Demo Lab | 2025-06-02 | Y Non-linearity | 0.195
| 단국대학교_융합반도체공학과 | 2025-06-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.219
| PS UK Demo Lab | 2025-06-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.276
| 단국대학교_융합반도체공학과 | 2025-06-02 | Y Detector noise (XY Servo on) | 0.274
| PS UK Demo Lab | 2025-06-02 | Y Detector noise (XY servo on) | 0.165
| 단국대학교_융합반도체공학과 | 2025-06-02 | Y Detector noise (XY Servo off) | 0.265
| PS UK Demo Lab | 2025-06-02 | Y Detector noise (XY servo off) | 0.159
| PS UK Demo Lab | 2025-06-02 | Y axis OPM (-52000,-52000) | 1.321
| PS UK Demo Lab | 2025-06-02 | Y axis OPM (-52000,52000) | 2.309
| PS UK Demo Lab | 2025-06-02 | Y axis OPM (52000,-52000) | 1.443
| PS UK Demo Lab | 2025-06-02 | Y axis OPM (52000,52000) | 2.304
| PS UK Demo Lab | 2025-06-02 | Y axis OPM (0,0) | 0.8402
| 단국대학교_융합반도체공학과 | 2025-06-02 | Y axis OPM | 0.4304
| PS UK Demo Lab | 2025-06-02 | XY Orthogonality (-52000,-52000) | 89.946
| PS UK Demo Lab | 2025-06-02 | XY Orthogonality (-52000,52000) | 89.95
| PS UK Demo Lab | 2025-06-02 | XY Orthogonality (52000,-52000) | 89.845
| PS UK Demo Lab | 2025-06-02 | XY Orthogonality (52000,52000) | 89.732
| PS UK Demo Lab | 2025-06-02 | XY Orthogonality (0,0) | 89.93
| PS UK Demo Lab | 2025-06-02 | X2 Detector noise (XY Servo on) | 0.199
| PS UK Demo Lab | 2025-06-02 | X2 Detector noise (XY Servo off) | 0.19
| PS UK Demo Lab | 2025-06-02 | X1 Detector noise (XY servo on) | 0.209
| PS UK Demo Lab | 2025-06-02 | X1 Detector noise (XY servo off) | 0.197
| 단국대학교_융합반도체공학과 | 2025-06-02 | X Scan Stroke | 55
| PS UK Demo Lab | 2025-06-02 | X Scan Stroke | 168
| 단국대학교_융합반도체공학과 | 2025-06-02 | X Non-linearity | 0.028
| PS UK Demo Lab | 2025-06-02 | X Non-linearity | 0.137
| 단국대학교_융합반도체공학과 | 2025-06-02 | X Detector stroke calibration error (nominal scale error X) | 0.014
| PS UK Demo Lab | 2025-06-02 | X Detector stroke calibration error (nominal scale error X) | 0.25
| 단국대학교_융합반도체공학과 | 2025-06-02 | X Detector noise (XY Servo on) | 0.271
| PS UK Demo Lab | 2025-06-02 | X Detector noise (XY servo on) | 0.153
| 단국대학교_융합반도체공학과 | 2025-06-02 | X Detector noise (XY Servo off) | 0.267
| PS UK Demo Lab | 2025-06-02 | X Detector noise (XY servo off) | 0.142
| PS UK Demo Lab | 2025-06-02 | X axis OPM (-52000,-52000) | 2.118
| PS UK Demo Lab | 2025-06-02 | X axis OPM (-52000,52000) | 2.181
| PS UK Demo Lab | 2025-06-02 | X axis OPM (52000,-52000) | 1.978
| PS UK Demo Lab | 2025-06-02 | X axis OPM (52000,52000) | 2.337
| PS UK Demo Lab | 2025-06-02 | X axis OPM (0,0) | 1.5894
| 단국대학교_융합반도체공학과 | 2025-06-02 | X axis OPM | 0.786
| PS UK Demo Lab | 2025-06-02 | SLD beam power at the end of the lens | 0.28
| 단국대학교_융합반도체공학과 | 2025-06-02 | Roughness Repeatability 1x1um(RSD%) | 0.516
| PS UK Demo Lab | 2025-06-02 | Roughness Repeatability 1x1um(RSD%) | 3.992
| 단국대학교_융합반도체공학과 | 2025-06-02 | Roughness Repeatability 1x1um(1σ) | 0.001
| PS UK Demo Lab | 2025-06-02 | Roughness Repeatability 1x1um(1σ) | 0.005
| 단국대학교_융합반도체공학과 | 2025-06-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| 단국대학교_융합반도체공학과 | 2025-06-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.13
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.14
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.14
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.13
| PS UK Demo Lab | 2025-06-02 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.12
| 단국대학교_융합반도체공학과 | 2025-06-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| PS UK Demo Lab | 2025-06-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| 단국대학교_융합반도체공학과 | 2025-06-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| PS UK Demo Lab | 2025-06-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| 단국대학교_융합반도체공학과 | 2025-06-02 | NCM driving% with AC160TS cantilever | 2.351
| PS UK Demo Lab | 2025-06-02 | NCM driving% with AC160TS cantilever | 1.736
| 단국대학교_융합반도체공학과 | 2025-06-02 | NCM Amplitude noise | 11.081
| PS UK Demo Lab | 2025-06-02 | NCM Amplitude noise | 4.353
| 단국대학교_융합반도체공학과 | 2025-06-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.986
| PS UK Demo Lab | 2025-06-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.249
| 단국대학교_융합반도체공학과 | 2025-06-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.055
| PS UK Demo Lab | 2025-06-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.612
| PS UK Demo Lab | 2025-06-02 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 235.7
| PS UK Demo Lab | 2025-06-02 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 233.3
| 단국대학교_융합반도체공학과 | 2025-06-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 109.4
| PS UK Demo Lab | 2025-06-02 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 232.2
| PS UK Demo Lab | 2025-06-02 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 234.4
| 단국대학교_융합반도체공학과 | 2025-06-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.8
| PS UK Demo Lab | 2025-06-02 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| 단국대학교_융합반도체공학과 | 2025-06-02 | Lateral(C-D) sensitivity | 31.8
| PS UK Demo Lab | 2025-06-02 | Lateral(C-D) sensitivity | 6.9
| 단국대학교_융합반도체공학과 | 2025-06-02 | Lateral(C-D) noise | 0.063
| PS UK Demo Lab | 2025-06-02 | Lateral(C-D) noise | 0.113
| 단국대학교_융합반도체공학과 | 2025-06-02 | Full range of Z Height during Z Scan sweep | 17.8
| PS UK Demo Lab | 2025-06-02 | Full range of Z Height during Z Scan sweep | 16.4
| 단국대학교_융합반도체공학과 | 2025-06-02 | Full range of Z Detector during Z Scan sweep | 13.1
| PS UK Demo Lab | 2025-06-02 | Full range of Z Detector during Z Scan sweep | 13.9
| PS UK Demo Lab | 2025-06-02 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| PS UK Demo Lab | 2025-06-02 | Full range of Y2 Detector during Y Scan sweep | 169.3
| PS UK Demo Lab | 2025-06-02 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| PS UK Demo Lab | 2025-06-02 | Full range of Y1 Detector during Y Scan sweep | 163.5
| 단국대학교_융합반도체공학과 | 2025-06-02 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| 단국대학교_융합반도체공학과 | 2025-06-02 | Full range of Y Detector during Y Scan sweep | 74
| PS UK Demo Lab | 2025-06-02 | Full range of Y Detector during Y Scan sweep | 166.4
| PS UK Demo Lab | 2025-06-02 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| PS UK Demo Lab | 2025-06-02 | Full range of X2 Detector during X Scan sweep | 165.6
| PS UK Demo Lab | 2025-06-02 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| PS UK Demo Lab | 2025-06-02 | Full range of X1 Detector during X Scan sweep | 168.3
| 단국대학교_융합반도체공학과 | 2025-06-02 | Full range of X Detector Raw during X Scan sweep | 14
| 단국대학교_융합반도체공학과 | 2025-06-02 | Full range of X Detector during X Scan sweep | 74.8
| PS UK Demo Lab | 2025-06-02 | Full range of X Detector during X Scan sweep | 166.9
| 단국대학교_융합반도체공학과 | 2025-06-02 | A-B noise | 8.366
| PS UK Demo Lab | 2025-06-02 | A-B noise | 15.323
| QMOT(3) | 2025-06-02 | Z-XY Orthogonality (Y direction-Right) | 54.96
| QMOT(3) | 2025-06-02 | Z-XY Orthogonality (Y direction-Left) | 54.47
| QMOT(3) | 2025-06-02 | Z-XY Orthogonality (X direction-Right) | 54.4
| QMOT(3) | 2025-06-02 | Z-XY Orthogonality (X direction-Left) | 54.73
| QMOT(3) | 2025-06-02 | Z Linearity by Z scanner position | 0.1
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Linearity by Z scanner position | 0.35638
| QMOT(3) | 2025-06-02 | Z Linearity by 4 different sample height | 1.0037
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Linearity by 4 different sample height | 0.99504
| QMOT(3) | 2025-06-02 | Z Height noise | 0.028
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Height noise | 0.026
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.0367
| QMOT(3) | 2025-06-02 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.1817
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.0172
| QMOT(3) | 2025-06-02 | Z Height calibration error | 0.076106
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Height calibration error | 0.00873
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12.1146
| QMOT(3) | 2025-06-02 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 13.0492
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.0328
| QMOT(3) | 2025-06-02 | Z Drive calibration error | 0.10668
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Drive calibration error | 0.003928
| QMOT(3) | 2025-06-02 | Z Drift | 16.755
| Dongguan Guangmao Technology(3) | 2025-06-02 | Z Drift | 14.9079
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y2 Detector noise (XY Servo on) | 0.466
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 275.6751
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y1 Detector noise (XY Servo on) | 0.471
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 304.7435
| QMOT(3) | 2025-06-02 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 63.3783
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 178.928
| QMOT(3) | 2025-06-02 | Y Ringing(Rpv) | 0.161
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y Ringing(Rpv) | 0.073
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y Non-linearity error | 3.2483
| QMOT(3) | 2025-06-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.485
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.16
| QMOT(3) | 2025-06-02 | Y detector noise (XY Servo on) | 0.163
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y Detector noise (XY Servo on) | 0.338
| QMOT(3) | 2025-06-02 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 86.6382
| QMOT(3) | 2025-06-02 | Y axis OPM | 3.405
| Dongguan Guangmao Technology(3) | 2025-06-02 | Y axis OPM | 1.448
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.4551
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y offset (LLC Translation) | 3.0713
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y offset (In-Die Align) | 1.6064
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y offset (Global Align) | 7.3437
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.2451
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y 1σ (LLC Translation) | 0.7175
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y 1σ (In-Die Align) | 0.2684
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning Y 1σ (Global Align) | 2.1723
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X offset (Vision Pattern Matching) | 1.0107
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X offset (LLC Translation) | 1.9287
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X offset (In-Die Align) | 0.8887
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X offset (Global Align) | 4.2285
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.1976
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X 1σ (LLC Translation) | 0.4299
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X 1σ (In-Die Align) | 0.2082
| Dongguan Guangmao Technology(3) | 2025-06-02 | XY Stage positioning X 1σ (Global Align) | 1.0284
| QMOT(3) | 2025-06-02 | XY Orthogonality | 90.242
| Dongguan Guangmao Technology(3) | 2025-06-02 | X2 Detector noise (XY Servo on) | 0.459
| Dongguan Guangmao Technology(3) | 2025-06-02 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 285.5654
| Dongguan Guangmao Technology(3) | 2025-06-02 | X1 Detector noise (XY Servo on) | 0.474
| Dongguan Guangmao Technology(3) | 2025-06-02 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 286.2456
| QMOT(3) | 2025-06-02 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 62.8244
| Dongguan Guangmao Technology(3) | 2025-06-02 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 178.121
| QMOT(3) | 2025-06-02 | X Ringing(Rpv) | 0.09
| Dongguan Guangmao Technology(3) | 2025-06-02 | X Ringing(Rpv) | 0.128
| Dongguan Guangmao Technology(3) | 2025-06-02 | X Non-linearity error | 3.0581
| QMOT(3) | 2025-06-02 | X Detector stroke calibration error (nominal scale error X) | 0.33
| Dongguan Guangmao Technology(3) | 2025-06-02 | X Detector stroke calibration error (nominal scale error X) | 0.017
| QMOT(3) | 2025-06-02 | X detector noise (XY Servo on) | 0.14
| Dongguan Guangmao Technology(3) | 2025-06-02 | X Detector noise (XY Servo on) | 0.34
| QMOT(3) | 2025-06-02 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 87.5108
| QMOT(3) | 2025-06-02 | X axis OPM | 1.639
| Dongguan Guangmao Technology(3) | 2025-06-02 | X axis OPM | 0.746
| Dongguan Guangmao Technology(3) | 2025-06-02 | Wafer placement(Loading) error on the load port2 (Y offset) | 19.2
| Dongguan Guangmao Technology(3) | 2025-06-02 | Wafer placement(Loading) error on the load port2 (X offset) | 33.95
| Dongguan Guangmao Technology(3) | 2025-06-02 | Wafer placement(Loading) error on the load port1 (Y offset) | 16.71
| Dongguan Guangmao Technology(3) | 2025-06-02 | Wafer placement(Loading) error on the load port1 (X offset) | 18.55
| Dongguan Guangmao Technology(3) | 2025-06-02 | Torque Profile (50mm, 20um/s) | 0.45586
| Dongguan Guangmao Technology(3) | 2025-06-02 | Torque Profile (50mm, 200um/s) | 0.41847
| Dongguan Guangmao Technology(3) | 2025-06-02 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.64115
| QMOT(3) | 2025-06-02 | Scan offset accuracy (Precision) | 0.757
| Dongguan Guangmao Technology(3) | 2025-06-02 | Roughness Repeatability 80x80um(RSD%) | 1.1001
| Dongguan Guangmao Technology(3) | 2025-06-02 | Roughness Repeatability 80x80um(1σ) | 0.002119
| Dongguan Guangmao Technology(3) | 2025-06-02 | Roughness Repeatability 1x1um(RSD%) | 0.88579
| Dongguan Guangmao Technology(3) | 2025-06-02 | Roughness Repeatability 1x1um(1σ) | 0.000527
| Dongguan Guangmao Technology(3) | 2025-06-02 | Roll motion to X+ full scan | 1.5
| QMOT(3) | 2025-06-02 | RMS of ZeroScan | 0.046
| Dongguan Guangmao Technology(3) | 2025-06-02 | RMS of ZeroScan | 0.045
| QMOT(3) | 2025-06-02 | RMS error after Stage Calibration | 0.86
| Dongguan Guangmao Technology(3) | 2025-06-02 | RMS error after Stage Calibration | 2.57
| Dongguan Guangmao Technology(3) | 2025-06-02 | Profiler noise @200um | 0.113
| Dongguan Guangmao Technology(3) | 2025-06-02 | Position repeatability(10repeats, Position Based) | 0.66
| QMOT(3) | 2025-06-02 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9.94e-05
| Dongguan Guangmao Technology(3) | 2025-06-02 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.000267
| QMOT(3) | 2025-06-02 | OpticalAxisShiftPerFocusY | 0.0006
| Dongguan Guangmao Technology(3) | 2025-06-02 | OpticalAxisShiftPerFocusY | -0.0032
| QMOT(3) | 2025-06-02 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8.76e-05
| Dongguan Guangmao Technology(3) | 2025-06-02 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0
| QMOT(3) | 2025-06-02 | OpticalAxisShiftPerFocusX | 0.0026
| Dongguan Guangmao Technology(3) | 2025-06-02 | OpticalAxisShiftPerFocusX | -0.0018
| Dongguan Guangmao Technology(3) | 2025-06-02 | OPM Repeatability @5mm | 0.676
| Dongguan Guangmao Technology(3) | 2025-06-02 | OPM Repeatability @25mm | 1.747
| Dongguan Guangmao Technology(3) | 2025-06-02 | OPM Repeatability @1mm | 0.471
| Dongguan Guangmao Technology(3) | 2025-06-02 | OPM Repeatability @10mm | 0.995
| QMOT(3) | 2025-06-02 | Min. Vibration Transmissbility (~100Hz) | 0.026
| Dongguan Guangmao Technology(3) | 2025-06-02 | Min. Vibration Transmissbility (~100Hz) | 0.052761
| Dongguan Guangmao Technology(3) | 2025-06-02 | Max OPM @5mm | 9.855
| Dongguan Guangmao Technology(3) | 2025-06-02 | Max OPM @25mm | 29.98
| Dongguan Guangmao Technology(3) | 2025-06-02 | Max OPM @1mm | 4.157
| Dongguan Guangmao Technology(3) | 2025-06-02 | Max OPM @10mm | 16.916
| QMOT(3) | 2025-06-02 | Mask Roughness Repeatability (10Repeats, σ ) | 0.010064
| QMOT(3) | 2025-06-02 | Mask Roughness Repeatability (10Repeats, RSD%) | 0.14215
| QMOT(3) | 2025-06-02 | Mask exchange time | 4'18"
| Dongguan Guangmao Technology(3) | 2025-06-02 | Lift pins synchronization offset check (Y offset) | 22.97
| Dongguan Guangmao Technology(3) | 2025-06-02 | Lift pins synchronization offset check (X offset) | 6.87
| Dongguan Guangmao Technology(3) | 2025-06-02 | Lateral(C-D) sensitivity | 9.1
| Dongguan Guangmao Technology(3) | 2025-06-02 | Lateral(C-D) noise | 0.16582
| QMOT(3) | 2025-06-02 | Image offset after moving from -Y to +Y (Y scan) | 161.1328
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from -Y to +Y (Y scan) | 119.3848
| QMOT(3) | 2025-06-02 | Image offset after moving from -Y to +Y (X scan) | 19.5312
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from -Y to +Y (X scan) | 14.6484
| QMOT(3) | 2025-06-02 | Image offset after moving from -X to +X (Y scan) | 83.0078
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from -X to +X (Y scan) | 57.1289
| QMOT(3) | 2025-06-02 | Image offset after moving from -X to +X (X scan) | 14.6484
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from -X to +X (X scan) | 27.0996
| QMOT(3) | 2025-06-02 | Image offset after moving from +Y to -Y (Y scan) | 170.8984
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from +Y to -Y (Y scan) | 27.0996
| QMOT(3) | 2025-06-02 | Image offset after moving from +Y to -Y (X scan) | 19.5312
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from +Y to -Y (X scan) | 16.1133
| QMOT(3) | 2025-06-02 | Image offset after moving from +X to -X (Y scan) | 58.5938
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from +X to -X (Y scan) | 13.916
| QMOT(3) | 2025-06-02 | Image offset after moving from +X to -X (X scan) | 19.5312
| Dongguan Guangmao Technology(3) | 2025-06-02 | Image offset after moving from +X to -X (X scan) | 35.1562
| QMOT(3) | 2025-06-02 | Imag offset after Lock stage (Y scan) | 146.4844
| QMOT(3) | 2025-06-02 | Imag offset after Lock stage (X scan) | 68.3594
| QMOT(3) | 2025-06-02 | Imag offset after EFEM loading (Y scan) | 122.0703
| QMOT(3) | 2025-06-02 | Imag offset after EFEM loading (X scan) | 170.8984
| QMOT(3) | 2025-06-02 | Full range of Z Height during Z Scan sweep | 19.9146
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Z Height during Z Scan sweep | 17.8308
| QMOT(3) | 2025-06-02 | Full range of Z Detector during Z Scan sweep | 14.12
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Z Detector during Z Scan sweep | 13.17
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Y2 Detector Raw during Y Scan sweep | 14.42
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Y2 Detector during Y Scan sweep | 204.3175
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Y1 Detector Raw during Y Scan sweep | 14.43
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Y1 Detector during Y Scan sweep | 219.0902
| QMOT(3) | 2025-06-02 | Full range of Y Detector Raw during Y Scan sweep | 14.99
| QMOT(3) | 2025-06-02 | Full range of Y Detector during Y Scan sweep | 63.6055
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of Y Detector during Y Scan sweep | 211.7039
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of X2 Detector Raw during X Scan sweep | 14.64
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of X2 Detector during X Scan sweep | 212.2338
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of X1 Detector Raw during X Scan sweep | 14.47
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of X1 Detector during X Scan sweep | 215.755
| QMOT(3) | 2025-06-02 | Full range of X Detector Raw during X Scan sweep | 14.58
| QMOT(3) | 2025-06-02 | Full range of X Detector during X Scan sweep | 62.2311
| Dongguan Guangmao Technology(3) | 2025-06-02 | Full range of X Detector during X Scan sweep | 213.9944
| Dongguan Guangmao Technology(3) | 2025-06-02 | Cycle time for 15-point measurement | 1956.856
| QMOT(3) | 2025-06-02 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.505
| QMOT(3) | 2025-06-02 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.505
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (Y offset-Point Align3) | 1.03
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (Y offset-Point Align2) | 1.6
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (Y offset-Point Align1) | -0.58
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (Y offset) | 0.77
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (X offset-Point Align3) | -0.38
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (X offset-Point Align2) | 0.78
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (X offset-Point Align1) | -0.54
| QMOT(3) | 2025-06-02 | Aligned stage accuracy (X offset) | 0.71
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.066783
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.35681
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37063
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.069358
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.032615
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17482
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.092431
| Dongguan Guangmao Technology(3) | 2025-06-02 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.017248
| QMOT(3) | 2025-05-30 | Z-XY Orthogonality (Y direction-Right) | 54.96
| QMOT(3) | 2025-05-30 | Z-XY Orthogonality (Y direction-Left) | 54.47
| QMOT(3) | 2025-05-30 | Z-XY Orthogonality (X direction-Right) | 54.4
| QMOT(3) | 2025-05-30 | Z-XY Orthogonality (X direction-Left) | 54.73
| QMOT(3) | 2025-05-30 | Z Linearity by Z scanner position | 0.1
| QMOT(3) | 2025-05-30 | Z Linearity by 4 different sample height | 1.0037
| QMOT(3) | 2025-05-30 | Z Height noise | 0.028
| QMOT(3) | 2025-05-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.1817
| QMOT(3) | 2025-05-30 | Z Height calibration error | 0.076106
| QMOT(3) | 2025-05-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 13.0492
| QMOT(3) | 2025-05-30 | Z Drive calibration error | 0.10668
| QMOT(3) | 2025-05-30 | Z Drift | 16.755
| QMOT(3) | 2025-05-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 63.3783
| QMOT(3) | 2025-05-30 | Y Ringing(Rpv) | 0.161
| QMOT(3) | 2025-05-30 | Y Detector stroke calibration error (nominal scale error Y) | 0.485
| QMOT(3) | 2025-05-30 | Y detector noise (XY Servo on) | 0.163
| QMOT(3) | 2025-05-30 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 86.6382
| QMOT(3) | 2025-05-30 | Y axis OPM | 3.405
| QMOT(3) | 2025-05-30 | XY Orthogonality | 90.242
| QMOT(3) | 2025-05-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 62.8244
| QMOT(3) | 2025-05-30 | X Ringing(Rpv) | 0.09
| QMOT(3) | 2025-05-30 | X Detector stroke calibration error (nominal scale error X) | 0.33
| QMOT(3) | 2025-05-30 | X detector noise (XY Servo on) | 0.14
| QMOT(3) | 2025-05-30 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 87.5108
| QMOT(3) | 2025-05-30 | X axis OPM | 1.639
| QMOT(3) | 2025-05-30 | Scan offset accuracy (Precision) | 0.757
| QMOT(3) | 2025-05-30 | RMS of ZeroScan | 0.046
| QMOT(3) | 2025-05-30 | RMS error after Stage Calibration | 0.86
| QMOT(3) | 2025-05-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9.94e-05
| QMOT(3) | 2025-05-30 | OpticalAxisShiftPerFocusY | 0.0006
| QMOT(3) | 2025-05-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8.76e-05
| QMOT(3) | 2025-05-30 | OpticalAxisShiftPerFocusX | 0.0026
| QMOT(3) | 2025-05-30 | Min. Vibration Transmissbility (~100Hz) | 0.026
| QMOT(3) | 2025-05-30 | Mask Roughness Repeatability (10Repeats, σ ) | 0.010064
| QMOT(3) | 2025-05-30 | Mask Roughness Repeatability (10Repeats, RSD%) | 0.14215
| QMOT(3) | 2025-05-30 | Mask exchange time | 4'18"
| QMOT(3) | 2025-05-30 | Image offset after moving from -Y to +Y (Y scan) | 161.1328
| QMOT(3) | 2025-05-30 | Image offset after moving from -Y to +Y (X scan) | 19.5312
| QMOT(3) | 2025-05-30 | Image offset after moving from -X to +X (Y scan) | 83.0078
| QMOT(3) | 2025-05-30 | Image offset after moving from -X to +X (X scan) | 14.6484
| QMOT(3) | 2025-05-30 | Image offset after moving from +Y to -Y (Y scan) | 170.8984
| QMOT(3) | 2025-05-30 | Image offset after moving from +Y to -Y (X scan) | 19.5312
| QMOT(3) | 2025-05-30 | Image offset after moving from +X to -X (Y scan) | 58.5938
| QMOT(3) | 2025-05-30 | Image offset after moving from +X to -X (X scan) | 19.5312
| QMOT(3) | 2025-05-30 | Imag offset after Lock stage (Y scan) | 146.4844
| QMOT(3) | 2025-05-30 | Imag offset after Lock stage (X scan) | 68.3594
| QMOT(3) | 2025-05-30 | Imag offset after EFEM loading (Y scan) | 122.0703
| QMOT(3) | 2025-05-30 | Imag offset after EFEM loading (X scan) | 170.8984
| QMOT(3) | 2025-05-30 | Full range of Z Height during Z Scan sweep | 19.9146
| QMOT(3) | 2025-05-30 | Full range of Z Detector during Z Scan sweep | 14.12
| QMOT(3) | 2025-05-30 | Full range of Y Detector Raw during Y Scan sweep | 14.99
| QMOT(3) | 2025-05-30 | Full range of Y Detector during Y Scan sweep | 63.6055
| QMOT(3) | 2025-05-30 | Full range of X Detector Raw during X Scan sweep | 14.58
| QMOT(3) | 2025-05-30 | Full range of X Detector during X Scan sweep | 62.2311
| QMOT(3) | 2025-05-30 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.505
| QMOT(3) | 2025-05-30 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.505
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (Y offset-Point Align3) | 1.03
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (Y offset-Point Align2) | 1.6
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (Y offset-Point Align1) | -0.58
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (Y offset) | 0.77
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (X offset-Point Align3) | -0.38
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (X offset-Point Align2) | 0.78
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (X offset-Point Align1) | -0.54
| QMOT(3) | 2025-05-30 | Aligned stage accuracy (X offset) | 0.71
| 고려대학교 융합대학원 | 2025-05-29 | Z Scanner position at -3μm | 0.001677
| The Second Research Institute of CAAC | 2025-05-29 | Z Scanner position at -3μm | 0.000183
| 고려대학교 융합대학원 | 2025-05-29 | Z Scanner position at +3μm | 0.005357
| The Second Research Institute of CAAC | 2025-05-29 | Z Scanner position at +3μm | 0.004571
| 고려대학교 융합대학원 | 2025-05-29 | Z linearity by sample height | 0.99914
| The Second Research Institute of CAAC | 2025-05-29 | Z linearity by sample height | 0.98776
| 고려대학교 융합대학원 | 2025-05-29 | Z Height noise | 0.035
| 고려대학교 융합대학원 | 2025-05-29 | Z Height calibration error | 0.047893
| The Second Research Institute of CAAC | 2025-05-29 | Z Height calibration error | 0.057471
| 고려대학교 융합대학원 | 2025-05-29 | Z Drive calibration error | 0.13793
| The Second Research Institute of CAAC | 2025-05-29 | Z Drive calibration error | 0.12069
| 고려대학교 융합대학원 | 2025-05-29 | Y Scan Stroke | 55
| The Second Research Institute of CAAC | 2025-05-29 | Y Scan Stroke | 55
| 고려대학교 융합대학원 | 2025-05-29 | Y Non-linearity | 0.06
| The Second Research Institute of CAAC | 2025-05-29 | Y Non-linearity | 0.101
| 고려대학교 융합대학원 | 2025-05-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.396
| The Second Research Institute of CAAC | 2025-05-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.49
| 고려대학교 융합대학원 | 2025-05-29 | Y Detector noise (XY Servo on) | 0.261
| The Second Research Institute of CAAC | 2025-05-29 | Y Detector noise (XY Servo on) | 0.265
| 고려대학교 융합대학원 | 2025-05-29 | Y Detector noise (XY Servo off) | 0.253
| The Second Research Institute of CAAC | 2025-05-29 | Y Detector noise (XY Servo off) | 0.258
| 고려대학교 융합대학원 | 2025-05-29 | Y axis OPM | 0.475
| The Second Research Institute of CAAC | 2025-05-29 | Y axis OPM | 0.6486
| 고려대학교 융합대학원 | 2025-05-29 | X Scan Stroke | 55
| The Second Research Institute of CAAC | 2025-05-29 | X Scan Stroke | 55
| 고려대학교 융합대학원 | 2025-05-29 | X Non-linearity | 0.07
| The Second Research Institute of CAAC | 2025-05-29 | X Non-linearity | 0.08
| 고려대학교 융합대학원 | 2025-05-29 | X Detector stroke calibration error (nominal scale error X) | 0.187
| The Second Research Institute of CAAC | 2025-05-29 | X Detector stroke calibration error (nominal scale error X) | 0.084
| 고려대학교 융합대학원 | 2025-05-29 | X Detector noise (XY Servo on) | 0.255
| The Second Research Institute of CAAC | 2025-05-29 | X Detector noise (XY Servo on) | 0.251
| 고려대학교 융합대학원 | 2025-05-29 | X Detector noise (XY Servo off) | 0.247
| The Second Research Institute of CAAC | 2025-05-29 | X Detector noise (XY Servo off) | 0.243
| 고려대학교 융합대학원 | 2025-05-29 | X axis OPM | 0.4542
| The Second Research Institute of CAAC | 2025-05-29 | X axis OPM | 0.7978
| 고려대학교 융합대학원 | 2025-05-29 | Roughness Repeatability 1x1um(RSD%) | 1.156
| The Second Research Institute of CAAC | 2025-05-29 | Roughness Repeatability 1x1um(RSD%) | 0.51
| 고려대학교 융합대학원 | 2025-05-29 | Roughness Repeatability 1x1um(1σ) | 0.002
| The Second Research Institute of CAAC | 2025-05-29 | Roughness Repeatability 1x1um(1σ) | 0.001
| 고려대학교 융합대학원 | 2025-05-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| The Second Research Institute of CAAC | 2025-05-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 고려대학교 융합대학원 | 2025-05-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| The Second Research Institute of CAAC | 2025-05-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| 고려대학교 융합대학원 | 2025-05-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| The Second Research Institute of CAAC | 2025-05-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| 고려대학교 융합대학원 | 2025-05-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| The Second Research Institute of CAAC | 2025-05-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 고려대학교 융합대학원 | 2025-05-29 | NCM driving% with AC160TS cantilever | 13.641
| The Second Research Institute of CAAC | 2025-05-29 | NCM driving% with AC160TS cantilever | 0.88
| 고려대학교 융합대학원 | 2025-05-29 | NCM Amplitude noise | 15.689
| The Second Research Institute of CAAC | 2025-05-29 | NCM Amplitude noise | 12.864
| 고려대학교 융합대학원 | 2025-05-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.712
| The Second Research Institute of CAAC | 2025-05-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.778
| 고려대학교 융합대학원 | 2025-05-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.22
| The Second Research Institute of CAAC | 2025-05-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.549
| 고려대학교 융합대학원 | 2025-05-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.8
| The Second Research Institute of CAAC | 2025-05-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 110.5
| 고려대학교 융합대학원 | 2025-05-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.1
| The Second Research Institute of CAAC | 2025-05-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.6
| 고려대학교 융합대학원 | 2025-05-29 | Lateral(C-D) sensitivity | 51.6
| The Second Research Institute of CAAC | 2025-05-29 | Lateral(C-D) sensitivity | 17.4
| 고려대학교 융합대학원 | 2025-05-29 | Lateral(C-D) noise | 0.069
| The Second Research Institute of CAAC | 2025-05-29 | Lateral(C-D) noise | 0.1
| 고려대학교 융합대학원 | 2025-05-29 | Full range of Z Height during Z Scan sweep | 17.2
| The Second Research Institute of CAAC | 2025-05-29 | Full range of Z Height during Z Scan sweep | 18.1
| 고려대학교 융합대학원 | 2025-05-29 | Full range of Z Detector during Z Scan sweep | 12.7
| The Second Research Institute of CAAC | 2025-05-29 | Full range of Z Detector during Z Scan sweep | 13.1
| 고려대학교 융합대학원 | 2025-05-29 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| The Second Research Institute of CAAC | 2025-05-29 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| 고려대학교 융합대학원 | 2025-05-29 | Full range of Y Detector during Y Scan sweep | 72.4
| The Second Research Institute of CAAC | 2025-05-29 | Full range of Y Detector during Y Scan sweep | 74.6
| 고려대학교 융합대학원 | 2025-05-29 | Full range of X Detector Raw during X Scan sweep | 14
| The Second Research Institute of CAAC | 2025-05-29 | Full range of X Detector Raw during X Scan sweep | 14.2
| 고려대학교 융합대학원 | 2025-05-29 | Full range of X Detector during X Scan sweep | 72.3
| The Second Research Institute of CAAC | 2025-05-29 | Full range of X Detector during X Scan sweep | 74.1
| 고려대학교 융합대학원 | 2025-05-29 | A-B noise | 8.495
| The Second Research Institute of CAAC | 2025-05-29 | A-B noise | 7.866
| Infineon Dresden(4) | 2025-05-29 | Z Linearity by Z scanner position | 0.44
| Infineon Dresden(4) | 2025-05-29 | Z Linearity by 4 different sample height | 1.0029
| Infineon Dresden(4) | 2025-05-29 | Z Height noise | 0.063
| Infineon Dresden(4) | 2025-05-29 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 64.6745
| Infineon Dresden(4) | 2025-05-29 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 64.6
| Infineon Dresden(4) | 2025-05-29 | Z Height calibration error | 0.017997
| Infineon Dresden(4) | 2025-05-29 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 28.1955
| Infineon Dresden(4) | 2025-05-29 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.9996
| Infineon Dresden(4) | 2025-05-29 | Z Drive calibration error | 0.17433
| Infineon Dresden(4) | 2025-05-29 | Z Drift | 9.2
| Infineon Dresden(4) | 2025-05-29 | Y2 Detector noise (XY Servo on) | 0.478
| Infineon Dresden(4) | 2025-05-29 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 265.0917
| Infineon Dresden(4) | 2025-05-29 | Y1 Detector noise (XY Servo on) | 0.364
| Infineon Dresden(4) | 2025-05-29 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 271.1529
| Infineon Dresden(4) | 2025-05-29 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 178.356
| Infineon Dresden(4) | 2025-05-29 | Y Ringing(Rpv) | 0.397
| Infineon Dresden(4) | 2025-05-29 | Y Non-linearity error | 2.47
| Infineon Dresden(4) | 2025-05-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.447
| Infineon Dresden(4) | 2025-05-29 | Y Detector noise (XY Servo on) | 0.346
| Infineon Dresden(4) | 2025-05-29 | Y axis OPM | 1.348
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.4355
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y offset (LLC Translation) | 2.0361
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y offset (In-Die Align) | 1.6943
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y offset (Global Align) | 4.3945
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.283
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y 1σ (LLC Translation) | 0.4389
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y 1σ (In-Die Align) | 0.3387
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning Y 1σ (Global Align) | 1.4231
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X offset (Vision Pattern Matching) | 0.8105
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X offset (LLC Translation) | 1.123
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X offset (In-Die Align) | 1.4404
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X offset (Global Align) | 3.0078
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.1648
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X 1σ (LLC Translation) | 0.2823
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X 1σ (In-Die Align) | 0.274
| Infineon Dresden(4) | 2025-05-29 | XY Stage positioning X 1σ (Global Align) | 0.6721
| Infineon Dresden(4) | 2025-05-29 | X2 Detector noise (XY Servo on) | 0.488
| Infineon Dresden(4) | 2025-05-29 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 274.5608
| Infineon Dresden(4) | 2025-05-29 | X1 Detector noise (XY Servo on) | 0.324
| Infineon Dresden(4) | 2025-05-29 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 281.0868
| Infineon Dresden(4) | 2025-05-29 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 177.349
| Infineon Dresden(4) | 2025-05-29 | X Ringing(Rpv) | 0.268
| Infineon Dresden(4) | 2025-05-29 | X Non-linearity error | 2.56
| Infineon Dresden(4) | 2025-05-29 | X Detector stroke calibration error (nominal scale error X) | 0.417
| Infineon Dresden(4) | 2025-05-29 | X Detector noise (XY Servo on) | 0.322
| Infineon Dresden(4) | 2025-05-29 | X axis OPM | 1.343
| Infineon Dresden(4) | 2025-05-29 | Wafer placement(Loading) error on the load port2 (Y offset) | 26.16
| Infineon Dresden(4) | 2025-05-29 | Wafer placement(Loading) error on the load port2 (X offset) | 50.2
| Infineon Dresden(4) | 2025-05-29 | Wafer placement(Loading) error on the load port1 (Y offset) | 39.48
| Infineon Dresden(4) | 2025-05-29 | Wafer placement(Loading) error on the load port1 (X offset) | 27.8
| Infineon Dresden(4) | 2025-05-29 | Torque Profile (50mm, 20um/s) | 0.68
| Infineon Dresden(4) | 2025-05-29 | Torque Profile (50mm, 200um/s) | 0.532
| Infineon Dresden(4) | 2025-05-29 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.415
| Infineon Dresden(4) | 2025-05-29 | Roughness Repeatability 80x80um(RSD%) | 1.0959
| Infineon Dresden(4) | 2025-05-29 | Roughness Repeatability 80x80um(1σ) | 0.001751
| Infineon Dresden(4) | 2025-05-29 | Roughness Repeatability 1x1um(RSD%) | 1.6896
| Infineon Dresden(4) | 2025-05-29 | Roughness Repeatability 1x1um(1σ) | 0.001059
| Infineon Dresden(4) | 2025-05-29 | Roll motion to X+ full scan | 24.6
| Infineon Dresden(4) | 2025-05-29 | RMS of ZeroScan | 0.029
| Infineon Dresden(4) | 2025-05-29 | RMS error after Stage Calibration | 2.61
| Infineon Dresden(4) | 2025-05-29 | Profiler noise @200um | 0.15
| Infineon Dresden(4) | 2025-05-29 | Position repeatability(10repeats, Position Based) | 0.24
| Infineon Dresden(4) | 2025-05-29 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Infineon Dresden(4) | 2025-05-29 | OpticalAxisShiftPerFocusY | -0.0014
| Infineon Dresden(4) | 2025-05-29 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| Infineon Dresden(4) | 2025-05-29 | OpticalAxisShiftPerFocusX | -0.0022
| Infineon Dresden(4) | 2025-05-29 | OPM Repeatability @5mm | 0.63
| Infineon Dresden(4) | 2025-05-29 | OPM Repeatability @25mm | 4.423
| Infineon Dresden(4) | 2025-05-29 | OPM Repeatability @1mm | 0.546
| Infineon Dresden(4) | 2025-05-29 | OPM Repeatability @10mm | 0.924
| Infineon Dresden(4) | 2025-05-29 | Min. Vibration Transmissbility (~100Hz) | 0.032
| Infineon Dresden(4) | 2025-05-29 | Max OPM @5mm | 25.491
| Infineon Dresden(4) | 2025-05-29 | Max OPM @25mm | 141.113
| Infineon Dresden(4) | 2025-05-29 | Max OPM @1mm | 4.513
| Infineon Dresden(4) | 2025-05-29 | Max OPM @10mm | 35.885
| Infineon Dresden(4) | 2025-05-29 | Lift pins synchronization offset check (Y offset) | 24.24
| Infineon Dresden(4) | 2025-05-29 | Lift pins synchronization offset check (X offset) | 26.71
| Infineon Dresden(4) | 2025-05-29 | Lateral(C-D) sensitivity | 16.9
| Infineon Dresden(4) | 2025-05-29 | Lateral(C-D) noise | 0.14
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from -Y to +Y (Y scan) | 32.96
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from -Y to +Y (X scan) | 30.03
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from -X to +X (Y scan) | 113.53
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from -X to +X (X scan) | 16.11
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from +Y to -Y (Y scan) | 71.04
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from +Y to -Y (X scan) | 26.37
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from +X to -X (Y scan) | 19.78
| Infineon Dresden(4) | 2025-05-29 | Image offset after moving from +X to -X (X scan) | 47.61
| Infineon Dresden(4) | 2025-05-29 | Full range of Z Height during Z Scan sweep | 44.3
| Infineon Dresden(4) | 2025-05-29 | Full range of Z Detector during Z Scan sweep | 13.4
| Infineon Dresden(4) | 2025-05-29 | Full range of Y2 Detector Raw during Y Scan sweep | 14.15
| Infineon Dresden(4) | 2025-05-29 | Full range of Y2 Detector during Y Scan sweep | 205.7
| Infineon Dresden(4) | 2025-05-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Infineon Dresden(4) | 2025-05-29 | Full range of Y1 Detector during Y Scan sweep | 208.1
| Infineon Dresden(4) | 2025-05-29 | Full range of Y Detector during Y Scan sweep | 206.9
| Infineon Dresden(4) | 2025-05-29 | Full range of X2 Detector Raw during X Scan sweep | 14.25
| Infineon Dresden(4) | 2025-05-29 | Full range of X2 Detector during X Scan sweep | 199.8
| Infineon Dresden(4) | 2025-05-29 | Full range of X1 Detector Raw during X Scan sweep | 13.99
| Infineon Dresden(4) | 2025-05-29 | Full range of X1 Detector during X Scan sweep | 204.9
| Infineon Dresden(4) | 2025-05-29 | Full range of X Detector during X Scan sweep | 202.4
| Infineon Dresden(4) | 2025-05-29 | Cycle time for 15-point measurement | 1852.166
| Infineon Dresden(4) | 2025-05-29 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Infineon Dresden(4) | 2025-05-29 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| Infineon Dresden(4) | 2025-05-29 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| Infineon Dresden(4) | 2025-05-29 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Infineon Dresden(4) | 2025-05-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Infineon Dresden(4) | 2025-05-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| Infineon Dresden(4) | 2025-05-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Infineon Dresden(4) | 2025-05-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| Shanxi Institute of Technology | 2025-05-28 | Z Scanner position at -3μm | 0.002412
| Shanxi Institute of Technology | 2025-05-28 | Z Scanner position at +3μm | 0.001889
| Shanxi Institute of Technology | 2025-05-28 | Z linearity by sample height | 1.0253
| Shanxi Institute of Technology | 2025-05-28 | Z Height noise | 0.03
| Shanxi Institute of Technology | 2025-05-28 | Z Height calibration error | 0.043103
| Shanxi Institute of Technology | 2025-05-28 | Z Drive calibration error | 0.078544
| Shanxi Institute of Technology | 2025-05-28 | Y Scan Stroke | 125.5
| Shanxi Institute of Technology | 2025-05-28 | Y Scan Stroke | 17.4
| Shanxi Institute of Technology | 2025-05-28 | Y Non-linearity | 0.043
| Shanxi Institute of Technology | 2025-05-28 | Y Non-linearity | 0.13
| Shanxi Institute of Technology | 2025-05-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.463
| Shanxi Institute of Technology | 2025-05-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.13
| Shanxi Institute of Technology | 2025-05-28 | Y Detector noise (XY Servo on) | 0.396
| Shanxi Institute of Technology | 2025-05-28 | Y Detector noise (XY Servo on) | 0.025
| Shanxi Institute of Technology | 2025-05-28 | Y Detector noise (XY Servo off) | 0.424
| Shanxi Institute of Technology | 2025-05-28 | Y Detector noise (XY servo off) | 0.028
| Shanxi Institute of Technology | 2025-05-28 | Y axis OPM | 1.13
| Shanxi Institute of Technology | 2025-05-28 | Y axis OPM | 1.8736
| Shanxi Institute of Technology | 2025-05-28 | X Scan Stroke | 114.8
| Shanxi Institute of Technology | 2025-05-28 | X Scan Stroke | 17
| Shanxi Institute of Technology | 2025-05-28 | X Non-linearity | 0.084
| Shanxi Institute of Technology | 2025-05-28 | X Non-linearity | 0.12
| Shanxi Institute of Technology | 2025-05-28 | X Detector stroke calibration error (nominal scale error X) | 0.111
| Shanxi Institute of Technology | 2025-05-28 | X Detector stroke calibration error (nominal scale error X) | 0.024
| Shanxi Institute of Technology | 2025-05-28 | X Detector noise (XY Servo on) | 0.409
| Shanxi Institute of Technology | 2025-05-28 | X Detector noise (XY Servo on) | 0.029
| Shanxi Institute of Technology | 2025-05-28 | X Detector noise (XY Servo off) | 0.437
| Shanxi Institute of Technology | 2025-05-28 | X Detector noise (XY servo off) | 0.03
| Shanxi Institute of Technology | 2025-05-28 | X axis OPM | 0.932
| Shanxi Institute of Technology | 2025-05-28 | X axis OPM | 0.9528
| Shanxi Institute of Technology | 2025-05-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Shanxi Institute of Technology | 2025-05-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Shanxi Institute of Technology | 2025-05-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Shanxi Institute of Technology | 2025-05-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| Shanxi Institute of Technology | 2025-05-28 | NCM driving% with AC160TS cantilever | 0.285
| Shanxi Institute of Technology | 2025-05-28 | NCM Amplitude noise | 9.794
| Shanxi Institute of Technology | 2025-05-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.882
| Shanxi Institute of Technology | 2025-05-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.399
| Shanxi Institute of Technology | 2025-05-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.6
| Shanxi Institute of Technology | 2025-05-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 24.4
| Shanxi Institute of Technology | 2025-05-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 165.2
| Shanxi Institute of Technology | 2025-05-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 26.2
| Shanxi Institute of Technology | 2025-05-28 | Full range of Z Height during Z Scan sweep | 18.1
| Shanxi Institute of Technology | 2025-05-28 | Full range of Z Detector during Z Scan sweep | 12.7
| Shanxi Institute of Technology | 2025-05-28 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| Shanxi Institute of Technology | 2025-05-28 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Shanxi Institute of Technology | 2025-05-28 | Full range of Y Detector during Y Scan sweep | 125.6
| Shanxi Institute of Technology | 2025-05-28 | Full range of Y Detector during Y Scan sweep | 17.4
| Shanxi Institute of Technology | 2025-05-28 | Full range of X Detector Raw during X Scan sweep | 13.9
| Shanxi Institute of Technology | 2025-05-28 | Full range of X Detector Raw during X Scan sweep | 12.9
| Shanxi Institute of Technology | 2025-05-28 | Full range of X Detector during X Scan sweep | 114.8
| Shanxi Institute of Technology | 2025-05-28 | Full range of X Detector during X Scan sweep | 17
| Shanxi Institute of Technology | 2025-05-28 | A-B noise | 5.659
| 광주과학기술원_반도체공학과 | 2025-05-20 | Z Scanner position at -3μm | 0.004314
| 광주과학기술원_반도체공학과 | 2025-05-20 | Z Scanner position at +3μm | 0.008801
| 광주과학기술원_반도체공학과 | 2025-05-20 | Z linearity by sample height | 0.99949
| 광주과학기술원_반도체공학과 | 2025-05-20 | Z Height noise | 0.03
| 광주과학기술원_반도체공학과 | 2025-05-20 | Z Height calibration error | 0.074713
| 광주과학기술원_반도체공학과 | 2025-05-20 | Z Drive calibration error | 0.28448
| 광주과학기술원_반도체공학과 | 2025-05-20 | Y Scan Stroke | 55
| 광주과학기술원_반도체공학과 | 2025-05-20 | Y Non-linearity | 0.09
| 광주과학기술원_반도체공학과 | 2025-05-20 | Y Detector stroke calibration error (nominal scale error Y) | 0.453
| 광주과학기술원_반도체공학과 | 2025-05-20 | Y Detector noise (XY Servo on) | 0.268
| 광주과학기술원_반도체공학과 | 2025-05-20 | Y Detector noise (XY Servo off) | 0.268
| 광주과학기술원_반도체공학과 | 2025-05-20 | Y axis OPM | 0.2244
| 광주과학기술원_반도체공학과 | 2025-05-20 | X Scan Stroke | 55
| 광주과학기술원_반도체공학과 | 2025-05-20 | X Non-linearity | 0.06
| 광주과학기술원_반도체공학과 | 2025-05-20 | X Detector stroke calibration error (nominal scale error X) | 0.127
| 광주과학기술원_반도체공학과 | 2025-05-20 | X Detector noise (XY Servo on) | 0.258
| 광주과학기술원_반도체공학과 | 2025-05-20 | X Detector noise (XY Servo off) | 0.257
| 광주과학기술원_반도체공학과 | 2025-05-20 | X axis OPM | 0.3096
| 광주과학기술원_반도체공학과 | 2025-05-20 | Roughness Repeatability 1x1um(RSD%) | 3.203
| 광주과학기술원_반도체공학과 | 2025-05-20 | Roughness Repeatability 1x1um(1σ) | 0.006
| 광주과학기술원_반도체공학과 | 2025-05-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| 광주과학기술원_반도체공학과 | 2025-05-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 광주과학기술원_반도체공학과 | 2025-05-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 광주과학기술원_반도체공학과 | 2025-05-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| 광주과학기술원_반도체공학과 | 2025-05-20 | NCM driving% with AC160TS cantilever | 0.838
| 광주과학기술원_반도체공학과 | 2025-05-20 | NCM Amplitude noise | 8.997
| 광주과학기술원_반도체공학과 | 2025-05-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.962
| 광주과학기술원_반도체공학과 | 2025-05-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.613
| 광주과학기술원_반도체공학과 | 2025-05-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 104
| 광주과학기술원_반도체공학과 | 2025-05-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.9
| 광주과학기술원_반도체공학과 | 2025-05-20 | Lateral(C-D) sensitivity | 48.1
| 광주과학기술원_반도체공학과 | 2025-05-20 | Lateral(C-D) noise | 0.087
| 광주과학기술원_반도체공학과 | 2025-05-20 | Full range of Z Height during Z Scan sweep | 17
| 광주과학기술원_반도체공학과 | 2025-05-20 | Full range of Z Detector during Z Scan sweep | 12.3
| 광주과학기술원_반도체공학과 | 2025-05-20 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| 광주과학기술원_반도체공학과 | 2025-05-20 | Full range of Y Detector during Y Scan sweep | 69.1
| 광주과학기술원_반도체공학과 | 2025-05-20 | Full range of X Detector Raw during X Scan sweep | 13.8
| 광주과학기술원_반도체공학과 | 2025-05-20 | Full range of X Detector during X Scan sweep | 70.2
| 광주과학기술원_반도체공학과 | 2025-05-20 | A-B noise | 11.617
| University of Siena | 2025-05-19 | Z Scanner position at -3μm | 0.004241
| University of Zurich | 2025-05-19 | Z Scanner position at -3μm | 0.0057
| University of Siena | 2025-05-19 | Z Scanner position at +3μm | 0.007382
| University of Zurich | 2025-05-19 | Z Scanner position at +3μm | 0.0057
| University of Siena | 2025-05-19 | Z linearity by sample height | 1.0184
| University of Zurich | 2025-05-19 | Z linearity by sample height | 0.98778
| University of Siena | 2025-05-19 | Z Height noise | 0.027
| University of Zurich | 2025-05-19 | Z Height noise | 0.027
| University of Siena | 2025-05-19 | Z Height calibration error | 0.16954
| University of Zurich | 2025-05-19 | Z Height calibration error | 0.098699
| University of Siena | 2025-05-19 | Z Drive calibration error | 0.10441
| University of Zurich | 2025-05-19 | Z Drive calibration error | 0
| University of Siena | 2025-05-19 | Y Scan Stroke | 55
| University of Zurich | 2025-05-19 | Y Scan Stroke | 55
| University of Siena | 2025-05-19 | Y Non-linearity | 0.093
| University of Zurich | 2025-05-19 | Y Non-linearity | 0.116
| University of Siena | 2025-05-19 | Y Detector stroke calibration error (nominal scale error Y) | 0.047
| University of Zurich | 2025-05-19 | Y Detector stroke calibration error (nominal scale error Y) | 0.254
| University of Siena | 2025-05-19 | Y Detector noise (XY Servo on) | 0.341
| University of Zurich | 2025-05-19 | Y Detector noise (XY Servo on) | 0.32
| University of Siena | 2025-05-19 | Y Detector noise (XY Servo off) | 0.36
| University of Zurich | 2025-05-19 | Y Detector noise (XY Servo off) | 0.344
| University of Siena | 2025-05-19 | Y axis OPM | 0.947
| University of Zurich | 2025-05-19 | Y axis OPM | 0.4722
| University of Siena | 2025-05-19 | X Scan Stroke | 55
| University of Zurich | 2025-05-19 | X Scan Stroke | 55
| University of Siena | 2025-05-19 | X Non-linearity | 0.035
| University of Zurich | 2025-05-19 | X Non-linearity | 0.076
| University of Siena | 2025-05-19 | X Detector stroke calibration error (nominal scale error X) | 0.021
| University of Zurich | 2025-05-19 | X Detector stroke calibration error (nominal scale error X) | 0.382
| University of Siena | 2025-05-19 | X Detector noise (XY Servo on) | 0.337
| University of Zurich | 2025-05-19 | X Detector noise (XY Servo on) | 0.314
| University of Siena | 2025-05-19 | X Detector noise (XY Servo off) | 0.348
| University of Zurich | 2025-05-19 | X Detector noise (XY Servo off) | 0.334
| University of Siena | 2025-05-19 | X axis OPM | 0.6
| University of Zurich | 2025-05-19 | X axis OPM | 1.2148
| University of Siena | 2025-05-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Zurich | 2025-05-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| University of Siena | 2025-05-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| University of Zurich | 2025-05-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| University of Siena | 2025-05-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of Zurich | 2025-05-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| University of Siena | 2025-05-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| University of Zurich | 2025-05-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.19
| University of Siena | 2025-05-19 | NCM driving% with AC160TS cantilever | 1.409
| University of Zurich | 2025-05-19 | NCM driving% with AC160TS cantilever | 0.68
| University of Siena | 2025-05-19 | NCM Amplitude noise | 10.288
| University of Zurich | 2025-05-19 | NCM Amplitude noise | 17.129
| University of Siena | 2025-05-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.557
| University of Zurich | 2025-05-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.246
| University of Siena | 2025-05-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.397
| University of Zurich | 2025-05-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.894
| University of Siena | 2025-05-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 86.4
| University of Zurich | 2025-05-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 89.3
| University of Siena | 2025-05-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 84.4
| University of Zurich | 2025-05-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 87.7
| University of Siena | 2025-05-19 | Full range of Z Height during Z Scan sweep | 18.7
| University of Zurich | 2025-05-19 | Full range of Z Height during Z Scan sweep | 17
| University of Siena | 2025-05-19 | Full range of Z Detector during Z Scan sweep | 14.1
| University of Zurich | 2025-05-19 | Full range of Z Detector during Z Scan sweep | 12.7
| University of Siena | 2025-05-19 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| University of Zurich | 2025-05-19 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| University of Siena | 2025-05-19 | Full range of Y Detector during Y Scan sweep | 59.5
| University of Zurich | 2025-05-19 | Full range of Y Detector during Y Scan sweep | 67.9
| University of Siena | 2025-05-19 | Full range of X Detector Raw during X Scan sweep | 15.8
| University of Zurich | 2025-05-19 | Full range of X Detector Raw during X Scan sweep | 15.2
| University of Siena | 2025-05-19 | Full range of X Detector during X Scan sweep | 66.7
| University of Zurich | 2025-05-19 | Full range of X Detector during X Scan sweep | 66.8
| University of Siena | 2025-05-19 | A-B noise | 5.864
| University of Zurich | 2025-05-19 | A-B noise | 7.804
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Z Scanner position at -3μm | 0.004183
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Z Scanner position at -3μm | 0.001658
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Z Scanner position at +3μm | 0.004637
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Z Scanner position at +3μm | 0.003384
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Z linearity by sample height | 1.0091
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Z linearity by sample height | 1.0116
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Z Height noise | 0.027
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Z Height noise | 0.033
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Z Height calibration error | 0.43219
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Z Height calibration error | 0.08046
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Z Drive calibration error | 0.18927
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Z Drive calibration error | 0.15613
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y2 Detector noise (XY Servo on) | 0.354
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y2 Detector noise (XY Servo on) | 0.226
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y2 Detector noise (XY Servo off) | 0.336
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y2 Detector noise (XY Servo off) | 0.223
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y1 Detector noise (XY servo on) | 0.288
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y1 Detector noise (XY servo on) | 0.311
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y1 Detector noise (XY servo off) | 0.28
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y1 Detector noise (XY servo off) | 0.311
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y Scan Stroke | 175.5
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y Scan Stroke | 168
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y Non-linearity | 0.069
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y Non-linearity | 0.176
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.28
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y Detector stroke calibration error (nominal scale error Y) | 0.066
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y Detector noise (XY servo on) | 0.234
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y Detector noise (XY servo on) | 0.197
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y Detector noise (XY servo off) | 0.223
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y Detector noise (XY servo off) | 0.197
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Y axis OPM | 0.9006
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Y axis OPM | 0.5282
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X2 Detector noise (XY Servo on) | 0.228
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X2 Detector noise (XY Servo on) | 0.242
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X2 Detector noise (XY Servo off) | 0.218
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X2 Detector noise (XY Servo off) | 0.235
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X1 Detector noise (XY servo on) | 0.242
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X1 Detector noise (XY servo on) | 0.242
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X1 Detector noise (XY servo off) | 0.236
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X1 Detector noise (XY servo off) | 0.236
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X Scan Stroke | 174.3
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X Scan Stroke | 174.8
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X Non-linearity | 0.143
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X Non-linearity | 0.178
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X Detector stroke calibration error (nominal scale error X) | 0.19
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X Detector stroke calibration error (nominal scale error X) | 0.299
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X Detector noise (XY servo on) | 0.176
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X Detector noise (XY servo on) | 0.178
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X Detector noise (XY servo off) | 0.169
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X Detector noise (XY servo off) | 0.172
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | X axis OPM | 0.675
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | X axis OPM | 0.8674
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Roughness Repeatability 1x1um(RSD%) | 1.232
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Roughness Repeatability 1x1um(RSD%) | 2.215
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Roughness Repeatability 1x1um(1σ) | 0.001
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Roughness Repeatability 1x1um(1σ) | 0.002
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.32
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.31
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.32
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.32
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.28
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | NCM driving% with AC160TS cantilever | 0.682
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | NCM driving% with AC160TS cantilever | 0.534
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | NCM Amplitude noise | 6.126
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | NCM Amplitude noise | 14.247
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.325
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.411
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.2
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.521
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 237.4
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 237
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.4
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 233
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 239.7
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 241.8
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238.6
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 254
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Lateral(C-D) sensitivity | 22.4
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Lateral(C-D) sensitivity | 27.6
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Lateral(C-D) noise | 0.113
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Lateral(C-D) noise | 0.105
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Z Height during Z Scan sweep | 16.7
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Z Height during Z Scan sweep | 16.7
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Z Detector during Z Scan sweep | 12.3
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Z Detector during Z Scan sweep | 12.6
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Y2 Detector during Y Scan sweep | 174.3
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Y2 Detector during Y Scan sweep | 169.9
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Y1 Detector Raw during Y Scan sweep | 14.3
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Y1 Detector during Y Scan sweep | 176.1
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Y1 Detector during Y Scan sweep | 166.3
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of Y Detector during Y Scan sweep | 175.2
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of Y Detector during Y Scan sweep | 168.1
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of X2 Detector during X Scan sweep | 174
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of X2 Detector during X Scan sweep | 171.7
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of X1 Detector Raw during X Scan sweep | 14.6
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of X1 Detector Raw during X Scan sweep | 14
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of X1 Detector during X Scan sweep | 173.9
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of X1 Detector during X Scan sweep | 177.8
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | Full range of X Detector during X Scan sweep | 173.9
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | Full range of X Detector during X Scan sweep | 174.8
| Hon Hai Technology Group (FOXCONN) (FOXCONN) | 2025-05-16 | A-B noise | 10.474
| Hangzhou Institute of Extremely-Weak Magnetic Field | 2025-05-16 | A-B noise | 7.485
| Walther-Meissner-Institute | 2025-05-15 | Z Scanner position at -3μm | 0.006157
| ATC Demo | 2025-05-15 | Z Scanner position at -3μm | 0.003984
| Walther-Meissner-Institute | 2025-05-15 | Z Scanner position at +3μm | 0.019289
| ATC Demo | 2025-05-15 | Z Scanner position at +3μm | 0.018466
| Walther-Meissner-Institute | 2025-05-15 | Z linearity by sample height | 1.0066
| ATC Demo | 2025-05-15 | Z linearity by sample height | 0.9931
| Walther-Meissner-Institute | 2025-05-15 | Z Height noise | 0.027
| ATC Demo | 2025-05-15 | Z Height noise | 0.017
| ATC Demo | 2025-05-15 | Z Height calibration error | 0.019157
| Walther-Meissner-Institute | 2025-05-15 | Z Height calibration error | 0.086207
| ATC Demo | 2025-05-15 | Z Drive calibration error | 0.18295
| Walther-Meissner-Institute | 2025-05-15 | Z Drive calibration error | 0.062261
| Walther-Meissner-Institute | 2025-05-15 | Y2 Detector noise (XY Servo on) | 0.186
| ATC Demo | 2025-05-15 | Y2 Detector noise (XY Servo on) | 0.208
| Walther-Meissner-Institute | 2025-05-15 | Y2 Detector noise (XY Servo off) | 0.182
| ATC Demo | 2025-05-15 | Y2 Detector noise (XY Servo off) | 0.187
| Walther-Meissner-Institute | 2025-05-15 | Y1 Detector noise (XY servo on) | 0.212
| ATC Demo | 2025-05-15 | Y1 Detector noise (XY servo on) | 0.203
| Walther-Meissner-Institute | 2025-05-15 | Y1 Detector noise (XY servo off) | 0.2
| ATC Demo | 2025-05-15 | Y1 Detector noise (XY servo off) | 0.182
| Walther-Meissner-Institute | 2025-05-15 | Y Scan Stroke | 149.2
| ATC Demo | 2025-05-15 | Y Scan Stroke | 167.5
| Walther-Meissner-Institute | 2025-05-15 | Y Non-linearity | 0.255
| ATC Demo | 2025-05-15 | Y Non-linearity | 0.096
| Walther-Meissner-Institute | 2025-05-15 | Y Detector stroke calibration error (nominal scale error Y) | 0.252
| ATC Demo | 2025-05-15 | Y Detector stroke calibration error (nominal scale error Y) | 0.026
| Walther-Meissner-Institute | 2025-05-15 | Y Detector noise (XY servo on) | 0.146
| ATC Demo | 2025-05-15 | Y Detector noise (XY servo on) | 0.157
| Walther-Meissner-Institute | 2025-05-15 | Y Detector noise (XY servo off) | 0.137
| ATC Demo | 2025-05-15 | Y Detector noise (XY servo off) | 0.137
| Walther-Meissner-Institute | 2025-05-15 | Y axis OPM (-52000,-52000) | 1.342
| ATC Demo | 2025-05-15 | Y axis OPM (-52000,-52000) | 1.253
| Walther-Meissner-Institute | 2025-05-15 | Y axis OPM (-52000,52000) | 2.936
| ATC Demo | 2025-05-15 | Y axis OPM (-52000,52000) | 2.824
| Walther-Meissner-Institute | 2025-05-15 | Y axis OPM (52000,-52000) | 2.55
| ATC Demo | 2025-05-15 | Y axis OPM (52000,-52000) | 2.475
| Walther-Meissner-Institute | 2025-05-15 | Y axis OPM (52000,52000) | 1.605
| ATC Demo | 2025-05-15 | Y axis OPM (52000,52000) | 1.92
| Walther-Meissner-Institute | 2025-05-15 | Y axis OPM (0,0) | 1.038
| ATC Demo | 2025-05-15 | Y axis OPM (0,0) | 0.815
| Walther-Meissner-Institute | 2025-05-15 | XY Orthogonality (-52000,-52000) | 89.995
| ATC Demo | 2025-05-15 | XY Orthogonality (-52000,-52000) | 89.774
| Walther-Meissner-Institute | 2025-05-15 | XY Orthogonality (-52000,52000) | 89.911
| ATC Demo | 2025-05-15 | XY Orthogonality (-52000,52000) | 90.143
| Walther-Meissner-Institute | 2025-05-15 | XY Orthogonality (52000,-52000) | 90.233
| ATC Demo | 2025-05-15 | XY Orthogonality (52000,-52000) | 90.018
| Walther-Meissner-Institute | 2025-05-15 | XY Orthogonality (52000,52000) | 90.092
| ATC Demo | 2025-05-15 | XY Orthogonality (52000,52000) | 90.034
| Walther-Meissner-Institute | 2025-05-15 | XY Orthogonality (0,0) | 90.041
| ATC Demo | 2025-05-15 | XY Orthogonality (0,0) | 90.064
| Walther-Meissner-Institute | 2025-05-15 | X2 Detector noise (XY Servo on) | 0.249
| ATC Demo | 2025-05-15 | X2 Detector noise (XY Servo on) | 0.205
| Walther-Meissner-Institute | 2025-05-15 | X2 Detector noise (XY Servo off) | 0.244
| ATC Demo | 2025-05-15 | X2 Detector noise (XY Servo off) | 0.191
| Walther-Meissner-Institute | 2025-05-15 | X1 Detector noise (XY servo on) | 0.222
| ATC Demo | 2025-05-15 | X1 Detector noise (XY servo on) | 0.195
| Walther-Meissner-Institute | 2025-05-15 | X1 Detector noise (XY servo off) | 0.212
| ATC Demo | 2025-05-15 | X1 Detector noise (XY servo off) | 0.178
| Walther-Meissner-Institute | 2025-05-15 | X Scan Stroke | 174.6
| ATC Demo | 2025-05-15 | X Scan Stroke | 169.5
| Walther-Meissner-Institute | 2025-05-15 | X Non-linearity | 0.207
| ATC Demo | 2025-05-15 | X Non-linearity | 0.166
| Walther-Meissner-Institute | 2025-05-15 | X Detector stroke calibration error (nominal scale error X) | 0.05
| ATC Demo | 2025-05-15 | X Detector stroke calibration error (nominal scale error X) | 0.029
| Walther-Meissner-Institute | 2025-05-15 | X Detector noise (XY servo on) | 0.179
| ATC Demo | 2025-05-15 | X Detector noise (XY servo on) | 0.151
| Walther-Meissner-Institute | 2025-05-15 | X Detector noise (XY servo off) | 0.172
| ATC Demo | 2025-05-15 | X Detector noise (XY servo off) | 0.137
| Walther-Meissner-Institute | 2025-05-15 | X axis OPM (-52000,-52000) | 2.603
| ATC Demo | 2025-05-15 | X axis OPM (-52000,-52000) | 2.815
| Walther-Meissner-Institute | 2025-05-15 | X axis OPM (-52000,52000) | 1.546
| ATC Demo | 2025-05-15 | X axis OPM (-52000,52000) | 3.154
| Walther-Meissner-Institute | 2025-05-15 | X axis OPM (52000,-52000) | 1.149
| ATC Demo | 2025-05-15 | X axis OPM (52000,-52000) | 1.637
| Walther-Meissner-Institute | 2025-05-15 | X axis OPM (52000,52000) | 1.977
| ATC Demo | 2025-05-15 | X axis OPM (52000,52000) | 2.601
| Walther-Meissner-Institute | 2025-05-15 | X axis OPM (0,0) | 1.415
| ATC Demo | 2025-05-15 | X axis OPM (0,0) | 1.0888
| Walther-Meissner-Institute | 2025-05-15 | SLD beam power at the end of the lens | 0.24
| ATC Demo | 2025-05-15 | SLD beam power at the end of the lens | 0.28
| Walther-Meissner-Institute | 2025-05-15 | Roughness Repeatability 1x1um(RSD%) | 1.535
| ATC Demo | 2025-05-15 | Roughness Repeatability 1x1um(RSD%) | 2.184
| Walther-Meissner-Institute | 2025-05-15 | Roughness Repeatability 1x1um(1σ) | 0.003
| ATC Demo | 2025-05-15 | Roughness Repeatability 1x1um(1σ) | 0.002
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| ATC Demo | 2025-05-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| ATC Demo | 2025-05-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.23
| ATC Demo | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.14
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.25
| ATC Demo | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.14
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.18
| ATC Demo | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.14
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.17
| ATC Demo | 2025-05-15 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.13
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.17
| ATC Demo | 2025-05-15 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.13
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| ATC Demo | 2025-05-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| Walther-Meissner-Institute | 2025-05-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| ATC Demo | 2025-05-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| Walther-Meissner-Institute | 2025-05-15 | NCM driving% with AC160TS cantilever | 0.783
| ATC Demo | 2025-05-15 | NCM driving% with AC160TS cantilever | 12.572
| Walther-Meissner-Institute | 2025-05-15 | NCM Amplitude noise | 4.636
| ATC Demo | 2025-05-15 | NCM Amplitude noise | 5.834
| Walther-Meissner-Institute | 2025-05-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.289
| ATC Demo | 2025-05-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.406
| Walther-Meissner-Institute | 2025-05-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.233
| ATC Demo | 2025-05-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.395
| Walther-Meissner-Institute | 2025-05-15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 234.7
| ATC Demo | 2025-05-15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 244.8
| Walther-Meissner-Institute | 2025-05-15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 246.6
| ATC Demo | 2025-05-15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 237.4
| Walther-Meissner-Institute | 2025-05-15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 247.9
| ATC Demo | 2025-05-15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 242.1
| Walther-Meissner-Institute | 2025-05-15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249.5
| ATC Demo | 2025-05-15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238.8
| Walther-Meissner-Institute | 2025-05-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| ATC Demo | 2025-05-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Walther-Meissner-Institute | 2025-05-15 | Lateral(C-D) sensitivity | 55.8
| ATC Demo | 2025-05-15 | Lateral(C-D) sensitivity | 8.4
| Walther-Meissner-Institute | 2025-05-15 | Lateral(C-D) noise | 0.118
| ATC Demo | 2025-05-15 | Lateral(C-D) noise | 0.108
| Walther-Meissner-Institute | 2025-05-15 | Full range of Z Height during Z Scan sweep | 16.5
| ATC Demo | 2025-05-15 | Full range of Z Height during Z Scan sweep | 16.3
| Walther-Meissner-Institute | 2025-05-15 | Full range of Z Detector during Z Scan sweep | 14.8
| ATC Demo | 2025-05-15 | Full range of Z Detector during Z Scan sweep | 14
| Walther-Meissner-Institute | 2025-05-15 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| ATC Demo | 2025-05-15 | Full range of Y2 Detector Raw during Y Scan sweep | 13.9
| Walther-Meissner-Institute | 2025-05-15 | Full range of Y2 Detector during Y Scan sweep | 166.1
| ATC Demo | 2025-05-15 | Full range of Y2 Detector during Y Scan sweep | 169.9
| Walther-Meissner-Institute | 2025-05-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| ATC Demo | 2025-05-15 | Full range of Y1 Detector Raw during Y Scan sweep | 13.9
| Walther-Meissner-Institute | 2025-05-15 | Full range of Y1 Detector during Y Scan sweep | 173.4
| ATC Demo | 2025-05-15 | Full range of Y1 Detector during Y Scan sweep | 165
| Walther-Meissner-Institute | 2025-05-15 | Full range of Y Detector during Y Scan sweep | 170.8
| ATC Demo | 2025-05-15 | Full range of Y Detector during Y Scan sweep | 167.4
| Walther-Meissner-Institute | 2025-05-15 | Full range of X2 Detector Raw during X Scan sweep | 14
| ATC Demo | 2025-05-15 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| Walther-Meissner-Institute | 2025-05-15 | Full range of X2 Detector during X Scan sweep | 175.5
| ATC Demo | 2025-05-15 | Full range of X2 Detector during X Scan sweep | 168.7
| Walther-Meissner-Institute | 2025-05-15 | Full range of X1 Detector Raw during X Scan sweep | 14
| ATC Demo | 2025-05-15 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| Walther-Meissner-Institute | 2025-05-15 | Full range of X1 Detector during X Scan sweep | 174.5
| ATC Demo | 2025-05-15 | Full range of X1 Detector during X Scan sweep | 169.1
| Walther-Meissner-Institute | 2025-05-15 | Full range of X Detector during X Scan sweep | 174.1
| ATC Demo | 2025-05-15 | Full range of X Detector during X Scan sweep | 168.9
| Walther-Meissner-Institute | 2025-05-15 | A-B noise | 15.697
| ATC Demo | 2025-05-15 | A-B noise | 4.822
| AGH University | 2025-05-12 | Z Scanner position at -3μm | 0.003573
| AGH University | 2025-05-12 | Z Scanner position at +3μm | 0.001547
| AGH University | 2025-05-12 | Z linearity by sample height | 0.99134
| AGH University | 2025-05-12 | Z Height noise | 0.028
| AGH University | 2025-05-12 | Z Height calibration error | 0.02682
| AGH University | 2025-05-12 | Z Drive calibration error | 0.14368
| AGH University | 2025-05-12 | Y Scan Stroke | 55
| AGH University | 2025-05-12 | Y Non-linearity | 0.071
| AGH University | 2025-05-12 | Y Detector stroke calibration error (nominal scale error Y) | 0.286
| AGH University | 2025-05-12 | Y Detector noise (XY Servo on) | 0.376
| AGH University | 2025-05-12 | Y Detector noise (XY Servo off) | 0.42
| AGH University | 2025-05-12 | Y axis OPM | 0.451
| AGH University | 2025-05-12 | X Scan Stroke | 55
| AGH University | 2025-05-12 | X Non-linearity | 0.09
| AGH University | 2025-05-12 | X Detector stroke calibration error (nominal scale error X) | 0.067
| AGH University | 2025-05-12 | X Detector noise (XY Servo on) | 0.377
| AGH University | 2025-05-12 | X Detector noise (XY Servo off) | 0.412
| AGH University | 2025-05-12 | X axis OPM | 0.7522
| AGH University | 2025-05-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| AGH University | 2025-05-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| AGH University | 2025-05-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| AGH University | 2025-05-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| AGH University | 2025-05-12 | NCM driving% with AC160TS cantilever | 0.743
| AGH University | 2025-05-12 | NCM Amplitude noise | 15.462
| AGH University | 2025-05-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.733
| AGH University | 2025-05-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.121
| AGH University | 2025-05-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 87.3
| AGH University | 2025-05-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 87.1
| AGH University | 2025-05-12 | Full range of Z Height during Z Scan sweep | 17.2
| AGH University | 2025-05-12 | Full range of Z Detector during Z Scan sweep | 12.7
| AGH University | 2025-05-12 | Full range of Y Detector Raw during Y Scan sweep | 15
| AGH University | 2025-05-12 | Full range of Y Detector during Y Scan sweep | 65.5
| AGH University | 2025-05-12 | Full range of X Detector Raw during X Scan sweep | 14.9
| AGH University | 2025-05-12 | Full range of X Detector during X Scan sweep | 64.7
| AGH University | 2025-05-12 | A-B noise | 7.751
| University of Siegen, Germany | 2025-05-09 | Z Scanner position at -3μm | 0.006808
| University of Siegen, Germany | 2025-05-09 | Z Scanner position at +3μm | 0.001934
| University of Siegen, Germany | 2025-05-09 | Z linearity by sample height | 0.98415
| University of Siegen, Germany | 2025-05-09 | Z Height noise | 0.029
| University of Siegen, Germany | 2025-05-09 | Z Height calibration error | 0.023946
| University of Siegen, Germany | 2025-05-09 | Z Drive calibration error | 0.19732
| University of Siegen, Germany | 2025-05-09 | Y Scan Stroke | 123.6
| University of Siegen, Germany | 2025-05-09 | Y Non-linearity | 0.223
| University of Siegen, Germany | 2025-05-09 | Y Detector stroke calibration error (nominal scale error Y) | 0.008
| University of Siegen, Germany | 2025-05-09 | Y Detector noise (XY Servo on) | 0.379
| University of Siegen, Germany | 2025-05-09 | Y Detector noise (XY Servo off) | 0.417
| University of Siegen, Germany | 2025-05-09 | Y axis OPM | 1.8146
| University of Siegen, Germany | 2025-05-09 | X Scan Stroke | 127.5
| University of Siegen, Germany | 2025-05-09 | X Non-linearity | 0.121
| University of Siegen, Germany | 2025-05-09 | X Detector stroke calibration error (nominal scale error X) | 0.063
| University of Siegen, Germany | 2025-05-09 | X Detector noise (XY Servo on) | 0.397
| University of Siegen, Germany | 2025-05-09 | X Detector noise (XY Servo off) | 0.438
| University of Siegen, Germany | 2025-05-09 | X axis OPM | 1.357
| University of Siegen, Germany | 2025-05-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Siegen, Germany | 2025-05-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Siegen, Germany | 2025-05-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| University of Siegen, Germany | 2025-05-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| University of Siegen, Germany | 2025-05-09 | NCM driving% with AC160TS cantilever | 1.902
| University of Siegen, Germany | 2025-05-09 | NCM Amplitude noise | 14.267
| University of Siegen, Germany | 2025-05-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.302
| University of Siegen, Germany | 2025-05-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.703
| University of Siegen, Germany | 2025-05-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.1
| University of Siegen, Germany | 2025-05-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 174.7
| University of Siegen, Germany | 2025-05-09 | Full range of Z Height during Z Scan sweep | 16.9
| University of Siegen, Germany | 2025-05-09 | Full range of Z Detector during Z Scan sweep | 12.7
| University of Siegen, Germany | 2025-05-09 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| University of Siegen, Germany | 2025-05-09 | Full range of Y Detector during Y Scan sweep | 123.6
| University of Siegen, Germany | 2025-05-09 | Full range of X Detector Raw during X Scan sweep | 14.6
| University of Siegen, Germany | 2025-05-09 | Full range of X Detector during X Scan sweep | 127.5
| University of Siegen, Germany | 2025-05-09 | A-B noise | 7.313
| Wuhan Cyanopto Technology | 2025-05-07 | Z Scanner position at -3μm | 0.005798
| 아이큐랩 | 2025-05-07 | Z Scanner position at -3μm | 0.004383
| Wuhan Cyanopto Technology | 2025-05-07 | Z Scanner position at +3μm | 0.013254
| 아이큐랩 | 2025-05-07 | Z Scanner position at +3μm | 0.001743
| Wuhan Cyanopto Technology | 2025-05-07 | Z linearity by sample height | 1.0299
| 아이큐랩 | 2025-05-07 | Z linearity by sample height | 1.007
| Wuhan Cyanopto Technology | 2025-05-07 | Z Height noise | 0.027
| 아이큐랩 | 2025-05-07 | Z Height noise | 0.03
| Wuhan Cyanopto Technology | 2025-05-07 | Z Height calibration error | 0.057561
| 아이큐랩 | 2025-05-07 | Z Height calibration error | 0.18341
| Wuhan Cyanopto Technology | 2025-05-07 | Z Drive calibration error | 0.087805
| 아이큐랩 | 2025-05-07 | Z Drive calibration error | 0.096585
| 아이큐랩 | 2025-05-07 | Y2 Detector noise (XY Servo on) | 0.345
| 아이큐랩 | 2025-05-07 | Y2 Detector noise (XY Servo off) | 0.325
| 아이큐랩 | 2025-05-07 | Y1 Detector noise (XY servo on) | 0.289
| 아이큐랩 | 2025-05-07 | Y1 Detector noise (XY servo off) | 0.273
| 아이큐랩 | 2025-05-07 | Y Scan Stroke | 171.7
| Wuhan Cyanopto Technology | 2025-05-07 | Y Non-linearity | 0.205
| 아이큐랩 | 2025-05-07 | Y Non-linearity | 0.218
| Wuhan Cyanopto Technology | 2025-05-07 | Y Detector stroke calibration error (nominal scale error Y) | 0.12
| 아이큐랩 | 2025-05-07 | Y Detector stroke calibration error (nominal scale error Y) | 0.095
| Wuhan Cyanopto Technology | 2025-05-07 | Y Detector noise (XY Servo on) | 0.282
| 아이큐랩 | 2025-05-07 | Y Detector noise (XY servo on) | 0.245
| Wuhan Cyanopto Technology | 2025-05-07 | Y Detector noise (XY servo off) | 0.266
| 아이큐랩 | 2025-05-07 | Y Detector noise (XY servo off) | 0.216
| Wuhan Cyanopto Technology | 2025-05-07 | Y axis OPM | 1.112
| 아이큐랩 | 2025-05-07 | Y axis OPM | 1.096
| 아이큐랩 | 2025-05-07 | X2 Detector noise (XY Servo on) | 0.303
| 아이큐랩 | 2025-05-07 | X2 Detector noise (XY Servo off) | 0.275
| 아이큐랩 | 2025-05-07 | X1 Detector noise (XY servo on) | 0.264
| 아이큐랩 | 2025-05-07 | X1 Detector noise (XY servo off) | 0.235
| 아이큐랩 | 2025-05-07 | X Scan Stroke | 180.1
| Wuhan Cyanopto Technology | 2025-05-07 | X Non-linearity | 0.082
| 아이큐랩 | 2025-05-07 | X Non-linearity | 0.272
| Wuhan Cyanopto Technology | 2025-05-07 | X Detector stroke calibration error (nominal scale error X) | 0.045
| 아이큐랩 | 2025-05-07 | X Detector stroke calibration error (nominal scale error X) | 0.014
| Wuhan Cyanopto Technology | 2025-05-07 | X Detector noise (XY Servo on) | 0.323
| 아이큐랩 | 2025-05-07 | X Detector noise (XY servo on) | 0.224
| Wuhan Cyanopto Technology | 2025-05-07 | X Detector noise (XY servo off) | 0.272
| 아이큐랩 | 2025-05-07 | X Detector noise (XY servo off) | 0.188
| Wuhan Cyanopto Technology | 2025-05-07 | X axis OPM | 1.216
| 아이큐랩 | 2025-05-07 | X axis OPM | 0.701
| Wuhan Cyanopto Technology | 2025-05-07 | Roughness Repeatability 1x1um(RSD%) | 1.673
| 아이큐랩 | 2025-05-07 | Roughness Repeatability 1x1um(RSD%) | 2.081
| Wuhan Cyanopto Technology | 2025-05-07 | Roughness Repeatability 1x1um(1σ) | 0.003
| 아이큐랩 | 2025-05-07 | Roughness Repeatability 1x1um(1σ) | 0.005
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 아이큐랩 | 2025-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| 아이큐랩 | 2025-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| 아이큐랩 | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.32
| 아이큐랩 | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.31
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.24
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.25
| 아이큐랩 | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| 아이큐랩 | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.35
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.25
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.23
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.24
| 아이큐랩 | 2025-05-07 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.34
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 아이큐랩 | 2025-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Wuhan Cyanopto Technology | 2025-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| 아이큐랩 | 2025-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Wuhan Cyanopto Technology | 2025-05-07 | NCM driving% with AC160TS cantilever | 2.966
| 아이큐랩 | 2025-05-07 | NCM driving% with AC160TS cantilever | 1.74
| Wuhan Cyanopto Technology | 2025-05-07 | NCM Amplitude noise | 6.583
| 아이큐랩 | 2025-05-07 | NCM Amplitude noise | 12.995
| Wuhan Cyanopto Technology | 2025-05-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.601
| 아이큐랩 | 2025-05-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.575
| Wuhan Cyanopto Technology | 2025-05-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.181
| 아이큐랩 | 2025-05-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.944
| 아이큐랩 | 2025-05-07 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 229.1
| 아이큐랩 | 2025-05-07 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 238.2
| Wuhan Cyanopto Technology | 2025-05-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 187.5
| 아이큐랩 | 2025-05-07 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 248.5
| 아이큐랩 | 2025-05-07 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238.3
| Wuhan Cyanopto Technology | 2025-05-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 184
| Wuhan Cyanopto Technology | 2025-05-07 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| 아이큐랩 | 2025-05-07 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| Wuhan Cyanopto Technology | 2025-05-07 | Lateral(C-D) sensitivity | 245.817
| 아이큐랩 | 2025-05-07 | Lateral(C-D) sensitivity | 9.9
| Wuhan Cyanopto Technology | 2025-05-07 | Lateral(C-D) noise | 0.127
| 아이큐랩 | 2025-05-07 | Lateral(C-D) noise | 0.134
| Wuhan Cyanopto Technology | 2025-05-07 | Full range of Z Height during Z Scan sweep | 17.2
| Wuhan Cyanopto Technology | 2025-05-07 | Full range of Z Detector during Z Scan sweep | 12.7
| 아이큐랩 | 2025-05-07 | Full range of Z Detector during Z Scan sweep | 17.6
| 아이큐랩 | 2025-05-07 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| 아이큐랩 | 2025-05-07 | Full range of Y2 Detector during Y Scan sweep | 168.9
| 아이큐랩 | 2025-05-07 | Full range of Y1 Detector Raw during Y Scan sweep | 14.7
| 아이큐랩 | 2025-05-07 | Full range of Y1 Detector during Y Scan sweep | 174.8
| Wuhan Cyanopto Technology | 2025-05-07 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Wuhan Cyanopto Technology | 2025-05-07 | Full range of Y Detector during Y Scan sweep | 133.1
| 아이큐랩 | 2025-05-07 | Full range of Y Detector during Y Scan sweep | 171.9
| 아이큐랩 | 2025-05-07 | Full range of X2 Detector Raw during X Scan sweep | 14.6
| 아이큐랩 | 2025-05-07 | Full range of X2 Detector during X Scan sweep | 180.9
| 아이큐랩 | 2025-05-07 | Full range of X1 Detector Raw during X Scan sweep | 15
| 아이큐랩 | 2025-05-07 | Full range of X1 Detector during X Scan sweep | 178.9
| Wuhan Cyanopto Technology | 2025-05-07 | Full range of X Detector Raw during X Scan sweep | 14.3
| Wuhan Cyanopto Technology | 2025-05-07 | Full range of X Detector during X Scan sweep | 133.9
| 아이큐랩 | 2025-05-07 | Full range of X Detector during X Scan sweep | 179.9
| Wuhan Cyanopto Technology | 2025-05-07 | A-B noise | 10.237
| 아이큐랩 | 2025-05-07 | A-B noise | 8.481
| Morgan State University | 2025-05-02 | Z Scanner position at -3μm | 0.001149
| Morgan State University | 2025-05-02 | Z Scanner position at +3μm | 0.003663
| Morgan State University | 2025-05-02 | Z linearity by sample height | 0.99497
| Morgan State University | 2025-05-02 | Z Height noise | 0.026
| Morgan State University | 2025-05-02 | Z Height calibration error | 0.24904
| Morgan State University | 2025-05-02 | Z Drive calibration error | 0.31609
| Morgan State University | 2025-05-02 | Y Scan Stroke | 55
| Morgan State University | 2025-05-02 | Y Non-linearity | 0.037
| Morgan State University | 2025-05-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.436
| Morgan State University | 2025-05-02 | Y Detector noise (XY Servo on) | 0.266
| Morgan State University | 2025-05-02 | Y Detector noise (XY Servo off) | 0.258
| Morgan State University | 2025-05-02 | Y axis OPM | 0.779
| Morgan State University | 2025-05-02 | X Scan Stroke | 55
| Morgan State University | 2025-05-02 | X Non-linearity | 0.063
| Morgan State University | 2025-05-02 | X Detector stroke calibration error (nominal scale error X) | 0.036
| Morgan State University | 2025-05-02 | X Detector noise (XY Servo on) | 0.263
| Morgan State University | 2025-05-02 | X Detector noise (XY Servo off) | 0.254
| Morgan State University | 2025-05-02 | X axis OPM | 1.306
| Morgan State University | 2025-05-02 | Roughness Repeatability 1x1um(RSD%) | 2.633
| Morgan State University | 2025-05-02 | Roughness Repeatability 1x1um(1σ) | 0.004
| Morgan State University | 2025-05-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Morgan State University | 2025-05-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Morgan State University | 2025-05-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Morgan State University | 2025-05-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Morgan State University | 2025-05-02 | NCM driving% with AC160TS cantilever | 4.53
| Morgan State University | 2025-05-02 | NCM Amplitude noise | 8.549
| Morgan State University | 2025-05-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.482
| Morgan State University | 2025-05-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.11
| Morgan State University | 2025-05-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.6
| Morgan State University | 2025-05-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.1
| Morgan State University | 2025-05-02 | Lateral(C-D) sensitivity | 40.7
| Morgan State University | 2025-05-02 | Lateral(C-D) noise | 0.1
| Morgan State University | 2025-05-02 | Full range of Z Height during Z Scan sweep | 18.3
| Morgan State University | 2025-05-02 | Full range of Z Detector during Z Scan sweep | 14
| Morgan State University | 2025-05-02 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Morgan State University | 2025-05-02 | Full range of Y Detector during Y Scan sweep | 75.6
| Morgan State University | 2025-05-02 | Full range of X Detector Raw during X Scan sweep | 14.4
| Morgan State University | 2025-05-02 | Full range of X Detector during X Scan sweep | 74.8
| Morgan State University | 2025-05-02 | A-B noise | 11.112
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Z Scanner position at -3μm | 0.008699
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Z Scanner position at -3μm | 0.00514
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Z Scanner position at +3μm | 0.014195
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Z Scanner position at +3μm | 0.017795
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Z linearity by sample height | 0.99836
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Z linearity by sample height | 0.99288
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Z Height noise | 0.028
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Z Height noise | 0.018
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Z Height calibration error | 0.2251
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Z Drive calibration error | 0.23946
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y2 Detector noise (XY Servo on) | 0.23
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y2 Detector noise (XY Servo off) | 0.209
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y1 Detector noise (XY servo on) | 0.214
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y1 Detector noise (XY servo off) | 0.192
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Y Scan Stroke | 55
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y Scan Stroke | 171.014
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Y Non-linearity | 0.032
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Y Detector stroke calibration error (nominal scale error Y) | 0.283
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Y Detector noise (XY Servo on) | 0.395
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y Detector noise (XY servo on) | 0.17
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Y Detector noise (XY Servo off) | 0.445
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y Detector noise (XY servo off) | 0.149
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Repeatability @1st position | 0.02
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Repeatability @ 4th position | 0.02
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Repeatability @ 3rd position | 0.02
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Repeatability @ 2nd position | 0.02
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Accuracy @1st position | 0.1771
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Accuracy @ 4th position | 0.27579
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Accuracy @ 3rd position | 0.35431
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis, Accuracy @ 2nd position | 0.21879
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis OPM (-52000,-52000) | 3.062
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis OPM (-52000,52000) | 3.313
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis OPM (52000,-52000) | 3.875
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis OPM (52000,52000) | 3.142
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Y axis OPM (0,0) | 0.5871
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Y axis OPM | 0.5
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | XY stage calibration | 2.56
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | XY Orthogonality (-52000,-52000) | 90.03
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | XY Orthogonality (-52000,52000) | 89.74
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | XY Orthogonality (52000,-52000) | 89.85
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | XY Orthogonality (52000,52000) | 89.94
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | XY Orthogonality (0,0) | 89.93
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X2 Detector noise (XY Servo on) | 0.22
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X2 Detector noise (XY Servo off) | 0.202
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X1 Detector noise (XY servo on) | 0.215
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X1 Detector noise (XY servo off) | 0.199
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | X Scan Stroke | 55
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X Scan Stroke | 171.432
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | X Non-linearity | 0.062
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | X Detector stroke calibration error (nominal scale error X) | 0.167
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | X Detector noise (XY Servo on) | 0.391
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X Detector noise (XY servo on) | 0.165
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | X Detector noise (XY Servo off) | 0.439
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X Detector noise (XY servo off) | 0.148
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Repeatability @1st position | 0.03
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Repeatability @ 4th position | 0.02
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Repeatability @ 3rd position | 0.04
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Repeatability @ 2nd position | 0.02
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Accuracy @1st position | 0.28634
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Accuracy @ 4th position | 0.2421
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Accuracy @ 3rd position | 0.30864
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis, Accuracy @ 2nd position | 0.28804
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis OPM (-52000,-52000) | 1.62
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis OPM (-52000,52000) | 3.052
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis OPM (52000,-52000) | 1.815
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis OPM (52000,52000) | 2.731
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | X axis OPM (0,0) | 1.6494
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | X axis OPM | 1.014
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | SLD beam power at the end of the lens | 0.256
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.284
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.18
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.18
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.16
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.15
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.15
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | NCM driving% with AC160TS cantilever | 0.361
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | NCM driving% with AC160TS cantilever | 2.812
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | NCM Amplitude noise | 13.528
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | NCM Amplitude noise | 6.192
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.826
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.744
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.34
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.995
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 231.5
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 230.7
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 88.8
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.5
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 232.7
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 87.2
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | LD beam power at the end of the lens | 1
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Lateral(C-D) sensitivity | 10.9
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Lateral(C-D) noise | 0.106
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Full range of Z Height during Z Scan sweep | 17.9
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Z Height during Z Scan sweep | 16.8
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Full range of Z Detector during Z Scan sweep | 13.1
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Z Detector during Z Scan sweep | 14.6
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Y2 Detector during Y Scan sweep | 170.5
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.8
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Y1 Detector during Y Scan sweep | 170.4
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Full range of Y Detector during Y Scan sweep | 60
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of Y Detector during Y Scan sweep | 170.5
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of X2 Detector during X Scan sweep | 174.6
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of X1 Detector during X Scan sweep | 167
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Full range of X Detector Raw during X Scan sweep | 14.7
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | Full range of X Detector during X Scan sweep | 64
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | Full range of X Detector during X Scan sweep | 170.8
| 울산과학기술_신소재공학과(UNIST) | 2025-04-29 | A-B noise | 11.566
| Taiwan Instrument Research Institute(TIRI)(TIRI) | 2025-04-29 | A-B noise | 7.408
| CNR-Nanotec(Lecce) | 2025-04-28 | Z Scanner position at -3μm | 0.002011
| BestBay Co.,Limited | 2025-04-28 | Z Scanner position at -3μm | 0.005028
| CNR-Nanotec(Lecce) | 2025-04-28 | Z Scanner position at +3μm | 0.00842
| BestBay Co.,Limited | 2025-04-28 | Z Scanner position at +3μm | 0.00077
| CNR-Nanotec(Lecce) | 2025-04-28 | Z linearity by sample height | 0.99055
| BestBay Co.,Limited | 2025-04-28 | Z linearity by sample height | 1.0066
| CNR-Nanotec(Lecce) | 2025-04-28 | Z Height noise | 0.029
| BestBay Co.,Limited | 2025-04-28 | Z Height noise | 0.028
| CNR-Nanotec(Lecce) | 2025-04-28 | Z Height calibration error | 0.16379
| BestBay Co.,Limited | 2025-04-28 | Z Height calibration error | 0.2864
| CNR-Nanotec(Lecce) | 2025-04-28 | Z Drive calibration error | 0.068966
| BestBay Co.,Limited | 2025-04-28 | Z Drive calibration error | 0.28927
| BestBay Co.,Limited | 2025-04-28 | Y2 Detector noise (XY Servo on) | 0.293
| BestBay Co.,Limited | 2025-04-28 | Y2 Detector noise (XY Servo off) | 0.246
| BestBay Co.,Limited | 2025-04-28 | Y1 Detector noise (XY servo on) | 0.292
| BestBay Co.,Limited | 2025-04-28 | Y1 Detector noise (XY servo off) | 0.252
| CNR-Nanotec(Lecce) | 2025-04-28 | Y Scan Stroke | 107.8
| BestBay Co.,Limited | 2025-04-28 | Y Scan Stroke | 188.5
| CNR-Nanotec(Lecce) | 2025-04-28 | Y Non-linearity | 0.163
| BestBay Co.,Limited | 2025-04-28 | Y Non-linearity | 0.179
| CNR-Nanotec(Lecce) | 2025-04-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.126
| BestBay Co.,Limited | 2025-04-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.276
| CNR-Nanotec(Lecce) | 2025-04-28 | Y Detector noise (XY Servo on) | 0.262
| BestBay Co.,Limited | 2025-04-28 | Y Detector noise (XY servo on) | 0.238
| CNR-Nanotec(Lecce) | 2025-04-28 | Y Detector noise (XY servo off) | 0.257
| BestBay Co.,Limited | 2025-04-28 | Y Detector noise (XY servo off) | 0.187
| CNR-Nanotec(Lecce) | 2025-04-28 | Y axis OPM | 0.9848
| BestBay Co.,Limited | 2025-04-28 | Y axis OPM | 1.1354
| BestBay Co.,Limited | 2025-04-28 | X2 Detector noise (XY Servo on) | 0.284
| BestBay Co.,Limited | 2025-04-28 | X2 Detector noise (XY Servo off) | 0.273
| BestBay Co.,Limited | 2025-04-28 | X1 Detector noise (XY servo on) | 0.282
| BestBay Co.,Limited | 2025-04-28 | X1 Detector noise (XY servo off) | 0.262
| CNR-Nanotec(Lecce) | 2025-04-28 | X Scan Stroke | 107.8
| BestBay Co.,Limited | 2025-04-28 | X Scan Stroke | 196.2
| CNR-Nanotec(Lecce) | 2025-04-28 | X Non-linearity | 0.196
| BestBay Co.,Limited | 2025-04-28 | X Non-linearity | 0.08
| CNR-Nanotec(Lecce) | 2025-04-28 | X Detector stroke calibration error (nominal scale error X) | 0.051
| BestBay Co.,Limited | 2025-04-28 | X Detector stroke calibration error (nominal scale error X) | 0.186
| CNR-Nanotec(Lecce) | 2025-04-28 | X Detector noise (XY Servo on) | 0.257
| BestBay Co.,Limited | 2025-04-28 | X Detector noise (XY servo on) | 0.217
| CNR-Nanotec(Lecce) | 2025-04-28 | X Detector noise (XY servo off) | 0.257
| BestBay Co.,Limited | 2025-04-28 | X Detector noise (XY servo off) | 0.201
| CNR-Nanotec(Lecce) | 2025-04-28 | X axis OPM | 0.9418
| BestBay Co.,Limited | 2025-04-28 | X axis OPM | 1.765
| CNR-Nanotec(Lecce) | 2025-04-28 | Roughness Repeatability 1x1um(RSD%) | 1.83
| BestBay Co.,Limited | 2025-04-28 | Roughness Repeatability 1x1um(RSD%) | 3.574
| CNR-Nanotec(Lecce) | 2025-04-28 | Roughness Repeatability 1x1um(1σ) | 0.007
| BestBay Co.,Limited | 2025-04-28 | Roughness Repeatability 1x1um(1σ) | 0.006
| CNR-Nanotec(Lecce) | 2025-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| CNR-Nanotec(Lecce) | 2025-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.29
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.27
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.29
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.25
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| CNR-Nanotec(Lecce) | 2025-04-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| CNR-Nanotec(Lecce) | 2025-04-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| BestBay Co.,Limited | 2025-04-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| CNR-Nanotec(Lecce) | 2025-04-28 | NCM driving% with AC160TS cantilever | 1.189
| BestBay Co.,Limited | 2025-04-28 | NCM driving% with AC160TS cantilever | 0.315
| CNR-Nanotec(Lecce) | 2025-04-28 | NCM Amplitude noise | 16.6
| BestBay Co.,Limited | 2025-04-28 | NCM Amplitude noise | 9.563
| CNR-Nanotec(Lecce) | 2025-04-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.846
| BestBay Co.,Limited | 2025-04-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.45
| CNR-Nanotec(Lecce) | 2025-04-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.862
| BestBay Co.,Limited | 2025-04-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.268
| BestBay Co.,Limited | 2025-04-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 248.1
| BestBay Co.,Limited | 2025-04-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 266.6
| CNR-Nanotec(Lecce) | 2025-04-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.9
| BestBay Co.,Limited | 2025-04-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 275.2
| BestBay Co.,Limited | 2025-04-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 269.9
| CNR-Nanotec(Lecce) | 2025-04-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 161.7
| BestBay Co.,Limited | 2025-04-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| CNR-Nanotec(Lecce) | 2025-04-28 | Lateral(C-D) sensitivity | 24.3
| BestBay Co.,Limited | 2025-04-28 | Lateral(C-D) sensitivity | 24.5
| CNR-Nanotec(Lecce) | 2025-04-28 | Lateral(C-D) noise | 0.113
| BestBay Co.,Limited | 2025-04-28 | Lateral(C-D) noise | 0.155
| CNR-Nanotec(Lecce) | 2025-04-28 | Full range of Z Height during Z Scan sweep | 17
| BestBay Co.,Limited | 2025-04-28 | Full range of Z Height during Z Scan sweep | 16.8
| CNR-Nanotec(Lecce) | 2025-04-28 | Full range of Z Detector during Z Scan sweep | 12.2
| BestBay Co.,Limited | 2025-04-28 | Full range of Z Detector during Z Scan sweep | 12.7
| BestBay Co.,Limited | 2025-04-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| BestBay Co.,Limited | 2025-04-28 | Full range of Y2 Detector during Y Scan sweep | 182.9
| BestBay Co.,Limited | 2025-04-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| BestBay Co.,Limited | 2025-04-28 | Full range of Y1 Detector during Y Scan sweep | 194.3
| CNR-Nanotec(Lecce) | 2025-04-28 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| CNR-Nanotec(Lecce) | 2025-04-28 | Full range of Y Detector during Y Scan sweep | 107.8
| BestBay Co.,Limited | 2025-04-28 | Full range of Y Detector during Y Scan sweep | 188.6
| BestBay Co.,Limited | 2025-04-28 | Full range of X2 Detector Raw during X Scan sweep | 14.6
| BestBay Co.,Limited | 2025-04-28 | Full range of X2 Detector during X Scan sweep | 200.9
| BestBay Co.,Limited | 2025-04-28 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| BestBay Co.,Limited | 2025-04-28 | Full range of X1 Detector during X Scan sweep | 191.2
| CNR-Nanotec(Lecce) | 2025-04-28 | Full range of X Detector Raw during X Scan sweep | 13.3
| CNR-Nanotec(Lecce) | 2025-04-28 | Full range of X Detector during X Scan sweep | 107.8
| BestBay Co.,Limited | 2025-04-28 | Full range of X Detector during X Scan sweep | 196
| CNR-Nanotec(Lecce) | 2025-04-28 | A-B noise | 7.91
| BestBay Co.,Limited | 2025-04-28 | A-B noise | 6.775
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Z Scanner position at -3μm | 0.001148
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Z Scanner position at +3μm | 0.002582
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Z linearity by sample height | 1.0134
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Z Height noise | 0.018
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Z Height calibration error | 0.14943
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Z Drive calibration error | 0.34004
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Y Scan Stroke | 118.3
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Y Non-lineairty | 0.055
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Y Detector stroke calibration error (nominal scale error Y) | 0.095
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Y Detector noise (XY Servo on) | 0.26
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Y Detector noise (XY Servo off) | 0.25
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Y axis OPM | 1.22
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | X Scan Stroke | 118.2
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | X Non-lineairty | 0.103
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | X Detector stroke calibration error (nominal scale error X) | 0.27
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | X Detector noise (XY Servo on) | 0.26
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | X Detector noise (XY Servo off) | 0.25
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | X axis OPM | 1.51
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | SLD beam power at the end of the lens | 0.232
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Roughness Repeatability 1x1um(RSD%) | 4.156
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Roughness Repeatability 1x1um(1σ) | 0.007
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | NCM driving% with AC160TS cantilever | 2.441
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | NCM Amplitude noise | 7.29
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.464
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.56
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 177.4
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 169.7
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Lateral(C-D) sensitivity | 26.3
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Lateral(C-D) noise | 0.08
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Full range of Z Height during Z Scan sweep | 16.8
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Full range of Z Detector during Z Scan sweep | 15.6
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Full range of Y Detector during Y Scan sweep | 118.4
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Full range of X Detector Raw during X Scan sweep | 14
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | Full range of X Detector during X Scan sweep | 118.7
| 금오공과대학교_신소재공학과(일신랩서비스) | 2025-04-23 | A-B noise | 8.41
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Z Scanner position at -3μm | 0.002698
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Z Scanner position at +3μm | 0.001533
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Z linearity by sample height | 1.0027
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Z Height noise | 0.027
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Z Height calibration error | 0.1341
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Z Drive calibration error | 0.034483
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Y Scan Stroke | 55
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Y Non-linearity | 0.067
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Y Detector stroke calibration error (nominal scale error Y) | 0.437
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Y Detector noise (XY Servo on) | 0.275
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Y Detector noise (XY Servo off) | 0.269
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Y axis OPM | 0.7958
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | X Scan Stroke | 55
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | X Non-linearity | 0.037
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | X Detector stroke calibration error (nominal scale error X) | 0.072
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | X Detector noise (XY Servo on) | 0.292
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | X Detector noise (XY Servo off) | 0.265
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | X axis OPM | 0.8306
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Roughness Repeatability 1x1um(RSD%) | 2.758
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Roughness Repeatability 1x1um(1σ) | 0.008
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.59
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.58
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | NCM driving% with AC160TS cantilever | 0.215
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | NCM Amplitude noise | 9.014
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.149
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.862
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 108.5
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.9
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Lateral(C-D) sensitivity | 25.1
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Lateral(C-D) noise | 0.098
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Full range of Z Height during Z Scan sweep | 18.6
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Full range of Z Detector during Z Scan sweep | 12.9
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Full range of Y Detector during Y Scan sweep | 76.5
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Full range of X Detector Raw during X Scan sweep | 14.4
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | Full range of X Detector during X Scan sweep | 76.8
| Taiwan Instrument Research Institute(TIRI) | 2025-04-22 | A-B noise | 10.642
| Purdue University | 2025-04-21 | Z Scanner position at -3μm | 0.004255
| PSE demo lab (1/2) | 2025-04-21 | Z Scanner position at -3μm | 0.002875
| Purdue University | 2025-04-21 | Z Scanner position at +3μm | 0.006255
| PSE demo lab (1/2) | 2025-04-21 | Z Scanner position at +3μm | 0.010998
| Purdue University | 2025-04-21 | Z linearity by sample height | 0.99083
| PSE demo lab (1/2) | 2025-04-21 | Z linearity by sample height | 1.0165
| Purdue University | 2025-04-21 | Z Height noise | 0.029
| PSE demo lab (1/2) | 2025-04-21 | Z Height noise | 0.024
| PSE demo lab (1/2) | 2025-04-21 | Z Height calibration error | 0.11415
| Purdue University | 2025-04-21 | Z Height calibration error | 0.11707
| PSE demo lab (1/2) | 2025-04-21 | Z Drive calibration error | 0.08
| Purdue University | 2025-04-21 | Z Drive calibration error | 0.37951
| Purdue University | 2025-04-21 | Y Scan Stroke | 120.9
| PSE demo lab (1/2) | 2025-04-21 | Y Scan Stroke | 114.8
| Purdue University | 2025-04-21 | Y Non-linearity | 0.087
| PSE demo lab (1/2) | 2025-04-21 | Y Non-linearity | 0.213
| Purdue University | 2025-04-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.274
| PSE demo lab (1/2) | 2025-04-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.153
| Purdue University | 2025-04-21 | Y Detector noise (XY Servo on) | 0.409
| PSE demo lab (1/2) | 2025-04-21 | Y Detector noise (XY Servo on) | 0.263
| Purdue University | 2025-04-21 | Y Detector noise (XY Servo off) | 0.433
| PSE demo lab (1/2) | 2025-04-21 | Y Detector noise (XY Servo off) | 0.255
| Purdue University | 2025-04-21 | Y axis OPM | 1.3224
| PSE demo lab (1/2) | 2025-04-21 | Y axis OPM | 0.784
| Purdue University | 2025-04-21 | X Scan Stroke | 126.6
| PSE demo lab (1/2) | 2025-04-21 | X Scan Stroke | 114.9
| Purdue University | 2025-04-21 | X Non-linearity | 0.053
| PSE demo lab (1/2) | 2025-04-21 | X Non-linearity | 0.247
| Purdue University | 2025-04-21 | X Detector stroke calibration error (nominal scale error X) | 0.068
| PSE demo lab (1/2) | 2025-04-21 | X Detector stroke calibration error (nominal scale error X) | 0.034
| Purdue University | 2025-04-21 | X Detector noise (XY Servo on) | 0.428
| PSE demo lab (1/2) | 2025-04-21 | X Detector noise (XY Servo on) | 0.269
| Purdue University | 2025-04-21 | X Detector noise (XY Servo off) | 0.449
| PSE demo lab (1/2) | 2025-04-21 | X Detector noise (XY Servo off) | 0.258
| Purdue University | 2025-04-21 | X axis OPM | 1.1366
| PSE demo lab (1/2) | 2025-04-21 | X axis OPM | 0.985
| PSE demo lab (1/2) | 2025-04-21 | SLD beam power at the end of the lens | 0.28
| PSE demo lab (1/2) | 2025-04-21 | Roughness Repeatability 1x1um(RSD%) | 1.157
| PSE demo lab (1/2) | 2025-04-21 | Roughness Repeatability 1x1um(1σ) | 0.004
| Purdue University | 2025-04-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| PSE demo lab (1/2) | 2025-04-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Purdue University | 2025-04-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| PSE demo lab (1/2) | 2025-04-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Purdue University | 2025-04-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| PSE demo lab (1/2) | 2025-04-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Purdue University | 2025-04-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| PSE demo lab (1/2) | 2025-04-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Purdue University | 2025-04-21 | NCM driving% with AC160TS cantilever | 0.455
| PSE demo lab (1/2) | 2025-04-21 | NCM driving% with AC160TS cantilever | 3.997
| Purdue University | 2025-04-21 | NCM Amplitude noise | 9.088
| PSE demo lab (1/2) | 2025-04-21 | NCM Amplitude noise | 5.912
| Purdue University | 2025-04-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.648
| PSE demo lab (1/2) | 2025-04-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.802
| Purdue University | 2025-04-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.546
| PSE demo lab (1/2) | 2025-04-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.492
| Purdue University | 2025-04-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.1
| PSE demo lab (1/2) | 2025-04-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.9
| Purdue University | 2025-04-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 169.5
| PSE demo lab (1/2) | 2025-04-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 169.5
| PSE demo lab (1/2) | 2025-04-21 | Lateral(C-D) sensitivity | 48.109
| PSE demo lab (1/2) | 2025-04-21 | Lateral(C-D) noise | 0.119
| Purdue University | 2025-04-21 | Full range of Z Height during Z Scan sweep | 17.3
| PSE demo lab (1/2) | 2025-04-21 | Full range of Z Height during Z Scan sweep | 16.9
| Purdue University | 2025-04-21 | Full range of Z Detector during Z Scan sweep | 12.6
| PSE demo lab (1/2) | 2025-04-21 | Full range of Z Detector during Z Scan sweep | 14.4
| Purdue University | 2025-04-21 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| PSE demo lab (1/2) | 2025-04-21 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Purdue University | 2025-04-21 | Full range of Y Detector during Y Scan sweep | 120.9
| PSE demo lab (1/2) | 2025-04-21 | Full range of Y Detector during Y Scan sweep | 114.8
| Purdue University | 2025-04-21 | Full range of X Detector Raw during X Scan sweep | 14.9
| PSE demo lab (1/2) | 2025-04-21 | Full range of X Detector Raw during X Scan sweep | 13.6
| Purdue University | 2025-04-21 | Full range of X Detector during X Scan sweep | 126.6
| PSE demo lab (1/2) | 2025-04-21 | Full range of X Detector during X Scan sweep | 114.9
| Purdue University | 2025-04-21 | A-B noise | 7.689
| PSE demo lab (1/2) | 2025-04-21 | A-B noise | 16.446
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Z Scanner position at -3μm | 0.002131
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Z Scanner position at +3μm | 0.00647
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Z linearity by sample height | 0.99564
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Z Height noise | 0.03
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Z Height calibration error | 0.2682
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Z Drive calibration error | 0.19061
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Y Scan Stroke | 55
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Y Non-linearity | 0.046
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Y Detector stroke calibration error (nominal scale error Y) | 0.132
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Y Detector noise (XY Servo on) | 0.256
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Y Detector noise (XY Servo off) | 0.249
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Y axis OPM | 0.6852
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | X Scan Stroke | 55
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | X Non-linearity | 0.057
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | X Detector stroke calibration error (nominal scale error X) | 0.043
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | X Detector noise (XY Servo on) | 0.249
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | X Detector noise (XY Servo off) | 0.243
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | X axis OPM | 0.633
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Roughness Repeatability 1x1um(RSD%) | 0.541
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Roughness Repeatability 1x1um(1σ) | 0.002
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | NCM driving% with AC160TS cantilever | 1.498
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | NCM Amplitude noise | 9.237
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.371
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.146
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 109.2
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.2
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Lateral(C-D) sensitivity | 21.8
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Lateral(C-D) noise | 0.075
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Full range of Z Height during Z Scan sweep | 16.9
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Full range of Z Detector during Z Scan sweep | 12
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Full range of Y Detector during Y Scan sweep | 73.7
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Full range of X Detector Raw during X Scan sweep | 14.2
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | Full range of X Detector during X Scan sweep | 73.5
| Institute of Macromolecular Chemistry Petru Poni | 2025-04-17 | A-B noise | 5.848
| Samsung (SEC) SRD | 2025-04-16 | Z Linearity by Z scanner position | 1.35
| Samsung (SEC) SRD | 2025-04-16 | Full range of Z Height during Z Scan sweep | 13.21
| Samsung (SEC) SRD | 2025-04-16 | Z Linearity by 3 different sample height | 0.99851
| Samsung (SEC) SRD | 2025-04-16 | Full range of Z Detector during Z Scan sweep | 16.39
| Samsung (SEC) SRD | 2025-04-16 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.0217
| Samsung (SEC) SRD | 2025-04-16 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.5747
| Samsung (SEC) SRD | 2025-04-16 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.0217
| Samsung (SEC) SRD | 2025-04-16 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5585
| Samsung (SEC) SRD | 2025-04-16 | Z Height noise | 0.025
| Samsung (SEC) SRD | 2025-04-16 | Wafer placement(Loading) error on the load port1 (X offset) | 39.75
| Samsung (SEC) SRD | 2025-04-16 | Wafer placement(Loading) error on the load port1 (Y offset) | 66.67
| Samsung (SEC) SRD | 2025-04-16 | Wafer placement(Loading) error on the load port2 (X offset) | 28.1
| Samsung (SEC) SRD | 2025-04-16 | Wafer placement(Loading) error on the load port2 (Y offset) | 31.01
| Samsung (SEC) SRD | 2025-04-16 | OpticalAxisShiftPerFocusX | 0.008
| Samsung (SEC) SRD | 2025-04-16 | OpticalAxisShiftPerFocusY | -0.007
| Samsung (SEC) SRD | 2025-04-16 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 2.86e-20
| Samsung (SEC) SRD | 2025-04-16 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5.68e-05
| Samsung (SEC) SRD | 2025-04-16 | Lift pins synchronization offset check (X offset) | 9.29
| Samsung (SEC) SRD | 2025-04-16 | Lift pins synchronization offset check (Y offset) | 6.41
| Samsung (SEC) SRD | 2025-04-16 | Max OPM @25mm | 160.605
| Samsung (SEC) SRD | 2025-04-16 | Max OPM @10mm | 45.578
| Samsung (SEC) SRD | 2025-04-16 | Max OPM @5mm | 25.246
| Samsung (SEC) SRD | 2025-04-16 | Max OPM @1mm | 12.426
| Samsung (SEC) SRD | 2025-04-16 | OPM Repeatability @25mm | 6.156
| Samsung (SEC) SRD | 2025-04-16 | OPM Repeatability @10mm | 1.926
| Samsung (SEC) SRD | 2025-04-16 | OPM Repeatability @5mm | 0.819
| Samsung (SEC) SRD | 2025-04-16 | OPM Repeatability @1mm | 1.032
| Samsung (SEC) SRD | 2025-04-16 | Torque Profile (50mm, 200um/s) | 0.841
| Samsung (SEC) SRD | 2025-04-16 | Torque Profile (50mm, 20um/s) | 0.837
| Samsung (SEC) SRD | 2025-04-16 | Profiler noise @200um | 0.173
| Samsung (SEC) SRD | 2025-04-16 | Roll motion to X+ full scan | 36.5
| Samsung (SEC) SRD | 2025-04-16 | Position repeatability(10repeats, Position Based) | 0.82
| Samsung (SEC) SRD | 2025-04-16 | Lateral(C-D) sensitivity | 25.6
| Samsung (SEC) SRD | 2025-04-16 | Lateral(C-D) noise | 0.331
| Samsung (SEC) SRD | 2025-04-16 | Min. Vibration Transmissbility (~100Hz) | 0.036
| Samsung (SEC) SRD | 2025-04-16 | RMS of ZeroScan | 0.057
| Samsung (SEC) SRD | 2025-04-16 | Full range of X Detector during X Scan sweep | 186.5565
| Samsung (SEC) SRD | 2025-04-16 | Full range of X1 Detector during X Scan sweep | 184.7737
| Samsung (SEC) SRD | 2025-04-16 | Full range of X2 Detector during X Scan sweep | 188.3393
| Samsung (SEC) SRD | 2025-04-16 | Full range of Y Detector during Y Scan sweep | 178.4636
| Samsung (SEC) SRD | 2025-04-16 | Full range of Y1 Detector during Y Scan sweep | 179.1754
| Samsung (SEC) SRD | 2025-04-16 | Full range of Y2 Detector during Y Scan sweep | 177.7518
| Samsung (SEC) SRD | 2025-04-16 | Full range of X1 Detector Raw during X Scan sweep | 13.84
| Samsung (SEC) SRD | 2025-04-16 | Full range of X2 Detector Raw during X Scan sweep | 13.89
| Samsung (SEC) SRD | 2025-04-16 | Full range of Y1 Detector Raw during Y Scan sweep | 13.56
| Samsung (SEC) SRD | 2025-04-16 | Full range of Y2 Detector Raw during Y Scan sweep | 13.54
| Samsung (SEC) SRD | 2025-04-16 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 270.921
| Samsung (SEC) SRD | 2025-04-16 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 269.48
| Samsung (SEC) SRD | 2025-04-16 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 265.835
| Samsung (SEC) SRD | 2025-04-16 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 262.698
| Samsung (SEC) SRD | 2025-04-16 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 168.917
| Samsung (SEC) SRD | 2025-04-16 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.298
| Samsung (SEC) SRD | 2025-04-16 | XY Orthogonality | 89.99
| Samsung (SEC) SRD | 2025-04-16 | X Non-linearity error | 3.46
| Samsung (SEC) SRD | 2025-04-16 | Y Non-linearity error | 3.48
| Samsung (SEC) SRD | 2025-04-16 | X axis OPM | 1.151
| Samsung (SEC) SRD | 2025-04-16 | Y axis OPM | 1.265
| Samsung (SEC) SRD | 2025-04-16 | X Ringing(Rpv) | 0.245
| Samsung (SEC) SRD | 2025-04-16 | Y Ringing(Rpv) | 0.336
| Samsung (SEC) SRD | 2025-04-16 | X Detector noise (XY Servo on) | 0.267
| Samsung (SEC) SRD | 2025-04-16 | Y Detector noise (XY Servo on) | 0.536
| Samsung (SEC) SRD | 2025-04-16 | X1 Detector noise (XY Servo on) | 0.361
| Samsung (SEC) SRD | 2025-04-16 | Y1 Detector noise (XY Servo on) | 0.614
| Samsung (SEC) SRD | 2025-04-16 | X2 Detector noise (XY Servo on) | 0.31
| Samsung (SEC) SRD | 2025-04-16 | Y2 Detector noise (XY Servo on) | 0.516
| Samsung (SEC) SRD | 2025-04-16 | RMS error after Stage Calibration | 3.02
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from -X to +X (X scan) | 27.0996
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from -Y to +Y (Y scan) | 31.4941
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from -X to +X (Y scan) | 88.6231
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from -Y to +Y (X scan) | 13.1836
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from +X to -X (X scan) | 43.9453
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from +Y to -Y (Y scan) | 47.6074
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from +X to -X (Y scan) | 76.1719
| Samsung (SEC) SRD | 2025-04-16 | Image offset after moving from +Y to -Y (X scan) | 15.3809
| Samsung (SEC) SRD | 2025-04-16 | X axis, Repeatability @1st position | 0.05
| Samsung (SEC) SRD | 2025-04-16 | X axis, Repeatability @ 2nd position | 0.05
| Samsung (SEC) SRD | 2025-04-16 | X axis, Repeatability @ 3rd position | 0.05
| Samsung (SEC) SRD | 2025-04-16 | X axis, Repeatability @ 4th position | 0.06
| Samsung (SEC) SRD | 2025-04-16 | X axis, Accuracy @1st position | 0.29
| Samsung (SEC) SRD | 2025-04-16 | X axis, Accuracy @ 2nd position | 0.28
| Samsung (SEC) SRD | 2025-04-16 | X axis, Accuracy @ 3rd position | 0.3
| Samsung (SEC) SRD | 2025-04-16 | X axis, Accuracy @ 4th position | 0.32
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Repeatability @1st position | 0.08
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Repeatability @ 2nd position | 0.09
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Repeatability @ 3rd position | 0.09
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Repeatability @ 4th position | 0.09
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Accuracy @1st position | 0.36
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Accuracy @ 2nd position | 0.39
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Accuracy @ 3rd position | 0.4
| Samsung (SEC) SRD | 2025-04-16 | Y axis, Accuracy @ 4th position | 0.41
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning X offset (Vision Pattern Matching) | 1.6211
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.0996
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning X offset (In-Die Align) | 2.3779
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning Y offset (In-Die Align) | 2.5586
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning X offset (LLC Translation) | 3.2715
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning Y offset (LLC Translation) | 2.5977
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning X 1σ (LLC Translation) | 0.741
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning Y 1σ (LLC Translation) | 0.472
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning X offset (Global Align) | 7.1973
| Samsung (SEC) SRD | 2025-04-16 | XY Stage positioning Y offset (Global Align) | 2.8711
| Samsung (SEC) SRD | 2025-04-16 | Z Drift | 6.396
| Samsung (SEC) SRD | 2025-04-16 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.44828
| Samsung (SEC) SRD | 2025-04-16 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08126
| Samsung (SEC) SRD | 2025-04-16 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.63204
| Samsung (SEC) SRD | 2025-04-16 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11428
| Samsung (SEC) SRD | 2025-04-16 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22518
| Samsung (SEC) SRD | 2025-04-16 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.040819
| Samsung (SEC) SRD | 2025-04-16 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.61579
| Samsung (SEC) SRD | 2025-04-16 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11132
| Samsung (SEC) SRD | 2025-04-16 | Cycle time for 15-point measurement | 2748.483
| Samsung (SEC) SRD | 2025-04-16 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.641
| Samsung (SEC) SRD | 2025-04-16 | Roughness Repeatability 1x1um(10Repeats, RSD%) | 0.728
| Samsung (SEC) SRD | 2025-04-16 | Roughness Repeatability 80x80um(10Repeats, RSD%) | 0.44
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Z Scanner position at -3μm | 0.003464
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Z Scanner position at -3μm | 0.002196
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Z Scanner position at +3μm | 0.003407
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Z Scanner position at +3μm | 0.002461
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Z linearity by sample height | 0.99787
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Z linearity by sample height | 0.99063
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Z Height noise | 0.029
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Z Height noise | 0.028
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Z Height calibration error | 0.18965
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Z Height calibration error | 0.04023
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Z Drive calibration error | 0.04023
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Z Drive calibration error | 0.24521
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Y Scan Stroke | 55
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Y Non-linearity | 0.029
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Y Non-linearity | 0.1
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Y Detector stroke calibration error (nominal scale error Y) | 0.162
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Y Detector stroke calibration error (nominal scale error Y) | 0.372
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Y Detector noise (XY Servo on) | 0.385
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Y Detector noise (XY Servo on) | 0.316
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Y Detector noise (XY Servo off) | 0.43
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Y Detector noise (XY servo off) | 0.318
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Y axis OPM | 0.6018
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Y axis OPM | 1.0386
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | X Scan Stroke | 55
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | X Non-linearity | 0.031
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | X Non-linearity | 0.159
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | X Detector stroke calibration error (nominal scale error X) | 0.014
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | X Detector stroke calibration error (nominal scale error X) | 0.111
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | X Detector noise (XY Servo on) | 0.385
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | X Detector noise (XY Servo on) | 0.271
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | X Detector noise (XY Servo off) | 0.44
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | X Detector noise (XY servo off) | 0.26
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | X axis OPM | 0.8636
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | X axis OPM | 1.0406
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Roughness Repeatability 1x1um(RSD%) | 2.458
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Roughness Repeatability 1x1um(1σ) | 0.004
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.25
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.25
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.25
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.09
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | NCM driving% with AC160TS cantilever | 0.896
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | NCM driving% with AC160TS cantilever | 0.88
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | NCM Amplitude noise | 13.264
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | NCM Amplitude noise | 9.701
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.718
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.633
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.977
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.604
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 89.1
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 180.4
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 89.2
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 192
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Lateral(C-D) sensitivity | 5.6
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Lateral(C-D) noise | 0.143
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Full range of Z Height during Z Scan sweep | 16.8
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Full range of Z Height during Z Scan sweep | 17.7
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Full range of Z Detector during Z Scan sweep | 11.6
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Full range of Z Detector during Z Scan sweep | 13.3
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Full range of Y Detector during Y Scan sweep | 63.9
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Full range of Y Detector during Y Scan sweep | 133.6
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Full range of X Detector Raw during X Scan sweep | 14.1
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Full range of X Detector Raw during X Scan sweep | 13.8
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | Full range of X Detector during X Scan sweep | 62.8
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | Full range of X Detector during X Scan sweep | 132.4
| UNIVERSITY FEDERAL MINAS GERAIS | 2025-04-15 | A-B noise | 7.539
| Guangdong Huicheng Vacuum Technology | 2025-04-15 | A-B noise | 7.225
| North China University of Technology | 2025-04-14 | Z Scanner position at -3μm | 0.000903
| University of California, Davis | 2025-04-14 | Z Scanner position at -3μm | 0.002571
| Kyocera Corporation | 2025-04-14 | Z Scanner position at -3μm | 0.006675
| North China University of Technology | 2025-04-14 | Z Scanner position at +3μm | 0.006968
| University of California, Davis | 2025-04-14 | Z Scanner position at +3μm | 0.004374
| Kyocera Corporation | 2025-04-14 | Z Scanner position at +3μm | 0.009309
| North China University of Technology | 2025-04-14 | Z linearity by sample height | 1.007
| University of California, Davis | 2025-04-14 | Z linearity by sample height | 0.99088
| Kyocera Corporation | 2025-04-14 | Z linearity by sample height | 1.0126
| North China University of Technology | 2025-04-14 | Z Height noise | 0.029
| University of California, Davis | 2025-04-14 | Z Height noise | 0.029
| Kyocera Corporation | 2025-04-14 | Z Height noise | 0.027
| North China University of Technology | 2025-04-14 | Z Height calibration error | 0.22605
| University of California, Davis | 2025-04-14 | Z Height calibration error | 0.011494
| Kyocera Corporation | 2025-04-14 | Z Height calibration error | 0.021073
| North China University of Technology | 2025-04-14 | Z Drive calibration error | 0.057471
| University of California, Davis | 2025-04-14 | Z Drive calibration error | 0.39655
| Kyocera Corporation | 2025-04-14 | Z Drive calibration error | 0.44349
| Kyocera Corporation | 2025-04-14 | Y2 Detector noise (XY Servo on) | 0.211
| Kyocera Corporation | 2025-04-14 | Y2 Detector noise (XY Servo off) | 0.21
| Kyocera Corporation | 2025-04-14 | Y1 Detector noise (XY servo on) | 0.189
| Kyocera Corporation | 2025-04-14 | Y1 Detector noise (XY servo off) | 0.2
| North China University of Technology | 2025-04-14 | Y Scan Stroke | 108.4
| University of California, Davis | 2025-04-14 | Y Scan Stroke | 55
| Kyocera Corporation | 2025-04-14 | Y Scan Stroke | 171.3
| North China University of Technology | 2025-04-14 | Y Non-linearity | 0.22
| University of California, Davis | 2025-04-14 | Y Non-linearity | 0.121
| Kyocera Corporation | 2025-04-14 | Y Non-lineairty | 0.2
| North China University of Technology | 2025-04-14 | Y Detector stroke calibration error (nominal scale error Y) | 0.006
| University of California, Davis | 2025-04-14 | Y Detector stroke calibration error (nominal scale error Y) | 0.289
| Kyocera Corporation | 2025-04-14 | Y Detector stroke calibration error (nominal scale error Y) | 0.119
| North China University of Technology | 2025-04-14 | Y Detector noise (XY Servo on) | 0.245
| University of California, Davis | 2025-04-14 | Y Detector noise (XY Servo on) | 0.374
| Kyocera Corporation | 2025-04-14 | Y Detector noise (XY servo on) | 0.144
| North China University of Technology | 2025-04-14 | Y Detector noise (XY servo off) | 0.244
| University of California, Davis | 2025-04-14 | Y Detector noise (XY Servo off) | 0.39
| Kyocera Corporation | 2025-04-14 | Y Detector noise (XY servo off) | 0.14
| North China University of Technology | 2025-04-14 | Y axis OPM | 0.8056
| University of California, Davis | 2025-04-14 | Y axis OPM | 1.0976
| Kyocera Corporation | 2025-04-14 | Y axis OPM | 1.259
| Kyocera Corporation | 2025-04-14 | X2 Detector noise (XY Servo on) | 0.211
| Kyocera Corporation | 2025-04-14 | X2 Detector noise (XY Servo off) | 0.22
| Kyocera Corporation | 2025-04-14 | X1 Detector noise (XY servo on) | 0.204
| Kyocera Corporation | 2025-04-14 | X1 Detector noise (XY servo off) | 0.206
| North China University of Technology | 2025-04-14 | X Scan Stroke | 108.2
| University of California, Davis | 2025-04-14 | X Scan Stroke | 55
| Kyocera Corporation | 2025-04-14 | X Scan Stroke | 175.2
| North China University of Technology | 2025-04-14 | X Non-linearity | 0.22
| University of California, Davis | 2025-04-14 | X Non-linearity | 0.091
| Kyocera Corporation | 2025-04-14 | X Non-lineairty | 0.104
| North China University of Technology | 2025-04-14 | X Detector stroke calibration error (nominal scale error X) | 0.11
| University of California, Davis | 2025-04-14 | X Detector stroke calibration error (nominal scale error X) | 0.156
| Kyocera Corporation | 2025-04-14 | X Detector stroke calibration error (nominal scale error X) | 0.139
| North China University of Technology | 2025-04-14 | X Detector noise (XY Servo on) | 0.264
| University of California, Davis | 2025-04-14 | X Detector noise (XY Servo on) | 0.37
| Kyocera Corporation | 2025-04-14 | X Detector noise (XY servo on) | 0.149
| North China University of Technology | 2025-04-14 | X Detector noise (XY servo off) | 0.263
| University of California, Davis | 2025-04-14 | X Detector noise (XY Servo off) | 0.387
| Kyocera Corporation | 2025-04-14 | X Detector noise (XY servo off) | 0.154
| North China University of Technology | 2025-04-14 | X axis OPM | 1.4232
| University of California, Davis | 2025-04-14 | X axis OPM | 0.7942
| Kyocera Corporation | 2025-04-14 | X axis OPM | 1.607
| North China University of Technology | 2025-04-14 | Roughness Repeatability 1x1um(RSD%) | 2.2
| Kyocera Corporation | 2025-04-14 | Roughness Repeatability 1x1um(RSD%) | 2.377
| North China University of Technology | 2025-04-14 | Roughness Repeatability 1x1um(1σ) | 0.006
| Kyocera Corporation | 2025-04-14 | Roughness Repeatability 1x1um(1σ) | 0.004
| North China University of Technology | 2025-04-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| University of California, Davis | 2025-04-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| North China University of Technology | 2025-04-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| University of California, Davis | 2025-04-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.24
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.28
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.24
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.24
| Kyocera Corporation | 2025-04-14 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.23
| North China University of Technology | 2025-04-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of California, Davis | 2025-04-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Kyocera Corporation | 2025-04-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| North China University of Technology | 2025-04-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| University of California, Davis | 2025-04-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Kyocera Corporation | 2025-04-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.19
| North China University of Technology | 2025-04-14 | NCM driving% with AC160TS cantilever | 0.86
| University of California, Davis | 2025-04-14 | NCM driving% with AC160TS cantilever | 0.197
| Kyocera Corporation | 2025-04-14 | NCM driving% with AC160TS cantilever | 0.174
| North China University of Technology | 2025-04-14 | NCM Amplitude noise | 6.604
| University of California, Davis | 2025-04-14 | NCM Amplitude noise | 17.182
| Kyocera Corporation | 2025-04-14 | NCM Amplitude noise | 11.353
| North China University of Technology | 2025-04-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.343
| University of California, Davis | 2025-04-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.948
| Kyocera Corporation | 2025-04-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.808
| North China University of Technology | 2025-04-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.937
| University of California, Davis | 2025-04-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.486
| Kyocera Corporation | 2025-04-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.744
| Kyocera Corporation | 2025-04-14 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 236.3
| Kyocera Corporation | 2025-04-14 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 232.9
| North China University of Technology | 2025-04-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 159.8
| University of California, Davis | 2025-04-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 88
| Kyocera Corporation | 2025-04-14 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 246.4
| Kyocera Corporation | 2025-04-14 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 243.7
| North China University of Technology | 2025-04-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163
| University of California, Davis | 2025-04-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 86.2
| Kyocera Corporation | 2025-04-14 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| North China University of Technology | 2025-04-14 | Lateral(C-D) sensitivity | 35.8
| Kyocera Corporation | 2025-04-14 | Lateral(C-D) sensitivity | 7.1
| North China University of Technology | 2025-04-14 | Lateral(C-D) noise | 0.112
| Kyocera Corporation | 2025-04-14 | Lateral(C-D) noise | 0.148
| North China University of Technology | 2025-04-14 | Full range of Z Height during Z Scan sweep | 16.8
| University of California, Davis | 2025-04-14 | Full range of Z Height during Z Scan sweep | 18.1
| Kyocera Corporation | 2025-04-14 | Full range of Z Height during Z Scan sweep | 16.8
| North China University of Technology | 2025-04-14 | Full range of Z Detector during Z Scan sweep | 12.1
| University of California, Davis | 2025-04-14 | Full range of Z Detector during Z Scan sweep | 13.2
| Kyocera Corporation | 2025-04-14 | Full range of Z Detector during Z Scan sweep | 13.6
| Kyocera Corporation | 2025-04-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| Kyocera Corporation | 2025-04-14 | Full range of Y2 Detector during Y Scan sweep | 173.3
| Kyocera Corporation | 2025-04-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| Kyocera Corporation | 2025-04-14 | Full range of Y1 Detector during Y Scan sweep | 169.2
| North China University of Technology | 2025-04-14 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| University of California, Davis | 2025-04-14 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| North China University of Technology | 2025-04-14 | Full range of Y Detector during Y Scan sweep | 108.4
| University of California, Davis | 2025-04-14 | Full range of Y Detector during Y Scan sweep | 63.3
| Kyocera Corporation | 2025-04-14 | Full range of Y Detector during Y Scan sweep | 171.3
| Kyocera Corporation | 2025-04-14 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Kyocera Corporation | 2025-04-14 | Full range of X2 Detector during X Scan sweep | 174.5
| Kyocera Corporation | 2025-04-14 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Kyocera Corporation | 2025-04-14 | Full range of X1 Detector during X Scan sweep | 174.2
| North China University of Technology | 2025-04-14 | Full range of X Detector Raw during X Scan sweep | 13.3
| University of California, Davis | 2025-04-14 | Full range of X Detector Raw during X Scan sweep | 14.9
| North China University of Technology | 2025-04-14 | Full range of X Detector during X Scan sweep | 108.2
| University of California, Davis | 2025-04-14 | Full range of X Detector during X Scan sweep | 64
| Kyocera Corporation | 2025-04-14 | Full range of X Detector during X Scan sweep | 174.3
| North China University of Technology | 2025-04-14 | A-B noise | 7.631
| University of California, Davis | 2025-04-14 | A-B noise | 6.593
| Kyocera Corporation | 2025-04-14 | A-B noise | 8.025
| JCET Microelectronics | 2025-04-14 | Z Linearity by Z scanner position | 0.13
| JCET Microelectronics | 2025-04-14 | Z Linearity by Z scanner position | 1.07
| JCET Microelectronics | 2025-04-14 | Z Linearity by 4 different sample height | 0.99505
| JCET Microelectronics | 2025-04-14 | Z Linearity by 4 different sample height | 0.99478
| JCET Microelectronics | 2025-04-14 | Z Height noise | 0.062
| JCET Microelectronics | 2025-04-14 | Z Height noise | 0.026
| JCET Microelectronics | 2025-04-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 65.5575
| JCET Microelectronics | 2025-04-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7843
| JCET Microelectronics | 2025-04-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 65.4641
| JCET Microelectronics | 2025-04-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7083
| JCET Microelectronics | 2025-04-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 28.3718
| JCET Microelectronics | 2025-04-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6938
| JCET Microelectronics | 2025-04-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 28.1954
| JCET Microelectronics | 2025-04-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.604
| JCET Microelectronics | 2025-04-14 | Z Drift | 18.023
| JCET Microelectronics | 2025-04-14 | Y2 Detector noise (XY Servo on) | 0.303
| JCET Microelectronics | 2025-04-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 247.9795
| JCET Microelectronics | 2025-04-14 | Y1 Detector noise (XY Servo on) | 0.306
| JCET Microelectronics | 2025-04-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 254.1173
| JCET Microelectronics | 2025-04-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 158.959
| JCET Microelectronics | 2025-04-14 | Y Ringing(Rpv) | 0.176
| JCET Microelectronics | 2025-04-14 | Y Non-linearity error | 3.15
| JCET Microelectronics | 2025-04-14 | Y Detector noise (XY Servo on) | 0.249
| JCET Microelectronics | 2025-04-14 | Y axis OPM | 0.503
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.0654
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y offset (LLC Translation) | 1.6357
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y offset (In-Die Align) | 1.0889
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y offset (Global Align) | 3.5693
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3012
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y 1σ (LLC Translation) | 0.2666
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y 1σ (In-Die Align) | 0.1978
| JCET Microelectronics | 2025-04-14 | XY Stage positioning Y 1σ (Global Align) | 0.8404
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X offset (Vision Pattern Matching) | 2.5
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X offset (LLC Translation) | 1.4941
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X offset (In-Die Align) | 1.0791
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X offset (Global Align) | 3.6914
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.5385
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X 1σ (LLC Translation) | 0.3192
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X 1σ (In-Die Align) | 0.2292
| JCET Microelectronics | 2025-04-14 | XY Stage positioning X 1σ (Global Align) | 0.7453
| JCET Microelectronics | 2025-04-14 | X2 Detector noise (XY Servo on) | 0.321
| JCET Microelectronics | 2025-04-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 271.2516
| JCET Microelectronics | 2025-04-14 | X1 Detector noise (XY Servo on) | 0.314
| JCET Microelectronics | 2025-04-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 267.7471
| JCET Microelectronics | 2025-04-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 173.869
| JCET Microelectronics | 2025-04-14 | X Ringing(Rpv) | 0.153
| JCET Microelectronics | 2025-04-14 | X Non-linearity error | 2.38
| JCET Microelectronics | 2025-04-14 | X Detector noise (XY Servo on) | 0.259
| JCET Microelectronics | 2025-04-14 | X axis OPM | 2.586
| JCET Microelectronics | 2025-04-14 | Wafer placement(Loading) error on the load port2 (Y offset) | 24.5
| JCET Microelectronics | 2025-04-14 | Wafer placement(Loading) error on the load port2 (X offset) | 36.85
| JCET Microelectronics | 2025-04-14 | Wafer placement(Loading) error on the load port1 (Y offset) | 13.5
| JCET Microelectronics | 2025-04-14 | Wafer placement(Loading) error on the load port1 (X offset) | 29.3
| JCET Microelectronics | 2025-04-14 | Torque Profile (50mm, 20um/s) | 0.452
| JCET Microelectronics | 2025-04-14 | Torque Profile (50mm, 200um/s) | 0.438
| JCET Microelectronics | 2025-04-14 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.863
| JCET Microelectronics | 2025-04-14 | Roughness Repeatability 80x80um(RSD%) | 0.219
| JCET Microelectronics | 2025-04-14 | Roughness Repeatability 1x1um(RSD%) | 7.988
| JCET Microelectronics | 2025-04-14 | Roll motion to X+ full scan | 7.4
| JCET Microelectronics | 2025-04-14 | RMS of ZeroScan | 0.038
| JCET Microelectronics | 2025-04-14 | RMS error after Stage Calibration | 2.28
| JCET Microelectronics | 2025-04-14 | Profiler noise @200um | 0.083
| JCET Microelectronics | 2025-04-14 | Position repeatability(10repeats, Position Based) | 0.38
| JCET Microelectronics | 2025-04-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| JCET Microelectronics | 2025-04-14 | OpticalAxisShiftPerFocusY | -0.0016
| JCET Microelectronics | 2025-04-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| JCET Microelectronics | 2025-04-14 | OpticalAxisShiftPerFocusX | -0.0038
| JCET Microelectronics | 2025-04-14 | OPM Repeatability @5mm | 1.078
| JCET Microelectronics | 2025-04-14 | OPM Repeatability @25mm | 2.56
| JCET Microelectronics | 2025-04-14 | OPM Repeatability @1mm | 0.701
| JCET Microelectronics | 2025-04-14 | OPM Repeatability @10mm | 1.604
| JCET Microelectronics | 2025-04-14 | Min. Vibration Transmissbility (~100Hz) | 0.057
| JCET Microelectronics | 2025-04-14 | Max OPM @5mm | 26.31
| JCET Microelectronics | 2025-04-14 | Max OPM @25mm | 132.213
| JCET Microelectronics | 2025-04-14 | Max OPM @1mm | 7.828
| JCET Microelectronics | 2025-04-14 | Max OPM @10mm | 44.972
| JCET Microelectronics | 2025-04-14 | Lift pins synchronization offset check (Y offset) | 17.59
| JCET Microelectronics | 2025-04-14 | Lift pins synchronization offset check (X offset) | 11.52
| JCET Microelectronics | 2025-04-14 | Lateral(C-D) sensitivity | 11.9
| JCET Microelectronics | 2025-04-14 | Lateral(C-D) noise | 0.126
| JCET Microelectronics | 2025-04-14 | Image offset after moving from -Y to +Y (Y scan) | 53.4668
| JCET Microelectronics | 2025-04-14 | Image offset after moving from -Y to +Y (X scan) | 15.38
| JCET Microelectronics | 2025-04-14 | Image offset after moving from -X to +X (Y scan) | 42.48
| JCET Microelectronics | 2025-04-14 | Image offset after moving from -X to +X (X scan) | 55.66
| JCET Microelectronics | 2025-04-14 | Image offset after moving from +Y to -Y (Y scan) | 41.0156
| JCET Microelectronics | 2025-04-14 | Image offset after moving from +Y to -Y (X scan) | 16.8457
| JCET Microelectronics | 2025-04-14 | Image offset after moving from +X to -X (Y scan) | 30.7617
| JCET Microelectronics | 2025-04-14 | Image offset after moving from +X to -X (X scan) | 65.1856
| JCET Microelectronics | 2025-04-14 | Full range of Z Height during Z Scan sweep | 16.5
| JCET Microelectronics | 2025-04-14 | Full range of Z Detector during Z Scan sweep | 12.51
| JCET Microelectronics | 2025-04-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| JCET Microelectronics | 2025-04-14 | Full range of Y2 Detector during Y Scan sweep | 197.141
| JCET Microelectronics | 2025-04-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.07
| JCET Microelectronics | 2025-04-14 | Full range of Y1 Detector during Y Scan sweep | 189.5671
| JCET Microelectronics | 2025-04-14 | Full range of Y Detector during Y Scan sweep | 193.3541
| JCET Microelectronics | 2025-04-14 | Full range of X2 Detector Raw during X Scan sweep | 14.79
| JCET Microelectronics | 2025-04-14 | Full range of X2 Detector during X Scan sweep | 208.0181
| JCET Microelectronics | 2025-04-14 | Full range of X1 Detector Raw during X Scan sweep | 14.78
| JCET Microelectronics | 2025-04-14 | Full range of X1 Detector during X Scan sweep | 214.0128
| JCET Microelectronics | 2025-04-14 | Full range of X Detector during X Scan sweep | 211.0154
| JCET Microelectronics | 2025-04-14 | Full range in voltage during Z Scan sweep | 13.56
| JCET Microelectronics | 2025-04-14 | Full range in um during Z Scan sweep | 44.4
| JCET Microelectronics | 2025-04-14 | Cycle time for 15-point measurement | 1749.192
| JCET Microelectronics | 2025-04-14 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| JCET Microelectronics | 2025-04-14 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| JCET Microelectronics | 2025-04-14 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| JCET Microelectronics | 2025-04-14 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| JCET Microelectronics | 2025-04-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| JCET Microelectronics | 2025-04-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| JCET Microelectronics | 2025-04-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.11
| JCET Microelectronics | 2025-04-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Z Scanner position at -3μm | 0.007238
| East China Normal University | 2025-04-10 | Z Scanner position at -3μm | 0.000154
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Z Scanner position at +3μm | 0.010431
| East China Normal University | 2025-04-10 | Z Scanner position at +3μm | 0.005974
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Z linearity by sample height | 1.0077
| East China Normal University | 2025-04-10 | Z linearity by sample height | 0.99965
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Z Height noise | 0.029
| East China Normal University | 2025-04-10 | Z Height noise | 0.029
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Z Height calibration error | 0.18487
| East China Normal University | 2025-04-10 | Z Height calibration error | 0.034483
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Z Drive calibration error | 0.063218
| East China Normal University | 2025-04-10 | Z Drive calibration error | 0.12165
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y2 Detector noise (XY Servo on) | 0.232
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y2 Detector noise (XY Servo off) | 0.23
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y1 Detector noise (XY servo on) | 0.239
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y1 Detector noise (XY servo off) | 0.24
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y Scan Stroke | 180.7
| East China Normal University | 2025-04-10 | Y Scan Stroke | 111.9
| East China Normal University | 2025-04-10 | Y Non-linearity | 0.244
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y Non-lineairty | 0.24
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y Detector stroke calibration error (nominal scale error Y) | 0.274
| East China Normal University | 2025-04-10 | Y Detector stroke calibration error (nominal scale error Y) | 0.405
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y Detector noise (XY servo on) | 0.173
| East China Normal University | 2025-04-10 | Y Detector noise (XY Servo on) | 0.248
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y Detector noise (XY servo off) | 0.173
| East China Normal University | 2025-04-10 | Y Detector noise (XY servo off) | 0.244
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Y axis OPM | 1.5934
| East China Normal University | 2025-04-10 | Y axis OPM | 0.712
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X2 Detector noise (XY Servo on) | 0.234
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X2 Detector noise (XY Servo off) | 0.231
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X1 Detector noise (XY servo on) | 0.23
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X1 Detector noise (XY servo off) | 0.226
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X Scan Stroke | 176
| East China Normal University | 2025-04-10 | X Scan Stroke | 109.3
| East China Normal University | 2025-04-10 | X Non-linearity | 0.292
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X Non-lineairty | 0.18
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X Detector stroke calibration error (nominal scale error X) | 0.253
| East China Normal University | 2025-04-10 | X Detector stroke calibration error (nominal scale error X) | 0.098
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X Detector noise (XY servo on) | 0.175
| East China Normal University | 2025-04-10 | X Detector noise (XY Servo on) | 0.25
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X Detector noise (XY servo off) | 0.171
| East China Normal University | 2025-04-10 | X Detector noise (XY servo off) | 0.246
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | X axis OPM | 1.2414
| East China Normal University | 2025-04-10 | X axis OPM | 0.9366
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Roughness Repeatability 1x1um(RSD%) | 2.282
| East China Normal University | 2025-04-10 | Roughness Repeatability 1x1um(RSD%) | 0.883
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Roughness Repeatability 1x1um(1σ) | 0.003
| East China Normal University | 2025-04-10 | Roughness Repeatability 1x1um(1σ) | 0.003
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| East China Normal University | 2025-04-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| East China Normal University | 2025-04-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.32
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.32
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.31
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.31
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| East China Normal University | 2025-04-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| East China Normal University | 2025-04-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | NCM driving% with AC160TS cantilever | 0.887
| East China Normal University | 2025-04-10 | NCM driving% with AC160TS cantilever | 0.775
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | NCM Amplitude noise | 13.239
| East China Normal University | 2025-04-10 | NCM Amplitude noise | 11.081
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.717
| East China Normal University | 2025-04-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.66
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.571
| East China Normal University | 2025-04-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.838
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 250.9
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 244.3
| East China Normal University | 2025-04-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.4
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 240.6
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 240
| East China Normal University | 2025-04-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.2
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Lateral(C-D) sensitivity | 4.7
| East China Normal University | 2025-04-10 | Lateral(C-D) sensitivity | 19.8
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Lateral(C-D) noise | 0.14
| East China Normal University | 2025-04-10 | Lateral(C-D) noise | 0.118
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Z Height during Z Scan sweep | 17.4
| East China Normal University | 2025-04-10 | Full range of Z Height during Z Scan sweep | 17.5
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Z Detector during Z Scan sweep | 13.1
| East China Normal University | 2025-04-10 | Full range of Z Detector during Z Scan sweep | 13.5
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Y2 Detector during Y Scan sweep | 183.4
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Y1 Detector during Y Scan sweep | 178.6
| East China Normal University | 2025-04-10 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of Y Detector during Y Scan sweep | 181
| East China Normal University | 2025-04-10 | Full range of Y Detector during Y Scan sweep | 111.8
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of X2 Detector Raw during X Scan sweep | 14.7
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of X2 Detector during X Scan sweep | 176.6
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of X1 Detector Raw during X Scan sweep | 14.6
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of X1 Detector during X Scan sweep | 175.3
| East China Normal University | 2025-04-10 | Full range of X Detector Raw during X Scan sweep | 13.4
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | Full range of X Detector during X Scan sweep | 175.9
| East China Normal University | 2025-04-10 | Full range of X Detector during X Scan sweep | 109.4
| Nanjing Pukou buyuexin integrated circuit Co., Ltd | 2025-04-10 | A-B noise | 7.478
| East China Normal University | 2025-04-10 | A-B noise | 5.808
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Z Scanner position at -3μm | 0.004401
| Accretech America Inc. | 2025-04-08 | Z Scanner position at -3μm | 0.002741
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Z Scanner position at +3μm | 0.011185
| Accretech America Inc. | 2025-04-08 | Z Scanner position at +3μm | 0.005337
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Z linearity by sample height | 0.99255
| Accretech America Inc. | 2025-04-08 | Z linearity by sample height | 1.0128
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Z Height noise | 0.028
| Accretech America Inc. | 2025-04-08 | Z Height noise | 0.033
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Z Height calibration error | 0.18487
| Accretech America Inc. | 2025-04-08 | Z Height calibration error | 0.000976
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Z Drive calibration error | 0.3841
| Accretech America Inc. | 2025-04-08 | Z Drive calibration error | 0.28195
| Accretech America Inc. | 2025-04-08 | Y2 Detector noise (XY Servo on) | 0.227
| Accretech America Inc. | 2025-04-08 | Y2 Detector noise (XY Servo off) | 0.226
| Accretech America Inc. | 2025-04-08 | Y1 Detector noise (XY servo on) | 0.222
| Accretech America Inc. | 2025-04-08 | Y1 Detector noise (XY servo off) | 0.221
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Y Scan Stroke | 107.8
| Accretech America Inc. | 2025-04-08 | Y Scan Stroke | 179.6
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Y Non-linearity | 0.211
| Accretech America Inc. | 2025-04-08 | Y Non-linearity | 0.166
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Y Detector stroke calibration error (nominal scale error Y) | 0.38
| Accretech America Inc. | 2025-04-08 | Y Detector stroke calibration error (nominal scale error Y) | 0.082
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Y Detector noise (XY Servo on) | 0.252
| Accretech America Inc. | 2025-04-08 | Y Detector noise (XY servo on) | 0.159
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Y Detector noise (XY servo off) | 0.248
| Accretech America Inc. | 2025-04-08 | Y Detector noise (XY servo off) | 0.159
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Y axis OPM | 0.716
| Accretech America Inc. | 2025-04-08 | Y axis OPM | 0.868
| Accretech America Inc. | 2025-04-08 | X2 Detector noise (XY Servo on) | 0.248
| Accretech America Inc. | 2025-04-08 | X2 Detector noise (XY Servo off) | 0.247
| Accretech America Inc. | 2025-04-08 | X1 Detector noise (XY servo on) | 0.234
| Accretech America Inc. | 2025-04-08 | X1 Detector noise (XY servo off) | 0.225
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | X Scan Stroke | 111.1
| Accretech America Inc. | 2025-04-08 | X Scan Stroke | 175
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | X Non-linearity | 0.125
| Accretech America Inc. | 2025-04-08 | X Non-linearity | 0.172
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | X Detector stroke calibration error (nominal scale error X) | 0.155
| Accretech America Inc. | 2025-04-08 | X Detector stroke calibration error (nominal scale error X) | 0.4
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | X Detector noise (XY Servo on) | 0.266
| Accretech America Inc. | 2025-04-08 | X Detector noise (XY servo on) | 0.181
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | X Detector noise (XY servo off) | 0.259
| Accretech America Inc. | 2025-04-08 | X Detector noise (XY servo off) | 0.174
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | X axis OPM | 1.549
| Accretech America Inc. | 2025-04-08 | X axis OPM | 0.85
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Roughness Repeatability 1x1um(RSD%) | 1.22
| Accretech America Inc. | 2025-04-08 | Roughness Repeatability 1x1um(RSD%) | 0.68
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Roughness Repeatability 1x1um(1σ) | 0.002
| Accretech America Inc. | 2025-04-08 | Roughness Repeatability 1x1um(1σ) | 0.002
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.32
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.34
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.35
| Accretech America Inc. | 2025-04-08 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Accretech America Inc. | 2025-04-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Accretech America Inc. | 2025-04-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | NCM driving% with AC160TS cantilever | 0.499
| Accretech America Inc. | 2025-04-08 | NCM driving% with AC160TS cantilever | 0.461
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | NCM Amplitude noise | 10.446
| Accretech America Inc. | 2025-04-08 | NCM Amplitude noise | 14.056
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.884
| Accretech America Inc. | 2025-04-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.925
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.234
| Accretech America Inc. | 2025-04-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.95
| Accretech America Inc. | 2025-04-08 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 250.7
| Accretech America Inc. | 2025-04-08 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 244.5
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 155.9
| Accretech America Inc. | 2025-04-08 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 246.6
| Accretech America Inc. | 2025-04-08 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 236.3
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160.8
| Accretech America Inc. | 2025-04-08 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Lateral(C-D) sensitivity | 23.8
| Accretech America Inc. | 2025-04-08 | Lateral(C-D) sensitivity | 23.8
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Lateral(C-D) noise | 0.103
| Accretech America Inc. | 2025-04-08 | Lateral(C-D) noise | 0.105
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Full range of Z Height during Z Scan sweep | 18
| Accretech America Inc. | 2025-04-08 | Full range of Z Height during Z Scan sweep | 17.6
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Full range of Z Detector during Z Scan sweep | 13.2
| Accretech America Inc. | 2025-04-08 | Full range of Z Detector during Z Scan sweep | 13.1
| Accretech America Inc. | 2025-04-08 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| Accretech America Inc. | 2025-04-08 | Full range of Y2 Detector during Y Scan sweep | 180.7
| Accretech America Inc. | 2025-04-08 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Accretech America Inc. | 2025-04-08 | Full range of Y1 Detector during Y Scan sweep | 178.5
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Full range of Y Detector during Y Scan sweep | 107.8
| Accretech America Inc. | 2025-04-08 | Full range of Y Detector during Y Scan sweep | 179.6
| Accretech America Inc. | 2025-04-08 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Accretech America Inc. | 2025-04-08 | Full range of X2 Detector during X Scan sweep | 178
| Accretech America Inc. | 2025-04-08 | Full range of X1 Detector Raw during X Scan sweep | 14.6
| Accretech America Inc. | 2025-04-08 | Full range of X1 Detector during X Scan sweep | 172.9
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Full range of X Detector Raw during X Scan sweep | 13.8
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | Full range of X Detector during X Scan sweep | 111.1
| Accretech America Inc. | 2025-04-08 | Full range of X Detector during X Scan sweep | 175.5
| Quantum Technology Yangtze River Delta Industrial Innovation Center | 2025-04-08 | A-B noise | 7.941
| Accretech America Inc. | 2025-04-08 | A-B noise | 7.456
| Hubei Dinghui | 2025-04-07 | Z Linearity by Z scanner position | 0.38
| Hubei Dinghui | 2025-04-07 | Z Linearity by 4 different sample height | 1.0056
| Hubei Dinghui | 2025-04-07 | Z Height noise | 0.027
| Hubei Dinghui | 2025-04-07 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.6535
| Hubei Dinghui | 2025-04-07 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.6039
| Hubei Dinghui | 2025-04-07 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4902
| Hubei Dinghui | 2025-04-07 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4536
| Hubei Dinghui | 2025-04-07 | Z Drift | 9.0953
| Hubei Dinghui | 2025-04-07 | Y2 Detector noise (XY Servo on) | 0.526
| Hubei Dinghui | 2025-04-07 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 244.5148
| Hubei Dinghui | 2025-04-07 | Y1 Detector noise (XY Servo on) | 0.55
| Hubei Dinghui | 2025-04-07 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 265.615
| Hubei Dinghui | 2025-04-07 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 164.261
| Hubei Dinghui | 2025-04-07 | Y Ringing(Rpv) | 0.127
| Hubei Dinghui | 2025-04-07 | Y Non-linearity error | 3.9454
| Hubei Dinghui | 2025-04-07 | Y Detector noise (XY Servo on) | 0.45
| Hubei Dinghui | 2025-04-07 | Y axis, Repeatability @1st position | 0.04
| Hubei Dinghui | 2025-04-07 | Y axis, Repeatability @ 4th position | 0.05
| Hubei Dinghui | 2025-04-07 | Y axis, Repeatability @ 3rd position | 0.03
| Hubei Dinghui | 2025-04-07 | Y axis, Repeatability @ 2nd position | 0.05
| Hubei Dinghui | 2025-04-07 | Y axis, Accuracy @1st position | 0.21
| Hubei Dinghui | 2025-04-07 | Y axis, Accuracy @ 4th position | 0.33
| Hubei Dinghui | 2025-04-07 | Y axis, Accuracy @ 3rd position | 0.21
| Hubei Dinghui | 2025-04-07 | Y axis, Accuracy @ 2nd position | 0.29
| Hubei Dinghui | 2025-04-07 | Y axis OPM | 0.847
| Hubei Dinghui | 2025-04-07 | X2 Detector noise (XY Servo on) | 0.494
| Hubei Dinghui | 2025-04-07 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 279.5774
| Hubei Dinghui | 2025-04-07 | X1 Detector noise (XY Servo on) | 0.498
| Hubei Dinghui | 2025-04-07 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 278.1135
| Hubei Dinghui | 2025-04-07 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 179.965
| Hubei Dinghui | 2025-04-07 | X Ringing(Rpv) | 0.172
| Hubei Dinghui | 2025-04-07 | X Non-linearity error | 3.1152
| Hubei Dinghui | 2025-04-07 | X Detector noise (XY Servo on) | 0.382
| Hubei Dinghui | 2025-04-07 | X axis, Repeatability @1st position | 0.04
| Hubei Dinghui | 2025-04-07 | X axis, Repeatability @ 4th position | 0.03
| Hubei Dinghui | 2025-04-07 | X axis, Repeatability @ 3rd position | 0.03
| Hubei Dinghui | 2025-04-07 | X axis, Repeatability @ 2nd position | 0.06
| Hubei Dinghui | 2025-04-07 | X axis, Accuracy @1st position | 0.25
| Hubei Dinghui | 2025-04-07 | X axis, Accuracy @ 4th position | 0.25
| Hubei Dinghui | 2025-04-07 | X axis, Accuracy @ 3rd position | 0.14
| Hubei Dinghui | 2025-04-07 | X axis, Accuracy @ 2nd position | 0.29
| Hubei Dinghui | 2025-04-07 | X axis OPM | 0.243
| Hubei Dinghui | 2025-04-07 | Torque Profile (50mm, 20um/s) | 0.57434
| Hubei Dinghui | 2025-04-07 | Torque Profile (50mm, 200um/s) | 0.35309
| Hubei Dinghui | 2025-04-07 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.68649
| Hubei Dinghui | 2025-04-07 | Roughness Repeatability 80x80um(RSD%) | 0.60652
| Hubei Dinghui | 2025-04-07 | Roughness Repeatability 1x1um(RSD%) | 1.9706
| Hubei Dinghui | 2025-04-07 | RMS of ZeroScan | 0.052
| Hubei Dinghui | 2025-04-07 | RMS error after Stage Calibration | 2.12
| Hubei Dinghui | 2025-04-07 | Profiler noise @200um | 0.183
| Hubei Dinghui | 2025-04-07 | Position repeatability(10repeats, Position Based) | 0.90332
| Hubei Dinghui | 2025-04-07 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4.83e-05
| Hubei Dinghui | 2025-04-07 | OpticalAxisShiftPerFocusY | -0.0003
| Hubei Dinghui | 2025-04-07 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000207
| Hubei Dinghui | 2025-04-07 | OpticalAxisShiftPerFocusX | -0.0005
| Hubei Dinghui | 2025-04-07 | OPM Repeatability @5mm | 1.587
| Hubei Dinghui | 2025-04-07 | OPM Repeatability @25mm | 2.769
| Hubei Dinghui | 2025-04-07 | OPM Repeatability @1mm | 0.46
| Hubei Dinghui | 2025-04-07 | OPM Repeatability @10mm | 0.578
| Hubei Dinghui | 2025-04-07 | Min. Vibration Transmissbility (~100Hz) | 0.038224
| Hubei Dinghui | 2025-04-07 | Max OPM @5mm | 10.874
| Hubei Dinghui | 2025-04-07 | Max OPM @25mm | 33.909
| Hubei Dinghui | 2025-04-07 | Max OPM @1mm | 5.688
| Hubei Dinghui | 2025-04-07 | Max OPM @10mm | 14.589
| Hubei Dinghui | 2025-04-07 | Lift pins synchronization offset check (Y offset) | 41.82
| Hubei Dinghui | 2025-04-07 | Lift pins synchronization offset check (X offset) | 39.73
| Hubei Dinghui | 2025-04-07 | Lateral(C-D) sensitivity | 62.1
| Hubei Dinghui | 2025-04-07 | Lateral(C-D) noise | 0.19144
| Hubei Dinghui | 2025-04-07 | Image offset after moving from -Y to +Y (Y scan) | 42.4805
| Hubei Dinghui | 2025-04-07 | Image offset after moving from -Y to +Y (X scan) | 27.832
| Hubei Dinghui | 2025-04-07 | Image offset after moving from -X to +X (Y scan) | 50.5371
| Hubei Dinghui | 2025-04-07 | Image offset after moving from -X to +X (X scan) | 21.2402
| Hubei Dinghui | 2025-04-07 | Image offset after moving from +Y to -Y (Y scan) | 59.3262
| Hubei Dinghui | 2025-04-07 | Image offset after moving from +Y to -Y (X scan) | 49.0723
| Hubei Dinghui | 2025-04-07 | Image offset after moving from +X to -X (Y scan) | 8.0566
| Hubei Dinghui | 2025-04-07 | Image offset after moving from +X to -X (X scan) | 60.0586
| Hubei Dinghui | 2025-04-07 | Full range of Z Height during Z Scan sweep | 16.8504
| Hubei Dinghui | 2025-04-07 | Full range of Z Detector during Z Scan sweep | 12.66
| Hubei Dinghui | 2025-04-07 | Full range of Y2 Detector Raw during Y Scan sweep | 13.53
| Hubei Dinghui | 2025-04-07 | Full range of Y2 Detector during Y Scan sweep | 168.5151
| Hubei Dinghui | 2025-04-07 | Full range of Y1 Detector Raw during Y Scan sweep | 13.97
| Hubei Dinghui | 2025-04-07 | Full range of Y1 Detector during Y Scan sweep | 199.5741
| Hubei Dinghui | 2025-04-07 | Full range of Y Detector during Y Scan sweep | 168.5151
| Hubei Dinghui | 2025-04-07 | Full range of X2 Detector Raw during X Scan sweep | 13.17
| Hubei Dinghui | 2025-04-07 | Full range of X2 Detector during X Scan sweep | 197.0639
| Hubei Dinghui | 2025-04-07 | Full range of X1 Detector Raw during X Scan sweep | 13.07
| Hubei Dinghui | 2025-04-07 | Full range of X1 Detector during X Scan sweep | 187.1891
| Hubei Dinghui | 2025-04-07 | Full range of X Detector during X Scan sweep | 197.0639
| Hubei Dinghui | 2025-04-07 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.17192
| Hubei Dinghui | 2025-04-07 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.95311
| Hubei Dinghui | 2025-04-07 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.17488
| Hubei Dinghui | 2025-04-07 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.96571
| PSI Demo | 2025-04-04 | Z Scanner position at -3μm | 0.00166
| PSI Demo | 2025-04-04 | Z Scanner position at +3μm | 0.017309
| PSI Demo | 2025-04-04 | Z linearity by sample height | 1.0052
| PSI Demo | 2025-04-04 | Z Height noise | 0.027
| PSI Demo | 2025-04-04 | Y Scan Stroke | 55
| PSI Demo | 2025-04-04 | Y Detector noise (XY Servo on) | 0.246
| PSI Demo | 2025-04-04 | Y Detector noise (XY Servo off) | 0.237
| PSI Demo | 2025-04-04 | Y axis OPM | 0.44
| PSI Demo | 2025-04-04 | X Scan Stroke | 55
| PSI Demo | 2025-04-04 | X Detector noise (XY Servo on) | 0.235
| PSI Demo | 2025-04-04 | X Detector noise (XY Servo off) | 0.243
| PSI Demo | 2025-04-04 | X axis OPM | 0.577
| PSI Demo | 2025-04-04 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 2.843
| PSI Demo | 2025-04-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| PSI Demo | 2025-04-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| PSI Demo | 2025-04-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| PSI Demo | 2025-04-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| PSI Demo | 2025-04-04 | NCM driving% with AC160TS cantilever | 0.444
| PSI Demo | 2025-04-04 | NCM Amplitude noise | 17.455
| PSI Demo | 2025-04-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.594
| PSI Demo | 2025-04-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.725
| PSI Demo | 2025-04-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.8
| PSI Demo | 2025-04-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.6
| PSI Demo | 2025-04-04 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0
| PSI Demo | 2025-04-04 | Lateral(C-D) sensitivity | 11.5
| PSI Demo | 2025-04-04 | Lateral(C-D) noise | 0.1
| PSI Demo | 2025-04-04 | Full range of Z Height during Z Scan sweep | 17
| PSI Demo | 2025-04-04 | Full range of Z Detector during Z Scan sweep | 13.2
| PSI Demo | 2025-04-04 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| PSI Demo | 2025-04-04 | Full range of Y Detector during Y Scan sweep | 75.7
| PSI Demo | 2025-04-04 | Full range of X Detector Raw during X Scan sweep | 14.4
| PSI Demo | 2025-04-04 | Full range of X Detector during X Scan sweep | 75.3
| PSI Demo | 2025-04-04 | A-B noise | 8.463
| Zhejiang ICTruth | 2025-04-04 | Z Linearity by Z scanner position | 1.0186
| Zhejiang ICTruth | 2025-04-04 | Z Linearity by Z scanner position | 0.4
| Zhejiang ICTruth | 2025-04-04 | Z Linearity by 4 different sample height | 0.9996
| Zhejiang ICTruth | 2025-04-04 | Z Linearity by 4 different sample height | 1.0004
| Zhejiang ICTruth | 2025-04-04 | Z Height noise | 0.026
| Zhejiang ICTruth | 2025-04-04 | Z Height noise | 0.061
| Zhejiang ICTruth | 2025-04-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.1612
| Zhejiang ICTruth | 2025-04-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 64.513
| Zhejiang ICTruth | 2025-04-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.6877
| Zhejiang ICTruth | 2025-04-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 64.441
| Zhejiang ICTruth | 2025-04-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.9906
| Zhejiang ICTruth | 2025-04-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.703
| Zhejiang ICTruth | 2025-04-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.1701
| Zhejiang ICTruth | 2025-04-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.5752
| Zhejiang ICTruth | 2025-04-04 | Z Drift | 12.7377
| Zhejiang ICTruth | 2025-04-04 | Y2 Detector noise (XY Servo on) | 0.316
| Zhejiang ICTruth | 2025-04-04 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 260.8257
| Zhejiang ICTruth | 2025-04-04 | Y1 Detector noise (XY Servo on) | 0.343
| Zhejiang ICTruth | 2025-04-04 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 240.9509
| Zhejiang ICTruth | 2025-04-04 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 180
| Zhejiang ICTruth | 2025-04-04 | Y Ringing(Rpv) | 0.129
| Zhejiang ICTruth | 2025-04-04 | Y Non-linearity error | 3.76
| Zhejiang ICTruth | 2025-04-04 | Y Detector noise (XY Servo on) | 0.234
| Zhejiang ICTruth | 2025-04-04 | Y axis, Repeatability @1st position | 0.03
| Zhejiang ICTruth | 2025-04-04 | Y axis, Repeatability @ 4th position | 0.02
| Zhejiang ICTruth | 2025-04-04 | Y axis, Repeatability @ 3rd position | 0.03
| Zhejiang ICTruth | 2025-04-04 | Y axis, Repeatability @ 2nd position | 0.03
| Zhejiang ICTruth | 2025-04-04 | Y axis, Accuracy @1st position | 0.16
| Zhejiang ICTruth | 2025-04-04 | Y axis, Accuracy @ 4th position | 0.12
| Zhejiang ICTruth | 2025-04-04 | Y axis, Accuracy @ 3rd position | 0.16
| Zhejiang ICTruth | 2025-04-04 | Y axis, Accuracy @ 2nd position | 0.14
| Zhejiang ICTruth | 2025-04-04 | Y axis OPM | 0.957
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.5527
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning Y offset (LLC Translation) | 2.6855
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning Y offset (In-Die Align) | 3.1689
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning Y offset (Global Align) | 3.6963
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning Y 1σ (LLC Translation) | 0.57905
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning X offset (Vision Pattern Matching) | 2.6465
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning X offset (LLC Translation) | 2.9004
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning X offset (In-Die Align) | 1.7236
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning X offset (Global Align) | 2.5928
| Zhejiang ICTruth | 2025-04-04 | XY Stage positioning X 1σ (LLC Translation) | 0.60296
| Zhejiang ICTruth | 2025-04-04 | X2 Detector noise (XY Servo on) | 0.36
| Zhejiang ICTruth | 2025-04-04 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 263.0974
| Zhejiang ICTruth | 2025-04-04 | X1 Detector noise (XY Servo on) | 0.549
| Zhejiang ICTruth | 2025-04-04 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 258.1111
| Zhejiang ICTruth | 2025-04-04 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 166.635
| Zhejiang ICTruth | 2025-04-04 | X Ringing(Rpv) | 0.054
| Zhejiang ICTruth | 2025-04-04 | X Non-linearity error | 3.79
| Zhejiang ICTruth | 2025-04-04 | X Detector noise (XY Servo on) | 0.335
| Zhejiang ICTruth | 2025-04-04 | X axis, Repeatability @1st position | 0.04
| Zhejiang ICTruth | 2025-04-04 | X axis, Repeatability @ 4th position | 0.05
| Zhejiang ICTruth | 2025-04-04 | X axis, Repeatability @ 3rd position | 0.07
| Zhejiang ICTruth | 2025-04-04 | X axis, Repeatability @ 2nd position | 0.03
| Zhejiang ICTruth | 2025-04-04 | X axis, Accuracy @1st position | 0.23
| Zhejiang ICTruth | 2025-04-04 | X axis, Accuracy @ 4th position | 0.27
| Zhejiang ICTruth | 2025-04-04 | X axis, Accuracy @ 3rd position | 0.37
| Zhejiang ICTruth | 2025-04-04 | X axis, Accuracy @ 2nd position | 0.15
| Zhejiang ICTruth | 2025-04-04 | X axis OPM | 0.485
| Zhejiang ICTruth | 2025-04-04 | Wafer placement(Loading) error on the load port2 (Y offset) | 45.78
| Zhejiang ICTruth | 2025-04-04 | Wafer placement(Loading) error on the load port2 (X offset) | 55.1
| Zhejiang ICTruth | 2025-04-04 | Wafer placement(Loading) error on the load port1 (Y offset) | 52.88
| Zhejiang ICTruth | 2025-04-04 | Wafer placement(Loading) error on the load port1 (X offset) | 25.9
| Zhejiang ICTruth | 2025-04-04 | Torque Profile (50mm, 20um/s) | 0.69717
| Zhejiang ICTruth | 2025-04-04 | Torque Profile (50mm, 200um/s) | 0.4572
| Zhejiang ICTruth | 2025-04-04 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.50098
| Zhejiang ICTruth | 2025-04-04 | Roughness Repeatability 80x80um(RSD%) | 1.4677
| Zhejiang ICTruth | 2025-04-04 | Roughness Repeatability 1x1um(RSD%) | 3.1368
| Zhejiang ICTruth | 2025-04-04 | Roll motion to X+ full scan | 9.2
| Zhejiang ICTruth | 2025-04-04 | RMS of ZeroScan | 0.038
| Zhejiang ICTruth | 2025-04-04 | RMS error after Stage Calibration | 2.99
| Zhejiang ICTruth | 2025-04-04 | Profiler noise @200um | 0.102
| Zhejiang ICTruth | 2025-04-04 | Position repeatability(10repeats, Position Based) | 0.48
| Zhejiang ICTruth | 2025-04-04 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7.89e-05
| Zhejiang ICTruth | 2025-04-04 | OpticalAxisShiftPerFocusY | -0.0031
| Zhejiang ICTruth | 2025-04-04 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 9.66e-05
| Zhejiang ICTruth | 2025-04-04 | OpticalAxisShiftPerFocusX | -0.0014
| Zhejiang ICTruth | 2025-04-04 | OPM Repeatability @5mm | 0.755
| Zhejiang ICTruth | 2025-04-04 | OPM Repeatability @25mm | 2.607
| Zhejiang ICTruth | 2025-04-04 | OPM Repeatability @1mm | 0.739
| Zhejiang ICTruth | 2025-04-04 | OPM Repeatability @10mm | 1.109
| Zhejiang ICTruth | 2025-04-04 | Min. Vibration Transmissbility (~100Hz) | 0.023
| Zhejiang ICTruth | 2025-04-04 | Max OPM @5mm | 24.717
| Zhejiang ICTruth | 2025-04-04 | Max OPM @25mm | 81.027
| Zhejiang ICTruth | 2025-04-04 | Max OPM @1mm | 7.804
| Zhejiang ICTruth | 2025-04-04 | Max OPM @10mm | 25.549
| Zhejiang ICTruth | 2025-04-04 | Lift pins synchronization offset check (Y offset) | 33.4
| Zhejiang ICTruth | 2025-04-04 | Lift pins synchronization offset check (X offset) | 18.25
| Zhejiang ICTruth | 2025-04-04 | Lateral(C-D) sensitivity | 37.1
| Zhejiang ICTruth | 2025-04-04 | Lateral(C-D) noise | 0.24798
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from -Y to +Y (Y scan) | 75.4394
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from -Y to +Y (X scan) | 9.5215
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from -X to +X (Y scan) | 54.9316
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from -X to +X (X scan) | 31.4941
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from +Y to -Y (Y scan) | 29.2969
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from +Y to -Y (X scan) | 8.7891
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from +X to -X (Y scan) | 7.3242
| Zhejiang ICTruth | 2025-04-04 | Image offset after moving from +X to -X (X scan) | 36.6211
| Zhejiang ICTruth | 2025-04-04 | Full range of Z Height during Z Scan sweep | 16.4
| Zhejiang ICTruth | 2025-04-04 | Full range of Z Detector during Z Scan sweep | 12.08
| Zhejiang ICTruth | 2025-04-04 | Full range of Y2 Detector Raw during Y Scan sweep | 14.63
| Zhejiang ICTruth | 2025-04-04 | Full range of Y2 Detector during Y Scan sweep | 192.7
| Zhejiang ICTruth | 2025-04-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.63
| Zhejiang ICTruth | 2025-04-04 | Full range of Y1 Detector during Y Scan sweep | 186.8
| Zhejiang ICTruth | 2025-04-04 | Full range of Y Detector during Y Scan sweep | 186.8
| Zhejiang ICTruth | 2025-04-04 | Full range of X2 Detector Raw during X Scan sweep | 14.41
| Zhejiang ICTruth | 2025-04-04 | Full range of X2 Detector during X Scan sweep | 191.8
| Zhejiang ICTruth | 2025-04-04 | Full range of X1 Detector Raw during X Scan sweep | 14.59
| Zhejiang ICTruth | 2025-04-04 | Full range of X1 Detector during X Scan sweep | 196.1
| Zhejiang ICTruth | 2025-04-04 | Full range of X Detector during X Scan sweep | 191.8
| Zhejiang ICTruth | 2025-04-04 | Full range in voltage during Z Scan sweep | 13.59
| Zhejiang ICTruth | 2025-04-04 | Full range in um during Z Scan sweep | 43.8
| Zhejiang ICTruth | 2025-04-04 | Cycle time for 15-point measurement | 2578.751
| Zhejiang ICTruth | 2025-04-04 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Zhejiang ICTruth | 2025-04-04 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.72
| Zhejiang ICTruth | 2025-04-04 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| Zhejiang ICTruth | 2025-04-04 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Zhejiang ICTruth | 2025-04-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Zhejiang ICTruth | 2025-04-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.73
| Zhejiang ICTruth | 2025-04-04 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.49
| Zhejiang ICTruth | 2025-04-04 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Z Scanner position at -3μm | 0.006312
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Z Scanner position at +3μm | 0.010485
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Z linearity by sample height | 1.0038
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Z Height noise | 0.032
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Z Height calibration error | 0.20098
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Z Drive calibration error | 0.14537
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Y Scan Stroke | 55
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Y Non-linearity | 0.033
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Y Detector stroke calibration error (nominal scale error Y) | 0.322
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Y Detector noise (XY Servo on) | 0.353
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Y Detector noise (XY Servo off) | 0.373
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Y axis OPM | 1.0012
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | X Scan Stroke | 55
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | X Non-linearity | 0.032
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | X Detector stroke calibration error (nominal scale error X) | 0.059
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | X Detector noise (XY Servo on) | 0.348
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | X Detector noise (XY Servo off) | 0.376
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | X axis OPM | 1.7042
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.16
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | NCM driving% with AC160TS cantilever | 0.743
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | NCM Amplitude noise | 10.552
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.418
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.053
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 93.3
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 91.9
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Full range of Z Height during Z Scan sweep | 16.9
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Full range of Z Detector during Z Scan sweep | 12.1
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Full range of Y Detector Raw during Y Scan sweep | 15
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Full range of Y Detector during Y Scan sweep | 69.8
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Full range of X Detector Raw during X Scan sweep | 13.4
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | Full range of X Detector during X Scan sweep | 61.6
| Zhejiang University_School of Mechanical Engineering | 2025-04-02 | A-B noise | 6.989
| University of Texas at Austin | 2025-04-01 | Z Scanner position at -3μm | 0.000994
| Zhejiang Sci-tech University | 2025-04-01 | Z Scanner position at -3μm | 0.00686
| App Team Demo | 2025-04-01 | Z Scanner position at -3μm | 0.009495
| University of Texas at Austin | 2025-04-01 | Z Scanner position at +3μm | 0.004087
| Zhejiang Sci-tech University | 2025-04-01 | Z Scanner position at +3μm | 0.009438
| App Team Demo | 2025-04-01 | Z Scanner position at +3μm | 0.002249
| University of Texas at Austin | 2025-04-01 | Z linearity by sample height | 1.0094
| Zhejiang Sci-tech University | 2025-04-01 | Z linearity by sample height | 0.98545
| App Team Demo | 2025-04-01 | Z linearity by sample height | 1.0125
| University of Texas at Austin | 2025-04-01 | Z Height noise | 0.028
| Zhejiang Sci-tech University | 2025-04-01 | Z Height noise | 0.027
| App Team Demo | 2025-04-01 | Z Height noise | 0.028
| University of Texas at Austin | 2025-04-01 | Z Height calibration error | 0.25479
| Zhejiang Sci-tech University | 2025-04-01 | Z Height calibration error | 0.15038
| App Team Demo | 2025-04-01 | Z Height calibration error | 0.28161
| University of Texas at Austin | 2025-04-01 | Z Drive calibration error | 0.33333
| Zhejiang Sci-tech University | 2025-04-01 | Z Drive calibration error | 0.13027
| App Team Demo | 2025-04-01 | Z Drive calibration error | 0.054598
| University of Texas at Austin | 2025-04-01 | Y Scan Stroke | 180
| Zhejiang Sci-tech University | 2025-04-01 | Y Scan Stroke | 116.1
| App Team Demo | 2025-04-01 | Y Scan Stroke | 185.4
| University of Texas at Austin | 2025-04-01 | Y Non-linearity | 0.171
| App Team Demo | 2025-04-01 | Y Non-linearity | 0.268
| Zhejiang Sci-tech University | 2025-04-01 | Y Non-lineairty | 0.215
| University of Texas at Austin | 2025-04-01 | Y Detector stroke calibration error (nominal scale error Y) | 0.029
| Zhejiang Sci-tech University | 2025-04-01 | Y Detector stroke calibration error (nominal scale error Y) | 0.366
| App Team Demo | 2025-04-01 | Y Detector stroke calibration error (nominal scale error Y) | 0.377
| University of Texas at Austin | 2025-04-01 | Y Detector noise (XY Servo on) | 0.405
| Zhejiang Sci-tech University | 2025-04-01 | Y Detector noise (XY Servo on) | 0.262
| App Team Demo | 2025-04-01 | Y Detector noise (XY Servo on) | 0.309
| University of Texas at Austin | 2025-04-01 | Y Detector noise (XY servo off) | 0.396
| Zhejiang Sci-tech University | 2025-04-01 | Y Detector noise (XY servo off) | 0.258
| App Team Demo | 2025-04-01 | Y Detector noise (XY servo off) | 0.316
| University of Texas at Austin | 2025-04-01 | Y axis OPM | 0.8356
| Zhejiang Sci-tech University | 2025-04-01 | Y axis OPM | 1.322
| App Team Demo | 2025-04-01 | Y axis OPM | 1.16
| University of Texas at Austin | 2025-04-01 | X Scan Stroke | 185.2
| Zhejiang Sci-tech University | 2025-04-01 | X Scan Stroke | 107.8
| App Team Demo | 2025-04-01 | X Scan Stroke | 188.8
| University of Texas at Austin | 2025-04-01 | X Non-linearity | 0.111
| App Team Demo | 2025-04-01 | X Non-linearity | 0.064
| Zhejiang Sci-tech University | 2025-04-01 | X Non-lineairty | 0.231
| University of Texas at Austin | 2025-04-01 | X Detector stroke calibration error (nominal scale error X) | 0.379
| Zhejiang Sci-tech University | 2025-04-01 | X Detector stroke calibration error (nominal scale error X) | 0.181
| App Team Demo | 2025-04-01 | X Detector stroke calibration error (nominal scale error X) | 0.338
| University of Texas at Austin | 2025-04-01 | X Detector noise (XY Servo on) | 0.394
| Zhejiang Sci-tech University | 2025-04-01 | X Detector noise (XY Servo on) | 0.253
| App Team Demo | 2025-04-01 | X Detector noise (XY Servo on) | 0.322
| University of Texas at Austin | 2025-04-01 | X Detector noise (XY servo off) | 0.384
| Zhejiang Sci-tech University | 2025-04-01 | X Detector noise (XY servo off) | 0.249
| App Team Demo | 2025-04-01 | X Detector noise (XY servo off) | 0.369
| University of Texas at Austin | 2025-04-01 | X axis OPM | 1.1424
| Zhejiang Sci-tech University | 2025-04-01 | X axis OPM | 1.7012
| App Team Demo | 2025-04-01 | X axis OPM | 1.1974
| University of Texas at Austin | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 0.9
| Zhejiang Sci-tech University | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 2.3
| App Team Demo | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 3.298
| University of Texas at Austin | 2025-04-01 | Roughness Repeatability 1x1um(1σ) | 0.001578
| Zhejiang Sci-tech University | 2025-04-01 | Roughness Repeatability 1x1um(1σ) | 0.004
| App Team Demo | 2025-04-01 | Roughness Repeatability 1x1um(1σ) | 0.019
| University of Texas at Austin | 2025-04-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| Zhejiang Sci-tech University | 2025-04-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.38
| App Team Demo | 2025-04-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| University of Texas at Austin | 2025-04-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.83
| Zhejiang Sci-tech University | 2025-04-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.61
| App Team Demo | 2025-04-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.57
| University of Texas at Austin | 2025-04-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Zhejiang Sci-tech University | 2025-04-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| App Team Demo | 2025-04-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| University of Texas at Austin | 2025-04-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.84
| Zhejiang Sci-tech University | 2025-04-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.6
| App Team Demo | 2025-04-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.54
| University of Texas at Austin | 2025-04-01 | NCM driving% with AC160TS cantilever | 0.899
| Zhejiang Sci-tech University | 2025-04-01 | NCM driving% with AC160TS cantilever | 0.47
| App Team Demo | 2025-04-01 | NCM driving% with AC160TS cantilever | 1.52
| University of Texas at Austin | 2025-04-01 | NCM Amplitude noise | 9.349
| Zhejiang Sci-tech University | 2025-04-01 | NCM Amplitude noise | 13.78
| App Team Demo | 2025-04-01 | NCM Amplitude noise | 9.053
| University of Texas at Austin | 2025-04-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.775
| Zhejiang Sci-tech University | 2025-04-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.551
| App Team Demo | 2025-04-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.115
| University of Texas at Austin | 2025-04-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.045
| Zhejiang Sci-tech University | 2025-04-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.334
| App Team Demo | 2025-04-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.051
| Zhejiang Sci-tech University | 2025-04-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.6
| App Team Demo | 2025-04-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 256
| University of Texas at Austin | 2025-04-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 240
| Zhejiang Sci-tech University | 2025-04-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157
| App Team Demo | 2025-04-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 259.8
| University of Texas at Austin | 2025-04-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 234
| University of Texas at Austin | 2025-04-01 | Lateral(C-D) sensitivity | 19.8
| Zhejiang Sci-tech University | 2025-04-01 | Lateral(C-D) sensitivity | 17
| App Team Demo | 2025-04-01 | Lateral(C-D) sensitivity | 18.1
| University of Texas at Austin | 2025-04-01 | Lateral(C-D) noise | 0.095
| Zhejiang Sci-tech University | 2025-04-01 | Lateral(C-D) noise | 0.111
| App Team Demo | 2025-04-01 | Lateral(C-D) noise | 0.109
| University of Texas at Austin | 2025-04-01 | Full range of Z Height during Z Scan sweep | 17.3
| Zhejiang Sci-tech University | 2025-04-01 | Full range of Z Height during Z Scan sweep | 17
| App Team Demo | 2025-04-01 | Full range of Z Height during Z Scan sweep | 18.2
| University of Texas at Austin | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 12.5
| Zhejiang Sci-tech University | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 12.5
| App Team Demo | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 13
| University of Texas at Austin | 2025-04-01 | Full range of Y Detector Raw during Y Scan sweep | 15
| Zhejiang Sci-tech University | 2025-04-01 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| App Team Demo | 2025-04-01 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| University of Texas at Austin | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 180
| Zhejiang Sci-tech University | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 116.1
| App Team Demo | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 185.4
| University of Texas at Austin | 2025-04-01 | Full range of X Detector Raw during X Scan sweep | 15.8
| Zhejiang Sci-tech University | 2025-04-01 | Full range of X Detector Raw during X Scan sweep | 13.7
| App Team Demo | 2025-04-01 | Full range of X Detector Raw during X Scan sweep | 14.5
| University of Texas at Austin | 2025-04-01 | Full range of X Detector during X Scan sweep | 185.1
| Zhejiang Sci-tech University | 2025-04-01 | Full range of X Detector during X Scan sweep | 107.9
| App Team Demo | 2025-04-01 | Full range of X Detector during X Scan sweep | 188.8
| University of Texas at Austin | 2025-04-01 | A-B noise | 13.422
| Zhejiang Sci-tech University | 2025-04-01 | A-B noise | 6.08
| App Team Demo | 2025-04-01 | A-B noise | 5.76
| SZMTC#4 | 2025-04-01 | Z Linearity by Z scanner position | 1.75
| SZMTC#4 | 2025-04-01 | Z Linearity by Z scanner position | 1.77
| SZMTC#5 | 2025-04-01 | Z Linearity by Z scanner position | 0.26
| SZMTC#5 | 2025-04-01 | Z Linearity by Z scanner position | 0.43
| SZMTC#4 | 2025-04-01 | Z Linearity by 3 different sample height | 1.001
| SZMTC#4 | 2025-04-01 | Z Linearity by 3 different sample height | 1.0015
| SZMTC#5 | 2025-04-01 | Z Linearity by 3 different sample height | 0.99516
| SZMTC#5 | 2025-04-01 | Z Linearity by 3 different sample height | 0.99806
| SZMTC#4 | 2025-04-01 | Z Height noise | 0.025
| SZMTC#4 | 2025-04-01 | Z Height noise | 0.025
| SZMTC#5 | 2025-04-01 | Z Height noise | 0.027
| SZMTC#5 | 2025-04-01 | Z Height noise | 0.027
| SZMTC#4 | 2025-04-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.9354
| SZMTC#4 | 2025-04-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.93
| SZMTC#5 | 2025-04-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.9354
| SZMTC#5 | 2025-04-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.9354
| SZMTC#4 | 2025-04-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.7695
| SZMTC#4 | 2025-04-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.76
| SZMTC#5 | 2025-04-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.8
| SZMTC#5 | 2025-04-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.7695
| SZMTC#4 | 2025-04-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.359
| SZMTC#4 | 2025-04-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.42
| SZMTC#5 | 2025-04-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4776
| SZMTC#5 | 2025-04-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4315
| SZMTC#4 | 2025-04-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2952
| SZMTC#4 | 2025-04-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.31
| SZMTC#5 | 2025-04-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3888
| SZMTC#5 | 2025-04-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.381
| SZMTC#4 | 2025-04-01 | Z Drift | 17.935
| SZMTC#4 | 2025-04-01 | Z Drift | 14.206
| SZMTC#5 | 2025-04-01 | Z Drift | 10.677
| SZMTC#5 | 2025-04-01 | Z Drift | 9.381
| SZMTC#4 | 2025-04-01 | Y2 Detector noise (XY Servo on) | 0.286
| SZMTC#4 | 2025-04-01 | Y2 Detector noise (XY Servo on) | 0.379
| SZMTC#5 | 2025-04-01 | Y2 Detector noise (XY Servo on) | 0.419
| SZMTC#5 | 2025-04-01 | Y2 Detector noise (XY Servo on) | 0.228
| SZMTC#4 | 2025-04-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 237.9019
| SZMTC#4 | 2025-04-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 242.1636
| SZMTC#5 | 2025-04-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 249.0872
| SZMTC#5 | 2025-04-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 245.5511
| SZMTC#4 | 2025-04-01 | Y1 Detector noise (XY Servo on) | 0.393
| SZMTC#4 | 2025-04-01 | Y1 Detector noise (XY Servo on) | 0.371
| SZMTC#5 | 2025-04-01 | Y1 Detector noise (XY Servo on) | 0.432
| SZMTC#5 | 2025-04-01 | Y1 Detector noise (XY Servo on) | 0.22
| SZMTC#4 | 2025-04-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 233.4439
| SZMTC#4 | 2025-04-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 240.0539
| SZMTC#5 | 2025-04-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 244.3068
| SZMTC#5 | 2025-04-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 240.9365
| SZMTC#4 | 2025-04-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 145.238
| SZMTC#4 | 2025-04-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 157.548
| SZMTC#5 | 2025-04-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 154.072
| SZMTC#5 | 2025-04-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 153.667
| SZMTC#4 | 2025-04-01 | Y Ringing(Rpv) | 0.084
| SZMTC#4 | 2025-04-01 | Y Ringing(Rpv) | 0.109
| SZMTC#5 | 2025-04-01 | Y Ringing(Rpv) | 0.22
| SZMTC#5 | 2025-04-01 | Y Ringing(Rpv) | 0.108
| SZMTC#4 | 2025-04-01 | Y Detector noise (XY Servo on) | 0.244
| SZMTC#4 | 2025-04-01 | Y Detector noise (XY Servo on) | 0.266
| SZMTC#5 | 2025-04-01 | Y Detector noise (XY Servo on) | 0.304
| SZMTC#5 | 2025-04-01 | Y Detector noise (XY Servo on) | 0.163
| SZMTC#4 | 2025-04-01 | Y axis OPM | 0.538
| SZMTC#4 | 2025-04-01 | Y axis OPM | 1.427
| SZMTC#5 | 2025-04-01 | Y axis OPM | 0.14
| SZMTC#5 | 2025-04-01 | Y axis OPM | 0.508
| SZMTC#4 | 2025-04-01 | XY Orthogonality | 90.1
| SZMTC#4 | 2025-04-01 | XY Orthogonality | 90.178
| SZMTC#5 | 2025-04-01 | XY Orthogonality | 90.287
| SZMTC#5 | 2025-04-01 | XY Orthogonality | 89.951
| SZMTC#4 | 2025-04-01 | X2 Detector noise (XY Servo on) | 0.413
| SZMTC#4 | 2025-04-01 | X2 Detector noise (XY Servo on) | 0.315
| SZMTC#5 | 2025-04-01 | X2 Detector noise (XY Servo on) | 0.435
| SZMTC#5 | 2025-04-01 | X2 Detector noise (XY Servo on) | 0.232
| SZMTC#4 | 2025-04-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 234.1829
| SZMTC#4 | 2025-04-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 240.8487
| SZMTC#5 | 2025-04-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 242.3083
| SZMTC#5 | 2025-04-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 248.3824
| SZMTC#4 | 2025-04-01 | X1 Detector noise (XY Servo on) | 0.401
| SZMTC#4 | 2025-04-01 | X1 Detector noise (XY Servo on) | 0.425
| SZMTC#5 | 2025-04-01 | X1 Detector noise (XY Servo on) | 0.3
| SZMTC#5 | 2025-04-01 | X1 Detector noise (XY Servo on) | 0.233
| SZMTC#4 | 2025-04-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 224.9836
| SZMTC#4 | 2025-04-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 251.7749
| SZMTC#5 | 2025-04-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 246.8302
| SZMTC#5 | 2025-04-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 250.3534
| SZMTC#4 | 2025-04-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 147.32
| SZMTC#4 | 2025-04-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 155.956
| SZMTC#5 | 2025-04-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 151.987
| SZMTC#5 | 2025-04-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 152.567
| SZMTC#4 | 2025-04-01 | X Ringing(Rpv) | 0.093
| SZMTC#4 | 2025-04-01 | X Ringing(Rpv) | 0.095
| SZMTC#5 | 2025-04-01 | X Ringing(Rpv) | 0.155
| SZMTC#5 | 2025-04-01 | X Ringing(Rpv) | 0.125
| SZMTC#4 | 2025-04-01 | X Detector noise (XY Servo on) | 0.296
| SZMTC#4 | 2025-04-01 | X Detector noise (XY Servo on) | 0.325
| SZMTC#5 | 2025-04-01 | X Detector noise (XY Servo on) | 0.271
| SZMTC#5 | 2025-04-01 | X Detector noise (XY Servo on) | 0.171
| SZMTC#4 | 2025-04-01 | X axis OPM | 0.527
| SZMTC#4 | 2025-04-01 | X axis OPM | 0.564
| SZMTC#5 | 2025-04-01 | X axis OPM | 0.69
| SZMTC#5 | 2025-04-01 | X axis OPM | 1.141
| SZMTC#4 | 2025-04-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.838
| SZMTC#4 | 2025-04-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.798
| SZMTC#5 | 2025-04-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.216
| SZMTC#5 | 2025-04-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.795
| SZMTC#4 | 2025-04-01 | Roughness Repeatability 80x80um(RSD%) | 1.79
| SZMTC#4 | 2025-04-01 | Roughness Repeatability 80x80um(RSD%) | 1.37
| SZMTC#5 | 2025-04-01 | Roughness Repeatability 80x80um(RSD%) | 2.31
| SZMTC#5 | 2025-04-01 | Roughness Repeatability 80x80um(RSD%) | 0.45
| SZMTC#4 | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 2.292
| SZMTC#4 | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 1.794
| SZMTC#5 | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 1.425
| SZMTC#5 | 2025-04-01 | Roughness Repeatability 1x1um(RSD%) | 2.241
| SZMTC#4 | 2025-04-01 | RMS of ZeroScan | 0.03
| SZMTC#4 | 2025-04-01 | RMS of ZeroScan | 0.031
| SZMTC#5 | 2025-04-01 | RMS of ZeroScan | 0.027
| SZMTC#5 | 2025-04-01 | RMS of ZeroScan | 0.028
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00012
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.0002
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00013
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusY | -0.0027
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusY | 0.0005
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusY | 0.0018
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusY | -0.0018
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 4e-05
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00035
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusX | -0.0002
| SZMTC#4 | 2025-04-01 | OpticalAxisShiftPerFocusX | 0.0045
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusX | 0.0054
| SZMTC#5 | 2025-04-01 | OpticalAxisShiftPerFocusX | 0.0032
| SZMTC#4 | 2025-04-01 | Lateral(C-D) sensitivity | 62.6
| SZMTC#4 | 2025-04-01 | Lateral(C-D) sensitivity | 69.6
| SZMTC#5 | 2025-04-01 | Lateral(C-D) sensitivity | 27
| SZMTC#5 | 2025-04-01 | Lateral(C-D) sensitivity | 36.3
| SZMTC#4 | 2025-04-01 | Lateral(C-D) noise | 0.168
| SZMTC#4 | 2025-04-01 | Lateral(C-D) noise | 0.159
| SZMTC#5 | 2025-04-01 | Lateral(C-D) noise | 0.236
| SZMTC#5 | 2025-04-01 | Lateral(C-D) noise | 0.228
| SZMTC#4 | 2025-04-01 | Full range of Z Height during Z Scan sweep | 17.47
| SZMTC#4 | 2025-04-01 | Full range of Z Height during Z Scan sweep | 17.45
| SZMTC#5 | 2025-04-01 | Full range of Z Height during Z Scan sweep | 16.28
| SZMTC#5 | 2025-04-01 | Full range of Z Height during Z Scan sweep | 15.9
| SZMTC#4 | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 13.57
| SZMTC#4 | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 13.55
| SZMTC#5 | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 12.63
| SZMTC#5 | 2025-04-01 | Full range of Z Detector during Z Scan sweep | 12.3
| SZMTC#4 | 2025-04-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| SZMTC#4 | 2025-04-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.53
| SZMTC#5 | 2025-04-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.51
| SZMTC#5 | 2025-04-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| SZMTC#4 | 2025-04-01 | Full range of Y2 Detector during Y Scan sweep | 188.9
| SZMTC#4 | 2025-04-01 | Full range of Y2 Detector during Y Scan sweep | 188.8
| SZMTC#5 | 2025-04-01 | Full range of Y2 Detector during Y Scan sweep | 189.2
| SZMTC#5 | 2025-04-01 | Full range of Y2 Detector during Y Scan sweep | 189.3
| SZMTC#4 | 2025-04-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| SZMTC#4 | 2025-04-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.49
| SZMTC#5 | 2025-04-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.22
| SZMTC#5 | 2025-04-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.32
| SZMTC#4 | 2025-04-01 | Full range of Y1 Detector during Y Scan sweep | 179.9
| SZMTC#4 | 2025-04-01 | Full range of Y1 Detector during Y Scan sweep | 179.8
| SZMTC#5 | 2025-04-01 | Full range of Y1 Detector during Y Scan sweep | 176.7
| SZMTC#5 | 2025-04-01 | Full range of Y1 Detector during Y Scan sweep | 178
| SZMTC#4 | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 184.4
| SZMTC#4 | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 184.3
| SZMTC#5 | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 182.9
| SZMTC#5 | 2025-04-01 | Full range of Y Detector during Y Scan sweep | 183.6
| SZMTC#4 | 2025-04-01 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| SZMTC#4 | 2025-04-01 | Full range of X2 Detector Raw during X Scan sweep | 14.37
| SZMTC#5 | 2025-04-01 | Full range of X2 Detector Raw during X Scan sweep | 14.39
| SZMTC#5 | 2025-04-01 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| SZMTC#4 | 2025-04-01 | Full range of X2 Detector during X Scan sweep | 183.4
| SZMTC#4 | 2025-04-01 | Full range of X2 Detector during X Scan sweep | 183.2
| SZMTC#5 | 2025-04-01 | Full range of X2 Detector during X Scan sweep | 183.8
| SZMTC#5 | 2025-04-01 | Full range of X2 Detector during X Scan sweep | 179.8
| SZMTC#4 | 2025-04-01 | Full range of X1 Detector Raw during X Scan sweep | 15.8
| SZMTC#4 | 2025-04-01 | Full range of X1 Detector Raw during X Scan sweep | 13.97
| SZMTC#5 | 2025-04-01 | Full range of X1 Detector Raw during X Scan sweep | 14.34
| SZMTC#5 | 2025-04-01 | Full range of X1 Detector Raw during X Scan sweep | 14.24
| SZMTC#4 | 2025-04-01 | Full range of X1 Detector during X Scan sweep | 200.8
| SZMTC#4 | 2025-04-01 | Full range of X1 Detector during X Scan sweep | 177.1
| SZMTC#5 | 2025-04-01 | Full range of X1 Detector during X Scan sweep | 182.3
| SZMTC#5 | 2025-04-01 | Full range of X1 Detector during X Scan sweep | 180.5
| SZMTC#4 | 2025-04-01 | Full range of X Detector during X Scan sweep | 192.1
| SZMTC#4 | 2025-04-01 | Full range of X Detector during X Scan sweep | 180.2
| SZMTC#5 | 2025-04-01 | Full range of X Detector during X Scan sweep | 183.1
| SZMTC#5 | 2025-04-01 | Full range of X Detector during X Scan sweep | 180.1
| SZMTC#4 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| SZMTC#4 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| SZMTC#5 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| SZMTC#5 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| SZMTC#4 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.32
| SZMTC#4 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.15
| SZMTC#5 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| SZMTC#5 | 2025-04-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.15
| 가천대학교 전자공학부 | 2025-03-28 | Z Scanner position at -3μm | 0.000721
| SZMTC | 2025-03-28 | Z Scanner position at -3μm | 0.013664
| 가천대학교_전자공학부 | 2025-03-28 | Z Scanner position at -3μm | 0.000721
| 가천대학교 전자공학부 | 2025-03-28 | Z Scanner position at +3μm | 0.005295
| SZMTC | 2025-03-28 | Z Scanner position at +3μm | 0.000253
| 가천대학교_전자공학부 | 2025-03-28 | Z Scanner position at +3μm | 0.005295
| 가천대학교 전자공학부 | 2025-03-28 | Z linearity by sample height | 0.99843
| SZMTC | 2025-03-28 | Z linearity by sample height | 1.0091
| 가천대학교_전자공학부 | 2025-03-28 | Z linearity by sample height | 0.99843
| 가천대학교 전자공학부 | 2025-03-28 | Z Height noise | 0.038
| SZMTC | 2025-03-28 | Z Height noise | 0.027
| 가천대학교_전자공학부 | 2025-03-28 | Z Height noise | 0.038
| SZMTC | 2025-03-28 | Z Height calibration error | 0.033525
| 가천대학교 전자공학부 | 2025-03-28 | Z Height calibration error | 0.094828
| 가천대학교_전자공학부 | 2025-03-28 | Z Height calibration error | 0.094828
| SZMTC | 2025-03-28 | Z Drive calibration error | 0.14847
| 가천대학교 전자공학부 | 2025-03-28 | Z Drive calibration error | 0.074713
| 가천대학교_전자공학부 | 2025-03-28 | Z Drive calibration error | 0.074713
| SZMTC | 2025-03-28 | Y2 Detector noise (XY Servo on) | 0.264
| SZMTC | 2025-03-28 | Y2 Detector noise (XY Servo off) | 0.27
| SZMTC | 2025-03-28 | Y1 Detector noise (XY servo on) | 0.244
| SZMTC | 2025-03-28 | Y1 Detector noise (XY servo off) | 0.246
| 가천대학교 전자공학부 | 2025-03-28 | Y Scan Stroke | 55
| SZMTC | 2025-03-28 | Y Scan Stroke | 175
| 가천대학교_전자공학부 | 2025-03-28 | Y Scan Stroke | 55
| 가천대학교 전자공학부 | 2025-03-28 | Y Non-linearity | 0.054
| 가천대학교_전자공학부 | 2025-03-28 | Y Non-linearity | 0.054
| SZMTC | 2025-03-28 | Y Non-lineairty | 0.0833
| 가천대학교 전자공학부 | 2025-03-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.393
| SZMTC | 2025-03-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.196
| 가천대학교_전자공학부 | 2025-03-28 | Y Detector stroke calibration error (nominal scale error Y) | 0.393
| 가천대학교 전자공학부 | 2025-03-28 | Y Detector noise (XY Servo on) | 0.258
| SZMTC | 2025-03-28 | Y Detector noise (XY servo on) | 0.203
| 가천대학교_전자공학부 | 2025-03-28 | Y Detector noise (XY Servo on) | 0.258
| 가천대학교 전자공학부 | 2025-03-28 | Y Detector noise (XY Servo off) | 0.258
| SZMTC | 2025-03-28 | Y Detector noise (XY servo off) | 0.207
| 가천대학교_전자공학부 | 2025-03-28 | Y Detector noise (XY Servo off) | 0.258
| 가천대학교 전자공학부 | 2025-03-28 | Y axis OPM | 0.7132
| SZMTC | 2025-03-28 | Y axis OPM | 0.8478
| 가천대학교_전자공학부 | 2025-03-28 | Y axis OPM | 0.7132
| SZMTC | 2025-03-28 | X2 Detector noise (XY Servo on) | 0.224
| SZMTC | 2025-03-28 | X2 Detector noise (XY Servo off) | 0.231
| SZMTC | 2025-03-28 | X1 Detector noise (XY servo on) | 0.221
| SZMTC | 2025-03-28 | X1 Detector noise (XY servo off) | 0.23
| 가천대학교 전자공학부 | 2025-03-28 | X Scan Stroke | 55
| SZMTC | 2025-03-28 | X Scan Stroke | 170.5
| 가천대학교_전자공학부 | 2025-03-28 | X Scan Stroke | 55
| 가천대학교 전자공학부 | 2025-03-28 | X Non-linearity | 0.074
| 가천대학교_전자공학부 | 2025-03-28 | X Non-linearity | 0.074
| SZMTC | 2025-03-28 | X Non-lineairty | 0.281
| 가천대학교 전자공학부 | 2025-03-28 | X Detector stroke calibration error (nominal scale error X) | 0.144
| SZMTC | 2025-03-28 | X Detector stroke calibration error (nominal scale error X) | 0.083
| 가천대학교_전자공학부 | 2025-03-28 | X Detector stroke calibration error (nominal scale error X) | 0.144
| 가천대학교 전자공학부 | 2025-03-28 | X Detector noise (XY Servo on) | 0.258
| SZMTC | 2025-03-28 | X Detector noise (XY servo on) | 0.166
| 가천대학교_전자공학부 | 2025-03-28 | X Detector noise (XY Servo on) | 0.258
| 가천대학교 전자공학부 | 2025-03-28 | X Detector noise (XY Servo off) | 0.257
| SZMTC | 2025-03-28 | X Detector noise (XY servo off) | 0.175
| 가천대학교_전자공학부 | 2025-03-28 | X Detector noise (XY Servo off) | 0.257
| 가천대학교 전자공학부 | 2025-03-28 | X axis OPM | 0.9478
| SZMTC | 2025-03-28 | X axis OPM | 1.257
| 가천대학교_전자공학부 | 2025-03-28 | X axis OPM | 0.9478
| 가천대학교 전자공학부 | 2025-03-28 | Roughness Repeatability 1x1um(RSD%) | 2.063
| SZMTC | 2025-03-28 | Roughness Repeatability 1x1um(RSD%) | 0.939
| 가천대학교_전자공학부 | 2025-03-28 | Roughness Repeatability 1x1um(RSD%) | 2.063
| 가천대학교 전자공학부 | 2025-03-28 | Roughness Repeatability 1x1um(1σ) | 0.007
| SZMTC | 2025-03-28 | Roughness Repeatability 1x1um(1σ) | 0.001
| 가천대학교_전자공학부 | 2025-03-28 | Roughness Repeatability 1x1um(1σ) | 0.007
| 가천대학교 전자공학부 | 2025-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| SZMTC | 2025-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| 가천대학교_전자공학부 | 2025-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 가천대학교 전자공학부 | 2025-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| SZMTC | 2025-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| 가천대학교_전자공학부 | 2025-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| SZMTC | 2025-03-28 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.33
| SZMTC | 2025-03-28 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.34
| SZMTC | 2025-03-28 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.36
| SZMTC | 2025-03-28 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.34
| SZMTC | 2025-03-28 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.38
| 가천대학교 전자공학부 | 2025-03-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| SZMTC | 2025-03-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 가천대학교_전자공학부 | 2025-03-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| 가천대학교 전자공학부 | 2025-03-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.16
| SZMTC | 2025-03-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| 가천대학교_전자공학부 | 2025-03-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.16
| 가천대학교 전자공학부 | 2025-03-28 | NCM driving% with AC160TS cantilever | 0.683
| SZMTC | 2025-03-28 | NCM driving% with AC160TS cantilever | 1.061
| 가천대학교_전자공학부 | 2025-03-28 | NCM driving% with AC160TS cantilever | 0.683
| 가천대학교 전자공학부 | 2025-03-28 | NCM Amplitude noise | 11.165
| SZMTC | 2025-03-28 | NCM Amplitude noise | 10.208
| 가천대학교_전자공학부 | 2025-03-28 | NCM Amplitude noise | 11.165
| 가천대학교 전자공학부 | 2025-03-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.615
| SZMTC | 2025-03-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.593
| 가천대학교_전자공학부 | 2025-03-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.615
| 가천대학교 전자공학부 | 2025-03-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.681
| SZMTC | 2025-03-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.808
| 가천대학교_전자공학부 | 2025-03-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.681
| SZMTC | 2025-03-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 239.4
| SZMTC | 2025-03-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 246.8
| 가천대학교 전자공학부 | 2025-03-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 111.4
| 가천대학교_전자공학부 | 2025-03-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 111.4
| SZMTC | 2025-03-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 233.4
| SZMTC | 2025-03-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 227.3
| 가천대학교 전자공학부 | 2025-03-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 109.1
| 가천대학교_전자공학부 | 2025-03-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 109.1
| SZMTC | 2025-03-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| 가천대학교 전자공학부 | 2025-03-28 | Lateral(C-D) sensitivity | 19.8
| SZMTC | 2025-03-28 | Lateral(C-D) sensitivity | 11.1
| 가천대학교_전자공학부 | 2025-03-28 | Lateral(C-D) sensitivity | 19.8
| 가천대학교 전자공학부 | 2025-03-28 | Lateral(C-D) noise | 0.084
| SZMTC | 2025-03-28 | Lateral(C-D) noise | 0.159
| 가천대학교_전자공학부 | 2025-03-28 | Lateral(C-D) noise | 0.084
| 가천대학교 전자공학부 | 2025-03-28 | Full range of Z Height during Z Scan sweep | 17.3
| SZMTC | 2025-03-28 | Full range of Z Height during Z Scan sweep | 17.3
| 가천대학교_전자공학부 | 2025-03-28 | Full range of Z Height during Z Scan sweep | 17.3
| 가천대학교 전자공학부 | 2025-03-28 | Full range of Z Detector during Z Scan sweep | 13
| SZMTC | 2025-03-28 | Full range of Z Detector during Z Scan sweep | 13
| 가천대학교_전자공학부 | 2025-03-28 | Full range of Z Detector during Z Scan sweep | 13
| SZMTC | 2025-03-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| SZMTC | 2025-03-28 | Full range of Y2 Detector during Y Scan sweep | 173.75
| SZMTC | 2025-03-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| SZMTC | 2025-03-28 | Full range of Y1 Detector during Y Scan sweep | 176.85
| 가천대학교 전자공학부 | 2025-03-28 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| 가천대학교_전자공학부 | 2025-03-28 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| 가천대학교 전자공학부 | 2025-03-28 | Full range of Y Detector during Y Scan sweep | 74.1
| SZMTC | 2025-03-28 | Full range of Y Detector during Y Scan sweep | 175.3
| 가천대학교_전자공학부 | 2025-03-28 | Full range of Y Detector during Y Scan sweep | 74.1
| SZMTC | 2025-03-28 | Full range of X2 Detector Raw during X Scan sweep | 14.7
| SZMTC | 2025-03-28 | Full range of X2 Detector during X Scan sweep | 171.8
| SZMTC | 2025-03-28 | Full range of X1 Detector Raw during X Scan sweep | 15
| SZMTC | 2025-03-28 | Full range of X1 Detector during X Scan sweep | 169.65
| 가천대학교 전자공학부 | 2025-03-28 | Full range of X Detector Raw during X Scan sweep | 13.6
| 가천대학교_전자공학부 | 2025-03-28 | Full range of X Detector Raw during X Scan sweep | 13.6
| 가천대학교 전자공학부 | 2025-03-28 | Full range of X Detector during X Scan sweep | 74.2
| SZMTC | 2025-03-28 | Full range of X Detector during X Scan sweep | 170.7
| 가천대학교_전자공학부 | 2025-03-28 | Full range of X Detector during X Scan sweep | 74.2
| 가천대학교 전자공학부 | 2025-03-28 | A-B noise | 6.111
| SZMTC | 2025-03-28 | A-B noise | 6.164
| 가천대학교_전자공학부 | 2025-03-28 | A-B noise | 6.111
| Northwest University | 2025-03-27 | Z Scanner position at -3μm | 0.000731
| Northwest University | 2025-03-27 | Z Scanner position at +3μm | 0.003991
| Northwest University | 2025-03-27 | Z linearity by sample height | 0.99249
| Northwest University | 2025-03-27 | Z Height noise | 0.027
| Northwest University | 2025-03-27 | Z Height calibration error | 0.18199
| Northwest University | 2025-03-27 | Z Drive calibration error | 0.28161
| Northwest University | 2025-03-27 | Y Scan Stroke | 128.5
| Northwest University | 2025-03-27 | Y Non-lineairty | 0.048
| Northwest University | 2025-03-27 | Y Detector stroke calibration error (nominal scale error Y) | 0.38
| Northwest University | 2025-03-27 | Y Detector noise (XY Servo on) | 0.366
| Northwest University | 2025-03-27 | Y Detector noise (XY Servo off) | 0.375
| Northwest University | 2025-03-27 | Y axis OPM | 0.7002
| Northwest University | 2025-03-27 | X Scan Stroke | 126.74
| Northwest University | 2025-03-27 | X Non-lineairty | 0.044
| Northwest University | 2025-03-27 | X Detector stroke calibration error (nominal scale error X) | 0.338
| Northwest University | 2025-03-27 | X Detector noise (XY Servo on) | 0.367
| Northwest University | 2025-03-27 | X Detector noise (XY Servo off) | 0.374
| Northwest University | 2025-03-27 | X axis OPM | 1.4106
| Northwest University | 2025-03-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Northwest University | 2025-03-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Northwest University | 2025-03-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Northwest University | 2025-03-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Northwest University | 2025-03-27 | NCM driving% with AC160TS cantilever | 0.44
| Northwest University | 2025-03-27 | NCM Amplitude noise | 11.03
| Northwest University | 2025-03-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.781
| Northwest University | 2025-03-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.308
| Northwest University | 2025-03-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 166
| Northwest University | 2025-03-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167
| Northwest University | 2025-03-27 | Full range of Z Height during Z Scan sweep | 17.6
| Northwest University | 2025-03-27 | Full range of Z Detector during Z Scan sweep | 13.3
| Northwest University | 2025-03-27 | Full range of Y Detector Raw during Y Scan sweep | 15.5
| Northwest University | 2025-03-27 | Full range of Y Detector during Y Scan sweep | 128.5
| Northwest University | 2025-03-27 | Full range of X Detector Raw during X Scan sweep | 15.2
| Northwest University | 2025-03-27 | Full range of X Detector during X Scan sweep | 126.74
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Z Scanner position at -3μm | 0.011575
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Z Scanner position at +3μm | 0.007881
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Z linearity by sample height | 1.0244
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Z Height noise | 0.029
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Z Height calibration error | 0.000958
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Z Drive calibration error | 0.13793
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Y Non-lineairty | 0.228
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Y Detector stroke calibration error (nominal scale error Y) | 0.222
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Y Detector noise (XY Servo on) | 0.281
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Y Detector noise (XY servo off) | 0.26
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Y axis OPM | 0.729
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | X Non-lineairty | 0.106
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | X Detector stroke calibration error (nominal scale error X) | 0.426
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | X Detector noise (XY Servo on) | 0.287
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | X Detector noise (XY servo off) | 0.27
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | X axis OPM | 0.511
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Roughness Repeatability 1x1um(RSD%) | 1.933
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Roughness Repeatability 1x1um(1σ) | 0.004
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.28
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.3
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.27
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000) | 0.27
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000) | 0.26
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000) | 0.24
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000) | 0.24
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0) | 0.25
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(60000, 0) | 0.25
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(0, -60000) | 0.24
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(0, 60000) | 0.24
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Height of Mask Sample Zeroscan(0, 0) | 0.25
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | NCM driving% with AC160TS cantilever | 2.704
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | NCM Amplitude noise | 16.891
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.925
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.586
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 178
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 178.2
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Max Wafer Zeroscan-Min Mask Sample Zeroscan | 0.02
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Lateral(C-D) sensitivity | 15.5
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Lateral(C-D) noise | 0.116
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Full range of Z Height during Z Scan sweep | 18.1
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Full range of Z Detector during Z Scan sweep | 13.1
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Full range of Y Detector during Y Scan sweep | 131.1
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Full range of X Detector Raw during X Scan sweep | 13.4
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | Full range of X Detector during X Scan sweep | 119
| Shaoguang Core Material Technology Co., Ltd | 2025-03-26 | A-B noise | 14.283
| SwaySure Technology | 2025-03-26 | Z Linearity by Z scanner position | 0.43
| SwaySure Technology | 2025-03-26 | Full range of Z Detector during Z Scan sweep | 12.4
| SwaySure Technology | 2025-03-26 | Z Linearity by 4 different sample height | 0.99517
| SwaySure Technology | 2025-03-26 | Full range of Z Height during Z Scan sweep | 15.5598
| SwaySure Technology | 2025-03-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.4266
| SwaySure Technology | 2025-03-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2881
| SwaySure Technology | 2025-03-26 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5426
| SwaySure Technology | 2025-03-26 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2846
| SwaySure Technology | 2025-03-26 | Z Height noise | 0.027
| SwaySure Technology | 2025-03-26 | Wafer placement(Loading) error on the load port1 (X offset) | 43.55
| SwaySure Technology | 2025-03-26 | Wafer placement(Loading) error on the load port1 (Y offset) | 35.52
| SwaySure Technology | 2025-03-26 | Wafer placement(Loading) error on the load port2 (X offset) | 41.2
| SwaySure Technology | 2025-03-26 | Wafer placement(Loading) error on the load port2 (Y offset) | 10.86
| SwaySure Technology | 2025-03-26 | OpticalAxisShiftPerFocusX | -0.0039
| SwaySure Technology | 2025-03-26 | OpticalAxisShiftPerFocusY | -0.0008
| SwaySure Technology | 2025-03-26 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00024
| SwaySure Technology | 2025-03-26 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| SwaySure Technology | 2025-03-26 | Lift pins synchronization offset check (X offset) | 72.57
| SwaySure Technology | 2025-03-26 | Lift pins synchronization offset check (Y offset) | 3.83
| SwaySure Technology | 2025-03-26 | Profiler noise @200um | 0.083
| SwaySure Technology | 2025-03-26 | Roll motion to X+ full scan | 9.6
| SwaySure Technology | 2025-03-26 | Position repeatability(10repeats, Position Based) | 0.22
| SwaySure Technology | 2025-03-26 | OPM Repeatability @25mm | 2.56
| SwaySure Technology | 2025-03-26 | OPM Repeatability @10mm | 1.76
| SwaySure Technology | 2025-03-26 | OPM Repeatability @5mm | 1.955
| SwaySure Technology | 2025-03-26 | OPM Repeatability @1mm | 1.288
| SwaySure Technology | 2025-03-26 | Max OPM @25mm | 73.409
| SwaySure Technology | 2025-03-26 | Max OPM @10mm | 23.085
| SwaySure Technology | 2025-03-26 | Max OPM @5mm | 36.638
| SwaySure Technology | 2025-03-26 | Max OPM @1mm | 10.362
| SwaySure Technology | 2025-03-26 | Torque Profile (50mm, 200um/s) | 0.336
| SwaySure Technology | 2025-03-26 | Torque Profile (50mm, 20um/s) | 0.448
| SwaySure Technology | 2025-03-26 | Lateral(C-D) sensitivity | 41.7
| SwaySure Technology | 2025-03-26 | Lateral(C-D) noise | 0.199
| SwaySure Technology | 2025-03-26 | Min. Vibration Transmissbility (~100Hz) | 0.034
| SwaySure Technology | 2025-03-26 | RMS of ZeroScan | 0.143
| SwaySure Technology | 2025-03-26 | Full range of X Detector during X Scan sweep | 197.8844
| SwaySure Technology | 2025-03-26 | Full range of X1 Detector during X Scan sweep | 197.8844
| SwaySure Technology | 2025-03-26 | Full range of X2 Detector during X Scan sweep | 202.2149
| SwaySure Technology | 2025-03-26 | Full range of Y Detector during Y Scan sweep | 205.9986
| SwaySure Technology | 2025-03-26 | Full range of Y1 Detector during Y Scan sweep | 207.0513
| SwaySure Technology | 2025-03-26 | Full range of Y2 Detector during Y Scan sweep | 205.9986
| SwaySure Technology | 2025-03-26 | Full range of X1 Detector Raw during X Scan sweep | 14.32
| SwaySure Technology | 2025-03-26 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| SwaySure Technology | 2025-03-26 | Full range of Y1 Detector Raw during Y Scan sweep | 15.04
| SwaySure Technology | 2025-03-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.62
| SwaySure Technology | 2025-03-26 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.4784
| SwaySure Technology | 2025-03-26 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 272.2556
| SwaySure Technology | 2025-03-26 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 264.3223
| SwaySure Technology | 2025-03-26 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 272.9863
| SwaySure Technology | 2025-03-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 167.66
| SwaySure Technology | 2025-03-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 176.072
| SwaySure Technology | 2025-03-26 | X Non-linearity error | 3.51
| SwaySure Technology | 2025-03-26 | Y Non-linearity error | 3.31
| SwaySure Technology | 2025-03-26 | X axis OPM | 1.805
| SwaySure Technology | 2025-03-26 | Y axis OPM | 1.624
| SwaySure Technology | 2025-03-26 | X Ringing(Rpv) | 0.882
| SwaySure Technology | 2025-03-26 | Y Ringing(Rpv) | 0.86
| SwaySure Technology | 2025-03-26 | X Detector noise (XY Servo on) | 0.311
| SwaySure Technology | 2025-03-26 | Y Detector noise (XY Servo on) | 0.352
| SwaySure Technology | 2025-03-26 | X1 Detector noise (XY Servo on) | 0.464
| SwaySure Technology | 2025-03-26 | Y1 Detector noise (XY Servo on) | 0.352
| SwaySure Technology | 2025-03-26 | X2 Detector noise (XY Servo on) | 0.328
| SwaySure Technology | 2025-03-26 | Y2 Detector noise (XY Servo on) | 0.516
| SwaySure Technology | 2025-03-26 | RMS error after Stage Calibration | 4.55
| SwaySure Technology | 2025-03-26 | Image offset after moving from -X to +X (X scan) | 44.6777
| SwaySure Technology | 2025-03-26 | Image offset after moving from -Y to +Y (Y scan) | 83.4961
| SwaySure Technology | 2025-03-26 | Image offset after moving from -X to +X (Y scan) | 49.8047
| SwaySure Technology | 2025-03-26 | Image offset after moving from -Y to +Y (X scan) | 14.6484
| SwaySure Technology | 2025-03-26 | Image offset after moving from +X to -X (X scan) | 45.4102
| SwaySure Technology | 2025-03-26 | Image offset after moving from +Y to -Y (Y scan) | 31.4941
| SwaySure Technology | 2025-03-26 | Image offset after moving from +X to -X (Y scan) | 27.832
| SwaySure Technology | 2025-03-26 | Image offset after moving from +Y to -Y (X scan) | 21.9727
| SwaySure Technology | 2025-03-26 | XY Stage positioning X offset (Vision Pattern Matching) | 4.1455
| SwaySure Technology | 2025-03-26 | XY Stage positioning Y offset (Vision Pattern Matching) | 9.7705
| SwaySure Technology | 2025-03-26 | XY Stage positioning X offset (In-Die Align) | 1.5234
| SwaySure Technology | 2025-03-26 | XY Stage positioning Y offset (In-Die Align) | 5.6396
| SwaySure Technology | 2025-03-26 | XY Stage positioning X offset (LLC Translation) | 2.6367
| SwaySure Technology | 2025-03-26 | XY Stage positioning Y offset (LLC Translation) | 2.0117
| SwaySure Technology | 2025-03-26 | XY Stage positioning X 1σ (LLC Translation) | 0.5421
| SwaySure Technology | 2025-03-26 | XY Stage positioning Y 1σ (LLC Translation) | 0.4052
| SwaySure Technology | 2025-03-26 | XY Stage positioning X offset (Global Align) | 2.5391
| SwaySure Technology | 2025-03-26 | XY Stage positioning Y offset (Global Align) | 2.7441
| SwaySure Technology | 2025-03-26 | Z Drift | 5.012
| SwaySure Technology | 2025-03-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| SwaySure Technology | 2025-03-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| SwaySure Technology | 2025-03-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.61
| SwaySure Technology | 2025-03-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| SwaySure Technology | 2025-03-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| SwaySure Technology | 2025-03-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| SwaySure Technology | 2025-03-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.69
| SwaySure Technology | 2025-03-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| SwaySure Technology | 2025-03-26 | Cycle time for 15-point measurement | 693.572
| SwaySure Technology | 2025-03-26 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.264
| SwaySure Technology | 2025-03-26 | Roughness Repeatability 1x1um(10Repeats, RSD%) | 1.603
| SwaySure Technology | 2025-03-26 | Roughness Repeatability 80x80um(10Repeats, RSD%) | 0.19
| Samsung NRD-K eAFM | 2025-03-25 | Z Linearity by Z scanner position | 0.74
| Samsung NRD-K eAFM | 2025-03-25 | Z Linearity by 4 different sample height | 1.0006
| Samsung NRD-K eAFM | 2025-03-25 | Z Height noise | 0.026
| Samsung NRD-K eAFM | 2025-03-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.69
| Samsung NRD-K eAFM | 2025-03-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.69
| Samsung NRD-K eAFM | 2025-03-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.36
| Samsung NRD-K eAFM | 2025-03-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.46
| Samsung NRD-K eAFM | 2025-03-25 | Z Drift | 8.96
| Samsung NRD-K eAFM | 2025-03-25 | Y2 Detector noise (XY Servo on) | 0.365
| Samsung NRD-K eAFM | 2025-03-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 277.91
| Samsung NRD-K eAFM | 2025-03-25 | Y1 Detector noise (XY Servo on) | 0.439
| Samsung NRD-K eAFM | 2025-03-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 279.63
| Samsung NRD-K eAFM | 2025-03-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 178.26
| Samsung NRD-K eAFM | 2025-03-25 | Y Ringing(Rpv) | 0.046
| Samsung NRD-K eAFM | 2025-03-25 | Y Non-linearity error | 2.88
| Samsung NRD-K eAFM | 2025-03-25 | Y Detector noise (XY Servo on) | 0.337
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Repeatability @1st position | 0.04
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Repeatability @ 4th position | 0.04
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Repeatability @ 3rd position | 0.04
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Repeatability @ 2nd position | 0.04
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Accuracy @1st position | 0.24
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Accuracy @ 4th position | 0.22
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Accuracy @ 3rd position | 0.19
| Samsung NRD-K eAFM | 2025-03-25 | Y axis, Accuracy @ 2nd position | 0.22
| Samsung NRD-K eAFM | 2025-03-25 | Y axis OPM | 1.039
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.2061
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning Y offset (LLC Translation) | 1.59
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning Y offset (In-Die Align) | 1.3135
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning Y offset (Global Align) | 7.4658
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning Y 1σ (LLC Translation) | 0.3258
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning X offset (Vision Pattern Matching) | 1.8896
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning X offset (LLC Translation) | 2.4658
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning X offset (In-Die Align) | 1.7432
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning X offset (Global Align) | 1.9287
| Samsung NRD-K eAFM | 2025-03-25 | XY Stage positioning X 1σ (LLC Translation) | 0.4735
| Samsung NRD-K eAFM | 2025-03-25 | X2 Detector noise (XY Servo on) | 0.342
| Samsung NRD-K eAFM | 2025-03-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 281.79
| Samsung NRD-K eAFM | 2025-03-25 | X1 Detector noise (XY Servo on) | 0.339
| Samsung NRD-K eAFM | 2025-03-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 279.39
| Samsung NRD-K eAFM | 2025-03-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 180.41
| Samsung NRD-K eAFM | 2025-03-25 | X Ringing(Rpv) | 0.11
| Samsung NRD-K eAFM | 2025-03-25 | X Non-linearity error | 3.46
| Samsung NRD-K eAFM | 2025-03-25 | X Detector noise (XY Servo on) | 0.251
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Repeatability @1st position | 0.04
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Repeatability @ 4th position | 0.05
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Repeatability @ 3rd position | 0.03
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Repeatability @ 2nd position | 0.05
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Accuracy @1st position | 0.23
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Accuracy @ 4th position | 0.21
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Accuracy @ 3rd position | 0.19
| Samsung NRD-K eAFM | 2025-03-25 | X axis, Accuracy @ 2nd position | 0.25
| Samsung NRD-K eAFM | 2025-03-25 | X axis OPM | 0.784
| Samsung NRD-K eAFM | 2025-03-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 17.81
| Samsung NRD-K eAFM | 2025-03-25 | Wafer placement(Loading) error on the load port2 (X offset) | 19.25
| Samsung NRD-K eAFM | 2025-03-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 30.21
| Samsung NRD-K eAFM | 2025-03-25 | Wafer placement(Loading) error on the load port1 (X offset) | 35.45
| Samsung NRD-K eAFM | 2025-03-25 | Torque Profile (50mm, 20um/s) | 0.469
| Samsung NRD-K eAFM | 2025-03-25 | Torque Profile (50mm, 200um/s) | 0.415
| Samsung NRD-K eAFM | 2025-03-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.568
| Samsung NRD-K eAFM | 2025-03-25 | Roughness Repeatability 80x80um(RSD%) | 0.44
| Samsung NRD-K eAFM | 2025-03-25 | Roughness Repeatability 1x1um(RSD%) | 1.741
| Samsung NRD-K eAFM | 2025-03-25 | Roll motion to X+ full scan | 3.9
| Samsung NRD-K eAFM | 2025-03-25 | RMS of ZeroScan | 0.05
| Samsung NRD-K eAFM | 2025-03-25 | RMS error after Stage Calibration | 4.16
| Samsung NRD-K eAFM | 2025-03-25 | Profiler noise @200um | 0.081
| Samsung NRD-K eAFM | 2025-03-25 | Position repeatability(10repeats, Position Based) | 0.26
| Samsung NRD-K eAFM | 2025-03-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| Samsung NRD-K eAFM | 2025-03-25 | OpticalAxisShiftPerFocusY | 0.0023
| Samsung NRD-K eAFM | 2025-03-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00054
| Samsung NRD-K eAFM | 2025-03-25 | OpticalAxisShiftPerFocusX | 0.0055
| Samsung NRD-K eAFM | 2025-03-25 | OPM Repeatability @5mm | 0.404
| Samsung NRD-K eAFM | 2025-03-25 | OPM Repeatability @25mm | 1.515
| Samsung NRD-K eAFM | 2025-03-25 | OPM Repeatability @1mm | 0.302
| Samsung NRD-K eAFM | 2025-03-25 | OPM Repeatability @10mm | 1.339
| Samsung NRD-K eAFM | 2025-03-25 | Min. Vibration Transmissbility (~100Hz) | 0.03
| Samsung NRD-K eAFM | 2025-03-25 | Max OPM @5mm | 9.036
| Samsung NRD-K eAFM | 2025-03-25 | Max OPM @25mm | 36.608
| Samsung NRD-K eAFM | 2025-03-25 | Max OPM @1mm | 4.08
| Samsung NRD-K eAFM | 2025-03-25 | Max OPM @10mm | 11.069
| Samsung NRD-K eAFM | 2025-03-25 | Lateral(C-D) sensitivity | 48.7
| Samsung NRD-K eAFM | 2025-03-25 | Lateral(C-D) noise | 0.195
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from -Y to +Y (Y scan) | 79.83
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from -Y to +Y (X scan) | 12.451
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from -X to +X (Y scan) | 44.677
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from -X to +X (X scan) | 35.156
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from +Y to -Y (Y scan) | 19.78
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from +Y to -Y (X scan) | 19.775
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from +X to -X (Y scan) | 14.648
| Samsung NRD-K eAFM | 2025-03-25 | Image offset after moving from +X to -X (X scan) | 13.916
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Z Height during Z Scan sweep | 16.6
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Z Detector during Z Scan sweep | 12.9
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.26
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Y2 Detector during Y Scan sweep | 217.8
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.08
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Y1 Detector during Y Scan sweep | 214.7
| Samsung NRD-K eAFM | 2025-03-25 | Full range of Y Detector during Y Scan sweep | 217.7
| Samsung NRD-K eAFM | 2025-03-25 | Full range of X2 Detector Raw during X Scan sweep | 14.51
| Samsung NRD-K eAFM | 2025-03-25 | Full range of X2 Detector during X Scan sweep | 219.5
| Samsung NRD-K eAFM | 2025-03-25 | Full range of X1 Detector Raw during X Scan sweep | 14.58
| Samsung NRD-K eAFM | 2025-03-25 | Full range of X1 Detector during X Scan sweep | 211.7
| Samsung NRD-K eAFM | 2025-03-25 | Full range of X Detector during X Scan sweep | 219.5
| Samsung NRD-K eAFM | 2025-03-25 | Cycle time for 15-point measurement | 1554.144
| Samsung NRD-K eAFM | 2025-03-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Samsung NRD-K eAFM | 2025-03-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.4
| Samsung NRD-K eAFM | 2025-03-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.78
| Samsung NRD-K eAFM | 2025-03-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| Samsung NRD-K eAFM | 2025-03-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Samsung NRD-K eAFM | 2025-03-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| Samsung NRD-K eAFM | 2025-03-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Samsung NRD-K eAFM | 2025-03-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| University of California-San Diego | 2025-03-21 | Z Scanner position at -3μm | 0.002061
| University of California-San Diego | 2025-03-21 | Z Scanner position at +3μm | 0.002879
| University of California-San Diego | 2025-03-21 | Z linearity by sample height | 0.99116
| University of California-San Diego | 2025-03-21 | Z Height noise | 0.034
| University of California-San Diego | 2025-03-21 | Z Height calibration error | 0.15134
| University of California-San Diego | 2025-03-21 | Z Drive calibration error | 0.27778
| University of California-San Diego | 2025-03-21 | Y Scan Stroke | 55
| University of California-San Diego | 2025-03-21 | Y Non-lineairty | 0.059
| University of California-San Diego | 2025-03-21 | Y Detector stroke calibration error (nominal scale error Y) | 0.162
| University of California-San Diego | 2025-03-21 | Y Detector noise (XY Servo on) | 0.377
| University of California-San Diego | 2025-03-21 | Y Detector noise (XY Servo off) | 0.434
| University of California-San Diego | 2025-03-21 | Y axis OPM | 0.7064
| University of California-San Diego | 2025-03-21 | X Scan Stroke | 55
| University of California-San Diego | 2025-03-21 | X Non-lineairty | 0.057
| University of California-San Diego | 2025-03-21 | X Detector stroke calibration error (nominal scale error X) | 0.153
| University of California-San Diego | 2025-03-21 | X Detector noise (XY Servo on) | 0.379
| University of California-San Diego | 2025-03-21 | X Detector noise (XY Servo off) | 0.434
| University of California-San Diego | 2025-03-21 | X axis OPM | 1.2686
| University of California-San Diego | 2025-03-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| University of California-San Diego | 2025-03-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| University of California-San Diego | 2025-03-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| University of California-San Diego | 2025-03-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of California-San Diego | 2025-03-21 | NCM driving% with AC160TS cantilever | 4.796
| University of California-San Diego | 2025-03-21 | NCM Amplitude noise | 10.183
| University of California-San Diego | 2025-03-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.435
| University of California-San Diego | 2025-03-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.688
| University of California-San Diego | 2025-03-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 85.5
| University of California-San Diego | 2025-03-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 86.5
| University of California-San Diego | 2025-03-21 | Full range of Z Height during Z Scan sweep | 16.9
| University of California-San Diego | 2025-03-21 | Full range of Z Detector during Z Scan sweep | 12.2
| University of California-San Diego | 2025-03-21 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| University of California-San Diego | 2025-03-21 | Full range of Y Detector during Y Scan sweep | 65.3
| University of California-San Diego | 2025-03-21 | Full range of X Detector Raw during X Scan sweep | 15.1
| University of California-San Diego | 2025-03-21 | Full range of X Detector during X Scan sweep | 65.4
| University of California-San Diego | 2025-03-21 | A-B noise | 9.491
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Z Scanner position at -3μm | 0.000793
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Z Scanner position at +3μm | 0.014421
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Z linearity by sample height | 1.0045
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Z Height noise | 0.025
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Z Height calibration error | 0.24234
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Z Drive calibration error | 0.1092
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y2 Detector noise (XY Servo on) | 0.237
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y2 Detector noise (XY Servo off) | 0.226
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y1 Detector noise (XY servo on) | 0.295
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y1 Detector noise (XY servo off) | 0.292
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y Scan Stroke | 176.3
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y Non-lineairty | 0.266
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y Detector stroke calibration error (nominal scale error Y) | 0.404
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y Detector noise (XY servo on) | 0.191
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y Detector noise (XY servo off) | 0.187
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Y axis OPM | 1.3294
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X2 Detector noise (XY Servo on) | 0.227
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X2 Detector noise (XY Servo off) | 0.219
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X1 Detector noise (XY servo on) | 0.234
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X1 Detector noise (XY servo off) | 0.226
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X Scan Stroke | 171.9
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X Non-lineairty | 0.175
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X Detector stroke calibration error (nominal scale error X) | 0.317
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X Detector noise (XY servo on) | 0.17
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X Detector noise (XY servo off) | 0.163
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | X axis OPM | 1.5646
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Roughness Repeatability 1x1um(RSD%) | 1.986
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Roughness Repeatability 1x1um(1σ) | 0.008
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.26
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.25
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.25
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | NCM driving% with AC160TS cantilever | 0.123
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | NCM Amplitude noise | 12.372
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.643
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.739
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.2
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 254
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 242.7
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 229.1
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Lateral(C-D) sensitivity | 6.5
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Lateral(C-D) noise | 0.119
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Z Height during Z Scan sweep | 17
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Z Detector during Z Scan sweep | 13.8
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Y2 Detector during Y Scan sweep | 175.7
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Y1 Detector Raw during Y Scan sweep | 13.9
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Y1 Detector during Y Scan sweep | 176.8
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of Y Detector during Y Scan sweep | 176.3
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of X2 Detector during X Scan sweep | 172.9
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of X1 Detector Raw during X Scan sweep | 14.9
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of X1 Detector during X Scan sweep | 170.9
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | Full range of X Detector during X Scan sweep | 171.9
| Xinping Semiconductor (Shenzhen) | 2025-03-20 | A-B noise | 6.388
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Z Scanner position at -3μm | 0.00673
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Z Scanner position at +3μm | 0.00122
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Z linearity by sample height | 1.0154
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Z Height noise | 0.028
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Y Detector noise (XY Servo on) | 0.31
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Y Detector noise (XY servo off) | 0.286
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Y axis OPM | 0.5526
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | X Detector noise (XY Servo on) | 0.288
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | X Detector noise (XY servo off) | 0.272
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | X axis OPM | 0.7676
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.161
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.28
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.27
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | NCM driving% with AC160TS cantilever | 1.008
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | NCM Amplitude noise | 10.254
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.39
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.749
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 185.2
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 182.4
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Lateral(C-D) sensitivity | 19.2
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Lateral(C-D) noise | 0.113
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Full range of Z Height during Z Scan sweep | 16.5
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Full range of Z Detector during Z Scan sweep | 11.9
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Full range of Y Detector during Y Scan sweep | 141.3
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Full range of X Detector Raw during X Scan sweep | 15.4
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | Full range of X Detector during X Scan sweep | 140.5
| Zhejiang Fuxi Opto-electronic Manufacturing Co.,Ltd | 2025-03-19 | A-B noise | 5.806
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Z Scanner position at -3μm | 0.000462
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Z Scanner position at +3μm | 0.007128
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Z linearity by sample height | 0.99449
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Z Height noise | 0.029
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Y Scan Stroke | 55
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Y Detector noise (XY Servo on) | 0.254
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Y Detector noise (XY Servo off) | 0.249
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Y axis OPM | 0.8308
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | X Scan Stroke | 55
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | X Detector noise (XY Servo on) | 0.249
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | X Detector noise (XY Servo off) | 0.241
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | X axis OPM | 1.2442
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 5.565
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | NCM driving% with AC160TS cantilever | 1.903
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | NCM Amplitude noise | 14.009
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.697
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.388
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.2
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.6
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Lateral(C-D) sensitivity | 18.1
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Lateral(C-D) noise | 0.107
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Full range of Z Height during Z Scan sweep | 16.7
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Full range of Z Detector during Z Scan sweep | 12.2
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Full range of Y Detector during Y Scan sweep | 73.4
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Full range of X Detector Raw during X Scan sweep | 13.7
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | Full range of X Detector during X Scan sweep | 72.4
| University of Malaya_Low Dimensional Materials Research Centre | 2025-03-18 | A-B noise | 11.163
| 서울대학교_농림생물자원학부 | 2025-03-13 | Z Scanner position at -3μm | 0.003568
| Thammasat University | 2025-03-13 | Z Scanner position at -3μm | 0.001318
| 서울대학교_농림생물자원학부 | 2025-03-13 | Z Scanner position at +3μm | 0.007107
| Thammasat University | 2025-03-13 | Z Scanner position at +3μm | 0.005665
| 서울대학교_농림생물자원학부 | 2025-03-13 | Z linearity by sample height | 0.99582
| Thammasat University | 2025-03-13 | Z linearity by sample height | 0.99327
| 서울대학교_농림생물자원학부 | 2025-03-13 | Z Height noise | 0.027
| Thammasat University | 2025-03-13 | Z Height noise | 0.03
| 서울대학교_농림생물자원학부 | 2025-03-13 | Y Scan Stroke | 55
| Thammasat University | 2025-03-13 | Y Scan Stroke | 55
| 서울대학교_농림생물자원학부 | 2025-03-13 | Y Detector noise (XY Servo on) | 0.258
| Thammasat University | 2025-03-13 | Y Detector noise (XY Servo on) | 0.387
| 서울대학교_농림생물자원학부 | 2025-03-13 | Y Detector noise (XY Servo off) | 0.25
| Thammasat University | 2025-03-13 | Y Detector noise (XY Servo off) | 0.405
| 서울대학교_농림생물자원학부 | 2025-03-13 | Y axis OPM | 0.5228
| Thammasat University | 2025-03-13 | Y axis OPM | 0.6944
| 서울대학교_농림생물자원학부 | 2025-03-13 | X Scan Stroke | 55
| Thammasat University | 2025-03-13 | X Scan Stroke | 55
| 서울대학교_농림생물자원학부 | 2025-03-13 | X Detector noise (XY Servo on) | 0.254
| Thammasat University | 2025-03-13 | X Detector noise (XY Servo on) | 0.415
| 서울대학교_농림생물자원학부 | 2025-03-13 | X Detector noise (XY Servo off) | 0.247
| Thammasat University | 2025-03-13 | X Detector noise (XY Servo off) | 0.443
| 서울대학교_농림생물자원학부 | 2025-03-13 | X axis OPM | 0.8474
| Thammasat University | 2025-03-13 | X axis OPM | 0.7774
| 서울대학교_농림생물자원학부 | 2025-03-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 4.21
| Thammasat University | 2025-03-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.44
| 서울대학교_농림생물자원학부 | 2025-03-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Thammasat University | 2025-03-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 서울대학교_농림생물자원학부 | 2025-03-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Thammasat University | 2025-03-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| 서울대학교_농림생물자원학부 | 2025-03-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Thammasat University | 2025-03-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| 서울대학교_농림생물자원학부 | 2025-03-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Thammasat University | 2025-03-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| 서울대학교_농림생물자원학부 | 2025-03-13 | NCM driving% with AC160TS cantilever | 0.988
| Thammasat University | 2025-03-13 | NCM driving% with AC160TS cantilever | 0.981
| 서울대학교_농림생물자원학부 | 2025-03-13 | NCM Amplitude noise | 18.04
| Thammasat University | 2025-03-13 | NCM Amplitude noise | 14.009
| 서울대학교_농림생물자원학부 | 2025-03-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.45
| Thammasat University | 2025-03-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.948
| 서울대학교_농림생물자원학부 | 2025-03-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.44
| Thammasat University | 2025-03-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.768
| 서울대학교_농림생물자원학부 | 2025-03-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 108.8
| Thammasat University | 2025-03-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 91
| 서울대학교_농림생물자원학부 | 2025-03-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.7
| Thammasat University | 2025-03-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.5
| 서울대학교_농림생물자원학부 | 2025-03-13 | Lateral(C-D) sensitivity | 26.2
| Thammasat University | 2025-03-13 | Lateral(C-D) sensitivity | 22.4
| 서울대학교_농림생물자원학부 | 2025-03-13 | Lateral(C-D) noise | 0.062
| Thammasat University | 2025-03-13 | Lateral(C-D) noise | 0.052
| 서울대학교_농림생물자원학부 | 2025-03-13 | Full range of Z Height during Z Scan sweep | 16.9
| Thammasat University | 2025-03-13 | Full range of Z Height during Z Scan sweep | 18
| 서울대학교_농림생물자원학부 | 2025-03-13 | Full range of Z Detector during Z Scan sweep | 11.9
| Thammasat University | 2025-03-13 | Full range of Z Detector during Z Scan sweep | 12.9
| 서울대학교_농림생물자원학부 | 2025-03-13 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| Thammasat University | 2025-03-13 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| 서울대학교_농림생물자원학부 | 2025-03-13 | Full range of Y Detector during Y Scan sweep | 72.3
| Thammasat University | 2025-03-13 | Full range of Y Detector during Y Scan sweep | 67.8
| 서울대학교_농림생물자원학부 | 2025-03-13 | Full range of X Detector Raw during X Scan sweep | 13.7
| Thammasat University | 2025-03-13 | Full range of X Detector Raw during X Scan sweep | 14.2
| 서울대학교_농림생물자원학부 | 2025-03-13 | Full range of X Detector during X Scan sweep | 72.4
| Thammasat University | 2025-03-13 | Full range of X Detector during X Scan sweep | 69
| 서울대학교_농림생물자원학부 | 2025-03-13 | A-B noise | 8.788
| Thammasat University | 2025-03-13 | A-B noise | 7.837
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Z Scanner position at -3μm | 0.003318
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Z Scanner position at +3μm | 0.012639
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Z linearity by sample height | 0.99635
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Z Height noise | 0.028
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Y Scan Stroke | 110
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Y Detector noise (XY Servo on) | 0.247
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Y Detector noise (XY servo off) | 0.248
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Y axis OPM | 0.4522
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | X Scan Stroke | 108.4
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | X Detector noise (XY Servo on) | 0.253
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | X Detector noise (XY servo off) | 0.254
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | X axis OPM | 1.0044
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 0.653
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.53
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | NCM driving% with AC160TS cantilever | 3.247
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | NCM Amplitude noise | 12.598
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.756
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.44
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 159
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.5
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Lateral(C-D) sensitivity | 18.7
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Lateral(C-D) noise | 0.105
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Full range of Z Height during Z Scan sweep | 17
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Full range of Z Detector during Z Scan sweep | 12.4
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Full range of Y Detector Raw during Y Scan sweep | 14
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Full range of Y Detector during Y Scan sweep | 109.8
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Full range of X Detector Raw during X Scan sweep | 13.7
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | Full range of X Detector during X Scan sweep | 108.8
| Swiss Federal Institute of TWuhan Polytechnic University | 2025-03-12 | A-B noise | 12.005
| TSMC F22(1) | 2025-03-09 | Z Linearity by Z scanner position | 0.94
| TSMC F22(1) | 2025-03-09 | Z Linearity by 4 different sample height | 0.99731
| TSMC F22(1) | 2025-03-09 | Z Height noise | 0.024
| TSMC F22(1) | 2025-03-09 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.5619
| TSMC F22(1) | 2025-03-09 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.5711
| TSMC F22(1) | 2025-03-09 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.3841
| TSMC F22(1) | 2025-03-09 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3316
| TSMC F22(1) | 2025-03-09 | Z Drift | 16.068
| TSMC F22(1) | 2025-03-09 | Y2 Detector noise (XY Servo on) | 0.37
| TSMC F22(1) | 2025-03-09 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 259.2084
| TSMC F22(1) | 2025-03-09 | Y1 Detector noise (XY Servo on) | 0.424
| TSMC F22(1) | 2025-03-09 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 260.3838
| TSMC F22(1) | 2025-03-09 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 170.933
| TSMC F22(1) | 2025-03-09 | Y Ringing(Rpv) | 0.202
| TSMC F22(1) | 2025-03-09 | Y Non-linearity error | 3.32
| TSMC F22(1) | 2025-03-09 | Y Detector noise (XY Servo on) | 0.296
| TSMC F22(1) | 2025-03-09 | Y axis, Repeatability @1st position | 0.03
| TSMC F22(1) | 2025-03-09 | Y axis, Repeatability @ 4th position | 0.04
| TSMC F22(1) | 2025-03-09 | Y axis, Repeatability @ 3rd position | 0.03
| TSMC F22(1) | 2025-03-09 | Y axis, Repeatability @ 2nd position | 0.04
| TSMC F22(1) | 2025-03-09 | Y axis, Accuracy @1st position | 0.27
| TSMC F22(1) | 2025-03-09 | Y axis, Accuracy @ 4th position | 0.32
| TSMC F22(1) | 2025-03-09 | Y axis, Accuracy @ 3rd position | 0.18
| TSMC F22(1) | 2025-03-09 | Y axis, Accuracy @ 2nd position | 0.22
| TSMC F22(1) | 2025-03-09 | Y axis OPM | 0.636
| TSMC F22(1) | 2025-03-09 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.7627
| TSMC F22(1) | 2025-03-09 | XY Stage positioning Y offset (LLC Translation) | 1.7969
| TSMC F22(1) | 2025-03-09 | XY Stage positioning Y offset (In-Die Align) | 0.9521
| TSMC F22(1) | 2025-03-09 | XY Stage positioning Y offset (Global Align) | 2.6318
| TSMC F22(1) | 2025-03-09 | XY Stage positioning Y 1σ (LLC Translation) | 0.3654
| TSMC F22(1) | 2025-03-09 | XY Stage positioning X offset (Vision Pattern Matching) | 2.334
| TSMC F22(1) | 2025-03-09 | XY Stage positioning X offset (LLC Translation) | 1.5576
| TSMC F22(1) | 2025-03-09 | XY Stage positioning X offset (In-Die Align) | 1.1279
| TSMC F22(1) | 2025-03-09 | XY Stage positioning X offset (Global Align) | 1.4551
| TSMC F22(1) | 2025-03-09 | XY Stage positioning X 1σ (LLC Translation) | 0.3022
| TSMC F22(1) | 2025-03-09 | X2 Detector noise (XY Servo on) | 0.46
| TSMC F22(1) | 2025-03-09 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 258.734
| TSMC F22(1) | 2025-03-09 | X1 Detector noise (XY Servo on) | 0.455
| TSMC F22(1) | 2025-03-09 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 257.773
| TSMC F22(1) | 2025-03-09 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 169.671
| TSMC F22(1) | 2025-03-09 | X Ringing(Rpv) | 0.259
| TSMC F22(1) | 2025-03-09 | X Non-linearity error | 2.94
| TSMC F22(1) | 2025-03-09 | X Detector noise (XY Servo on) | 0.339
| TSMC F22(1) | 2025-03-09 | X axis, Repeatability @1st position | 0.04
| TSMC F22(1) | 2025-03-09 | X axis, Repeatability @ 4th position | 0.03
| TSMC F22(1) | 2025-03-09 | X axis, Repeatability @ 3rd position | 0.02
| TSMC F22(1) | 2025-03-09 | X axis, Repeatability @ 2nd position | 0.03
| TSMC F22(1) | 2025-03-09 | X axis, Accuracy @1st position | 0.23
| TSMC F22(1) | 2025-03-09 | X axis, Accuracy @ 4th position | 0.15
| TSMC F22(1) | 2025-03-09 | X axis, Accuracy @ 3rd position | 0.1
| TSMC F22(1) | 2025-03-09 | X axis, Accuracy @ 2nd position | 0.22
| TSMC F22(1) | 2025-03-09 | X axis OPM | 0.486
| TSMC F22(1) | 2025-03-09 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.65
| TSMC F22(1) | 2025-03-09 | Wafer placement(Loading) error on the load port2 (X offset) | 20
| TSMC F22(1) | 2025-03-09 | Wafer placement(Loading) error on the load port1 (Y offset) | 23.01
| TSMC F22(1) | 2025-03-09 | Wafer placement(Loading) error on the load port1 (X offset) | 17.65
| TSMC F22(1) | 2025-03-09 | Torque Profile (50mm, 20um/s) | 0.477
| TSMC F22(1) | 2025-03-09 | Torque Profile (50mm, 200um/s) | 0.438
| TSMC F22(1) | 2025-03-09 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.891
| TSMC F22(1) | 2025-03-09 | Roughness Repeatability 80x80um(RSD%) | 1.78
| TSMC F22(1) | 2025-03-09 | Roughness Repeatability 1x1um(RSD%) | 1.001
| TSMC F22(1) | 2025-03-09 | Roll motion to X+ full scan | 7.3
| TSMC F22(1) | 2025-03-09 | RMS of ZeroScan | 0.05
| TSMC F22(1) | 2025-03-09 | RMS error after Stage Calibration | 3.21
| TSMC F22(1) | 2025-03-09 | Profiler noise @200um | 0.227
| TSMC F22(1) | 2025-03-09 | Position repeatability(10repeats, Position Based) | 0.78
| TSMC F22(1) | 2025-03-09 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| TSMC F22(1) | 2025-03-09 | OpticalAxisShiftPerFocusY | -0.0018
| TSMC F22(1) | 2025-03-09 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| TSMC F22(1) | 2025-03-09 | OpticalAxisShiftPerFocusX | -0.0001
| TSMC F22(1) | 2025-03-09 | OPM Repeatability @5mm | 0.561
| TSMC F22(1) | 2025-03-09 | OPM Repeatability @25mm | 1.894
| TSMC F22(1) | 2025-03-09 | OPM Repeatability @1mm | 0.562
| TSMC F22(1) | 2025-03-09 | OPM Repeatability @10mm | 1.083
| TSMC F22(1) | 2025-03-09 | Min. Vibration Transmissbility (~100Hz) | 0.059
| TSMC F22(1) | 2025-03-09 | Max OPM @5mm | 12.295
| TSMC F22(1) | 2025-03-09 | Max OPM @25mm | 40.078
| TSMC F22(1) | 2025-03-09 | Max OPM @1mm | 6.246
| TSMC F22(1) | 2025-03-09 | Max OPM @10mm | 21.318
| TSMC F22(1) | 2025-03-09 | Lateral(C-D) sensitivity | 5
| TSMC F22(1) | 2025-03-09 | Lateral(C-D) noise | 0.175
| TSMC F22(1) | 2025-03-09 | Image offset after moving from -Y to +Y (Y scan) | 27.0996
| TSMC F22(1) | 2025-03-09 | Image offset after moving from -Y to +Y (X scan) | 63.7207
| TSMC F22(1) | 2025-03-09 | Image offset after moving from -X to +X (Y scan) | 57.1289
| TSMC F22(1) | 2025-03-09 | Image offset after moving from -X to +X (X scan) | 145.0195
| TSMC F22(1) | 2025-03-09 | Image offset after moving from +Y to -Y (Y scan) | 99.6094
| TSMC F22(1) | 2025-03-09 | Image offset after moving from +Y to -Y (X scan) | 22.7051
| TSMC F22(1) | 2025-03-09 | Image offset after moving from +X to -X (Y scan) | 33.6914
| TSMC F22(1) | 2025-03-09 | Image offset after moving from +X to -X (X scan) | 107.666
| TSMC F22(1) | 2025-03-09 | Full range of Z Height during Z Scan sweep | 16.4
| TSMC F22(1) | 2025-03-09 | Full range of Z Detector during Z Scan sweep | 12.84
| TSMC F22(1) | 2025-03-09 | Full range of Y2 Detector Raw during Y Scan sweep | 14.42
| TSMC F22(1) | 2025-03-09 | Full range of Y2 Detector during Y Scan sweep | 211
| TSMC F22(1) | 2025-03-09 | Full range of Y1 Detector Raw during Y Scan sweep | 14.43
| TSMC F22(1) | 2025-03-09 | Full range of Y1 Detector during Y Scan sweep | 203.9
| TSMC F22(1) | 2025-03-09 | Full range of Y Detector during Y Scan sweep | 207.4
| TSMC F22(1) | 2025-03-09 | Full range of X2 Detector Raw during X Scan sweep | 14.14
| TSMC F22(1) | 2025-03-09 | Full range of X2 Detector during X Scan sweep | 196.4
| TSMC F22(1) | 2025-03-09 | Full range of X1 Detector Raw during X Scan sweep | 14.38
| TSMC F22(1) | 2025-03-09 | Full range of X1 Detector during X Scan sweep | 204.4
| TSMC F22(1) | 2025-03-09 | Full range of X Detector during X Scan sweep | 200.4
| TSMC F22(1) | 2025-03-09 | Cycle time for 15-point measurement | 1853.557
| TSMC F22(1) | 2025-03-09 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| TSMC F22(1) | 2025-03-09 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.21
| TSMC F22(1) | 2025-03-09 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.33
| TSMC F22(1) | 2025-03-09 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| TSMC F22(1) | 2025-03-09 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| TSMC F22(1) | 2025-03-09 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| TSMC F22(1) | 2025-03-09 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| TSMC F22(1) | 2025-03-09 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Linearity by Z scanner position | 1.78
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Linearity by 4 different sample height | 1.001
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Height noise | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.2598
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.5615
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4631
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.348
| Dongguan Guangmao Technology(2) | 2025-03-07 | Z Drift | 10.5777
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y2 Detector noise (XY Servo on) | 0.636
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.6671
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y1 Detector noise (XY Servo on) | 0.656
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 277.9736
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 181.443
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y Ringing(Rpv) | 0.123
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y Non-linearity error | 3.18
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y Detector noise (XY Servo on) | 0.599
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Repeatability @1st position | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Repeatability @ 4th position | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Repeatability @ 3rd position | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Repeatability @ 2nd position | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Accuracy @1st position | 0.26
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Accuracy @ 4th position | 0.28
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Accuracy @ 3rd position | 0.22
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis, Accuracy @ 2nd position | 0.23
| Dongguan Guangmao Technology(2) | 2025-03-07 | Y axis OPM | 1.126
| Dongguan Guangmao Technology(2) | 2025-03-07 | X2 Detector noise (XY Servo on) | 0.323
| Dongguan Guangmao Technology(2) | 2025-03-07 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 241.5042
| Dongguan Guangmao Technology(2) | 2025-03-07 | X1 Detector noise (XY Servo on) | 0.373
| Dongguan Guangmao Technology(2) | 2025-03-07 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 244.8048
| Dongguan Guangmao Technology(2) | 2025-03-07 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 156.406
| Dongguan Guangmao Technology(2) | 2025-03-07 | X Ringing(Rpv) | 0.089
| Dongguan Guangmao Technology(2) | 2025-03-07 | X Non-linearity error | 2.14
| Dongguan Guangmao Technology(2) | 2025-03-07 | X Detector noise (XY Servo on) | 0.269
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Repeatability @1st position | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Repeatability @ 4th position | 0.04
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Repeatability @ 3rd position | 0.04
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Repeatability @ 2nd position | 0.03
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Accuracy @1st position | 0.12
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Accuracy @ 4th position | 0.26
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Accuracy @ 3rd position | 0.21
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis, Accuracy @ 2nd position | 0.15
| Dongguan Guangmao Technology(2) | 2025-03-07 | X axis OPM | 2.051
| Dongguan Guangmao Technology(2) | 2025-03-07 | Torque Profile (50mm, 20um/s) | 0.578
| Dongguan Guangmao Technology(2) | 2025-03-07 | Torque Profile (50mm, 200um/s) | 0.42
| Dongguan Guangmao Technology(2) | 2025-03-07 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.673
| Dongguan Guangmao Technology(2) | 2025-03-07 | Roughness Repeatability 80x80um(RSD%) | 3.6233
| Dongguan Guangmao Technology(2) | 2025-03-07 | Roughness Repeatability 1x1um(RSD%) | 5.9626
| Dongguan Guangmao Technology(2) | 2025-03-07 | Roll motion to X+ full scan | 36.7
| Dongguan Guangmao Technology(2) | 2025-03-07 | RMS of ZeroScan | 0.052
| Dongguan Guangmao Technology(2) | 2025-03-07 | RMS error after Stage Calibration | 3.51
| Dongguan Guangmao Technology(2) | 2025-03-07 | Profiler noise @200um | 0.123
| Dongguan Guangmao Technology(2) | 2025-03-07 | Position repeatability(10repeats, Position Based) | 0.92
| Dongguan Guangmao Technology(2) | 2025-03-07 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| Dongguan Guangmao Technology(2) | 2025-03-07 | OpticalAxisShiftPerFocusY | -0.0024
| Dongguan Guangmao Technology(2) | 2025-03-07 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00015
| Dongguan Guangmao Technology(2) | 2025-03-07 | OpticalAxisShiftPerFocusX | -0.0043
| Dongguan Guangmao Technology(2) | 2025-03-07 | OPM Repeatability @5mm | 0.584
| Dongguan Guangmao Technology(2) | 2025-03-07 | OPM Repeatability @25mm | 2.263
| Dongguan Guangmao Technology(2) | 2025-03-07 | OPM Repeatability @1mm | 0.955
| Dongguan Guangmao Technology(2) | 2025-03-07 | OPM Repeatability @10mm | 1.365
| Dongguan Guangmao Technology(2) | 2025-03-07 | Min. Vibration Transmissbility (~100Hz) | 0.027
| Dongguan Guangmao Technology(2) | 2025-03-07 | Max OPM @5mm | 15.996
| Dongguan Guangmao Technology(2) | 2025-03-07 | Max OPM @25mm | 99.189
| Dongguan Guangmao Technology(2) | 2025-03-07 | Max OPM @1mm | 7.131
| Dongguan Guangmao Technology(2) | 2025-03-07 | Max OPM @10mm | 20.65
| Dongguan Guangmao Technology(2) | 2025-03-07 | Lift pins synchronization offset check (Y offset) | 34.9
| Dongguan Guangmao Technology(2) | 2025-03-07 | Lift pins synchronization offset check (X offset) | 37.83
| Dongguan Guangmao Technology(2) | 2025-03-07 | Lateral(C-D) sensitivity | 48.3
| Dongguan Guangmao Technology(2) | 2025-03-07 | Lateral(C-D) noise | 0.228
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from -Y to +Y (Y scan) | 83.4961
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from -Y to +Y (X scan) | 14.6484
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from -X to +X (Y scan) | 49.8047
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from -X to +X (X scan) | 44.6777
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from +Y to -Y (Y scan) | 31.4941
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from +Y to -Y (X scan) | 21.9727
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from +X to -X (Y scan) | 27.832
| Dongguan Guangmao Technology(2) | 2025-03-07 | Image offset after moving from +X to -X (X scan) | 45.4102
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Z Height during Z Scan sweep | 16.5062
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Z Detector during Z Scan sweep | 11.56
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Y2 Detector Raw during Y Scan sweep | 14.34
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Y2 Detector during Y Scan sweep | 201.5479
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Y1 Detector Raw during Y Scan sweep | 14.51
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Y1 Detector during Y Scan sweep | 204.6906
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of Y Detector during Y Scan sweep | 204.6906
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of X2 Detector Raw during X Scan sweep | 14.29
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of X2 Detector during X Scan sweep | 172.5484
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of X1 Detector Raw during X Scan sweep | 14.39
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of X1 Detector during X Scan sweep | 178.6137
| Dongguan Guangmao Technology(2) | 2025-03-07 | Full range of X Detector during X Scan sweep | 178.6137
| Dongguan Guangmao Technology(2) | 2025-03-07 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13773
| Dongguan Guangmao Technology(2) | 2025-03-07 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.76389
| Dongguan Guangmao Technology(2) | 2025-03-07 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.079103
| Dongguan Guangmao Technology(2) | 2025-03-07 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.4372
| TSMC AP2 Mini Line (4) | 2025-02-28 | Z Linearity by Z scanner position | 0.78
| TSMC F22(2) | 2025-02-28 | Z Linearity by Z scanner position | 0.9
| TSMC AP2 Mini Line (4) | 2025-02-28 | Z Linearity by 4 different sample height | 0.997
| TSMC F22(2) | 2025-02-28 | Z Linearity by 4 different sample height | 0.99866
| TSMC AP2 Mini Line (4) | 2025-02-28 | Z Height noise | 0.027
| TSMC F22(2) | 2025-02-28 | Z Height noise | 0.025
| TSMC F22(2) | 2025-02-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.1679
| TSMC F22(2) | 2025-02-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.1562
| TSMC F22(2) | 2025-02-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.363
| TSMC AP2 Mini Line (4) | 2025-02-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.2685
| TSMC F22(2) | 2025-02-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3001
| TSMC F22(2) | 2025-02-28 | Z Drift | 5.94
| TSMC AP2 Mini Line (4) | 2025-02-28 | Z Detector offset | 161.4
| TSMC F22(2) | 2025-02-28 | Y2 Detector noise (XY Servo on) | 0.362
| TSMC F22(2) | 2025-02-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 239.1739
| TSMC F22(2) | 2025-02-28 | Y1 Detector noise (XY Servo on) | 0.378
| TSMC F22(2) | 2025-02-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 228.1682
| TSMC F22(2) | 2025-02-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 152.397
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y Ringing(Rpv) | 0.098
| TSMC F22(2) | 2025-02-28 | Y Ringing(Rpv) | 0.106
| TSMC F22(2) | 2025-02-28 | Y Non-linearity error | 3.41
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y detector noise (XY Servo on) | 0.224
| TSMC F22(2) | 2025-02-28 | Y Detector noise (XY Servo on) | 0.315
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 151.778
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Repeatability @1st position | 0.06
| TSMC F22(2) | 2025-02-28 | Y axis, Repeatability @1st position | 0.04
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Repeatability @ 4th position | 0.07
| TSMC F22(2) | 2025-02-28 | Y axis, Repeatability @ 4th position | 0.05
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Repeatability @ 3rd position | 0.05
| TSMC F22(2) | 2025-02-28 | Y axis, Repeatability @ 3rd position | 0.04
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Repeatability @ 2nd position | 0.06
| TSMC F22(2) | 2025-02-28 | Y axis, Repeatability @ 2nd position | 0.04
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Accuracy @1st position | 0.36
| TSMC F22(2) | 2025-02-28 | Y axis, Accuracy @1st position | 0.25
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Accuracy @ 4th position | 0.46
| TSMC F22(2) | 2025-02-28 | Y axis, Accuracy @ 4th position | 0.26
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Accuracy @ 3rd position | 0.36
| TSMC F22(2) | 2025-02-28 | Y axis, Accuracy @ 3rd position | 0.26
| TSMC AP2 Mini Line (4) | 2025-02-28 | Y axis, Accuracy @ 2nd position | 0.37
| TSMC F22(2) | 2025-02-28 | Y axis, Accuracy @ 2nd position | 0.23
| TSMC F22(2) | 2025-02-28 | Y axis OPM | 0.572
| TSMC AP2 Mini Line (4) | 2025-02-28 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.3965
| TSMC F22(2) | 2025-02-28 | XY Stage positioning Y offset (Vision Pattern Matching) | 9.7705
| TSMC F22(2) | 2025-02-28 | XY Stage positioning Y offset (LLC Translation) | 2.0117
| TSMC AP2 Mini Line (4) | 2025-02-28 | XY Stage positioning Y offset (In-Die Align) | 1.7285
| TSMC F22(2) | 2025-02-28 | XY Stage positioning Y offset (In-Die Align) | 5.6396
| TSMC F22(2) | 2025-02-28 | XY Stage positioning Y offset (Global Align) | 2.3242
| TSMC F22(2) | 2025-02-28 | XY Stage positioning Y 1σ (LLC Translation) | 0.4223
| TSMC AP2 Mini Line (4) | 2025-02-28 | XY Stage positioning X offset (Vision Pattern Matching) | 1.8701
| TSMC F22(2) | 2025-02-28 | XY Stage positioning X offset (Vision Pattern Matching) | 4.1455
| TSMC F22(2) | 2025-02-28 | XY Stage positioning X offset (LLC Translation) | 2.6367
| TSMC AP2 Mini Line (4) | 2025-02-28 | XY Stage positioning X offset (In-Die Align) | 1.416
| TSMC F22(2) | 2025-02-28 | XY Stage positioning X offset (In-Die Align) | 1.5234
| TSMC F22(2) | 2025-02-28 | XY Stage positioning X offset (Global Align) | 2.5391
| TSMC F22(2) | 2025-02-28 | XY Stage positioning X 1σ (LLC Translation) | 0.5421
| TSMC F22(2) | 2025-02-28 | X2 Detector noise (XY Servo on) | 0.354
| TSMC F22(2) | 2025-02-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 248.506
| TSMC F22(2) | 2025-02-28 | X1 Detector noise (XY Servo on) | 0.337
| TSMC F22(2) | 2025-02-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 246.9103
| TSMC F22(2) | 2025-02-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 164.919
| TSMC AP2 Mini Line (4) | 2025-02-28 | X Ringing(Rpv) | 0.162
| TSMC F22(2) | 2025-02-28 | X Ringing(Rpv) | 0.162
| TSMC F22(2) | 2025-02-28 | X Non-linearity error | 3.42
| TSMC F22(2) | 2025-02-28 | X Detector noise (XY Servo on) | 0.281
| TSMC AP2 Mini Line (4) | 2025-02-28 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 150.157
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Repeatability @1st position | 0.04
| TSMC F22(2) | 2025-02-28 | X axis, Repeatability @1st position | 0.04
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Repeatability @ 4th position | 0.06
| TSMC F22(2) | 2025-02-28 | X axis, Repeatability @ 4th position | 0.03
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Repeatability @ 3rd position | 0.03
| TSMC F22(2) | 2025-02-28 | X axis, Repeatability @ 3rd position | 0.04
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Repeatability @ 2nd position | 0.04
| TSMC F22(2) | 2025-02-28 | X axis, Repeatability @ 2nd position | 0.04
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Accuracy @1st position | 0.35
| TSMC F22(2) | 2025-02-28 | X axis, Accuracy @1st position | 0.21
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Accuracy @ 4th position | 0.33
| TSMC F22(2) | 2025-02-28 | X axis, Accuracy @ 4th position | 0.19
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Accuracy @ 3rd position | 0.16
| TSMC F22(2) | 2025-02-28 | X axis, Accuracy @ 3rd position | 0.22
| TSMC AP2 Mini Line (4) | 2025-02-28 | X axis, Accuracy @ 2nd position | 0.25
| TSMC F22(2) | 2025-02-28 | X axis, Accuracy @ 2nd position | 0.25
| TSMC F22(2) | 2025-02-28 | X axis OPM | 0.342
| TSMC AP2 Mini Line (4) | 2025-02-28 | Wafer placement(Loading) error on the load port2 (Y offset) | 26.61
| TSMC F22(2) | 2025-02-28 | Wafer placement(Loading) error on the load port2 (Y offset) | 30.56
| TSMC AP2 Mini Line (4) | 2025-02-28 | Wafer placement(Loading) error on the load port2 (X offset) | 23.35
| TSMC F22(2) | 2025-02-28 | Wafer placement(Loading) error on the load port2 (X offset) | 18.05
| TSMC AP2 Mini Line (4) | 2025-02-28 | Wafer placement(Loading) error on the load port1 (Y offset) | 61.07
| TSMC F22(2) | 2025-02-28 | Wafer placement(Loading) error on the load port1 (Y offset) | 24.91
| TSMC AP2 Mini Line (4) | 2025-02-28 | Wafer placement(Loading) error on the load port1 (X offset) | 44.66
| TSMC F22(2) | 2025-02-28 | Wafer placement(Loading) error on the load port1 (X offset) | 13.45
| TSMC AP2 Mini Line (4) | 2025-02-28 | UpperLimit | 29000
| TSMC AP2 Mini Line (4) | 2025-02-28 | Under 50nm Step Height Sample (10 Repeats) | 15.9867
| TSMC F22(2) | 2025-02-28 | Torque Profile (50mm, 20um/s) | 0.847
| TSMC F22(2) | 2025-02-28 | Torque Profile (50mm, 200um/s) | 0.444
| TSMC F22(2) | 2025-02-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.65
| TSMC F22(2) | 2025-02-28 | Roughness Repeatability 80x80um(RSD%) | 1.74
| TSMC F22(2) | 2025-02-28 | Roughness Repeatability 1x1um(RSD%) | 1.482
| TSMC F22(2) | 2025-02-28 | Roll motion to X+ full scan | 17
| TSMC F22(2) | 2025-02-28 | RMS of ZeroScan | 0.05
| TSMC F22(2) | 2025-02-28 | RMS error after Stage Calibration | 1.75
| TSMC F22(2) | 2025-02-28 | Profiler noise @200um | 0.125
| TSMC F22(2) | 2025-02-28 | Position repeatability(10repeats, Position Based) | 0.22
| TSMC AP2 Mini Line (4) | 2025-02-28 | Overlay Camera X Position(ΔX) | -99965.6
| TSMC AP2 Mini Line (4) | 2025-02-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7.07e-05
| TSMC F22(2) | 2025-02-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8e-05
| TSMC AP2 Mini Line (4) | 2025-02-28 | OpticalAxisShiftPerFocusY | 0.0006
| TSMC F22(2) | 2025-02-28 | OpticalAxisShiftPerFocusY | 0.0009
| TSMC AP2 Mini Line (4) | 2025-02-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000779
| TSMC F22(2) | 2025-02-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| TSMC AP2 Mini Line (4) | 2025-02-28 | OpticalAxisShiftPerFocusX | -0.0083
| TSMC F22(2) | 2025-02-28 | OpticalAxisShiftPerFocusX | -0.0042
| TSMC F22(2) | 2025-02-28 | OPM Repeatability @5mm | 1.794
| TSMC F22(2) | 2025-02-28 | OPM Repeatability @25mm | 8.102
| TSMC F22(2) | 2025-02-28 | OPM Repeatability @1mm | 1.105
| TSMC F22(2) | 2025-02-28 | OPM Repeatability @10mm | 2.322
| TSMC F22(2) | 2025-02-28 | Min. Vibration Transmissbility (~100Hz) | 0.06
| TSMC F22(2) | 2025-02-28 | Max OPM @5mm | 16.141
| TSMC F22(2) | 2025-02-28 | Max OPM @25mm | 135.092
| TSMC F22(2) | 2025-02-28 | Max OPM @1mm | 10.322
| TSMC F22(2) | 2025-02-28 | Max OPM @10mm | 25.057
| TSMC AP2 Mini Line (4) | 2025-02-28 | LowerLimit | 5880
| TSMC F22(2) | 2025-02-28 | Lateral(C-D) sensitivity | 18.8
| TSMC F22(2) | 2025-02-28 | Lateral(C-D) noise | 0.162
| TSMC AP2 Mini Line (4) | 2025-02-28 | Initial Z | 15500
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from -Y to +Y (Y scan) | 136.9629
| TSMC F22(2) | 2025-02-28 | Image offset after moving from -Y to +Y (Y scan) | 43.2129
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from -Y to +Y (X scan) | 17.5781
| TSMC F22(2) | 2025-02-28 | Image offset after moving from -Y to +Y (X scan) | 6.5918
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from -X to +X (Y scan) | 32.959
| TSMC F22(2) | 2025-02-28 | Image offset after moving from -X to +X (Y scan) | 48.3398
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from -X to +X (X scan) | 147.9492
| TSMC F22(2) | 2025-02-28 | Image offset after moving from -X to +X (X scan) | 28.5645
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from +Y to -Y (Y scan) | 23.4375
| TSMC F22(2) | 2025-02-28 | Image offset after moving from +Y to -Y (Y scan) | 30.7617
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from +Y to -Y (X scan) | 82.7637
| TSMC F22(2) | 2025-02-28 | Image offset after moving from +Y to -Y (X scan) | 8.7891
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from +X to -X (Y scan) | 11.7188
| TSMC F22(2) | 2025-02-28 | Image offset after moving from +X to -X (Y scan) | 14.6484
| TSMC AP2 Mini Line (4) | 2025-02-28 | Image offset after moving from +X to -X (X scan) | 5.127
| TSMC F22(2) | 2025-02-28 | Image offset after moving from +X to -X (X scan) | 30.7617
| TSMC AP2 Mini Line (4) | 2025-02-28 | FullPixelHeight | 1828
| TSMC AP2 Mini Line (4) | 2025-02-28 | Full range of Z Height during Z Scan sweep | 18.6693
| TSMC F22(2) | 2025-02-28 | Full range of Z Height during Z Scan sweep | 16.8907
| TSMC AP2 Mini Line (4) | 2025-02-28 | Full range of Z Detector during Z Scan sweep | 15.22
| TSMC F22(2) | 2025-02-28 | Full range of Z Detector during Z Scan sweep | 12.62
| TSMC F22(2) | 2025-02-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.42
| TSMC F22(2) | 2025-02-28 | Full range of Y2 Detector during Y Scan sweep | 187.3695
| TSMC F22(2) | 2025-02-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.41
| TSMC F22(2) | 2025-02-28 | Full range of Y1 Detector during Y Scan sweep | 178.405
| TSMC AP2 Mini Line (4) | 2025-02-28 | Full range of Y Detector Raw during Y Scan sweep | 15.4
| TSMC F22(2) | 2025-02-28 | Full range of Y Detector during Y Scan sweep | 182.8873
| TSMC F22(2) | 2025-02-28 | Full range of X2 Detector Raw during X Scan sweep | 14.21
| TSMC F22(2) | 2025-02-28 | Full range of X2 Detector during X Scan sweep | 185.252
| TSMC F22(2) | 2025-02-28 | Full range of X1 Detector Raw during X Scan sweep | 14.28
| TSMC F22(2) | 2025-02-28 | Full range of X1 Detector during X Scan sweep | 184.7349
| TSMC AP2 Mini Line (4) | 2025-02-28 | Full range of X Detector Raw during X Scan sweep | 15.57
| TSMC F22(2) | 2025-02-28 | Full range of X Detector during X Scan sweep | 184.9934
| TSMC F22(2) | 2025-02-28 | Cycle time for 15-point measurement | 1756.717
| TSMC AP2 Mini Line (4) | 2025-02-28 | Check the full stroke image(servo on) | Yes
| TSMC F22(2) | 2025-02-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| TSMC F22(2) | 2025-02-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.61
| TSMC F22(2) | 2025-02-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.69
| TSMC F22(2) | 2025-02-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| TSMC AP2 Mini Line (4) | 2025-02-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1222
| TSMC F22(2) | 2025-02-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| TSMC AP2 Mini Line (4) | 2025-02-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.67282
| TSMC F22(2) | 2025-02-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| TSMC F22(2) | 2025-02-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| TSMC F22(2) | 2025-02-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| TSMC AP2 Mini Line (4) | 2025-02-28 | A-B after Auto PSPD Alignment | 0.104
| TSMC AP2 Mini Line (4) | 2025-02-28 | 940nm Step Height Sample (10 Repeats) | 945.7094
| TSMC AP2 Mini Line (4) | 2025-02-28 | 50~200nm Step Height Sample (10 Repeats) | 184.4086
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Z Scanner position at -3μm | 0.00066
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Z Scanner position at +3μm | 0.012049
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Z linearity by sample height | 0.98251
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Z Height noise | 0.017
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y2 Detector noise (XY Servo on) | 0.429
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y2 Detector noise (XY Servo off) | 0.385
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y1 Detector noise (XY servo on) | 0.44
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y1 Detector noise (XY servo off) | 0.377
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y Scan Stroke | 177.9
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y Detector noise (XY servo on) | 0.317
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y Detector noise (XY servo off) | 0.277
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y axis OPM (-52000,-52000) | 2.184
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y axis OPM (-52000,52000) | 1.999
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y axis OPM (52000,-52000) | 2.084
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y axis OPM (52000,52000) | 2.89
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Y axis OPM (0,0) | 1.297
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | XY Orthogonality (-52000,-52000) | 89.98
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | XY Orthogonality (-52000,52000) | 89.95
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | XY Orthogonality (52000,-52000) | 90.14
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | XY Orthogonality (52000,52000) | 89.88
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | XY Orthogonality (0,0) | 90.01
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X2 Detector noise (XY Servo on) | 0.365
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X2 Detector noise (XY Servo off) | 0.355
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X1 Detector noise (XY servo on) | 0.388
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X1 Detector noise (XY servo off) | 0.389
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X Scan Stroke | 175.7
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X Detector noise (XY servo on) | 0.276
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X Detector noise (XY servo off) | 0.27
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X axis OPM (-52000,-52000) | 1.861
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X axis OPM (-52000,52000) | 3.106
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X axis OPM (52000,-52000) | 3.781
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X axis OPM (52000,52000) | 1.981
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | X axis OPM (0,0) | 1.8168
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | SLD beam power at the end of the lens | 0.256
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 9.96
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.14
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.14
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.15
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.15
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.14
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | NCM driving% with AC160TS cantilever | 12.365
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | NCM Amplitude noise | 6.013
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.835
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.479
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 246.8
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 254.4
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 247.1
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249.5
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Lateral(C-D) sensitivity | 11.8
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Lateral(C-D) noise | 0.093
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Z Height during Z Scan sweep | 17.1
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Z Detector during Z Scan sweep | 15.1
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Y2 Detector during Y Scan sweep | 176
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Y1 Detector during Y Scan sweep | 179.7
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of Y Detector during Y Scan sweep | 177.9
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of X2 Detector during X Scan sweep | 174.4
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of X1 Detector during X Scan sweep | 176.2
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | Full range of X Detector during X Scan sweep | 175.3
| Swiss Federal Institute of Technology Lausanne_EPFL | 2025-02-26 | A-B noise | 6.63
| 충북대학교 신소재공학과 | 2025-02-25 | Z Scanner position at -3μm | 0.001693
| 충북대학교 신소재공학과 | 2025-02-25 | Z Scanner position at +3μm | 0.000606
| 충북대학교 신소재공학과 | 2025-02-25 | Z linearity by sample height | 0.99759
| 충북대학교 신소재공학과 | 2025-02-25 | Z Height noise | 0.027
| 충북대학교 신소재공학과 | 2025-02-25 | Y Scan Stroke | 55
| 충북대학교 신소재공학과 | 2025-02-25 | Y Detector noise (XY Servo on) | 0.349
| 충북대학교 신소재공학과 | 2025-02-25 | Y Detector noise (XY Servo off) | 0.381
| 충북대학교 신소재공학과 | 2025-02-25 | Y axis OPM | 0.4394
| 충북대학교 신소재공학과 | 2025-02-25 | X Scan Stroke | 55
| 충북대학교 신소재공학과 | 2025-02-25 | X Detector noise (XY Servo on) | 0.345
| 충북대학교 신소재공학과 | 2025-02-25 | X Detector noise (XY Servo off) | 0.38
| 충북대학교 신소재공학과 | 2025-02-25 | X axis OPM | 0.8246
| 충북대학교 신소재공학과 | 2025-02-25 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 0.501
| 충북대학교 신소재공학과 | 2025-02-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 충북대학교 신소재공학과 | 2025-02-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 충북대학교 신소재공학과 | 2025-02-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| 충북대학교 신소재공학과 | 2025-02-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| 충북대학교 신소재공학과 | 2025-02-25 | NCM driving% with AC160TS cantilever | 0.43
| 충북대학교 신소재공학과 | 2025-02-25 | NCM Amplitude noise | 15.141
| 충북대학교 신소재공학과 | 2025-02-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.435
| 충북대학교 신소재공학과 | 2025-02-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.335
| 충북대학교 신소재공학과 | 2025-02-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 89.3
| 충북대학교 신소재공학과 | 2025-02-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 86.6
| 충북대학교 신소재공학과 | 2025-02-25 | Lateral(C-D) sensitivity | 15.3
| 충북대학교 신소재공학과 | 2025-02-25 | Lateral(C-D) noise | 0.052
| 충북대학교 신소재공학과 | 2025-02-25 | Full range of Z Height during Z Scan sweep | 17.2
| 충북대학교 신소재공학과 | 2025-02-25 | Full range of Z Detector during Z Scan sweep | 12.6
| 충북대학교 신소재공학과 | 2025-02-25 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| 충북대학교 신소재공학과 | 2025-02-25 | Full range of Y Detector during Y Scan sweep | 64
| 충북대학교 신소재공학과 | 2025-02-25 | Full range of X Detector Raw during X Scan sweep | 15.5
| 충북대학교 신소재공학과 | 2025-02-25 | Full range of X Detector during X Scan sweep | 67.3
| 충북대학교 신소재공학과 | 2025-02-25 | A-B noise | 6.809
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Z Scanner position at -3μm | 0.002118
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Z Scanner position at +3μm | 0.003236
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Z linearity by sample height | 1.0125
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Z Height noise | 0.027
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y2 Detector noise (XY Servo on) | 0.255
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y2 Detector noise (XY Servo off) | 0.249
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y1 Detector noise (XY servo on) | 0.26
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y1 Detector noise (XY servo off) | 0.24
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y Scan Stroke | 178.5
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y Detector noise (XY servo on) | 0.189
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y Detector noise (XY servo off) | 0.175
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Y axis OPM | 1.169
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X2 Detector noise (XY Servo on) | 0.249
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X2 Detector noise (XY Servo off) | 0.238
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X1 Detector noise (XY servo on) | 0.333
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X1 Detector noise (XY servo off) | 0.322
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X Scan Stroke | 174
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X Detector noise (XY servo on) | 0.217
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X Detector noise (XY servo off) | 0.206
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | X axis OPM | 1.508
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.93
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.37
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.35
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.32
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | NCM driving% with AC160TS cantilever | 3.504
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | NCM Amplitude noise | 12.945
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.604
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.2
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 251.6
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 242.6
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.8
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 244.6
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Lateral(C-D) sensitivity | 19.2
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Lateral(C-D) noise | 0.098
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Z Height during Z Scan sweep | 17.2
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Z Detector during Z Scan sweep | 12.7
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Y2 Detector during Y Scan sweep | 179.7
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Y1 Detector during Y Scan sweep | 176.9
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of Y Detector during Y Scan sweep | 178.3
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of X2 Detector during X Scan sweep | 173.8
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of X1 Detector during X Scan sweep | 173.6
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | Full range of X Detector during X Scan sweep | 173.7
| Anhui Jinglong Semiconductor Technology Co., Ltd | 2025-02-24 | A-B noise | 10.396
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Z Scanner position at -3μm | 0.006047
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Z Scanner position at +3μm | 0.007093
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Z linearity by sample height | 0.99356
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Z Height noise | 0.023
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Y Scan Stroke | 114.7
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Y Detector noise (XY Servo on) | 0.272
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Y Detector noise (XY Servo off) | 0.265
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Y axis OPM | 1.0236
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | X Scan Stroke | 113.1
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | X Detector noise (XY Servo on) | 0.271
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | X Detector noise (XY Servo off) | 0.269
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | X axis OPM | 0.9988
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | SLD beam power at the end of the lens | 0.213
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | NCM driving% with AC160TS cantilever | 1.665
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | NCM Amplitude noise | 6.813
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.153
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.509
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.3
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Lateral(C-D) sensitivity | 11.4
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Lateral(C-D) noise | 0.109
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Full range of Z Height during Z Scan sweep | 16
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Full range of Z Detector during Z Scan sweep | 14.2
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Full range of Y Detector during Y Scan sweep | 114
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Full range of X Detector Raw during X Scan sweep | 13.8
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | Full range of X Detector during X Scan sweep | 112.1
| The Hong kong University of Science & Technology(HKUST) | 2025-02-20 | A-B noise | 8.281
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Z Scanner position at -3μm | 0.0072
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Z Scanner position at +3μm | 0.013989
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Z linearity by sample height | 0.99855
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Z Height noise | 0.026
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y2 Detector noise (XY Servo on) | 0.272
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y2 Detector noise (XY Servo off) | 0.251
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y1 Detector noise (XY servo on) | 0.277
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y1 Detector noise (XY servo off) | 0.261
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y Scan Stroke | 198.8
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y Detector noise (XY servo on) | 0.203
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y Detector noise (XY servo off) | 0.186
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Y axis OPM | 2.6322
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X2 Detector noise (XY Servo on) | 0.302
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X2 Detector noise (XY Servo off) | 0.289
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X1 Detector noise (XY servo on) | 0.325
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X1 Detector noise (XY servo off) | 0.315
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X Scan Stroke | 205.8
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X Detector noise (XY servo on) | 0.253
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X Detector noise (XY servo off) | 0.242
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | X axis OPM | 3.5354
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 10.367
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.4
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.41
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.34
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.33
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.33
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | NCM driving% with AC160TS cantilever | 0.152
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | NCM Amplitude noise | 14.711
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.533
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.302
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 269.9
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 287.2
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 280
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 283.8
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Lateral(C-D) sensitivity | 3.9
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Lateral(C-D) noise | 0.173
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Z Height during Z Scan sweep | 16.6
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Z Detector during Z Scan sweep | 13.1
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Y2 Detector during Y Scan sweep | 194.3
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Y1 Detector during Y Scan sweep | 203.3
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of Y Detector during Y Scan sweep | 198.8
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of X2 Detector Raw during X Scan sweep | 14.6
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of X2 Detector during X Scan sweep | 204
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of X1 Detector Raw during X Scan sweep | 14.6
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of X1 Detector during X Scan sweep | 207.6
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | Full range of X Detector during X Scan sweep | 205.8
| BEIJING YAN DONG MICROELECTRONIC | 2025-02-17 | A-B noise | 6.584
| Beijing Institute of Technology | 2025-02-14 | Z Scanner position at -3μm | 0.000561
| Beijing Institute of Technology | 2025-02-14 | Z Scanner position at +3μm | 0.000621
| Beijing Institute of Technology | 2025-02-14 | Z linearity by sample height | 1.0094
| Beijing Institute of Technology | 2025-02-14 | Z Height noise | 0.026
| Beijing Institute of Technology | 2025-02-14 | Y Scan Stroke | 113.1
| Beijing Institute of Technology | 2025-02-14 | Y Detector noise (XY Servo on) | 0.273
| Beijing Institute of Technology | 2025-02-14 | Y Detector noise (XY servo off) | 0.284
| Beijing Institute of Technology | 2025-02-14 | Y axis OPM | 1.4458
| Beijing Institute of Technology | 2025-02-14 | X Scan Stroke | 112.9
| Beijing Institute of Technology | 2025-02-14 | X Detector noise (XY Servo on) | 0.281
| Beijing Institute of Technology | 2025-02-14 | X Detector noise (XY servo off) | 0.276
| Beijing Institute of Technology | 2025-02-14 | X axis OPM | 1.4842
| Beijing Institute of Technology | 2025-02-14 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 3.864
| Beijing Institute of Technology | 2025-02-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Beijing Institute of Technology | 2025-02-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Beijing Institute of Technology | 2025-02-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Beijing Institute of Technology | 2025-02-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Beijing Institute of Technology | 2025-02-14 | NCM driving% with AC160TS cantilever | 1.192
| Beijing Institute of Technology | 2025-02-14 | NCM Amplitude noise | 15.001
| Beijing Institute of Technology | 2025-02-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.541
| Beijing Institute of Technology | 2025-02-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.053
| Beijing Institute of Technology | 2025-02-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 162.8
| Beijing Institute of Technology | 2025-02-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167.2
| Beijing Institute of Technology | 2025-02-14 | Lateral(C-D) sensitivity | 7.9
| Beijing Institute of Technology | 2025-02-14 | Lateral(C-D) noise | 0.112
| Beijing Institute of Technology | 2025-02-14 | Full range of Z Height during Z Scan sweep | 16.9
| Beijing Institute of Technology | 2025-02-14 | Full range of Z Detector during Z Scan sweep | 13.5
| Beijing Institute of Technology | 2025-02-14 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Beijing Institute of Technology | 2025-02-14 | Full range of Y Detector during Y Scan sweep | 113.6
| Beijing Institute of Technology | 2025-02-14 | Full range of X Detector Raw during X Scan sweep | 13.6
| Beijing Institute of Technology | 2025-02-14 | Full range of X Detector during X Scan sweep | 113.9
| Beijing Institute of Technology | 2025-02-14 | A-B noise | 8.066
| 한양대학교_물리학과 | 2025-02-13 | Z Scanner position at -3μm | 0.000809
| Bosch Reutilngen | 2025-02-13 | Z Scanner position at -3μm | 0.000412
| INRIM Torino | 2025-02-13 | Z Scanner position at -3μm | 0.002302
| 한양대학교_물리학과 | 2025-02-13 | Z Scanner position at +3μm | 0.003615
| Bosch Reutilngen | 2025-02-13 | Z Scanner position at +3μm | 0.007715
| INRIM Torino | 2025-02-13 | Z Scanner position at +3μm | 0.009654
| 한양대학교_물리학과 | 2025-02-13 | Z linearity by sample height | 0.99158
| Bosch Reutilngen | 2025-02-13 | Z linearity by sample height | 1.0078
| INRIM Torino | 2025-02-13 | Z linearity by sample height | 0.99752
| 한양대학교_물리학과 | 2025-02-13 | Z Height noise | 0.029
| Bosch Reutilngen | 2025-02-13 | Z Height noise | 0.023
| INRIM Torino | 2025-02-13 | Z Height noise | 0.02
| Bosch Reutilngen | 2025-02-13 | Y2 Detector noise (XY Servo on) | 0.212
| Bosch Reutilngen | 2025-02-13 | Y2 Detector noise (XY Servo off) | 0.207
| Bosch Reutilngen | 2025-02-13 | Y1 Detector noise (XY servo on) | 0.219
| Bosch Reutilngen | 2025-02-13 | Y1 Detector noise (XY servo off) | 0.207
| 한양대학교_물리학과 | 2025-02-13 | Y Scan Stroke | 72.6
| Bosch Reutilngen | 2025-02-13 | Y Scan Stroke | 176.9
| INRIM Torino | 2025-02-13 | Y Scan Stroke | 55
| 한양대학교_물리학과 | 2025-02-13 | Y Detector noise (XY Servo on) | 0.276
| Bosch Reutilngen | 2025-02-13 | Y Detector noise (XY servo on) | 0.159
| INRIM Torino | 2025-02-13 | Y Detector noise (XY Servo on) | 0.252
| 한양대학교_물리학과 | 2025-02-13 | Y Detector noise (XY Servo off) | 0.283
| Bosch Reutilngen | 2025-02-13 | Y Detector noise (XY servo off) | 0.15
| INRIM Torino | 2025-02-13 | Y Detector noise (XY Servo off) | 0.246
| Bosch Reutilngen | 2025-02-13 | Y axis OPM (-52000,-52000) | 1.238
| Bosch Reutilngen | 2025-02-13 | Y axis OPM (-52000,52000) | 0.831
| Bosch Reutilngen | 2025-02-13 | Y axis OPM (52000,-52000) | 0.877
| Bosch Reutilngen | 2025-02-13 | Y axis OPM (52000,52000) | 0.642
| Bosch Reutilngen | 2025-02-13 | Y axis OPM (0,0) | 0.744
| 한양대학교_물리학과 | 2025-02-13 | Y axis OPM | 0.5496
| INRIM Torino | 2025-02-13 | Y axis OPM | 0.4892
| Bosch Reutilngen | 2025-02-13 | XY Orthogonality (-52000,-52000) | 90.11
| Bosch Reutilngen | 2025-02-13 | XY Orthogonality (-52000,52000) | 90.23
| Bosch Reutilngen | 2025-02-13 | XY Orthogonality (52000,-52000) | 90.29
| Bosch Reutilngen | 2025-02-13 | XY Orthogonality (52000,52000) | 89.74
| Bosch Reutilngen | 2025-02-13 | XY Orthogonality (0,0) | 89.86
| Bosch Reutilngen | 2025-02-13 | X2 Detector noise (XY Servo on) | 0.222
| Bosch Reutilngen | 2025-02-13 | X2 Detector noise (XY Servo off) | 0.214
| Bosch Reutilngen | 2025-02-13 | X1 Detector noise (XY servo on) | 0.233
| Bosch Reutilngen | 2025-02-13 | X1 Detector noise (XY servo off) | 0.232
| 한양대학교_물리학과 | 2025-02-13 | X Scan Stroke | 71.8
| Bosch Reutilngen | 2025-02-13 | X Scan Stroke | 174.2
| INRIM Torino | 2025-02-13 | X Scan Stroke | 55
| 한양대학교_물리학과 | 2025-02-13 | X Detector noise (XY Servo on) | 0.266
| Bosch Reutilngen | 2025-02-13 | X Detector noise (XY servo on) | 0.172
| INRIM Torino | 2025-02-13 | X Detector noise (XY Servo on) | 0.251
| 한양대학교_물리학과 | 2025-02-13 | X Detector noise (XY Servo off) | 0.279
| Bosch Reutilngen | 2025-02-13 | X Detector noise (XY servo off) | 0.17
| INRIM Torino | 2025-02-13 | X Detector noise (XY Servo off) | 0.245
| Bosch Reutilngen | 2025-02-13 | X axis OPM (-52000,-52000) | 3.423
| Bosch Reutilngen | 2025-02-13 | X axis OPM (-52000,52000) | 3.332
| Bosch Reutilngen | 2025-02-13 | X axis OPM (52000,-52000) | 3.099
| Bosch Reutilngen | 2025-02-13 | X axis OPM (52000,52000) | 2.778
| Bosch Reutilngen | 2025-02-13 | X axis OPM (0,0) | 0.4246
| 한양대학교_물리학과 | 2025-02-13 | X axis OPM | 0.3594
| INRIM Torino | 2025-02-13 | X axis OPM | 0.6226
| Bosch Reutilngen | 2025-02-13 | SLD beam power at the end of the lens | 0.32
| 한양대학교_물리학과 | 2025-02-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.728
| Bosch Reutilngen | 2025-02-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 4.414
| INRIM Torino | 2025-02-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.051
| 한양대학교_물리학과 | 2025-02-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| INRIM Torino | 2025-02-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 한양대학교_물리학과 | 2025-02-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| INRIM Torino | 2025-02-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.22
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.22
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.2
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.17
| Bosch Reutilngen | 2025-02-13 | RMS of Z Height of C122 Zeroscan(0,-90000) | 0.17
| 한양대학교_물리학과 | 2025-02-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Bosch Reutilngen | 2025-02-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| INRIM Torino | 2025-02-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 한양대학교_물리학과 | 2025-02-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Bosch Reutilngen | 2025-02-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| INRIM Torino | 2025-02-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 한양대학교_물리학과 | 2025-02-13 | NCM driving% with AC160TS cantilever | 0.744
| Bosch Reutilngen | 2025-02-13 | NCM driving% with AC160TS cantilever | 8.186
| INRIM Torino | 2025-02-13 | NCM driving% with AC160TS cantilever | 0.763
| 한양대학교_물리학과 | 2025-02-13 | NCM Amplitude noise | 13.042
| Bosch Reutilngen | 2025-02-13 | NCM Amplitude noise | 5.335
| INRIM Torino | 2025-02-13 | NCM Amplitude noise | 17.003
| 한양대학교_물리학과 | 2025-02-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.887
| Bosch Reutilngen | 2025-02-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.48
| INRIM Torino | 2025-02-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.974
| 한양대학교_물리학과 | 2025-02-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.815
| Bosch Reutilngen | 2025-02-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.137
| INRIM Torino | 2025-02-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.61
| Bosch Reutilngen | 2025-02-13 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 251.4
| Bosch Reutilngen | 2025-02-13 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 247
| 한양대학교_물리학과 | 2025-02-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 109.2
| INRIM Torino | 2025-02-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.2
| Bosch Reutilngen | 2025-02-13 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 242.9
| Bosch Reutilngen | 2025-02-13 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 247.7
| 한양대학교_물리학과 | 2025-02-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.9
| INRIM Torino | 2025-02-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100.6
| Bosch Reutilngen | 2025-02-13 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| 한양대학교_물리학과 | 2025-02-13 | Lateral(C-D) sensitivity | 21.9
| Bosch Reutilngen | 2025-02-13 | Lateral(C-D) sensitivity | 12.5
| INRIM Torino | 2025-02-13 | Lateral(C-D) sensitivity | 9.1
| 한양대학교_물리학과 | 2025-02-13 | Lateral(C-D) noise | 0.083
| Bosch Reutilngen | 2025-02-13 | Lateral(C-D) noise | 0.114
| INRIM Torino | 2025-02-13 | Lateral(C-D) noise | 0.097
| 한양대학교_물리학과 | 2025-02-13 | Full range of Z Height during Z Scan sweep | 17.2
| Bosch Reutilngen | 2025-02-13 | Full range of Z Height during Z Scan sweep | 16.7
| INRIM Torino | 2025-02-13 | Full range of Z Height during Z Scan sweep | 17.4
| 한양대학교_물리학과 | 2025-02-13 | Full range of Z Detector during Z Scan sweep | 12.4
| Bosch Reutilngen | 2025-02-13 | Full range of Z Detector during Z Scan sweep | 14.4
| INRIM Torino | 2025-02-13 | Full range of Z Detector during Z Scan sweep | 12.6
| Bosch Reutilngen | 2025-02-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Bosch Reutilngen | 2025-02-13 | Full range of Y2 Detector during Y Scan sweep | 178.8
| Bosch Reutilngen | 2025-02-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Bosch Reutilngen | 2025-02-13 | Full range of Y1 Detector during Y Scan sweep | 174.95
| 한양대학교_물리학과 | 2025-02-13 | Full range of Y Detector Raw during Y Scan sweep | 13.3
| INRIM Torino | 2025-02-13 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| 한양대학교_물리학과 | 2025-02-13 | Full range of Y Detector during Y Scan sweep | 72.6
| Bosch Reutilngen | 2025-02-13 | Full range of Y Detector during Y Scan sweep | 176.85
| INRIM Torino | 2025-02-13 | Full range of Y Detector during Y Scan sweep | 68.4
| Bosch Reutilngen | 2025-02-13 | Full range of X2 Detector Raw during X Scan sweep | 14
| Bosch Reutilngen | 2025-02-13 | Full range of X2 Detector during X Scan sweep | 170.25
| Bosch Reutilngen | 2025-02-13 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Bosch Reutilngen | 2025-02-13 | Full range of X1 Detector during X Scan sweep | 176.15
| 한양대학교_물리학과 | 2025-02-13 | Full range of X Detector Raw during X Scan sweep | 13.8
| INRIM Torino | 2025-02-13 | Full range of X Detector Raw during X Scan sweep | 13.3
| 한양대학교_물리학과 | 2025-02-13 | Full range of X Detector during X Scan sweep | 71.8
| Bosch Reutilngen | 2025-02-13 | Full range of X Detector during X Scan sweep | 173.2
| INRIM Torino | 2025-02-13 | Full range of X Detector during X Scan sweep | 66.7
| 한양대학교_물리학과 | 2025-02-13 | A-B noise | 8.809
| Bosch Reutilngen | 2025-02-13 | A-B noise | 5.039
| INRIM Torino | 2025-02-13 | A-B noise | 10.032
| Huntersun Huzhou | 2025-02-12 | Z Scanner position at -3μm | 0.004075
| Huntersun Huzhou | 2025-02-12 | Z Scanner position at -3μm | 0.006075
| Huntersun Huzhou | 2025-02-12 | Z Scanner position at +3μm | 0.00684
| Huntersun Huzhou | 2025-02-12 | Z Scanner position at +3μm | 0.005405
| Huntersun Huzhou | 2025-02-12 | Z linearity by sample height | 1.0176
| Huntersun Huzhou | 2025-02-12 | Z linearity by sample height | 1.0148
| Huntersun Huzhou | 2025-02-12 | Z Height noise | 0.027
| Huntersun Huzhou | 2025-02-12 | Z Height noise | 0.028
| Huntersun Huzhou | 2025-02-12 | Y2 Detector noise (XY Servo on) | 0.245
| Huntersun Huzhou | 2025-02-12 | Y2 Detector noise (XY Servo off) | 0.241
| Huntersun Huzhou | 2025-02-12 | Y1 Detector noise (XY servo on) | 0.223
| Huntersun Huzhou | 2025-02-12 | Y1 Detector noise (XY servo off) | 0.218
| Huntersun Huzhou | 2025-02-12 | Y Scan Stroke | 167.5
| Huntersun Huzhou | 2025-02-12 | Y Detector noise (XY servo on) | 0.17
| Huntersun Huzhou | 2025-02-12 | Y Detector noise (XY servo off) | 0.164
| Huntersun Huzhou | 2025-02-12 | Y axis OPM | 1.443
| Huntersun Huzhou | 2025-02-12 | X2 Detector noise (XY Servo on) | 0.25
| Huntersun Huzhou | 2025-02-12 | X2 Detector noise (XY Servo off) | 0.251
| Huntersun Huzhou | 2025-02-12 | X1 Detector noise (XY servo on) | 0.255
| Huntersun Huzhou | 2025-02-12 | X1 Detector noise (XY servo off) | 0.255
| Huntersun Huzhou | 2025-02-12 | X Scan Stroke | 178
| Huntersun Huzhou | 2025-02-12 | X Detector noise (XY servo on) | 0.194
| Huntersun Huzhou | 2025-02-12 | X Detector noise (XY servo off) | 0.195
| Huntersun Huzhou | 2025-02-12 | X axis OPM | 1.3514
| Huntersun Huzhou | 2025-02-12 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 0.69
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.28
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.28
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.27
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.27
| Huntersun Huzhou | 2025-02-12 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| Huntersun Huzhou | 2025-02-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Huntersun Huzhou | 2025-02-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Huntersun Huzhou | 2025-02-12 | NCM driving% with AC160TS cantilever | 0.613
| Huntersun Huzhou | 2025-02-12 | NCM driving% with AC160TS cantilever | 0.441
| Huntersun Huzhou | 2025-02-12 | NCM Amplitude noise | 10.6
| Huntersun Huzhou | 2025-02-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.839
| Huntersun Huzhou | 2025-02-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.816
| Huntersun Huzhou | 2025-02-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.778
| Huntersun Huzhou | 2025-02-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.082
| Huntersun Huzhou | 2025-02-12 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 221.9
| Huntersun Huzhou | 2025-02-12 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 233.9
| Huntersun Huzhou | 2025-02-12 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 248.2
| Huntersun Huzhou | 2025-02-12 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 246.3
| Huntersun Huzhou | 2025-02-12 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Huntersun Huzhou | 2025-02-12 | Lateral(C-D) sensitivity | 23.3
| Huntersun Huzhou | 2025-02-12 | Lateral(C-D) noise | 0.14
| Huntersun Huzhou | 2025-02-12 | Full range of Z Height during Z Scan sweep | 16.9
| Huntersun Huzhou | 2025-02-12 | Full range of Z Height during Z Scan sweep | 17.1
| Huntersun Huzhou | 2025-02-12 | Full range of Z Detector during Z Scan sweep | 12.6
| Huntersun Huzhou | 2025-02-12 | Full range of Z Detector during Z Scan sweep | 12.6
| Huntersun Huzhou | 2025-02-12 | Full range of Y2 Detector Raw during Y Scan sweep | 14.9
| Huntersun Huzhou | 2025-02-12 | Full range of Y2 Detector during Y Scan sweep | 165.7
| Huntersun Huzhou | 2025-02-12 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Huntersun Huzhou | 2025-02-12 | Full range of Y1 Detector during Y Scan sweep | 168.7
| Huntersun Huzhou | 2025-02-12 | Full range of Y Detector during Y Scan sweep | 167.1
| Huntersun Huzhou | 2025-02-12 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Huntersun Huzhou | 2025-02-12 | Full range of X2 Detector during X Scan sweep | 178.4
| Huntersun Huzhou | 2025-02-12 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| Huntersun Huzhou | 2025-02-12 | Full range of X1 Detector during X Scan sweep | 177.6
| Huntersun Huzhou | 2025-02-12 | Full range of X Detector during X Scan sweep | 178.1
| Huntersun Huzhou | 2025-02-12 | A-B noise | 5.227
| Federal University of Minas Gerais | 2025-02-11 | Z Scanner position at -3μm | 0.003808
| Federal University of Minas Gerais | 2025-02-11 | Z Scanner position at +3μm | 0.007107
| Federal University of Minas Gerais | 2025-02-11 | Z linearity by sample height | 1.0055
| Federal University of Minas Gerais | 2025-02-11 | Z Height noise | 0.029
| Federal University of Minas Gerais | 2025-02-11 | Y Scan Stroke | 55
| Federal University of Minas Gerais | 2025-02-11 | Y Detector noise (XY Servo on) | 0.394
| Federal University of Minas Gerais | 2025-02-11 | Y Detector noise (XY Servo off) | 0.406
| Federal University of Minas Gerais | 2025-02-11 | Y axis OPM | 0.4314
| Federal University of Minas Gerais | 2025-02-11 | X Scan Stroke | 55
| Federal University of Minas Gerais | 2025-02-11 | X Detector noise (XY Servo on) | 0.407
| Federal University of Minas Gerais | 2025-02-11 | X Detector noise (XY Servo off) | 0.416
| Federal University of Minas Gerais | 2025-02-11 | X axis OPM | 0.398
| Federal University of Minas Gerais | 2025-02-11 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 0.99
| Federal University of Minas Gerais | 2025-02-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Federal University of Minas Gerais | 2025-02-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| Federal University of Minas Gerais | 2025-02-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Federal University of Minas Gerais | 2025-02-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| Federal University of Minas Gerais | 2025-02-11 | NCM driving% with AC160TS cantilever | 0.154
| Federal University of Minas Gerais | 2025-02-11 | NCM Amplitude noise | 11.591
| Federal University of Minas Gerais | 2025-02-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.837
| Federal University of Minas Gerais | 2025-02-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.476
| Federal University of Minas Gerais | 2025-02-11 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.2
| Federal University of Minas Gerais | 2025-02-11 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98
| Federal University of Minas Gerais | 2025-02-11 | Lateral(C-D) sensitivity | 18
| Federal University of Minas Gerais | 2025-02-11 | Lateral(C-D) noise | 0.063
| Federal University of Minas Gerais | 2025-02-11 | Full range of Z Height during Z Scan sweep | 17.4
| Federal University of Minas Gerais | 2025-02-11 | Full range of Z Detector during Z Scan sweep | 12.7
| Federal University of Minas Gerais | 2025-02-11 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Federal University of Minas Gerais | 2025-02-11 | Full range of Y Detector during Y Scan sweep | 68
| Federal University of Minas Gerais | 2025-02-11 | Full range of X Detector Raw during X Scan sweep | 14
| Federal University of Minas Gerais | 2025-02-11 | Full range of X Detector during X Scan sweep | 68.9
| Federal University of Minas Gerais | 2025-02-11 | A-B noise | 6.992
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Linearity by Z scanner position | 0.74
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Linearity by 4 different sample height | 0.998
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Height noise | 0.027
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7953
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7627
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6148
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6065
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Z Drift | 18.387
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y2 Detector noise (XY Servo on) | 0.372
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 235.277
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y1 Detector noise (XY Servo on) | 0.456
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 233.464
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 155.034
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y Ringing(Rpv) | 0.176
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y Non-linearity error | 1.73
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y Detector noise (XY Servo on) | 0.346
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Repeatability @1st position | 0.04
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Repeatability @ 4th position | 0.05
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Repeatability @ 3rd position | 0.04
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Repeatability @ 2nd position | 0.04
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Accuracy @1st position | 0.24
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Accuracy @ 4th position | 0.3
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Accuracy @ 3rd position | 0.29
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis, Accuracy @ 2nd position | 0.24
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Y axis OPM | 0.567
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.168
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning Y offset (LLC Translation) | 1.9385
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning Y offset (In-Die Align) | 4.4971
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning Y offset (Global Align) | 3.2373
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning Y 1σ (LLC Translation) | 0.4166
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning X offset (Vision Pattern Matching) | 2.0117
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning X offset (LLC Translation) | 1.6797
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning X offset (In-Die Align) | 3.3105
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning X offset (Global Align) | 1.7822
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | XY Stage positioning X 1σ (LLC Translation) | 0.3568
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X2 Detector noise (XY Servo on) | 0.336
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 249.999
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X1 Detector noise (XY Servo on) | 0.321
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 254.667
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 164.93
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X Ringing(Rpv) | 0.121
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X Non-linearity error | 3.09
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X Detector noise (XY Servo on) | 0.227
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Repeatability @1st position | 0.03
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Repeatability @ 4th position | 0.04
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Repeatability @ 3rd position | 0.04
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Repeatability @ 2nd position | 0.04
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Accuracy @1st position | 0.13
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Accuracy @ 4th position | 0.27
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Accuracy @ 3rd position | 0.3
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis, Accuracy @ 2nd position | 0.19
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | X axis OPM | 0.448
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Wafer placement(Loading) error on the load port2 (Y offset) | 40.77
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Wafer placement(Loading) error on the load port2 (X offset) | 38.7
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Wafer placement(Loading) error on the load port1 (Y offset) | 17.66
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Wafer placement(Loading) error on the load port1 (X offset) | 42.7
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Torque Profile (50mm, 20um/s) | 0.873
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Torque Profile (50mm, 200um/s) | 0.717
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7604
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Roughness Repeatability 80x80um(RSD%) | 0.614
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Roughness Repeatability 1x1um(RSD%) | 0.5
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Roll motion to X+ full scan | 5.8
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | RMS of ZeroScan | 0.064
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | RMS error after Stage Calibration | 2.57
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Profiler noise @200um | 0.112
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Position repeatability(10repeats, Position Based) | 0.68
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00039
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OpticalAxisShiftPerFocusY | -0.0017
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00019
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OpticalAxisShiftPerFocusX | -0.0048
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OPM Repeatability @5mm | 1.182
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OPM Repeatability @25mm | 3.041
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OPM Repeatability @1mm | 0.839
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | OPM Repeatability @10mm | 1.585
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Min. Vibration Transmissbility (~100Hz) | 0.02
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Max OPM @5mm | 29.863
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Max OPM @25mm | 153.567
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Max OPM @1mm | 5.296
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Max OPM @10mm | 32.67
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Lift pins synchronization offset check (Y offset) | 44.45
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Lift pins synchronization offset check (X offset) | 2.86
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Lateral(C-D) sensitivity | 24
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Lateral(C-D) noise | 0.154
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from -Y to +Y (Y scan) | 46.875
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from -Y to +Y (X scan) | 30.0293
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from -X to +X (Y scan) | 82.0313
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from -X to +X (X scan) | 40.2832
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from +Y to -Y (Y scan) | 49.0723
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from +Y to -Y (X scan) | 48.3398
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from +X to -X (Y scan) | 48.3398
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Image offset after moving from +X to -X (X scan) | 37.3535
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Z Height during Z Scan sweep | 17.4594
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Z Detector during Z Scan sweep | 13.2
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Y2 Detector Raw during Y Scan sweep | 14.59
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Y2 Detector during Y Scan sweep | 189.4132
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Y1 Detector Raw during Y Scan sweep | 14.67
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Y1 Detector during Y Scan sweep | 179.832
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of Y Detector during Y Scan sweep | 184.6226
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of X2 Detector during X Scan sweep | 194.3621
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of X1 Detector Raw during X Scan sweep | 14.61
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of X1 Detector during X Scan sweep | 205.5267
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Full range of X Detector during X Scan sweep | 200.1697
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | Cycle time for 15-point measurement | 1617.116
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.76
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.61
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| Samsung (SEC) NRD-K (DRAM) | 2025-02-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Innolume GmbH | 2025-02-07 | Z Scanner position at -3μm | 0.003384
| Innolume GmbH | 2025-02-07 | Z Scanner position at +3μm | 0.009589
| Innolume GmbH | 2025-02-07 | Z linearity by sample height | 1.0011
| Innolume GmbH | 2025-02-07 | Y2 Detector noise (XY Servo on) | 0.226
| Innolume GmbH | 2025-02-07 | Y2 Detector noise (XY Servo off) | 0.224
| Innolume GmbH | 2025-02-07 | Y1 Detector noise (XY servo on) | 0.227
| Innolume GmbH | 2025-02-07 | Y1 Detector noise (XY servo off) | 0.224
| Innolume GmbH | 2025-02-07 | Y Scan Stroke | 168.9
| Innolume GmbH | 2025-02-07 | Y Detector noise (XY servo on) | 0.171
| Innolume GmbH | 2025-02-07 | Y Detector noise (XY servo off) | 0.167
| Innolume GmbH | 2025-02-07 | Y axis OPM | 1.4088
| Innolume GmbH | 2025-02-07 | X2 Detector noise (XY Servo on) | 0.24
| Innolume GmbH | 2025-02-07 | X2 Detector noise (XY Servo off) | 0.236
| Innolume GmbH | 2025-02-07 | X1 Detector noise (XY servo on) | 0.234
| Innolume GmbH | 2025-02-07 | X1 Detector noise (XY servo off) | 0.23
| Innolume GmbH | 2025-02-07 | X Scan Stroke | 175.7
| Innolume GmbH | 2025-02-07 | X Detector noise (XY servo on) | 0.179
| Innolume GmbH | 2025-02-07 | X Detector noise (XY servo off) | 0.172
| Innolume GmbH | 2025-02-07 | X axis OPM | 0.5228
| Innolume GmbH | 2025-02-07 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 0.332
| Innolume GmbH | 2025-02-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| Innolume GmbH | 2025-02-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Innolume GmbH | 2025-02-07 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.36
| Innolume GmbH | 2025-02-07 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.36
| Innolume GmbH | 2025-02-07 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.35
| Innolume GmbH | 2025-02-07 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.36
| Innolume GmbH | 2025-02-07 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.35
| Innolume GmbH | 2025-02-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Innolume GmbH | 2025-02-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Innolume GmbH | 2025-02-07 | NCM driving% with AC160TS cantilever | 2.18
| Innolume GmbH | 2025-02-07 | NCM Amplitude noise | 15.016
| Innolume GmbH | 2025-02-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.625
| Innolume GmbH | 2025-02-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.603
| Innolume GmbH | 2025-02-07 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 238.8
| Innolume GmbH | 2025-02-07 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 235.4
| Innolume GmbH | 2025-02-07 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 249.2
| Innolume GmbH | 2025-02-07 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 243.9
| Innolume GmbH | 2025-02-07 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Innolume GmbH | 2025-02-07 | Lateral(C-D) sensitivity | 25.1
| Innolume GmbH | 2025-02-07 | Lateral(C-D) noise | 0.166
| Innolume GmbH | 2025-02-07 | Full range of Z Height during Z Scan sweep | 17
| Innolume GmbH | 2025-02-07 | Full range of Z Detector during Z Scan sweep | 13.3
| Innolume GmbH | 2025-02-07 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| Innolume GmbH | 2025-02-07 | Full range of Y2 Detector during Y Scan sweep | 168.1
| Innolume GmbH | 2025-02-07 | Full range of Y1 Detector Raw during Y Scan sweep | 14.7
| Innolume GmbH | 2025-02-07 | Full range of Y1 Detector during Y Scan sweep | 173
| Innolume GmbH | 2025-02-07 | Full range of Y Detector during Y Scan sweep | 170.6
| Innolume GmbH | 2025-02-07 | Full range of X2 Detector Raw during X Scan sweep | 14.6
| Innolume GmbH | 2025-02-07 | Full range of X2 Detector during X Scan sweep | 181.6
| Innolume GmbH | 2025-02-07 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| Innolume GmbH | 2025-02-07 | Full range of X1 Detector during X Scan sweep | 175.2
| Innolume GmbH | 2025-02-07 | Full range of X Detector during X Scan sweep | 178.4
| Innolume GmbH | 2025-02-07 | A-B noise | 10.405
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.3965
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y offset (LLC Translation) | 1.9775
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y offset (In-Die Align) | 1.7285
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y offset (Global Align) | 3.3838
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.29035
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y 1σ (LLC Translation) | 0.43799
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y 1σ (In-Die Align) | 0.33573
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning Y 1σ (Global Align) | 0.99024
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X offset (Vision Pattern Matching) | 1.8701
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X offset (LLC Translation) | 1.2256
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X offset (In-Die Align) | 1.416
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X offset (Global Align) | 2.0361
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.30442
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X 1σ (LLC Translation) | 0.23753
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X 1σ (In-Die Align) | 0.23716
| TSMC AP2 Mini Line (4) | 2025-02-07 | XY Stage positioning X 1σ (Global Align) | 0.53008
| TIE Cerium Laboratories | 2025-02-05 | Z Scanner position at -3μm | 0.00514
| TIE Cerium Laboratories | 2025-02-05 | Z Scanner position at +3μm | 0.017795
| TIE Cerium Laboratories | 2025-02-05 | Z linearity by sample height | 0.99288
| TIE Cerium Laboratories | 2025-02-05 | Z Height noise | 0.018
| TIE Cerium Laboratories | 2025-02-05 | Y2 Detector noise (XY Servo on) | 0.23
| TIE Cerium Laboratories | 2025-02-05 | Y2 Detector noise (XY Servo off) | 0.209
| TIE Cerium Laboratories | 2025-02-05 | Y1 Detector noise (XY servo on) | 0.214
| TIE Cerium Laboratories | 2025-02-05 | Y1 Detector noise (XY servo off) | 0.192
| TIE Cerium Laboratories | 2025-02-05 | Y Scan Stroke | 171.014
| TIE Cerium Laboratories | 2025-02-05 | Y Detector noise (XY servo on) | 0.17
| TIE Cerium Laboratories | 2025-02-05 | Y Detector noise (XY servo off) | 0.149
| TIE Cerium Laboratories | 2025-02-05 | Y axis OPM (-52000,-52000) | 3.062
| TIE Cerium Laboratories | 2025-02-05 | Y axis OPM (-52000,52000) | 3.313
| TIE Cerium Laboratories | 2025-02-05 | Y axis OPM (52000,-52000) | 3.875
| TIE Cerium Laboratories | 2025-02-05 | Y axis OPM (52000,52000) | 3.142
| TIE Cerium Laboratories | 2025-02-05 | Y axis OPM (0,0) | 0.5871
| TIE Cerium Laboratories | 2025-02-05 | XY Orthogonality (-52000,-52000) | 90.03
| TIE Cerium Laboratories | 2025-02-05 | XY Orthogonality (-52000,52000) | 89.74
| TIE Cerium Laboratories | 2025-02-05 | XY Orthogonality (52000,-52000) | 89.85
| TIE Cerium Laboratories | 2025-02-05 | XY Orthogonality (52000,52000) | 89.94
| TIE Cerium Laboratories | 2025-02-05 | XY Orthogonality (0,0) | 89.93
| TIE Cerium Laboratories | 2025-02-05 | X2 Detector noise (XY Servo on) | 0.22
| TIE Cerium Laboratories | 2025-02-05 | X2 Detector noise (XY Servo off) | 0.202
| TIE Cerium Laboratories | 2025-02-05 | X1 Detector noise (XY servo on) | 0.215
| TIE Cerium Laboratories | 2025-02-05 | X1 Detector noise (XY servo off) | 0.199
| TIE Cerium Laboratories | 2025-02-05 | X Scan Stroke | 171.432
| TIE Cerium Laboratories | 2025-02-05 | X Detector noise (XY servo on) | 0.165
| TIE Cerium Laboratories | 2025-02-05 | X Detector noise (XY servo off) | 0.148
| TIE Cerium Laboratories | 2025-02-05 | X axis OPM (-52000,-52000) | 1.62
| TIE Cerium Laboratories | 2025-02-05 | X axis OPM (-52000,52000) | 3.052
| TIE Cerium Laboratories | 2025-02-05 | X axis OPM (52000,-52000) | 1.815
| TIE Cerium Laboratories | 2025-02-05 | X axis OPM (52000,52000) | 2.731
| TIE Cerium Laboratories | 2025-02-05 | X axis OPM (0,0) | 1.6494
| TIE Cerium Laboratories | 2025-02-05 | SLD beam power at the end of the lens | 0.256
| TIE Cerium Laboratories | 2025-02-05 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.284
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.18
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.18
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.16
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.15
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.15
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| TIE Cerium Laboratories | 2025-02-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| TIE Cerium Laboratories | 2025-02-05 | NCM driving% with AC160TS cantilever | 2.812
| TIE Cerium Laboratories | 2025-02-05 | NCM Amplitude noise | 6.192
| TIE Cerium Laboratories | 2025-02-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.744
| TIE Cerium Laboratories | 2025-02-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.995
| TIE Cerium Laboratories | 2025-02-05 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 231.5
| TIE Cerium Laboratories | 2025-02-05 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 230.7
| TIE Cerium Laboratories | 2025-02-05 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.5
| TIE Cerium Laboratories | 2025-02-05 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 232.7
| TIE Cerium Laboratories | 2025-02-05 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| TIE Cerium Laboratories | 2025-02-05 | Lateral(C-D) sensitivity | 10.9
| TIE Cerium Laboratories | 2025-02-05 | Lateral(C-D) noise | 0.106
| TIE Cerium Laboratories | 2025-02-05 | Full range of Z Height during Z Scan sweep | 16.8
| TIE Cerium Laboratories | 2025-02-05 | Full range of Z Detector during Z Scan sweep | 14.6
| TIE Cerium Laboratories | 2025-02-05 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| TIE Cerium Laboratories | 2025-02-05 | Full range of Y2 Detector during Y Scan sweep | 170.5
| TIE Cerium Laboratories | 2025-02-05 | Full range of Y1 Detector Raw during Y Scan sweep | 14.8
| TIE Cerium Laboratories | 2025-02-05 | Full range of Y1 Detector during Y Scan sweep | 170.4
| TIE Cerium Laboratories | 2025-02-05 | Full range of Y Detector during Y Scan sweep | 170.5
| TIE Cerium Laboratories | 2025-02-05 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| TIE Cerium Laboratories | 2025-02-05 | Full range of X2 Detector during X Scan sweep | 174.6
| TIE Cerium Laboratories | 2025-02-05 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| TIE Cerium Laboratories | 2025-02-05 | Full range of X1 Detector during X Scan sweep | 167
| TIE Cerium Laboratories | 2025-02-05 | Full range of X Detector during X Scan sweep | 170.8
| TIE Cerium Laboratories | 2025-02-05 | A-B noise | 7.408
| PSE demo lab | 2025-02-03 | Z Scanner position at -3μm | 0.008251
| PSE demo lab | 2025-02-03 | Z Scanner position at +3μm | 0.001754
| PSE demo lab | 2025-02-03 | Z linearity by sample height | 0.98174
| PSE demo lab | 2025-02-03 | Z Height noise | 0.028
| PSE demo lab | 2025-02-03 | Y Scan Stroke | 55
| PSE demo lab | 2025-02-03 | Y Detector noise (XY Servo on) | 0.238
| PSE demo lab | 2025-02-03 | Y Detector noise (XY Servo off) | 0.232
| PSE demo lab | 2025-02-03 | Y axis OPM | 0.3078
| PSE demo lab | 2025-02-03 | X Scan Stroke | 55
| PSE demo lab | 2025-02-03 | X Detector noise (XY Servo on) | 0.245
| PSE demo lab | 2025-02-03 | X Detector noise (XY Servo off) | 0.237
| PSE demo lab | 2025-02-03 | X axis OPM | 0.8688
| PSE demo lab | 2025-02-03 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 2.371
| PSE demo lab | 2025-02-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| PSE demo lab | 2025-02-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| PSE demo lab | 2025-02-03 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| PSE demo lab | 2025-02-03 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| PSE demo lab | 2025-02-03 | NCM driving% with AC160TS cantilever | 0.749
| PSE demo lab | 2025-02-03 | NCM Amplitude noise | 16.666
| PSE demo lab | 2025-02-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.692
| PSE demo lab | 2025-02-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.103
| PSE demo lab | 2025-02-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.7
| PSE demo lab | 2025-02-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.3
| PSE demo lab | 2025-02-03 | Lateral(C-D) sensitivity | 7.4
| PSE demo lab | 2025-02-03 | Lateral(C-D) noise | 0.073
| PSE demo lab | 2025-02-03 | Full range of Z Height during Z Scan sweep | 17.5
| PSE demo lab | 2025-02-03 | Full range of Z Detector during Z Scan sweep | 12.9
| PSE demo lab | 2025-02-03 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| PSE demo lab | 2025-02-03 | Full range of Y Detector during Y Scan sweep | 64.5
| PSE demo lab | 2025-02-03 | Full range of X Detector Raw during X Scan sweep | 14.4
| PSE demo lab | 2025-02-03 | Full range of X Detector during X Scan sweep | 68.6
| PSE demo lab | 2025-02-03 | A-B noise | 8.867
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Linearity by Z scanner position | 0.97
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Height noise | 0.025
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Drift | 14.01
| SK하이닉스 M11 (Flash) | 2025-02-03 | Z Detector offset | 3.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y2 Detector noise (XY Servo on) | 0.46
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 261.3
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y1 Detector noise (XY Servo on) | 0.44
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 249.6
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 165.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y Ringing(Rpv) | 0.12
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y Non-linearity error | 2.43
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y Detector noise (XY Servo on) | 0.33
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Repeatability @1st position | 0.05
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Repeatability @ 4th position | 0.04
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Repeatability @ 3rd position | 0.05
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Repeatability @ 2nd position | 0.04
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Accuracy @1st position | 0.22
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Accuracy @ 4th position | 0.22
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Accuracy @ 3rd position | 0.24
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis, Accuracy @ 2nd position | 0.22
| SK하이닉스 M11 (Flash) | 2025-02-03 | Y axis OPM | 0.83
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.6758
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y offset (LLC Translation) | 1.9531
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y offset (In-Die Align) | 1.1084
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y offset (Global Align) | 4.6826
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.49478
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y 1σ (LLC Translation) | 0.4268
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y 1σ (In-Die Align) | 0.22293
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning Y 1σ (Global Align) | 1.189
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X offset (Vision Pattern Matching) | 2.6514
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X offset (LLC Translation) | 1.6162
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X offset (In-Die Align) | 1.3428
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X offset (Global Align) | 1.9824
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.47812
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X 1σ (LLC Translation) | 0.3458
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X 1σ (In-Die Align) | 0.25808
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Stage positioning X 1σ (Global Align) | 0.46854
| SK하이닉스 M11 (Flash) | 2025-02-03 | XY Orthogonality | 90
| SK하이닉스 M11 (Flash) | 2025-02-03 | X2 Detector noise (XY Servo on) | 0.45
| SK하이닉스 M11 (Flash) | 2025-02-03 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 258.5
| SK하이닉스 M11 (Flash) | 2025-02-03 | X1 Detector noise (XY Servo on) | 0.44
| SK하이닉스 M11 (Flash) | 2025-02-03 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 267.5
| SK하이닉스 M11 (Flash) | 2025-02-03 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 174.1
| SK하이닉스 M11 (Flash) | 2025-02-03 | X Ringing(Rpv) | 0.11
| SK하이닉스 M11 (Flash) | 2025-02-03 | X Non-linearity error | 3.11
| SK하이닉스 M11 (Flash) | 2025-02-03 | X Detector noise (XY Servo on) | 0.33
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Repeatability @1st position | 0.05
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Repeatability @ 4th position | 0.05
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Repeatability @ 3rd position | 0.09
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Repeatability @ 2nd position | 0.05
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Accuracy @1st position | 0.28
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Accuracy @ 4th position | 0.28
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Accuracy @ 3rd position | 0.46
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis, Accuracy @ 2nd position | 0.23
| SK하이닉스 M11 (Flash) | 2025-02-03 | X axis OPM | 0.77
| SK하이닉스 M11 (Flash) | 2025-02-03 | Wafer placement(Loading) error on the load port2 (Y offset) | 12.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Wafer placement(Loading) error on the load port2 (X offset) | 10.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Wafer placement(Loading) error on the load port1 (Y offset) | 15.7
| SK하이닉스 M11 (Flash) | 2025-02-03 | Wafer placement(Loading) error on the load port1 (X offset) | 14.6
| SK하이닉스 M11 (Flash) | 2025-02-03 | Torque Profile (50mm, 20um/s) | 0.953
| SK하이닉스 M11 (Flash) | 2025-02-03 | Torque Profile (50mm, 200um/s) | 0.916
| SK하이닉스 M11 (Flash) | 2025-02-03 | StrokeYMin | -153043
| SK하이닉스 M11 (Flash) | 2025-02-03 | StrokeYMax | 150295
| SK하이닉스 M11 (Flash) | 2025-02-03 | StrokeY | 303338
| SK하이닉스 M11 (Flash) | 2025-02-03 | StrokeXMin | -204573
| SK하이닉스 M11 (Flash) | 2025-02-03 | StrokeXMax | 203434
| SK하이닉스 M11 (Flash) | 2025-02-03 | StrokeX | 408007
| SK하이닉스 M11 (Flash) | 2025-02-03 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.75
| SK하이닉스 M11 (Flash) | 2025-02-03 | Roughness Repeatability 80x80um(1σ) | 0.002333
| SK하이닉스 M11 (Flash) | 2025-02-03 | Roughness Repeatability 80x80um(RSD%) | 1.47
| SK하이닉스 M11 (Flash) | 2025-02-03 | Roughness Repeatability 1x1um(1σ) | 0.004333
| SK하이닉스 M11 (Flash) | 2025-02-03 | Roughness Repeatability 1x1um(RSD%) | 3.137
| SK하이닉스 M11 (Flash) | 2025-02-03 | Roll motion to X+ full scan | 9.6
| SK하이닉스 M11 (Flash) | 2025-02-03 | RMS of ZeroScan | 0.051
| SK하이닉스 M11 (Flash) | 2025-02-03 | Position repeatability(10repeats, Position Based) | 0.48
| SK하이닉스 M11 (Flash) | 2025-02-03 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00018
| SK하이닉스 M11 (Flash) | 2025-02-03 | OpticalAxisShiftPerFocusY | -0.001
| SK하이닉스 M11 (Flash) | 2025-02-03 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00029
| SK하이닉스 M11 (Flash) | 2025-02-03 | OpticalAxisShiftPerFocusX | -0.002
| SK하이닉스 M11 (Flash) | 2025-02-03 | OPM Repeatability @5mm | 0.75
| SK하이닉스 M11 (Flash) | 2025-02-03 | OPM Repeatability @25mm | 1.33
| SK하이닉스 M11 (Flash) | 2025-02-03 | OPM Repeatability @1mm | 0.6
| SK하이닉스 M11 (Flash) | 2025-02-03 | OPM Repeatability @10mm | 1.17
| SK하이닉스 M11 (Flash) | 2025-02-03 | NCM driving% with AC160TS cantilever | 0.22
| SK하이닉스 M11 (Flash) | 2025-02-03 | Min. Vibration Transmissbility (~100Hz) | 0.034
| SK하이닉스 M11 (Flash) | 2025-02-03 | Max OPM @5mm | 35.7
| SK하이닉스 M11 (Flash) | 2025-02-03 | Max OPM @25mm | 81.82
| SK하이닉스 M11 (Flash) | 2025-02-03 | Max OPM @1mm | 11.47
| SK하이닉스 M11 (Flash) | 2025-02-03 | Max OPM @10mm | 52.83
| SK하이닉스 M11 (Flash) | 2025-02-03 | LowerLimit | 2200.94
| SK하이닉스 M11 (Flash) | 2025-02-03 | LowerLimit | 4632
| SK하이닉스 M11 (Flash) | 2025-02-03 | Lateral(C-D) sensitivity | 16
| SK하이닉스 M11 (Flash) | 2025-02-03 | Lateral(C-D) noise | 0.199
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from -Y to +Y (Y scan) | 52.73
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from -Y to +Y (X scan) | 5.86
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from -X to +X (Y scan) | 65.92
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from -X to +X (X scan) | 59.33
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from +Y to -Y (Y scan) | 57.13
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from +Y to -Y (X scan) | 18.31
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from +X to -X (Y scan) | 38.09
| SK하이닉스 M11 (Flash) | 2025-02-03 | Image offset after moving from +X to -X (X scan) | 43.95
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Z Height during Z Scan sweep | 17.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Z Detector during Z Scan sweep | 12.9
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Y2 Detector Raw during Y Scan sweep | 14.59
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Y2 Detector during Y Scan sweep | 199.8
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Y1 Detector Raw during Y Scan sweep | 14.44
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Y1 Detector during Y Scan sweep | 191.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of Y Detector during Y Scan sweep | 191.2
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of X2 Detector Raw during X Scan sweep | 14.33
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of X2 Detector during X Scan sweep | 193.9
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of X1 Detector Raw during X Scan sweep | 14.53
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of X1 Detector during X Scan sweep | 209.4
| SK하이닉스 M11 (Flash) | 2025-02-03 | Full range of X Detector during X Scan sweep | 193.9
| SK하이닉스 M11 (Flash) | 2025-02-03 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.7
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.01
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| SK하이닉스 M11 (Flash) | 2025-02-03 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Linearity by Z scanner position | 1.1
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Height noise | 0.028
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 115.8
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 163.8
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 114.6
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 160.9
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Drift | 28.2
| TSMC AP2 Mini Line(3) | 2025-01-31 | Z Detector offset | 118.1
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y2 Detector noise (XY Servo on)
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y1 Detector noise (XY Servo on)
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 114.6
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y Ringing(Rpv) | 0.43
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y Non-linearity error
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y Detector noise (XY Servo on) | 0.31
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 160.9
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Repeatability @1st position | 0.05
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Repeatability @ 4th position | 0.08
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Repeatability @ 3rd position | 0.04
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Repeatability @ 2nd position | 0.05
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Accuracy @1st position | 0.43
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Accuracy @ 4th position | 0.68
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Accuracy @ 3rd position | 0.42
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis, Accuracy @ 2nd position | 0.37
| TSMC AP2 Mini Line(3) | 2025-01-31 | Y axis OPM | 1.62
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.8125
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y offset (LLC Translation) | 2.5732
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y offset (In-Die Align) | 2.3291
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y offset (Global Align) | 7.1973
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.5972
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y 1σ (LLC Translation) | 0.6871
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y 1σ (In-Die Align) | 0.47848
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning Y 1σ (Global Align) | 2.5024
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X offset (Vision Pattern Matching) | 2.4707
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X offset (LLC Translation) | 1.6748
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X offset (In-Die Align) | 2.9541
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X offset (Global Align) | 4.1553
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.53613
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X 1σ (LLC Translation) | 0.3251
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X 1σ (In-Die Align) | 0.78468
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Stage positioning X 1σ (Global Align) | 1.13
| TSMC AP2 Mini Line(3) | 2025-01-31 | XY Orthogonality | 90.01
| TSMC AP2 Mini Line(3) | 2025-01-31 | X2 Detector noise (XY Servo on)
| TSMC AP2 Mini Line(3) | 2025-01-31 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
| TSMC AP2 Mini Line(3) | 2025-01-31 | X1 Detector noise (XY Servo on)
| TSMC AP2 Mini Line(3) | 2025-01-31 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 115.8
| TSMC AP2 Mini Line(3) | 2025-01-31 | X Ringing(Rpv) | 0.3
| TSMC AP2 Mini Line(3) | 2025-01-31 | X Non-linearity error
| TSMC AP2 Mini Line(3) | 2025-01-31 | X Detector noise (XY Servo on) | 0.28
| TSMC AP2 Mini Line(3) | 2025-01-31 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 163.8
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Repeatability @1st position | 0.09
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Repeatability @ 4th position | 0.03
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Repeatability @ 3rd position | 0.05
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Repeatability @ 2nd position | 0.04
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Accuracy @1st position | 0.42
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Accuracy @ 4th position | 0.2
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Accuracy @ 3rd position | 0.31
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis, Accuracy @ 2nd position | 0.2
| TSMC AP2 Mini Line(3) | 2025-01-31 | X axis OPM | 0.96
| TSMC AP2 Mini Line(3) | 2025-01-31 | Wafer placement(Loading) error on the load port2 (Y offset) | 39.3
| TSMC AP2 Mini Line(3) | 2025-01-31 | Wafer placement(Loading) error on the load port2 (X offset) | 10.9
| TSMC AP2 Mini Line(3) | 2025-01-31 | Wafer placement(Loading) error on the load port1 (Y offset) | 51.5
| TSMC AP2 Mini Line(3) | 2025-01-31 | Wafer placement(Loading) error on the load port1 (X offset) | 14.1
| TSMC AP2 Mini Line(3) | 2025-01-31 | Torque Profile (50mm, 20um/s)
| TSMC AP2 Mini Line(3) | 2025-01-31 | Torque Profile (50mm, 200um/s)
| TSMC AP2 Mini Line(3) | 2025-01-31 | StrokeYMin | -152987
| TSMC AP2 Mini Line(3) | 2025-01-31 | StrokeYMax | 150600.4
| TSMC AP2 Mini Line(3) | 2025-01-31 | StrokeY | 303587
| TSMC AP2 Mini Line(3) | 2025-01-31 | StrokeXMin | -226589
| TSMC AP2 Mini Line(3) | 2025-01-31 | StrokeXMax | 225823.6
| TSMC AP2 Mini Line(3) | 2025-01-31 | StrokeX | 452412.1
| TSMC AP2 Mini Line(3) | 2025-01-31 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.87
| TSMC AP2 Mini Line(3) | 2025-01-31 | Roughness Repeatability 80x80um(1σ) | 0.002333
| TSMC AP2 Mini Line(3) | 2025-01-31 | Roughness Repeatability 80x80um(RSD%) | 0.57
| TSMC AP2 Mini Line(3) | 2025-01-31 | Roughness Repeatability 1x1um(1σ) | 0.002333
| TSMC AP2 Mini Line(3) | 2025-01-31 | Roughness Repeatability 1x1um(RSD%) | 1.168
| TSMC AP2 Mini Line(3) | 2025-01-31 | Roll motion to X+ full scan
| TSMC AP2 Mini Line(3) | 2025-01-31 | RMS of ZeroScan | 0.103
| TSMC AP2 Mini Line(3) | 2025-01-31 | Position repeatability(10repeats, Position Based)
| TSMC AP2 Mini Line(3) | 2025-01-31 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00014
| TSMC AP2 Mini Line(3) | 2025-01-31 | OpticalAxisShiftPerFocusY | 0.0032
| TSMC AP2 Mini Line(3) | 2025-01-31 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00042
| TSMC AP2 Mini Line(3) | 2025-01-31 | OpticalAxisShiftPerFocusX | -0.0028
| TSMC AP2 Mini Line(3) | 2025-01-31 | OPM Repeatability @5mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | OPM Repeatability @25mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | OPM Repeatability @1mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | OPM Repeatability @10mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | NCM driving% with AC160TS cantilever | 0.6
| TSMC AP2 Mini Line(3) | 2025-01-31 | Min. Vibration Transmissbility (~100Hz) | 0.115
| TSMC AP2 Mini Line(3) | 2025-01-31 | Max OPM @5mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | Max OPM @25mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | Max OPM @1mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | Max OPM @10mm
| TSMC AP2 Mini Line(3) | 2025-01-31 | LowerLimit | 2380.56
| TSMC AP2 Mini Line(3) | 2025-01-31 | LowerLimit | 6352
| TSMC AP2 Mini Line(3) | 2025-01-31 | Lateral(C-D) sensitivity
| TSMC AP2 Mini Line(3) | 2025-01-31 | Lateral(C-D) noise
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from -Y to +Y (Y scan) | 167.73
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from -Y to +Y (X scan) | 136.96
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from -X to +X (Y scan) | 172.12
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from -X to +X (X scan) | 25.63
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from +Y to -Y (Y scan) | 41.75
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from +Y to -Y (X scan) | 161.13
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from +X to -X (Y scan) | 18.31
| TSMC AP2 Mini Line(3) | 2025-01-31 | Image offset after moving from +X to -X (X scan) | 80.57
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Z Height during Z Scan sweep | 18.9
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Z Detector during Z Scan sweep | 13.32
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Y2 Detector Raw during Y Scan sweep
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Y2 Detector during Y Scan sweep
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Y1 Detector during Y Scan sweep
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of Y Detector during Y Scan sweep | 112.7
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of X2 Detector Raw during X Scan sweep
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of X2 Detector during X Scan sweep
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of X1 Detector during X Scan sweep
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of X Detector Raw during X Scan sweep | 14
| TSMC AP2 Mini Line(3) | 2025-01-31 | Full range of X Detector during X Scan sweep | 115.3
| TSMC AP2 Mini Line(3) | 2025-01-31 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.16
| TSMC AP2 Mini Line(3) | 2025-01-31 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Tianjin Medical University | 2025-01-21 | Z Scanner position at -3μm | 0.000408
| Tianjin Medical University | 2025-01-21 | Z Scanner position at +3μm | 0.00266
| Tianjin Medical University | 2025-01-21 | Z linearity by sample height | 0.99463
| Tianjin Medical University | 2025-01-21 | Z Height noise | 0.028
| Tianjin Medical University | 2025-01-21 | Y Scan Stroke | 185.6
| Tianjin Medical University | 2025-01-21 | Y Detector noise (XY Servo on) | 0.378
| Tianjin Medical University | 2025-01-21 | Y Detector noise (XY servo off) | 0.372
| Tianjin Medical University | 2025-01-21 | Y axis OPM | 0.9494
| Tianjin Medical University | 2025-01-21 | X Scan Stroke | 192.6
| Tianjin Medical University | 2025-01-21 | X Detector noise (XY Servo on) | 0.364
| Tianjin Medical University | 2025-01-21 | X Detector noise (XY servo off) | 0.352
| Tianjin Medical University | 2025-01-21 | X axis OPM | 0.4602
| Tianjin Medical University | 2025-01-21 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 4.013
| Tianjin Medical University | 2025-01-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.43
| Tianjin Medical University | 2025-01-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| Tianjin Medical University | 2025-01-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.39
| Tianjin Medical University | 2025-01-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| Tianjin Medical University | 2025-01-21 | NCM driving% with AC160TS cantilever | 0.41
| Tianjin Medical University | 2025-01-21 | NCM Amplitude noise | 13.78
| Tianjin Medical University | 2025-01-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.651
| Tianjin Medical University | 2025-01-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.7
| Tianjin Medical University | 2025-01-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 254.1
| Tianjin Medical University | 2025-01-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 250.4
| Tianjin Medical University | 2025-01-21 | Lateral(C-D) sensitivity | 46.2
| Tianjin Medical University | 2025-01-21 | Lateral(C-D) noise | 0.134
| Tianjin Medical University | 2025-01-21 | Full range of Z Height during Z Scan sweep | 18.8
| Tianjin Medical University | 2025-01-21 | Full range of Z Detector during Z Scan sweep | 14.1
| Tianjin Medical University | 2025-01-21 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| Tianjin Medical University | 2025-01-21 | Full range of Y Detector during Y Scan sweep | 185.6
| Tianjin Medical University | 2025-01-21 | Full range of X Detector Raw during X Scan sweep | 15.4
| Tianjin Medical University | 2025-01-21 | Full range of X Detector during X Scan sweep | 192.6
| Tianjin Medical University | 2025-01-21 | A-B noise | 7.62
| PSI Demo | 2025-01-15 | Z Scanner position at -3μm | 0.006193
| PSI Demo | 2025-01-15 | Z Scanner position at +3μm | 0.005846
| PSI Demo | 2025-01-15 | Z linearity by sample height | 1.0081
| PSI Demo | 2025-01-15 | Z Height noise | 0.028
| PSI Demo | 2025-01-15 | Y Scan Stroke | 111.9
| PSI Demo | 2025-01-15 | Y Detector noise (XY Servo on) | 0.276
| PSI Demo | 2025-01-15 | Y Detector noise (XY servo off) | 0.272
| PSI Demo | 2025-01-15 | Y axis OPM | 0.7168
| PSI Demo | 2025-01-15 | X Scan Stroke | 108.7
| PSI Demo | 2025-01-15 | X Detector noise (XY Servo on) | 0.273
| PSI Demo | 2025-01-15 | X Detector noise (XY servo off) | 0.275
| PSI Demo | 2025-01-15 | X axis OPM | 0.9864
| PSI Demo | 2025-01-15 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 0.53
| PSI Demo | 2025-01-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| PSI Demo | 2025-01-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| PSI Demo | 2025-01-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| PSI Demo | 2025-01-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| PSI Demo | 2025-01-15 | NCM driving% with AC160TS cantilever | 0.214
| PSI Demo | 2025-01-15 | NCM Amplitude noise | 13.285
| PSI Demo | 2025-01-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.289
| PSI Demo | 2025-01-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.198
| PSI Demo | 2025-01-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 155.8
| PSI Demo | 2025-01-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 155
| PSI Demo | 2025-01-15 | Lateral(C-D) sensitivity | 23.2
| PSI Demo | 2025-01-15 | Lateral(C-D) noise | 0.112
| PSI Demo | 2025-01-15 | Full range of Z Height during Z Scan sweep | 16.1
| PSI Demo | 2025-01-15 | Full range of Z Detector during Z Scan sweep | 12.3
| PSI Demo | 2025-01-15 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| PSI Demo | 2025-01-15 | Full range of Y Detector during Y Scan sweep | 111.8
| PSI Demo | 2025-01-15 | Full range of X Detector Raw during X Scan sweep | 14
| PSI Demo | 2025-01-15 | Full range of X Detector during X Scan sweep | 108.7
| PSI Demo | 2025-01-15 | A-B noise | 6.905
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Z Scanner position at -3μm | 0.003114
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Z Scanner position at -3μm | 0.005153
| Free University of Berlin | 2025-01-13 | Z Scanner position at -3μm | 0.003013
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Z Scanner position at +3μm | 0.006296
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Z Scanner position at +3μm | 0.017008
| Free University of Berlin | 2025-01-13 | Z Scanner position at +3μm | 0.005175
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Z linearity by sample height | 0.9969
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Z linearity by sample height | 0.99068
| Free University of Berlin | 2025-01-13 | Z linearity by sample height | 0.99588
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Z Height noise | 0.029
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Z Height noise | 0.03
| Free University of Berlin | 2025-01-13 | Z Height noise | 0.029
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y2 Detector noise (XY Servo on) | 0.238
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y2 Detector noise (XY Servo off) | 0.238
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y1 Detector noise (XY servo on) | 0.241
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y1 Detector noise (XY servo off) | 0.239
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Y Scan Stroke | 115.9
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y Scan Stroke | 184.1
| Free University of Berlin | 2025-01-13 | Y Scan Stroke | 19.2
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Y Detector noise (XY Servo on) | 0.26
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y Detector noise (XY servo on) | 0.169
| Free University of Berlin | 2025-01-13 | Y Detector noise (XY Servo on) | 0.022
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Y Detector noise (XY servo off) | 0.258
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y Detector noise (XY servo off) | 0.168
| Free University of Berlin | 2025-01-13 | Y Detector noise (XY servo off) | 0.024
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Y axis OPM | 1.4416
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Y axis OPM | 1.333
| Free University of Berlin | 2025-01-13 | Y axis OPM | 1.4392
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X2 Detector noise (XY Servo on) | 0.26
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X2 Detector noise (XY Servo off) | 0.253
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X1 Detector noise (XY servo on) | 0.264
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X1 Detector noise (XY servo off) | 0.273
| China Aluminum Science and Technology Research Institute | 2025-01-13 | X Scan Stroke | 112.3
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X Scan Stroke | 186.4
| Free University of Berlin | 2025-01-13 | X Scan Stroke | 20.1
| China Aluminum Science and Technology Research Institute | 2025-01-13 | X Detector noise (XY Servo on) | 0.262
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X Detector noise (XY servo on) | 0.197
| Free University of Berlin | 2025-01-13 | X Detector noise (XY Servo on) | 0.023
| China Aluminum Science and Technology Research Institute | 2025-01-13 | X Detector noise (XY servo off) | 0.259
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X Detector noise (XY servo off) | 0.198
| Free University of Berlin | 2025-01-13 | X Detector noise (XY servo off) | 0.025
| China Aluminum Science and Technology Research Institute | 2025-01-13 | X axis OPM | 1.483
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | X axis OPM | 1.193
| Free University of Berlin | 2025-01-13 | X axis OPM | 1.5562
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.382
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 3.82
| Free University of Berlin | 2025-01-13 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 4.401
| China Aluminum Science and Technology Research Institute | 2025-01-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Free University of Berlin | 2025-01-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| China Aluminum Science and Technology Research Institute | 2025-01-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.61
| Free University of Berlin | 2025-01-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.33
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.33
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| China Aluminum Science and Technology Research Institute | 2025-01-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Free University of Berlin | 2025-01-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| China Aluminum Science and Technology Research Institute | 2025-01-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.6
| Free University of Berlin | 2025-01-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| China Aluminum Science and Technology Research Institute | 2025-01-13 | NCM driving% with AC160TS cantilever | 0.273
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | NCM driving% with AC160TS cantilever | 2.338
| Free University of Berlin | 2025-01-13 | NCM driving% with AC160TS cantilever | 0.58
| China Aluminum Science and Technology Research Institute | 2025-01-13 | NCM Amplitude noise | 13.918
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | NCM Amplitude noise | 10.416
| Free University of Berlin | 2025-01-13 | NCM Amplitude noise | 9.503
| China Aluminum Science and Technology Research Institute | 2025-01-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.857
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.066
| Free University of Berlin | 2025-01-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.225
| China Aluminum Science and Technology Research Institute | 2025-01-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.005
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.118
| Free University of Berlin | 2025-01-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.674
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 252.3
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 253.9
| China Aluminum Science and Technology Research Institute | 2025-01-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 158.6
| Free University of Berlin | 2025-01-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 30
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 255.6
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 258.5
| China Aluminum Science and Technology Research Institute | 2025-01-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157
| Free University of Berlin | 2025-01-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 31.4
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Lateral(C-D) sensitivity | 62.9
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Lateral(C-D) sensitivity | 85.292
| Free University of Berlin | 2025-01-13 | Lateral(C-D) sensitivity | 24.6
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Lateral(C-D) noise | 0.103
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Lateral(C-D) noise | 0.174
| Free University of Berlin | 2025-01-13 | Lateral(C-D) noise | 0.054
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Full range of Z Height during Z Scan sweep | 17.8
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Z Height during Z Scan sweep | 17.4
| Free University of Berlin | 2025-01-13 | Full range of Z Height during Z Scan sweep | 17.9
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Full range of Z Detector during Z Scan sweep | 12.7
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Z Detector during Z Scan sweep | 12.4
| Free University of Berlin | 2025-01-13 | Full range of Z Detector during Z Scan sweep | 13.4
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Y2 Detector during Y Scan sweep | 183.9
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Y1 Detector during Y Scan sweep | 184.2
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Full range of Y Detector Raw during Y Scan sweep | 14
| Free University of Berlin | 2025-01-13 | Full range of Y Detector Raw during Y Scan sweep | 12.8
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Full range of Y Detector during Y Scan sweep | 111.3
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of Y Detector during Y Scan sweep | 184.1
| Free University of Berlin | 2025-01-13 | Full range of Y Detector during Y Scan sweep | 19.2
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of X2 Detector during X Scan sweep | 185.8
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of X1 Detector Raw during X Scan sweep | 14.5
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of X1 Detector during X Scan sweep | 188.1
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Full range of X Detector Raw during X Scan sweep | 13.9
| Free University of Berlin | 2025-01-13 | Full range of X Detector Raw during X Scan sweep | 12.8
| China Aluminum Science and Technology Research Institute | 2025-01-13 | Full range of X Detector during X Scan sweep | 109.1
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | Full range of X Detector during X Scan sweep | 186.9
| Free University of Berlin | 2025-01-13 | Full range of X Detector during X Scan sweep | 20.1
| China Aluminum Science and Technology Research Institute | 2025-01-13 | A-B noise | 10.451
| NILT Malaysia Sdn. Bhd. | 2025-01-13 | A-B noise | 14.937
| Free University of Berlin | 2025-01-13 | A-B noise | 9.458
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Linearity by Z scanner position | 0.8
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Linearity by 3 different sample height | 0.999
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Height noise | 0.025
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.1
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.1
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.8
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.8
| Xuzhou Meixing OE Technology | 2025-01-10 | Z Drift (Short Traveling) | 18.3
| Xuzhou Meixing OE Technology | 2025-01-10 | Y2 Detector noise (XY Servo on) | 0.25
| Xuzhou Meixing OE Technology | 2025-01-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 236.1
| Xuzhou Meixing OE Technology | 2025-01-10 | Y1 Detector noise (XY Servo on) | 0.25
| Xuzhou Meixing OE Technology | 2025-01-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 237.1
| Xuzhou Meixing OE Technology | 2025-01-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 151.3
| Xuzhou Meixing OE Technology | 2025-01-10 | Y Ringing(Rpv) | 0.66
| Xuzhou Meixing OE Technology | 2025-01-10 | Y Detector noise (XY Servo on) | 0.19
| Xuzhou Meixing OE Technology | 2025-01-10 | Y axis OPM | 1.31
| Xuzhou Meixing OE Technology | 2025-01-10 | XY Orthogonality | 89.79
| Xuzhou Meixing OE Technology | 2025-01-10 | X2 Detector noise (XY Servo on) | 0.24
| Xuzhou Meixing OE Technology | 2025-01-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 243.6
| Xuzhou Meixing OE Technology | 2025-01-10 | X1 Detector noise (XY Servo on) | 0.24
| Xuzhou Meixing OE Technology | 2025-01-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 241.7
| Xuzhou Meixing OE Technology | 2025-01-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 153.3
| Xuzhou Meixing OE Technology | 2025-01-10 | X Detector noise (XY Servo on) | 0.18
| Xuzhou Meixing OE Technology | 2025-01-10 | X axis OPM | 0.71
| Xuzhou Meixing OE Technology | 2025-01-10 | StrokeYMin | -102425
| Xuzhou Meixing OE Technology | 2025-01-10 | StrokeYMax | 103858
| Xuzhou Meixing OE Technology | 2025-01-10 | StrokeY | 206283
| Xuzhou Meixing OE Technology | 2025-01-10 | StrokeXMin | -116611
| Xuzhou Meixing OE Technology | 2025-01-10 | StrokeXMax | 116038.1
| Xuzhou Meixing OE Technology | 2025-01-10 | StrokeX | 232649.4
| Xuzhou Meixing OE Technology | 2025-01-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.3
| Xuzhou Meixing OE Technology(NX-PSS) | 2025-01-10 | Roughness Repeatability 80x80um(1σ) | 0.001333
| Xuzhou Meixing OE Technology(NX-PSS) | 2025-01-10 | Roughness Repeatability 80x80um(RSD%) | 0.37
| Xuzhou Meixing OE Technology(NX-PSS) | 2025-01-10 | Roughness Repeatability 1x1um(1σ) | 0.003667
| Xuzhou Meixing OE Technology(NX-PSS) | 2025-01-10 | Roughness Repeatability 1x1um(RSD%) | 1.792
| Xuzhou Meixing OE Technology | 2025-01-10 | RMS of ZeroScan | 0.03
| Xuzhou Meixing OE Technology | 2025-01-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4e-05
| Xuzhou Meixing OE Technology | 2025-01-10 | OpticalAxisShiftPerFocusY | -0.0006
| Xuzhou Meixing OE Technology | 2025-01-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 6e-05
| Xuzhou Meixing OE Technology | 2025-01-10 | OpticalAxisShiftPerFocusX | -0.0002
| Xuzhou Meixing OE Technology | 2025-01-10 | NCM driving% with AC160TS cantilever
| Xuzhou Meixing OE Technology | 2025-01-10 | LowerLimit | 4481.92
| Xuzhou Meixing OE Technology | 2025-01-10 | Lateral(C-D) sensitivity | 46.6
| Xuzhou Meixing OE Technology | 2025-01-10 | Lateral(C-D) noise | 0.217
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of Z Height during Z Scan sweep | 17.3
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of Z Detector during Z Scan sweep | 12.8
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14.54
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of Y2 Detector during Y Scan sweep | 176.9
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.64
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of Y1 Detector during Y Scan sweep | 187.9
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of X2 Detector Raw during X Scan sweep | 14.54
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of X2 Detector during X Scan sweep | 169.7
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of X1 Detector Raw during X Scan sweep | 14.7
| Xuzhou Meixing OE Technology | 2025-01-10 | Full range of X1 Detector during X Scan sweep | 169.6
| Xuzhou Meixing OE Technology | 2025-01-10 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Xuzhou Meixing OE Technology | 2025-01-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| Xuzhou Meixing OE Technology | 2025-01-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| Xuzhou Meixing OE Technology | 2025-01-10 | 940nm Step Height Sample (10 Repeats) | 947.79
| Xuzhou Meixing OE Technology | 2025-01-10 | 50~200nm Step Height Sample (10 Repeats) | 183.3
| Xuzhou Meixing OE Technology | 2025-01-10 | 200~500nm Step Height Sample (10 Repeats) | 461.86
| HEPIA | 2025-01-09 | Z Scanner position at -3μm | 0.009209
| HEPIA | 2025-01-09 | Z Scanner position at +3μm | 0.017628
| HEPIA | 2025-01-09 | Z linearity by sample height | 1.0082
| HEPIA | 2025-01-09 | Z Height noise | 0.029
| HEPIA | 2025-01-09 | Y Scan Stroke | 55
| HEPIA | 2025-01-09 | Y Detector noise (XY Servo on) | 0.33
| HEPIA | 2025-01-09 | Y Detector noise (XY Servo off) | 0.333
| HEPIA | 2025-01-09 | Y axis OPM | 0.9634
| HEPIA | 2025-01-09 | X Scan Stroke | 55
| HEPIA | 2025-01-09 | X Detector noise (XY Servo on) | 0.327
| HEPIA | 2025-01-09 | X Detector noise (XY Servo off) | 0.335
| HEPIA | 2025-01-09 | X axis OPM | 1.1504
| HEPIA | 2025-01-09 | Roughness Repeatability 1x1um(10 Repeats, RSD% ) | 1.3
| HEPIA | 2025-01-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| HEPIA | 2025-01-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| HEPIA | 2025-01-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| HEPIA | 2025-01-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| HEPIA | 2025-01-09 | NCM driving% with AC160TS cantilever | 0.363
| HEPIA | 2025-01-09 | NCM Amplitude noise | 8.169
| HEPIA | 2025-01-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.92
| HEPIA | 2025-01-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.093
| HEPIA | 2025-01-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 91.2
| HEPIA | 2025-01-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 90.5
| HEPIA | 2025-01-09 | Lateral(C-D) sensitivity | 24
| HEPIA | 2025-01-09 | Lateral(C-D) noise | 0.082
| HEPIA | 2025-01-09 | Full range of Z Height during Z Scan sweep | 17.1
| HEPIA | 2025-01-09 | Full range of Z Detector during Z Scan sweep | 13.1
| HEPIA | 2025-01-09 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| HEPIA | 2025-01-09 | Full range of Y Detector during Y Scan sweep | 66
| HEPIA | 2025-01-09 | Full range of X Detector Raw during X Scan sweep | 14.1
| HEPIA | 2025-01-09 | Full range of X Detector during X Scan sweep | 63.6
| HEPIA | 2025-01-09 | A-B noise | 3.138
| UMC Singapore | 2025-01-03 | Z Scanner position at -3μm | 0.001
| UMC Singapore | 2025-01-03 | Z Scanner position at +3μm | 0.0138
| UMC Singapore | 2025-01-03 | Z linearity by sample height | 1.004
| UMC Singapore | 2025-01-03 | Z Height noise | 0.031
| UMC Singapore | 2025-01-03 | Y Scan Stroke | 122.3
| UMC Singapore | 2025-01-03 | Y Detector noise (XY Servo on) | 0.28
| UMC Singapore | 2025-01-03 | Y Detector noise (XY Servo off) | 0.27
| UMC Singapore | 2025-01-03 | Y axis OPM | 0.75
| UMC Singapore | 2025-01-03 | X Scan Stroke | 124.3
| UMC Singapore | 2025-01-03 | X Detector noise (XY Servo on) | 0.28
| UMC Singapore | 2025-01-03 | X Detector noise (XY Servo off) | 0.28
| UMC Singapore | 2025-01-03 | X axis OPM | 1.38
| UMC Singapore | 2025-01-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| UMC Singapore | 2025-01-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| UMC Singapore | 2025-01-03 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| UMC Singapore | 2025-01-03 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| UMC Singapore | 2025-01-03 | NCM driving% with AC160TS cantilever | 0.97
| UMC Singapore | 2025-01-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.514
| UMC Singapore | 2025-01-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.137
| UMC Singapore | 2025-01-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 175.7
| UMC Singapore | 2025-01-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 173.7
| UMC Singapore | 2025-01-03 | Lateral(C-D) sensitivity | 16.1
| UMC Singapore | 2025-01-03 | Lateral(C-D) noise | 0.109
| UMC Singapore | 2025-01-03 | Full range of Z Height during Z Scan sweep | 18.4
| UMC Singapore | 2025-01-03 | Full range of Z Detector during Z Scan sweep | 15.3
| UMC Singapore | 2025-01-03 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| UMC Singapore | 2025-01-03 | Full range of Y Detector during Y Scan sweep | 122.3
| UMC Singapore | 2025-01-03 | Full range of X Detector Raw during X Scan sweep | 14.3
| UMC Singapore | 2025-01-03 | Full range of X Detector during X Scan sweep | 124.3
| Shanghai Polytechnic University | 2024-12-31 | Z Scanner position at -3μm | 0.004
| Shanghai Polytechnic University | 2024-12-31 | Z Scanner position at +3μm | 0.012
| Shanghai Polytechnic University | 2024-12-31 | Z linearity by sample height | 0.995
| Shanghai Polytechnic University | 2024-12-31 | Z Height noise | 0.029
| Shanghai Polytechnic University | 2024-12-31 | Y Scan Stroke | 121.5
| Shanghai Polytechnic University | 2024-12-31 | Y Detector noise (XY Servo on) | 0.426
| Shanghai Polytechnic University | 2024-12-31 | Y Detector noise (XY Servo off) | 0.423
| Shanghai Polytechnic University | 2024-12-31 | Y axis OPM | 1.31
| Shanghai Polytechnic University | 2024-12-31 | XY orthogonality | 89.53
| Shanghai Polytechnic University | 2024-12-31 | X Scan Stroke | 120.4
| Shanghai Polytechnic University | 2024-12-31 | X Detector noise (XY Servo on) | 0.433
| Shanghai Polytechnic University | 2024-12-31 | X Detector noise (XY Servo off) | 0.431
| Shanghai Polytechnic University | 2024-12-31 | X axis OPM | 0.43
| Shanghai Polytechnic University | 2024-12-31 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.094
| Shanghai Polytechnic University | 2024-12-31 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.463
| Shanghai Polytechnic University | 2024-12-31 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 5.152
| Shanghai Polytechnic University | 2024-12-31 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.176
| Shanghai Polytechnic University | 2024-12-31 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.354
| Shanghai Polytechnic University | 2024-12-31 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.772
| Shanghai Polytechnic University | 2024-12-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Shanghai Polytechnic University | 2024-12-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Shanghai Polytechnic University | 2024-12-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Shanghai Polytechnic University | 2024-12-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Shanghai Polytechnic University | 2024-12-31 | NCM Amplitude noise | 14.75
| Shanghai Polytechnic University | 2024-12-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.06
| Shanghai Polytechnic University | 2024-12-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.835
| Shanghai Polytechnic University | 2024-12-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.9
| Shanghai Polytechnic University | 2024-12-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167.1
| Shanghai Polytechnic University | 2024-12-31 | Lateral(C-D) sensitivity | 15.1
| Shanghai Polytechnic University | 2024-12-31 | Lateral(C-D) noise | 0.108
| Shanghai Polytechnic University | 2024-12-31 | Full range of Z Height during Z Scan sweep | 17.8
| Shanghai Polytechnic University | 2024-12-31 | Full range of Z Detector during Z Scan sweep | 13.3
| Shanghai Polytechnic University | 2024-12-31 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Shanghai Polytechnic University | 2024-12-31 | Full range of Y Detector during Y Scan sweep | 121.5
| Shanghai Polytechnic University | 2024-12-31 | Full range of X Detector Raw during X Scan sweep | 14.4
| Shanghai Polytechnic University | 2024-12-31 | Full range of X Detector during X Scan sweep | 120.4
| Shanghai Polytechnic University | 2024-12-31 | A-B noise | 9.18
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Linearity by Z scanner position | 0.67
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Height noise | 0.025
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.8
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.7
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Drift | 10.91
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Z Detector offset | 58.3
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y2 Detector noise (XY Servo on) | 0.58
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 261.2
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y1 Detector noise (XY Servo on) | 0.56
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 249.4
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 165.2
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y Ringing(Rpv) | 0.1
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y Non-linearity error | 1.97
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y Detector noise (XY Servo on) | 0.47
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Repeatability @1st position | 0.05
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Repeatability @ 4th position | 0.05
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Repeatability @ 3rd position | 0.05
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Repeatability @ 2nd position | 0.05
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Accuracy @1st position | 0.25
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Accuracy @ 4th position | 0.22
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Accuracy @ 3rd position | 0.28
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis, Accuracy @ 2nd position | 0.25
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Y axis OPM | 0.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.6758
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y offset (LLC Translation) | 1.9531
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y offset (In-Die Align) | 1.1084
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y offset (Global Align) | 4.6826
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.49478
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y 1σ (LLC Translation) | 0.4268
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y 1σ (In-Die Align) | 0.22293
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning Y 1σ (Global Align) | 1.189
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X offset (Vision Pattern Matching) | 2.6514
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X offset (LLC Translation) | 1.6162
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X offset (In-Die Align) | 1.3428
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X offset (Global Align) | 1.9824
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.47812
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X 1σ (LLC Translation) | 0.34582
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X 1σ (In-Die Align) | 0.25808
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Stage positioning X 1σ (Global Align) | 0.46854
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | XY Orthogonality | 90
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X2 Detector noise (XY Servo on) | 0.47
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 258.5
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X1 Detector noise (XY Servo on) | 0.48
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 267.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 174.1
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X Ringing(Rpv) | 0.11
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X Non-linearity error | 2.31
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X Detector noise (XY Servo on) | 0.36
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Repeatability @1st position | 0.03
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Repeatability @ 4th position | 0.04
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Repeatability @ 3rd position | 0.03
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Repeatability @ 2nd position | 0.03
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Accuracy @1st position | 0.18
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Accuracy @ 4th position | 0.16
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Accuracy @ 3rd position | 0.16
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis, Accuracy @ 2nd position | 0.19
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | X axis OPM | 0.86
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Wafer placement(Loading) error on the load port2 (Y offset) | 12.2
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Wafer placement(Loading) error on the load port2 (X offset) | 10.2
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Wafer placement(Loading) error on the load port1 (Y offset) | 15.7
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Wafer placement(Loading) error on the load port1 (X offset) | 14.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Torque Profile (50mm, 20um/s) | 0.493
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Torque Profile (50mm, 200um/s) | 0.353
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | StrokeYMin | -152608
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | StrokeYMax | 150680.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | StrokeY | 303288.2
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | StrokeXMin | -203649
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | StrokeXMax | 203068.9
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | StrokeX | 406717.4
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.53
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Roughness Repeatability 80x80um(1σ) | 0.000667
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Roughness Repeatability 80x80um(RSD%) | 0.478
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Roughness Repeatability 1x1um(1σ) | 0.000333
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Roughness Repeatability 1x1um(RSD%) | 0.355
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Roll motion to X+ full scan | 1.9
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | RMS of ZeroScan | 0.054
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Position repeatability(10repeats, Position Based) | 0.16
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OpticalAxisShiftPerFocusY | 0
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.0001
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OpticalAxisShiftPerFocusX | -0.003
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OPM Repeatability @5mm | 1.28
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OPM Repeatability @25mm | 3.97
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OPM Repeatability @1mm | 0.7
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | OPM Repeatability @10mm | 1.77
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | NCM driving% with AC160TS cantilever | 9.78
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Min. Vibration Transmissbility (~100Hz) | 0.034
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Max OPM @5mm | 21.97
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Max OPM @25mm | 68.34
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Max OPM @1mm | 6.35
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Max OPM @10mm | 31.56
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | LowerLimit | 2290.69
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | LowerLimit | 4520
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Lateral(C-D) sensitivity | 16
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Lateral(C-D) noise | 0.199
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from -Y to +Y (Y scan) | 52.73
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from -Y to +Y (X scan) | 5.86
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from -X to +X (Y scan) | 65.92
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from -X to +X (X scan) | 59.33
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from +Y to -Y (Y scan) | 57.13
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from +Y to -Y (X scan) | 18.31
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from +X to -X (Y scan) | 38.09
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Image offset after moving from +X to -X (X scan) | 43.95
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Z Height during Z Scan sweep | 16.8
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Z Detector during Z Scan sweep | 13.1
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.01
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Y2 Detector during Y Scan sweep | 205.1
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Y1 Detector during Y Scan sweep | 193.6
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of Y Detector during Y Scan sweep | 205.1
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of X2 Detector Raw during X Scan sweep | 13.96
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of X2 Detector during X Scan sweep | 201.5
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of X1 Detector Raw during X Scan sweep | 14.06
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of X1 Detector during X Scan sweep | 204.8
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Full range of X Detector during X Scan sweep | 201.5
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.7
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.01
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Samsung (SEC) NRD-K (Flash) | 2024-12-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Z Scanner position at -3μm | 0.0033
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Z Scanner position at +3μm | 0.0152
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Z linearity by sample height | 0.998
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Z Height noise | 0.027
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y2 Detector noise (XY Servo on) | 0.25
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y2 Detector noise (XY Servo off) | 0.23
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y1 Detector noise (XY Servo on) | 0.25
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y1 Detector noise (XY Servo off) | 0.25
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y Scan Stroke | 181
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y Detector noise (XY Servo on) | 0.18
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y Detector noise (XY Servo off) | 0.17
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Y axis OPM | 0.65
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X2 Detector noise (XY Servo on) | 0.25
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X2 Detector noise (XY Servo off) | 0.25
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X1 Detector noise (XY Servo on) | 0.26
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X1 Detector noise (XY Servo off) | 0.25
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X Scan Stroke | 181.6
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X Detector noise (XY Servo on) | 0.19
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X Detector noise (XY Servo off) | 0.19
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | X axis OPM | 1.07
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.017
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.332
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.473
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.2
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.299
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 8.647
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.33
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.28
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | NCM Amplitude noise | 17
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.738
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.074
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.3
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 252.6
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 252.3
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 252.4
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Maximum Y Scan range (Full Stroke) | 181.1
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Maximum X Scan range (Full Stroke) | 181.7
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Lateral(C-D) sensitivity | 17.5
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Lateral(C-D) noise | 0.136
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Full range of Z Height during Z Scan sweep | 17.1
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Full range of Z Detector during Z Scan sweep | 12.6
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Full range of Y2 Detector raw | 14.6
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Full range of Y1 Detector raw | 14.4
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Full range of X2 Detector raw | 14.4
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | Full range of X1 Detector raw | 14.4
| Shenzhen Jingweikaiwu Instruments | 2024-12-27 | A-B noise | 12.94
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.6895
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y offset (LLC Translation) | 2.251
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y offset (In-Die Align) | 1.4795
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y offset (Global Align) | 2.2852
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.19192
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y 1σ (LLC Translation) | 0.418
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y 1σ (In-Die Align) | 0.30081
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning Y 1σ (Global Align) | 0.70278
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X offset (Vision Pattern Matching) | 2.8174
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X offset (LLC Translation) | 1.6113
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X offset (In-Die Align) | 1.7285
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X offset (Global Align) | 1.3574
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.62028
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X 1σ (LLC Translation) | 0.34422
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X 1σ (In-Die Align) | 0.30429
| Huahong Wuxi #6 | 2024-12-27 | XY Stage positioning X 1σ (Global Align) | 0.32098
| Huahong Wuxi #6 | 2024-12-27 | Roughness Repeatability 80x80um(1σ) | 0.005
| Huahong Wuxi #6 | 2024-12-27 | Roughness Repeatability 80x80um(RSD%) | 2.31
| Huahong Wuxi #6 | 2024-12-27 | Roughness Repeatability 1x1um(1σ) | 0.001
| Huahong Wuxi #6 | 2024-12-27 | Roughness Repeatability 1x1um(RSD%) | 0.99
| Shenzhen JWKW(2) | 2024-12-24 | Z Linearity by Z scanner position | 1.4
| Shenzhen JWKW(2) | 2024-12-24 | Z Height noise | 0.026
| Shenzhen JWKW(2) | 2024-12-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.4
| Shenzhen JWKW(2) | 2024-12-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.4
| Shenzhen JWKW(2) | 2024-12-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| Shenzhen JWKW(2) | 2024-12-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.7
| Shenzhen JWKW(2) | 2024-12-24 | Z Drift | 4.5
| Shenzhen JWKW(2) | 2024-12-24 | Z Detector offset | 6.6
| Shenzhen JWKW(2) | 2024-12-24 | Y2 Detector noise (XY Servo on) | 0.44
| Shenzhen JWKW(2) | 2024-12-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 248
| Shenzhen JWKW(2) | 2024-12-24 | Y1 Detector noise (XY Servo on) | 0.45
| Shenzhen JWKW(2) | 2024-12-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 243.4
| Shenzhen JWKW(2) | 2024-12-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 151.1
| Shenzhen JWKW(2) | 2024-12-24 | Y Ringing(Rpv) | 0.18
| Shenzhen JWKW(2) | 2024-12-24 | Y Non-linearity error | 2.76
| Shenzhen JWKW(2) | 2024-12-24 | Y Detector noise (XY Servo on) | 0.32
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Repeatability @1st position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Repeatability @ 4th position | 0.02
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Repeatability @ 3rd position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Repeatability @ 2nd position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Accuracy @1st position | 0.14
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Accuracy @ 4th position | 0.13
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Accuracy @ 3rd position | 0.2
| Shenzhen JWKW(2) | 2024-12-24 | Y axis, Accuracy @ 2nd position | 0.14
| Shenzhen JWKW(2) | 2024-12-24 | Y axis OPM | 0.37
| Shenzhen JWKW(2) | 2024-12-24 | XY Orthogonality | 90.1
| Shenzhen JWKW(2) | 2024-12-24 | X2 Detector noise (XY Servo on) | 0.44
| Shenzhen JWKW(2) | 2024-12-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 251.6
| Shenzhen JWKW(2) | 2024-12-24 | X1 Detector noise (XY Servo on) | 0.44
| Shenzhen JWKW(2) | 2024-12-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 237.3
| Shenzhen JWKW(2) | 2024-12-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 158.7
| Shenzhen JWKW(2) | 2024-12-24 | X Ringing(Rpv) | 0.21
| Shenzhen JWKW(2) | 2024-12-24 | X Non-linearity error | 2.95
| Shenzhen JWKW(2) | 2024-12-24 | X Detector noise (XY Servo on) | 0.32
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Repeatability @1st position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Repeatability @ 4th position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Repeatability @ 3rd position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Repeatability @ 2nd position | 0.03
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Accuracy @1st position | 0.13
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Accuracy @ 4th position | 0.21
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Accuracy @ 3rd position | 0.22
| Shenzhen JWKW(2) | 2024-12-24 | X axis, Accuracy @ 2nd position | 0.17
| Shenzhen JWKW(2) | 2024-12-24 | X axis OPM | 0.96
| Shenzhen JWKW(2) | 2024-12-24 | Wafer placement(Loading) error on the load port2 (Y offset) | 50.8
| Shenzhen JWKW(2) | 2024-12-24 | Wafer placement(Loading) error on the load port2 (X offset) | 48.4
| Shenzhen JWKW(2) | 2024-12-24 | Wafer placement(Loading) error on the load port1 (Y offset) | 28
| Shenzhen JWKW(2) | 2024-12-24 | Wafer placement(Loading) error on the load port1 (X offset) | 22.1
| Shenzhen JWKW(2) | 2024-12-24 | Torque Profile (50mm, 20um/s)
| Shenzhen JWKW(2) | 2024-12-24 | Torque Profile (50mm, 200um/s)
| Shenzhen JWKW(2) | 2024-12-24 | StrokeYMin | -104711
| Shenzhen JWKW(2) | 2024-12-24 | StrokeYMax | 101764.8
| Shenzhen JWKW(2) | 2024-12-24 | StrokeY | 206475.4
| Shenzhen JWKW(2) | 2024-12-24 | StrokeXMin | -153387
| Shenzhen JWKW(2) | 2024-12-24 | StrokeXMax | 152922.8
| Shenzhen JWKW(2) | 2024-12-24 | StrokeX | 306309.6
| Shenzhen JWKW(2) | 2024-12-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7
| Shenzhen JWKW(2) | 2024-12-24 | Roughness Repeatability 80x80um(1σ) | 0.002667
| Shenzhen JWICW #1 | 2024-12-24 | Roughness Repeatability 80x80um(1σ) | 0.001333
| Shenzhen JWKW (NX-HDM) | 2024-12-24 | Roughness Repeatability 80x80um(1σ) | 0.001667
| Shenzhen JWKW(2) | 2024-12-24 | Roughness Repeatability 80x80um(RSD%) | 0.92
| Shenzhen JWICW #1 | 2024-12-24 | Roughness Repeatability 80x80um(RSD%) | 0.58
| Shenzhen JWKW (NX-HDM) | 2024-12-24 | Roughness Repeatability 80x80um(RSD%) | 0.55
| Shenzhen JWKW(2) | 2024-12-24 | Roughness Repeatability 1x1um(1σ) | 0.004667
| Shenzhen JWICW #1 | 2024-12-24 | Roughness Repeatability 1x1um(1σ) | 0.001333
| Shenzhen JWKW (NX-HDM) | 2024-12-24 | Roughness Repeatability 1x1um(1σ) | 0.003667
| Shenzhen JWKW(2) | 2024-12-24 | Roughness Repeatability 1x1um(RSD%) | 4.13
| Shenzhen JWICW #1 | 2024-12-24 | Roughness Repeatability 1x1um(RSD%) | 1.032
| Shenzhen JWKW (NX-HDM) | 2024-12-24 | Roughness Repeatability 1x1um(RSD%) | 2.589
| Shenzhen JWKW(2) | 2024-12-24 | Roll motion to X+ full scan
| Shenzhen JWKW(2) | 2024-12-24 | RMS of ZeroScan | 0.035
| Shenzhen JWKW(2) | 2024-12-24 | Position repeatability(10repeats, Position Based)
| Shenzhen JWKW(2) | 2024-12-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| Shenzhen JWKW(2) | 2024-12-24 | OpticalAxisShiftPerFocusY | -0.001
| Shenzhen JWKW(2) | 2024-12-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 3e-05
| Shenzhen JWKW(2) | 2024-12-24 | OpticalAxisShiftPerFocusX | -0.0002
| Shenzhen JWKW(2) | 2024-12-24 | OPM Repeatability @5mm
| Shenzhen JWKW(2) | 2024-12-24 | OPM Repeatability @25mm
| Shenzhen JWKW(2) | 2024-12-24 | OPM Repeatability @1mm
| Shenzhen JWKW(2) | 2024-12-24 | OPM Repeatability @10mm
| Shenzhen JWKW(2) | 2024-12-24 | NCM driving% with AC160TS cantilever | 10.4
| Shenzhen JWKW(2) | 2024-12-24 | Min. Vibration Transmissbility (~100Hz) | 0.02
| Shenzhen JWKW(2) | 2024-12-24 | Max OPM @5mm
| Shenzhen JWKW(2) | 2024-12-24 | Max OPM @25mm
| Shenzhen JWKW(2) | 2024-12-24 | Max OPM @1mm
| Shenzhen JWKW(2) | 2024-12-24 | Max OPM @10mm
| Shenzhen JWKW(2) | 2024-12-24 | LowerLimit | 2720.25
| Shenzhen JWKW(2) | 2024-12-24 | LowerLimit | 12389
| Shenzhen JWKW(2) | 2024-12-24 | Lateral(C-D) sensitivity | 36.7
| Shenzhen JWKW(2) | 2024-12-24 | Lateral(C-D) noise | 0.177
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from -Y to +Y (Y scan) | 50.54
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from -Y to +Y (X scan) | 30.76
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from -X to +X (Y scan) | 36.62
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from -X to +X (X scan) | 106.93
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from +Y to -Y (Y scan) | 60.06
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from +Y to -Y (X scan) | 9.52
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from +X to -X (Y scan) | 41.02
| Shenzhen JWKW(2) | 2024-12-24 | Image offset after moving from +X to -X (X scan) | 48.34
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Z Height during Z Scan sweep | 17
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Z Detector during Z Scan sweep | 12.5
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.52
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Y2 Detector during Y Scan sweep | 183.1
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.38
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Y1 Detector during Y Scan sweep | 179.4
| Shenzhen JWKW(2) | 2024-12-24 | Full range of Y Detector during Y Scan sweep | 179.4
| Shenzhen JWKW(2) | 2024-12-24 | Full range of X2 Detector Raw during X Scan sweep | 14.62
| Shenzhen JWKW(2) | 2024-12-24 | Full range of X2 Detector during X Scan sweep | 188.8
| Shenzhen JWKW(2) | 2024-12-24 | Full range of X1 Detector Raw during X Scan sweep | 14.84
| Shenzhen JWKW(2) | 2024-12-24 | Full range of X1 Detector during X Scan sweep | 195.2
| Shenzhen JWKW(2) | 2024-12-24 | Full range of X Detector during X Scan sweep | 188.8
| Shenzhen JWKW(2) | 2024-12-24 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Shenzhen JWKW(2) | 2024-12-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
| Shenzhen JWKW(2) | 2024-12-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| Shenzhen JWKW(2) | 2024-12-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| Shenzhen JWKW(2) | 2024-12-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
| Shenzhen JWKW(2) | 2024-12-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Shenzhen JWKW(2) | 2024-12-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| Shenzhen JWKW(2) | 2024-12-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| Shenzhen JWKW(2) | 2024-12-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| Fudan University | 2024-12-23 | Z Scanner position at -3μm | 0.0108
| NASA Kennedy Space Center KSC | 2024-12-23 | Z Scanner position at -3μm | 0.0001
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Z Scanner position at -3μm | 0.0028
| Fudan University | 2024-12-23 | Z Scanner position at +3μm | 0.014
| NASA Kennedy Space Center KSC | 2024-12-23 | Z Scanner position at +3μm | 0.0124
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Z Scanner position at +3μm | 0.0107
| Fudan University | 2024-12-23 | Z Linearity by Z Scanner position
| NASA Kennedy Space Center KSC | 2024-12-23 | Z Linearity by Z Scanner position
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Z linearity by sample height | 1.013
| Fudan University | 2024-12-23 | Z Linearity by 3 different sample height | 1.016
| NASA Kennedy Space Center KSC | 2024-12-23 | Z Linearity by 3 different sample height | 1.012
| Fudan University | 2024-12-23 | Z Height noise | 0.029
| NASA Kennedy Space Center KSC | 2024-12-23 | Z Height noise | 0.032
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Z Height noise | 0.028
| Fudan University | 2024-12-23 | Y Scan Stroke | 139.2
| NASA Kennedy Space Center KSC | 2024-12-23 | Y Scan Stroke | 122.5
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Y Scan Stroke | 55
| Fudan University | 2024-12-23 | -Y Direction software Limit | -75736.9
| Fudan University | 2024-12-23 | Y Detector noise (XY Servo on) | 0.28
| NASA Kennedy Space Center KSC | 2024-12-23 | Y Detector noise (XY Servo on) | 0.27
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Y Detector noise (XY Servo on) | 0.38
| NASA Kennedy Space Center KSC | 2024-12-23 | Y Detector noise (XY Servo off) | 0.27
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Y Detector noise (XY Servo off) | 0.42
| Fudan University | 2024-12-23 | Y Detector noise (XY Servo off) | 0.28
| Fudan University | 2024-12-23 | Y axis OPM | 0.76
| NASA Kennedy Space Center KSC | 2024-12-23 | Y axis OPM | 0.9
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Y axis OPM | 0.75
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | XY orthogonality | 90.2
| Fudan University | 2024-12-23 | X Scan Stroke | 137.6
| NASA Kennedy Space Center KSC | 2024-12-23 | X Scan Stroke | 119.7
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | X Scan Stroke | 55
| Fudan University | 2024-12-23 | -X Direction software Limit | -75633.1
| Fudan University | 2024-12-23 | X Detector noise (XY Servo on) | 0.3
| NASA Kennedy Space Center KSC | 2024-12-23 | X Detector noise (XY Servo on) | 0.26
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | X Detector noise (XY Servo on) | 0.37
| NASA Kennedy Space Center KSC | 2024-12-23 | X Detector noise (XY Servo off) | 0.26
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | X Detector noise (XY Servo off) | 0.4
| Fudan University | 2024-12-23 | X Detector noise (XY Servo off) | 0.31
| Fudan University | 2024-12-23 | X axis OPM | 0.77
| NASA Kennedy Space Center KSC | 2024-12-23 | X axis OPM | 1
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | X axis OPM | 0.76
| Fudan University | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.023
| NASA Kennedy Space Center KSC | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.109
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.432
| Fudan University | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.408
| NASA Kennedy Space Center KSC | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.987
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.265
| Fudan University | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.651
| NASA Kennedy Space Center KSC | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.184
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.97
| Fudan University | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.198
| NASA Kennedy Space Center KSC | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.186
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.238
| Fudan University | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.348
| NASA Kennedy Space Center KSC | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.313
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.098
| Fudan University | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.03
| NASA Kennedy Space Center KSC | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.673
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.051
| Fudan University | 2024-12-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| NASA Kennedy Space Center KSC | 2024-12-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Fudan University | 2024-12-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| NASA Kennedy Space Center KSC | 2024-12-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Fudan University | 2024-12-23 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.23
| Fudan University | 2024-12-23 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.23
| Fudan University | 2024-12-23 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.22
| Fudan University | 2024-12-23 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.23
| Fudan University | 2024-12-23 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.23
| Fudan University | 2024-12-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| NASA Kennedy Space Center KSC | 2024-12-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Fudan University | 2024-12-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| NASA Kennedy Space Center KSC | 2024-12-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.14
| Fudan University | 2024-12-23 | NCM driving% with AC160TS cantilever | 0.937
| NASA Kennedy Space Center KSC | 2024-12-23 | NCM driving% with AC160TS cantilever | 0.97
| Fudan University | 2024-12-23 | NCM Amplitude noise | 10.2
| NASA Kennedy Space Center KSC | 2024-12-23 | NCM Amplitude noise | 16.85
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | NCM Amplitude noise | 11.28
| Fudan University | 2024-12-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.756
| NASA Kennedy Space Center KSC | 2024-12-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.858
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.955
| Fudan University | 2024-12-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.371
| NASA Kennedy Space Center KSC | 2024-12-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.49
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.4
| Fudan University | 2024-12-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 206.1
| NASA Kennedy Space Center KSC | 2024-12-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 179.9
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.5
| Fudan University | 2024-12-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 200.9
| NASA Kennedy Space Center KSC | 2024-12-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 175.5
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 96.2
| Fudan University | 2024-12-23 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Fudan University | 2024-12-23 | Lateral(C-D) sensitivity | 17.2
| NASA Kennedy Space Center KSC | 2024-12-23 | Lateral(C-D) sensitivity | 16.8
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Lateral(C-D) sensitivity | 59.3
| Fudan University | 2024-12-23 | Lateral(C-D) noise | 0.154
| NASA Kennedy Space Center KSC | 2024-12-23 | Lateral(C-D) noise | 0.084
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Lateral(C-D) noise | 0.14
| Fudan University | 2024-12-23 | Full range of Z Height during Z Scan sweep | 17.3
| NASA Kennedy Space Center KSC | 2024-12-23 | Full range of Z Height during Z Scan sweep | 17.2
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Full range of Z Height during Z Scan sweep | 17.5
| Fudan University | 2024-12-23 | Full range of Z Detector during Z Scan sweep | 12.2
| NASA Kennedy Space Center KSC | 2024-12-23 | Full range of Z Detector during Z Scan sweep | 15.3
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Full range of Z Detector during Z Scan sweep | 12.8
| Fudan University | 2024-12-23 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| NASA Kennedy Space Center KSC | 2024-12-23 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Fudan University | 2024-12-23 | Full range of Y Detector during Y Scan sweep | 139.2
| NASA Kennedy Space Center KSC | 2024-12-23 | Full range of Y Detector during Y Scan sweep | 122.5
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Full range of Y Detector during Y Scan sweep | 68
| Fudan University | 2024-12-23 | Full range of X Detector Raw during X Scan sweep | 13.8
| NASA Kennedy Space Center KSC | 2024-12-23 | Full range of X Detector Raw during X Scan sweep | 13.6
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Full range of X Detector Raw during X Scan sweep | 14.1
| Fudan University | 2024-12-23 | Full range of X Detector during X Scan sweep | 137.6
| NASA Kennedy Space Center KSC | 2024-12-23 | Full range of X Detector during X Scan sweep | 119.7
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | Full range of X Detector during X Scan sweep | 68.4
| Fudan University | 2024-12-23 | A-B noise | 13.68
| NASA Kennedy Space Center KSC | 2024-12-23 | A-B noise | 12.74
| 순천향대학교_디스플레이신소재공학과 | 2024-12-23 | A-B noise | 8.7
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Linearity by Z scanner position | 0.55
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Height noise | 0.027
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.8
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Drift | 4
| Micron F16(3) (Taiwan) | 2024-12-23 | Z Detector offset | 46.4
| Micron F16(3) (Taiwan) | 2024-12-23 | Y2 Detector noise (XY Servo on) | 0.48
| Micron F16(3) (Taiwan) | 2024-12-23 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 275.3
| Micron F16(3) (Taiwan) | 2024-12-23 | Y1 Detector noise (XY Servo on) | 0.43
| Micron F16(3) (Taiwan) | 2024-12-23 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 269.7
| Micron F16(3) (Taiwan) | 2024-12-23 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 174.2
| Micron F16(3) (Taiwan) | 2024-12-23 | Y Ringing(Rpv) | 0.23
| Micron F16(3) (Taiwan) | 2024-12-23 | Y Non-linearity error | 2.56
| Micron F16(3) (Taiwan) | 2024-12-23 | Y Detector noise (XY Servo on) | 0.34
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Repeatability @1st position | 0.05
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Repeatability @ 4th position | 0.04
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Repeatability @ 3rd position | 0.03
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Repeatability @ 2nd position | 0.03
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Accuracy @1st position | 0.32
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Accuracy @ 4th position | 0.34
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Accuracy @ 3rd position | 0.31
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis, Accuracy @ 2nd position | 0.25
| Micron F16(3) (Taiwan) | 2024-12-23 | Y axis OPM | 0.26
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.6455
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning Y offset (LLC Translation) | 1.7676
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning Y offset (Global Align) | 4.9268
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.31238
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning Y 1σ (LLC Translation) | 0.4184
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning Y 1σ (Global Align) | 1.343
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning X offset (Vision Pattern Matching) | 2.2656
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning X offset (LLC Translation) | 1.4209
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning X offset (Global Align) | 1.958
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.29859
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning X 1σ (LLC Translation) | 0.3134
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Stage positioning X 1σ (Global Align) | 0.36262
| Micron F16(3) (Taiwan) | 2024-12-23 | XY Orthogonality | 89.9
| Micron F16(3) (Taiwan) | 2024-12-23 | X2 Detector noise (XY Servo on) | 0.4
| Micron F16(3) (Taiwan) | 2024-12-23 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 266.2
| Micron F16(3) (Taiwan) | 2024-12-23 | X1 Detector noise (XY Servo on) | 0.46
| Micron F16(3) (Taiwan) | 2024-12-23 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 257.6
| Micron F16(3) (Taiwan) | 2024-12-23 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 174.7
| Micron F16(3) (Taiwan) | 2024-12-23 | X Ringing(Rpv) | 0.1
| Micron F16(3) (Taiwan) | 2024-12-23 | X Non-linearity error | 2.3
| Micron F16(3) (Taiwan) | 2024-12-23 | X Detector noise (XY Servo on) | 0.31
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Repeatability @1st position | 0.03
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Repeatability @ 4th position | 0.03
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Repeatability @ 3rd position | 0.04
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Repeatability @ 2nd position | 0.03
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Accuracy @1st position | 0.17
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Accuracy @ 4th position | 0.2
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Accuracy @ 3rd position | 0.22
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis, Accuracy @ 2nd position | 0.19
| Micron F16(3) (Taiwan) | 2024-12-23 | X axis OPM | 1.51
| Micron F16(3) (Taiwan) | 2024-12-23 | Wafer placement(Loading) error on the load port2 (Y offset) | 39.47
| Micron F16(3) (Taiwan) | 2024-12-23 | Wafer placement(Loading) error on the load port2 (X offset) | 25.95
| Micron F16(3) (Taiwan) | 2024-12-23 | Wafer placement(Loading) error on the load port1 (Y offset) | 23.81
| Micron F16(3) (Taiwan) | 2024-12-23 | Wafer placement(Loading) error on the load port1 (X offset) | 19.9
| Micron F16(3) (Taiwan) | 2024-12-23 | Torque Profile (50mm, 20um/s) | 0.767
| Micron F16(3) (Taiwan) | 2024-12-23 | Torque Profile (50mm, 200um/s) | 0.525
| Micron F16(3) (Taiwan) | 2024-12-23 | StrokeYMin | -150739
| Micron F16(3) (Taiwan) | 2024-12-23 | StrokeYMax | 152384.4
| Micron F16(3) (Taiwan) | 2024-12-23 | StrokeY | 303122.9
| Micron F16(3) (Taiwan) | 2024-12-23 | StrokeXMin | -204065
| Micron F16(3) (Taiwan) | 2024-12-23 | StrokeXMax | 203624.8
| Micron F16(3) (Taiwan) | 2024-12-23 | StrokeX | 407689.5
| Micron F16(3) (Taiwan) | 2024-12-23 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.36
| Micron F16(3) (Taiwan) | 2024-12-23 | Roughness Repeatability 80x80um(1σ) | 0.004667
| Micron F16(3) (Taiwan) | 2024-12-23 | Roughness Repeatability 80x80um(RSD%) | 0.93
| Micron F16(3) (Taiwan) | 2024-12-23 | Roughness Repeatability 1x1um(1σ) | 0.003333
| Micron F16(3) (Taiwan) | 2024-12-23 | Roughness Repeatability 1x1um(RSD%) | 2.619
| Micron F16(3) (Taiwan) | 2024-12-23 | Roll motion to X+ full scan | 19.2
| Micron F16(3) (Taiwan) | 2024-12-23 | RMS of ZeroScan | 0.042
| Micron F16(3) (Taiwan) | 2024-12-23 | Position repeatability(10repeats, Position Based) | 0.58
| Micron F16(3) (Taiwan) | 2024-12-23 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0
| Micron F16(3) (Taiwan) | 2024-12-23 | OpticalAxisShiftPerFocusY | 0.0006
| Micron F16(3) (Taiwan) | 2024-12-23 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| Micron F16(3) (Taiwan) | 2024-12-23 | OpticalAxisShiftPerFocusX | -0.003
| Micron F16(3) (Taiwan) | 2024-12-23 | OPM Repeatability @5mm | 0.74
| Micron F16(3) (Taiwan) | 2024-12-23 | OPM Repeatability @25mm | 2.15
| Micron F16(3) (Taiwan) | 2024-12-23 | OPM Repeatability @1mm | 0.95
| Micron F16(3) (Taiwan) | 2024-12-23 | OPM Repeatability @10mm | 1.23
| Micron F16(3) (Taiwan) | 2024-12-23 | NCM driving% with AC160TS cantilever | 25
| Micron F16(3) (Taiwan) | 2024-12-23 | Min. Vibration Transmissbility (~100Hz) | 0.062
| Micron F16(3) (Taiwan) | 2024-12-23 | Max OPM @5mm | 5.58
| Micron F16(3) (Taiwan) | 2024-12-23 | Max OPM @25mm | 21.56
| Micron F16(3) (Taiwan) | 2024-12-23 | Max OPM @1mm | 3.63
| Micron F16(3) (Taiwan) | 2024-12-23 | Max OPM @10mm | 20.04
| Micron F16(3) (Taiwan) | 2024-12-23 | LowerLimit | 2321.3
| Micron F16(3) (Taiwan) | 2024-12-23 | LowerLimit | 4515
| Micron F16(3) (Taiwan) | 2024-12-23 | Lateral(C-D) sensitivity | 14.6
| Micron F16(3) (Taiwan) | 2024-12-23 | Lateral(C-D) noise | 0.122
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from -Y to +Y (Y scan) | 105.47
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from -Y to +Y (X scan) | 11.72
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from -X to +X (Y scan) | 100.34
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from -X to +X (X scan) | 52
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from +Y to -Y (Y scan) | 67.38
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from +Y to -Y (X scan) | 16.85
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from +X to -X (Y scan) | 28.56
| Micron F16(3) (Taiwan) | 2024-12-23 | Image offset after moving from +X to -X (X scan) | 27.83
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Z Height during Z Scan sweep | 17.1
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Z Detector during Z Scan sweep | 12.9
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Y2 Detector Raw during Y Scan sweep | 14.34
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Y2 Detector during Y Scan sweep | 200.6
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Y1 Detector during Y Scan sweep | 194.1
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of Y Detector during Y Scan sweep | 194.1
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of X2 Detector Raw during X Scan sweep | 14.15
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of X2 Detector during X Scan sweep | 197.5
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of X1 Detector Raw during X Scan sweep | 14.37
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of X1 Detector during X Scan sweep | 193.4
| Micron F16(3) (Taiwan) | 2024-12-23 | Full range of X Detector during X Scan sweep | 197.5
| Micron F16(3) (Taiwan) | 2024-12-23 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Micron F16(3) (Taiwan) | 2024-12-23 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Micron F16(3) (Taiwan) | 2024-12-23 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| Micron F16(3) (Taiwan) | 2024-12-23 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.36
| Micron F16(3) (Taiwan) | 2024-12-23 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Micron F16(3) (Taiwan) | 2024-12-23 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.01
| Micron F16(3) (Taiwan) | 2024-12-23 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| Micron F16(3) (Taiwan) | 2024-12-23 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Micron F16(3) (Taiwan) | 2024-12-23 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| TSMC AP2 Mini Line #2 | 2024-12-20 | Roughness Repeatability 80x80um(1σ) | 0.001
| TSMC AP2 Mini Line #2 | 2024-12-20 | Roughness Repeatability 80x80um(RSD%) | 0.33
| TSMC AP2 Mini Line #2 | 2024-12-20 | Roughness Repeatability 1x1um(1σ) | 0.002
| TSMC AP2 Mini Line #2 | 2024-12-20 | Roughness Repeatability 1x1um(RSD%) | 1.228
| 서강대학교_화공생명공학과 | 2024-12-19 | Z Scanner position at -3μm | 0.0079
| 서강대학교_화공생명공학과 | 2024-12-19 | Z Scanner position at +3μm | 0.0139
| 서강대학교_화공생명공학과 | 2024-12-19 | Z linearity by sample height | 0.999
| 서강대학교_화공생명공학과 | 2024-12-19 | Z Height noise | 0.029
| 서강대학교_화공생명공학과 | 2024-12-19 | Y Scan Stroke | 55
| 서강대학교_화공생명공학과 | 2024-12-19 | Y Detector noise (XY Servo on) | 0.27
| 서강대학교_화공생명공학과 | 2024-12-19 | Y Detector noise (XY Servo off) | 0.26
| 서강대학교_화공생명공학과 | 2024-12-19 | Y axis OPM | 0.38
| 서강대학교_화공생명공학과 | 2024-12-19 | X Scan Stroke | 55
| 서강대학교_화공생명공학과 | 2024-12-19 | X Detector noise (XY Servo on) | 0.25
| 서강대학교_화공생명공학과 | 2024-12-19 | X Detector noise (XY Servo off) | 0.24
| 서강대학교_화공생명공학과 | 2024-12-19 | X axis OPM | 0.42
| 서강대학교_화공생명공학과 | 2024-12-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.048
| 서강대학교_화공생명공학과 | 2024-12-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.547
| 서강대학교_화공생명공학과 | 2024-12-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.904
| 서강대학교_화공생명공학과 | 2024-12-19 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.191
| 서강대학교_화공생명공학과 | 2024-12-19 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.222
| 서강대학교_화공생명공학과 | 2024-12-19 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.67
| 서강대학교_화공생명공학과 | 2024-12-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 서강대학교_화공생명공학과 | 2024-12-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| 서강대학교_화공생명공학과 | 2024-12-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| 서강대학교_화공생명공학과 | 2024-12-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| 서강대학교_화공생명공학과 | 2024-12-19 | NCM driving% with AC160TS cantilever | 1.291
| 서강대학교_화공생명공학과 | 2024-12-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.751
| 서강대학교_화공생명공학과 | 2024-12-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.997
| 서강대학교_화공생명공학과 | 2024-12-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 104.6
| 서강대학교_화공생명공학과 | 2024-12-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 96.3
| 서강대학교_화공생명공학과 | 2024-12-19 | Lateral(C-D) sensitivity | 8.7
| 서강대학교_화공생명공학과 | 2024-12-19 | Lateral(C-D) noise | 0.069
| 서강대학교_화공생명공학과 | 2024-12-19 | Full range of Z Height during Z Scan sweep | 17
| 서강대학교_화공생명공학과 | 2024-12-19 | Full range of Z Detector during Z Scan sweep | 12.2
| 서강대학교_화공생명공학과 | 2024-12-19 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| 서강대학교_화공생명공학과 | 2024-12-19 | Full range of Y Detector during Y Scan sweep | 68.4
| 서강대학교_화공생명공학과 | 2024-12-19 | Full range of X Detector Raw during X Scan sweep | 14
| 서강대학교_화공생명공학과 | 2024-12-19 | Full range of X Detector during X Scan sweep | 67.5
| 인하대학교_물리학과 | 2024-12-17 | Z Scanner position at -3μm | 0.0103
| 인하대학교_물리학과 | 2024-12-17 | Z Scanner position at +3μm | 0.0026
| 인하대학교_물리학과 | 2024-12-17 | Z linearity by sample height | 1.014
| 인하대학교_물리학과 | 2024-12-17 | Z Height noise | 0.026
| 인하대학교_물리학과 | 2024-12-17 | Y Scan Stroke | 114.6
| 인하대학교_물리학과 | 2024-12-17 | Y Detector noise (XY Servo on) | 0.27
| 인하대학교_물리학과 | 2024-12-17 | Y Detector noise (XY Servo off) | 0.26
| 인하대학교_물리학과 | 2024-12-17 | Y axis OPM | 0.75
| 인하대학교_물리학과 | 2024-12-17 | X Scan Stroke | 110.5
| 인하대학교_물리학과 | 2024-12-17 | X Detector noise (XY Servo on) | 0.26
| 인하대학교_물리학과 | 2024-12-17 | X Detector noise (XY Servo off) | 0.26
| 인하대학교_물리학과 | 2024-12-17 | X axis OPM | 1.12
| 인하대학교_물리학과 | 2024-12-17 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.046
| 인하대학교_물리학과 | 2024-12-17 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.691
| 인하대학교_물리학과 | 2024-12-17 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.589
| 인하대학교_물리학과 | 2024-12-17 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.199
| 인하대학교_물리학과 | 2024-12-17 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.054
| 인하대학교_물리학과 | 2024-12-17 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.222
| 인하대학교_물리학과 | 2024-12-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 인하대학교_물리학과 | 2024-12-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 인하대학교_물리학과 | 2024-12-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| 인하대학교_물리학과 | 2024-12-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 인하대학교_물리학과 | 2024-12-17 | NCM driving% with AC160TS cantilever | 0.66
| 인하대학교_물리학과 | 2024-12-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.87
| 인하대학교_물리학과 | 2024-12-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.494
| 인하대학교_물리학과 | 2024-12-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 161.6
| 인하대학교_물리학과 | 2024-12-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156.9
| 인하대학교_물리학과 | 2024-12-17 | Lateral(C-D) sensitivity | 21.6
| 인하대학교_물리학과 | 2024-12-17 | Lateral(C-D) noise | 0.105
| 인하대학교_물리학과 | 2024-12-17 | Full range of Z Height during Z Scan sweep | 17.3
| 인하대학교_물리학과 | 2024-12-17 | Full range of Z Detector during Z Scan sweep | 13.6
| 인하대학교_물리학과 | 2024-12-17 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| 인하대학교_물리학과 | 2024-12-17 | Full range of Y Detector during Y Scan sweep | 114.6
| 인하대학교_물리학과 | 2024-12-17 | Full range of X Detector Raw during X Scan sweep | 14.1
| 인하대학교_물리학과 | 2024-12-17 | Full range of X Detector during X Scan sweep | 110.5
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Z Scanner position at -3μm | 0.0099
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Z Scanner position at +3μm | 0.0108
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Z linearity by sample height | 1.01
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Z Height noise | 0.028
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Y Scan Stroke | 190.7
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Y Detector noise (XY Servo on) | 0.37
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Y Detector noise (XY Servo off) | 0.37
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Y axis OPM | 1.88
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | X Scan Stroke | 186.9
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | X Detector noise (XY Servo on) | 0.36
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | X Detector noise (XY Servo off) | 0.35
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | X axis OPM | 0.86
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.311
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 4.742
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 10.321
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.162
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.257
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.855
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.51
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.88
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.49
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.87
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | NCM Amplitude noise | 14.68
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.426
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.747
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 237.4
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 235.5
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Lateral(C-D) sensitivity | 17.2
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Lateral(C-D) noise | 0.144
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Full range of Z Height during Z Scan sweep | 16.6
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Full range of Z Detector during Z Scan sweep | 11.9
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Full range of Y Detector Raw during Y Scan sweep | 15.9
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Full range of Y Detector during Y Scan sweep | 190
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Full range of X Detector Raw during X Scan sweep | 15.9
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | Full range of X Detector during X Scan sweep | 187
| University of Bologna (Graziosi) (Gembetti) | 2024-12-16 | A-B noise | 9.95
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.5869
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y offset (LLC Translation) | 2.583
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y offset (In-Die Align) | 1.6748
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y offset (Global Align) | 2.26
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.28696
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y 1σ (LLC Translation) | 0.583
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y 1σ (In-Die Align) | 0.49902
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning Y 1σ (Global Align) | 0.49902
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X offset (Vision Pattern Matching) | 2.168
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X offset (LLC Translation) | 1.69
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X offset (In-Die Align) | 1.6895
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X offset (Global Align) | 2.67
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.54418
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X 1σ (LLC Translation) | 0.3418
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X 1σ (In-Die Align) | 0.55781
| TSMC AP2 Mini Line #2 | 2024-12-16 | XY Stage positioning X 1σ (Global Align) | 0.55781
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Linearity by Z scanner position | 0.86
| EVGroup | 2024-12-13 | Z Linearity by Z scanner position | 0.8
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Height noise | 0.027
| EVGroup | 2024-12-13 | Z Height noise | 0.025
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.3
| EVGroup | 2024-12-13 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.4
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.95
| EVGroup | 2024-12-13 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.6
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.57
| EVGroup | 2024-12-13 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.9
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.29
| EVGroup | 2024-12-13 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.8
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Drift | 15.459
| EVGroup | 2024-12-13 | Z Drift | 14.01
| Fraunhofer IPMS Dresden | 2024-12-13 | Z Detector offset | 157.6
| EVGroup | 2024-12-13 | Z Detector offset | 75.1
| Fraunhofer IPMS Dresden | 2024-12-13 | Y2 Detector noise (XY Servo on) | 0.411
| EVGroup | 2024-12-13 | Y2 Detector noise (XY Servo on) | 0.46
| Fraunhofer IPMS Dresden | 2024-12-13 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 238.1151
| EVGroup | 2024-12-13 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 258.3
| Fraunhofer IPMS Dresden | 2024-12-13 | Y1 Detector noise (XY Servo on) | 0.406
| EVGroup | 2024-12-13 | Y1 Detector noise (XY Servo on) | 0.44
| Fraunhofer IPMS Dresden | 2024-12-13 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 234.8785
| EVGroup | 2024-12-13 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 260
| Fraunhofer IPMS Dresden | 2024-12-13 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 150.371
| EVGroup | 2024-12-13 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 187.3
| Fraunhofer IPMS Dresden | 2024-12-13 | Y Ringing(Rpv) | 0.093
| EVGroup | 2024-12-13 | Y Ringing(Rpv) | 0.19
| Fraunhofer IPMS Dresden | 2024-12-13 | Y Non-linearity error | 1.59
| EVGroup | 2024-12-13 | Y Non-linearity error | 2.32
| Fraunhofer IPMS Dresden | 2024-12-13 | Y Detector noise (XY Servo on) | 0.294
| EVGroup | 2024-12-13 | Y Detector noise (XY Servo on) | 0.33
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Repeatability @1st position | 0.04
| EVGroup | 2024-12-13 | Y axis, Repeatability @1st position | 0.04
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Repeatability @ 4th position | 0.04
| EVGroup | 2024-12-13 | Y axis, Repeatability @ 4th position | 0.05
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Repeatability @ 3rd position | 0.03
| EVGroup | 2024-12-13 | Y axis, Repeatability @ 3rd position | 0.09
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Repeatability @ 2nd position | 0.04
| EVGroup | 2024-12-13 | Y axis, Repeatability @ 2nd position | 0.04
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Accuracy @1st position | 0.25
| EVGroup | 2024-12-13 | Y axis, Accuracy @1st position | 0.2
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Accuracy @ 4th position | 0.19
| EVGroup | 2024-12-13 | Y axis, Accuracy @ 4th position | 0.26
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Accuracy @ 3rd position | 0.14
| EVGroup | 2024-12-13 | Y axis, Accuracy @ 3rd position | 0.46
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis, Accuracy @ 2nd position | 0.22
| EVGroup | 2024-12-13 | Y axis, Accuracy @ 2nd position | 0.21
| Fraunhofer IPMS Dresden | 2024-12-13 | Y axis OPM | 1.174
| EVGroup | 2024-12-13 | Y axis OPM | 1.94
| Fraunhofer IPMS Dresden | 2024-12-13 | XY Orthogonality | 89.93
| EVGroup | 2024-12-13 | XY Orthogonality | 90
| Fraunhofer IPMS Dresden | 2024-12-13 | X2 Detector noise (XY Servo on) | 0.43
| EVGroup | 2024-12-13 | X2 Detector noise (XY Servo on) | 0.45
| Fraunhofer IPMS Dresden | 2024-12-13 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 238.1151
| EVGroup | 2024-12-13 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 252.9
| Fraunhofer IPMS Dresden | 2024-12-13 | X1 Detector noise (XY Servo on) | 0.425
| EVGroup | 2024-12-13 | X1 Detector noise (XY Servo on) | 0.44
| Fraunhofer IPMS Dresden | 2024-12-13 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 234.8785
| EVGroup | 2024-12-13 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 250.5
| Fraunhofer IPMS Dresden | 2024-12-13 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 154.224
| EVGroup | 2024-12-13 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.9
| Fraunhofer IPMS Dresden | 2024-12-13 | X Ringing(Rpv) | 0.749
| EVGroup | 2024-12-13 | X Ringing(Rpv) | 0.08
| Fraunhofer IPMS Dresden | 2024-12-13 | X Non-linearity error | 1.68
| EVGroup | 2024-12-13 | X Non-linearity error | 3.11
| Fraunhofer IPMS Dresden | 2024-12-13 | X Detector noise (XY Servo on) | 0.31
| EVGroup | 2024-12-13 | X Detector noise (XY Servo on) | 0.33
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Repeatability @1st position | 0.04
| EVGroup | 2024-12-13 | X axis, Repeatability @1st position | 0.02
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Repeatability @ 4th position | 0.03
| EVGroup | 2024-12-13 | X axis, Repeatability @ 4th position | 0.07
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Repeatability @ 3rd position | 0.04
| EVGroup | 2024-12-13 | X axis, Repeatability @ 3rd position | 0.08
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Repeatability @ 2nd position | 0.04
| EVGroup | 2024-12-13 | X axis, Repeatability @ 2nd position | 0.04
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Accuracy @1st position | 0.17
| EVGroup | 2024-12-13 | X axis, Accuracy @1st position | 0.17
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Accuracy @ 4th position | 0.17
| EVGroup | 2024-12-13 | X axis, Accuracy @ 4th position | 0.36
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Accuracy @ 3rd position | 0.18
| EVGroup | 2024-12-13 | X axis, Accuracy @ 3rd position | 0.37
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis, Accuracy @ 2nd position | 0.18
| EVGroup | 2024-12-13 | X axis, Accuracy @ 2nd position | 0.24
| Fraunhofer IPMS Dresden | 2024-12-13 | X axis OPM | 0.765
| EVGroup | 2024-12-13 | X axis OPM | 0.3
| Fraunhofer IPMS Dresden | 2024-12-13 | Wafer placement(Loading) error on the load port2 (Y offset) | 41.52
| EVGroup | 2024-12-13 | Wafer placement(Loading) error on the load port2 (Y offset)
| Fraunhofer IPMS Dresden | 2024-12-13 | Wafer placement(Loading) error on the load port2 (X offset) | 29.8
| EVGroup | 2024-12-13 | Wafer placement(Loading) error on the load port2 (X offset)
| Fraunhofer IPMS Dresden | 2024-12-13 | Wafer placement(Loading) error on the load port1 (Y offset) | 45.87
| EVGroup | 2024-12-13 | Wafer placement(Loading) error on the load port1 (Y offset)
| Fraunhofer IPMS Dresden | 2024-12-13 | Wafer placement(Loading) error on the load port1 (X offset) | 18.2
| EVGroup | 2024-12-13 | Wafer placement(Loading) error on the load port1 (X offset)
| Fraunhofer IPMS Dresden | 2024-12-13 | Torque Profile (50mm, 20um/s) | 0.758
| EVGroup | 2024-12-13 | Torque Profile (50mm, 20um/s) | 0.703
| Fraunhofer IPMS Dresden | 2024-12-13 | Torque Profile (50mm, 200um/s) | 0.908
| EVGroup | 2024-12-13 | Torque Profile (50mm, 200um/s) | 0.597
| Fraunhofer IPMS Dresden | 2024-12-13 | StrokeYMin | -101542
| EVGroup | 2024-12-13 | StrokeYMin | -153383
| Fraunhofer IPMS Dresden | 2024-12-13 | StrokeYMax | 104073
| EVGroup | 2024-12-13 | StrokeYMax | 150584.7
| Fraunhofer IPMS Dresden | 2024-12-13 | StrokeY | 205615
| EVGroup | 2024-12-13 | StrokeY | 303967.8
| Fraunhofer IPMS Dresden | 2024-12-13 | StrokeXMin | -152810
| EVGroup | 2024-12-13 | StrokeXMin | -205288
| Fraunhofer IPMS Dresden | 2024-12-13 | StrokeXMax | 153261
| EVGroup | 2024-12-13 | StrokeXMax | 204678.6
| Fraunhofer IPMS Dresden | 2024-12-13 | StrokeX | 306071
| EVGroup | 2024-12-13 | StrokeX | 409966.5
| Fraunhofer IPMS Dresden | 2024-12-13 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.537
| EVGroup | 2024-12-13 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.24
| Fraunhofer IPMS Dresden | 2024-12-13 | Roughness Repeatability 80x80um(1σ) | 0.000667
| EVGroup | 2024-12-13 | Roughness Repeatability 80x80um(1σ) | 0.001
| Fraunhofer IPMS Dresden | 2024-12-13 | Roughness Repeatability 80x80um(RSD%) | 0.18
| EVGroup | 2024-12-13 | Roughness Repeatability 80x80um(RSD%) | 1.13
| Fraunhofer IPMS Dresden | 2024-12-13 | Roughness Repeatability 1x1um(1σ) | 0.004667
| EVGroup | 2024-12-13 | Roughness Repeatability 1x1um(1σ) | 0.001667
| Fraunhofer IPMS Dresden | 2024-12-13 | Roughness Repeatability 1x1um(RSD%) | 3.214
| EVGroup | 2024-12-13 | Roughness Repeatability 1x1um(RSD%) | 4.37
| Fraunhofer IPMS Dresden | 2024-12-13 | Roll motion to X+ full scan | 3.5
| EVGroup | 2024-12-13 | Roll motion to X+ full scan
| Fraunhofer IPMS Dresden | 2024-12-13 | RMS of ZeroScan | 0.032
| EVGroup | 2024-12-13 | RMS of ZeroScan | 0.044
| Fraunhofer IPMS Dresden | 2024-12-13 | Position repeatability(10repeats, Position Based) | 0.38
| EVGroup | 2024-12-13 | Position repeatability(10repeats, Position Based) | 0.44
| Fraunhofer IPMS Dresden | 2024-12-13 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.000103
| EVGroup | 2024-12-13 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00018
| Fraunhofer IPMS Dresden | 2024-12-13 | OpticalAxisShiftPerFocusY | -0.0013
| EVGroup | 2024-12-13 | OpticalAxisShiftPerFocusY | -0.0009
| Fraunhofer IPMS Dresden | 2024-12-13 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00014
| EVGroup | 2024-12-13 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00029
| Fraunhofer IPMS Dresden | 2024-12-13 | OpticalAxisShiftPerFocusX | -0.0016
| EVGroup | 2024-12-13 | OpticalAxisShiftPerFocusX | -0.002
| Fraunhofer IPMS Dresden | 2024-12-13 | OPM Repeatability @5mm | 1.384
| EVGroup | 2024-12-13 | OPM Repeatability @5mm | 2.28
| Fraunhofer IPMS Dresden | 2024-12-13 | OPM Repeatability @25mm | 2.606
| EVGroup | 2024-12-13 | OPM Repeatability @25mm | 3.16
| Fraunhofer IPMS Dresden | 2024-12-13 | OPM Repeatability @1mm | 1.237
| EVGroup | 2024-12-13 | OPM Repeatability @1mm | 1.16
| Fraunhofer IPMS Dresden | 2024-12-13 | OPM Repeatability @10mm | 1.712
| EVGroup | 2024-12-13 | OPM Repeatability @10mm | 2.8
| Fraunhofer IPMS Dresden | 2024-12-13 | NCM driving% with AC160TS cantilever | 5.7
| EVGroup | 2024-12-13 | NCM driving% with AC160TS cantilever | 10.5
| Fraunhofer IPMS Dresden | 2024-12-13 | Min. Vibration Transmissbility (~100Hz) | 0.016
| EVGroup | 2024-12-13 | Min. Vibration Transmissbility (~100Hz) | 0.024
| Fraunhofer IPMS Dresden | 2024-12-13 | Max OPM @5mm | 22.37
| EVGroup | 2024-12-13 | Max OPM @5mm | 39
| Fraunhofer IPMS Dresden | 2024-12-13 | Max OPM @25mm | 79.356
| EVGroup | 2024-12-13 | Max OPM @25mm | 120.16
| Fraunhofer IPMS Dresden | 2024-12-13 | Max OPM @1mm | 16.447
| EVGroup | 2024-12-13 | Max OPM @1mm | 12.3
| Fraunhofer IPMS Dresden | 2024-12-13 | Max OPM @10mm | 22.679
| EVGroup | 2024-12-13 | Max OPM @10mm | 58.97
| Fraunhofer IPMS Dresden | 2024-12-13 | LowerLimit | 2234.19
| Fraunhofer IPMS Dresden | 2024-12-13 | LowerLimit | 4770
| EVGroup | 2024-12-13 | LowerLimit | 2296.5
| EVGroup | 2024-12-13 | LowerLimit | 3955
| Fraunhofer IPMS Dresden | 2024-12-13 | Lateral(C-D) sensitivity | 9.8
| EVGroup | 2024-12-13 | Lateral(C-D) sensitivity | 14.8
| Fraunhofer IPMS Dresden | 2024-12-13 | Lateral(C-D) noise | 0.156
| EVGroup | 2024-12-13 | Lateral(C-D) noise | 0.19
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from -Y to +Y (Y scan) | 109.863
| EVGroup | 2024-12-13 | Image offset after moving from -Y to +Y (Y scan) | 147.22
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from -Y to +Y (X scan) | 46.875
| EVGroup | 2024-12-13 | Image offset after moving from -Y to +Y (X scan) | 32.96
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from -X to +X (Y scan) | 19.043
| EVGroup | 2024-12-13 | Image offset after moving from -X to +X (Y scan) | 44.68
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from -X to +X (X scan) | 81.2988
| EVGroup | 2024-12-13 | Image offset after moving from -X to +X (X scan) | 67.38
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from +Y to -Y (Y scan) | 111.328
| EVGroup | 2024-12-13 | Image offset after moving from +Y to -Y (Y scan) | 23.44
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from +Y to -Y (X scan) | 18.3105
| EVGroup | 2024-12-13 | Image offset after moving from +Y to -Y (X scan) | 12.45
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from +X to -X (Y scan) | 174.316
| EVGroup | 2024-12-13 | Image offset after moving from +X to -X (Y scan) | 35.16
| Fraunhofer IPMS Dresden | 2024-12-13 | Image offset after moving from +X to -X (X scan) | 65.1855
| EVGroup | 2024-12-13 | Image offset after moving from +X to -X (X scan) | 84.96
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Z Height during Z Scan sweep | 16.4
| EVGroup | 2024-12-13 | Full range of Z Height during Z Scan sweep | 15.9
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Z Detector during Z Scan sweep | 12.81
| EVGroup | 2024-12-13 | Full range of Z Detector during Z Scan sweep | 12.9
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.35
| EVGroup | 2024-12-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.56
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Y2 Detector during Y Scan sweep | 178.39
| EVGroup | 2024-12-13 | Full range of Y2 Detector during Y Scan sweep | 191.3
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.15
| EVGroup | 2024-12-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.31
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Y1 Detector during Y Scan sweep | 174.7
| EVGroup | 2024-12-13 | Full range of Y1 Detector during Y Scan sweep | 168.1
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of Y Detector during Y Scan sweep | 176.54
| EVGroup | 2024-12-13 | Full range of Y Detector during Y Scan sweep | 179.7
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of X2 Detector Raw during X Scan sweep | 14.15
| EVGroup | 2024-12-13 | Full range of X2 Detector Raw during X Scan sweep | 14.61
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of X2 Detector during X Scan sweep | 183.29
| EVGroup | 2024-12-13 | Full range of X2 Detector during X Scan sweep | 185.7
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of X1 Detector Raw during X Scan sweep | 14.51
| EVGroup | 2024-12-13 | Full range of X1 Detector Raw during X Scan sweep | 14.61
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of X1 Detector during X Scan sweep | 181.37
| EVGroup | 2024-12-13 | Full range of X1 Detector during X Scan sweep | 203.9
| Fraunhofer IPMS Dresden | 2024-12-13 | Full range of X Detector during X Scan sweep | 182.33
| EVGroup | 2024-12-13 | Full range of X Detector during X Scan sweep | 194.8
| Fraunhofer IPMS Dresden | 2024-12-13 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| EVGroup | 2024-12-13 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Fraunhofer IPMS Dresden | 2024-12-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| EVGroup | 2024-12-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.18
| Fraunhofer IPMS Dresden | 2024-12-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.45
| EVGroup | 2024-12-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.97
| Fraunhofer IPMS Dresden | 2024-12-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.34
| EVGroup | 2024-12-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| Fraunhofer IPMS Dresden | 2024-12-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| EVGroup | 2024-12-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
| Fraunhofer IPMS Dresden | 2024-12-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| EVGroup | 2024-12-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Fraunhofer IPMS Dresden | 2024-12-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.49
| EVGroup | 2024-12-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.5
| Fraunhofer IPMS Dresden | 2024-12-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| EVGroup | 2024-12-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| Fraunhofer IPMS Dresden | 2024-12-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| EVGroup | 2024-12-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
| Omron | 2024-12-12 | Z Scanner position at -3μm | 0.0081
| Omron | 2024-12-12 | Z Scanner position at +3μm | 0.0106
| Omron | 2024-12-12 | Z linearity by sample height | 0.999
| Omron | 2024-12-12 | Z Height noise | 0.026
| Omron | 2024-12-12 | Y2 Detector noise (XY Servo on) | 0.23
| Omron | 2024-12-12 | Y2 Detector noise (XY Servo off) | 0.22
| Omron | 2024-12-12 | Y1 Detector noise (XY Servo on) | 0.28
| Omron | 2024-12-12 | Y1 Detector noise (XY Servo off) | 0.26
| Omron | 2024-12-12 | Y Scan Stroke | 180.6
| Omron | 2024-12-12 | Y Detector noise (XY Servo on) | 0.18
| Omron | 2024-12-12 | Y Detector noise (XY Servo off) | 0.17
| Omron | 2024-12-12 | Y axis OPM | 1.71
| Omron | 2024-12-12 | X2 Detector noise (XY Servo on) | 0.24
| Omron | 2024-12-12 | X2 Detector noise (XY Servo off) | 0.24
| Omron | 2024-12-12 | X1 Detector noise (XY Servo on) | 0.26
| Omron | 2024-12-12 | X1 Detector noise (XY Servo off) | 0.24
| Omron | 2024-12-12 | X Scan Stroke | 178.1
| Omron | 2024-12-12 | X Detector noise (XY Servo on) | 0.18
| Omron | 2024-12-12 | X Detector noise (XY Servo off) | 0.18
| Omron | 2024-12-12 | X axis OPM | 1.55
| Omron | 2024-12-12 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.036
| Omron | 2024-12-12 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.018
| Omron | 2024-12-12 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.96
| Omron | 2024-12-12 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.182
| Omron | 2024-12-12 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 1.913
| Omron | 2024-12-12 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.125
| Omron | 2024-12-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Omron | 2024-12-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Omron | 2024-12-12 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.27
| Omron | 2024-12-12 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| Omron | 2024-12-12 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.26
| Omron | 2024-12-12 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| Omron | 2024-12-12 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| Omron | 2024-12-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Omron | 2024-12-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Omron | 2024-12-12 | NCM Amplitude noise | 11.05
| Omron | 2024-12-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.901
| Omron | 2024-12-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.357
| Omron | 2024-12-12 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 245.7
| Omron | 2024-12-12 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 255.3
| Omron | 2024-12-12 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 250
| Omron | 2024-12-12 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 242.2
| Omron | 2024-12-12 | Maximum Y Scan range (Full Stroke) | 180.8
| Omron | 2024-12-12 | Maximum X Scan range (Full Stroke) | 178.1
| Omron | 2024-12-12 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Omron | 2024-12-12 | Lateral(C-D) sensitivity | 22.1
| Omron | 2024-12-12 | Lateral(C-D) noise | 0.17
| Omron | 2024-12-12 | Full range of Z Height during Z Scan sweep | 17
| Omron | 2024-12-12 | Full range of Z Detector during Z Scan sweep | 14.6
| Omron | 2024-12-12 | Full range of Y2 Detector raw | 14.4
| Omron | 2024-12-12 | Full range of Y1 Detector raw | 14.4
| Omron | 2024-12-12 | Full range of X2 Detector raw | 14.4
| Omron | 2024-12-12 | Full range of X1 Detector raw | 14.6
| Omron | 2024-12-12 | A-B noise | 7.18
| CXT #9 | 2024-12-10 | XY Stage positioning Y offset (Vision Pattern Matching) | 0.9424
| CXT #8 | 2024-12-10 | XY Stage positioning Y offset (Vision Pattern Matching)
| CXT #9 | 2024-12-10 | XY Stage positioning Y offset (LLC Translation) | 2.9932
| CXT #8 | 2024-12-10 | XY Stage positioning Y offset (LLC Translation) | 2.6416
| CXT #9 | 2024-12-10 | XY Stage positioning Y offset (In-Die Align) | 2.7832
| CXT #8 | 2024-12-10 | XY Stage positioning Y offset (In-Die Align) | 1.5479
| CXT #9 | 2024-12-10 | XY Stage positioning Y offset (Global Align) | 4.7949
| CXT #8 | 2024-12-10 | XY Stage positioning Y offset (Global Align) | 5.7373
| CXT #9 | 2024-12-10 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.17063
| CXT #8 | 2024-12-10 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.56146
| CXT #9 | 2024-12-10 | XY Stage positioning Y 1σ (LLC Translation) | 0.6581
| CXT #8 | 2024-12-10 | XY Stage positioning Y 1σ (LLC Translation) | 0.5662
| CXT #9 | 2024-12-10 | XY Stage positioning Y 1σ (In-Die Align) | 0.46308
| CXT #8 | 2024-12-10 | XY Stage positioning Y 1σ (In-Die Align) | 0.26238
| CXT #9 | 2024-12-10 | XY Stage positioning Y 1σ (Global Align) | 1.3913
| CXT #8 | 2024-12-10 | XY Stage positioning Y 1σ (Global Align) | 1.8067
| CXT #9 | 2024-12-10 | XY Stage positioning X offset (Vision Pattern Matching) | 1.1133
| CXT #8 | 2024-12-10 | XY Stage positioning X offset (Vision Pattern Matching)
| CXT #9 | 2024-12-10 | XY Stage positioning X offset (LLC Translation) | 1.3037
| CXT #8 | 2024-12-10 | XY Stage positioning X offset (LLC Translation) | 1.9287
| CXT #9 | 2024-12-10 | XY Stage positioning X offset (In-Die Align) | 1.5381
| CXT #8 | 2024-12-10 | XY Stage positioning X offset (In-Die Align) | 1.1328
| CXT #9 | 2024-12-10 | XY Stage positioning X offset (Global Align) | 2.0801
| CXT #8 | 2024-12-10 | XY Stage positioning X offset (Global Align) | 3.418
| CXT #9 | 2024-12-10 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.24029
| CXT #8 | 2024-12-10 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.69144
| CXT #9 | 2024-12-10 | XY Stage positioning X 1σ (LLC Translation) | 0.2578
| CXT #8 | 2024-12-10 | XY Stage positioning X 1σ (LLC Translation) | 0.4237
| CXT #9 | 2024-12-10 | XY Stage positioning X 1σ (In-Die Align) | 0.31286
| CXT #8 | 2024-12-10 | XY Stage positioning X 1σ (In-Die Align) | 0.23555
| CXT #9 | 2024-12-10 | XY Stage positioning X 1σ (Global Align) | 0.4951
| CXT #8 | 2024-12-10 | XY Stage positioning X 1σ (Global Align) | 0.89075
| CXT #8 | 2024-12-10 | Roughness Repeatability 80x80um(1σ) | 0.003667
| CXT #9 | 2024-12-10 | Roughness Repeatability 80x80um(1σ) | 0.002
| CXT #8 | 2024-12-10 | Roughness Repeatability 80x80um(RSD%) | 1.96
| CXT #9 | 2024-12-10 | Roughness Repeatability 80x80um(RSD%) | 1.66
| CXT #8 | 2024-12-10 | Roughness Repeatability 1x1um(1σ) | 0.000667
| CXT #9 | 2024-12-10 | Roughness Repeatability 1x1um(1σ) | 0.001
| CXT #8 | 2024-12-10 | Roughness Repeatability 1x1um(RSD%) | 0.652
| CXT #9 | 2024-12-10 | Roughness Repeatability 1x1um(RSD%) | 1.024
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Z Scanner position at -3μm | 0.0013
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Z Scanner position at +3μm | 0.003
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Z Linearity by Z Scanner position
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Z Linearity by 3 different sample height | 1.01
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Z Height noise | 0.031
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Y Scan Stroke | 136.7
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | -Y Direction software Limit | -100758
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Y Detector noise (XY Servo on) | 0.39
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Y Detector noise (XY Servo off) | 0.27
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Y axis OPM | 0.55
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | X Scan Stroke | 137.9
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | -X Direction software Limit | -100659
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | X Detector noise (XY Servo on) | 0.3
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | X Detector noise (XY Servo off) | 0.28
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | X axis OPM | 0.83
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.009
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.216
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.477
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.255
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.353
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.261
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.23
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.25
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.24
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.23
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.23
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | NCM driving% with AC160TS cantilever | 1.843
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | NCM Amplitude noise | 15.49
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.37
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.113
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 197.5
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 201
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Lateral(C-D) sensitivity | 20.5
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Lateral(C-D) noise | 0.107
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Full range of Z Height during Z Scan sweep | 18.5
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Full range of Z Detector during Z Scan sweep | 15.4
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Full range of Y Detector during Y Scan sweep | 136.7
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Full range of X Detector Raw during X Scan sweep | 14
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | Full range of X Detector during X Scan sweep | 137.9
| University of Shanghai for Science and Technology_College of Optoelectronics | 2024-12-06 | A-B noise | 8.57
| JSR Corp. | 2024-12-06 | Z Linearity by Z scanner position | 0.36
| JSR Corp. | 2024-12-06 | Z Height noise | 0.026
| JSR Corp. | 2024-12-06 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.5019
| JSR Corp. | 2024-12-06 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.9555
| JSR Corp. | 2024-12-06 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12.1036
| JSR Corp. | 2024-12-06 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.0924
| JSR Corp. | 2024-12-06 | Z Drift | 4.176
| JSR Corp. | 2024-12-06 | Z Detector offset | 3.7
| JSR Corp. | 2024-12-06 | Y2 Detector noise (XY Servo on) | 0.482
| JSR Corp. | 2024-12-06 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 256.8246
| JSR Corp. | 2024-12-06 | Y1 Detector noise (XY Servo on) | 0.449
| JSR Corp. | 2024-12-06 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 253.4652
| JSR Corp. | 2024-12-06 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 169.581
| JSR Corp. | 2024-12-06 | Y Ringing(Rpv) | 0.1
| JSR Corp. | 2024-12-06 | Y Non-linearity error | 1.73
| JSR Corp. | 2024-12-06 | Y Detector noise (XY Servo on) | 0.34
| JSR Corp. | 2024-12-06 | Y axis, Repeatability @1st position | 0.07
| JSR Corp. | 2024-12-06 | Y axis, Repeatability @ 4th position | 0.07
| JSR Corp. | 2024-12-06 | Y axis, Repeatability @ 3rd position | 0.06
| JSR Corp. | 2024-12-06 | Y axis, Repeatability @ 2nd position | 0.05
| JSR Corp. | 2024-12-06 | Y axis, Accuracy @1st position | 0.74
| JSR Corp. | 2024-12-06 | Y axis, Accuracy @ 4th position | 1.12
| JSR Corp. | 2024-12-06 | Y axis, Accuracy @ 3rd position | 0.79
| JSR Corp. | 2024-12-06 | Y axis, Accuracy @ 2nd position | 0.54
| JSR Corp. | 2024-12-06 | Y axis OPM | 0.913
| JSR Corp. | 2024-12-06 | XY Stage positioning Y offset (LLC Translation) | 2.3535
| JSR Corp. | 2024-12-06 | XY Stage positioning Y 1σ (LLC Translation) | 0.41326
| JSR Corp. | 2024-12-06 | XY Stage positioning X offset (LLC Translation) | 1.5723
| JSR Corp. | 2024-12-06 | XY Stage positioning X 1σ (LLC Translation) | 0.3657
| JSR Corp. | 2024-12-06 | XY Orthogonality | 89.83
| JSR Corp. | 2024-12-06 | X2 Detector noise (XY Servo on) | 0.438
| JSR Corp. | 2024-12-06 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 253.059
| JSR Corp. | 2024-12-06 | X1 Detector noise (XY Servo on) | 0.432
| JSR Corp. | 2024-12-06 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 242.0154
| JSR Corp. | 2024-12-06 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 169.028
| JSR Corp. | 2024-12-06 | X Ringing(Rpv) | 0.172
| JSR Corp. | 2024-12-06 | X Non-linearity error | 2.31
| JSR Corp. | 2024-12-06 | X Detector noise (XY Servo on) | 0.318
| JSR Corp. | 2024-12-06 | X axis, Repeatability @1st position | 0.07
| JSR Corp. | 2024-12-06 | X axis, Repeatability @ 4th position | 0.07
| JSR Corp. | 2024-12-06 | X axis, Repeatability @ 3rd position | 0.06
| JSR Corp. | 2024-12-06 | X axis, Repeatability @ 2nd position | 0.04
| JSR Corp. | 2024-12-06 | X axis, Accuracy @1st position | 0.58
| JSR Corp. | 2024-12-06 | X axis, Accuracy @ 4th position | 0.34
| JSR Corp. | 2024-12-06 | X axis, Accuracy @ 3rd position | 0.26
| JSR Corp. | 2024-12-06 | X axis, Accuracy @ 2nd position | 0.29
| JSR Corp. | 2024-12-06 | X axis OPM | 1.197
| JSR Corp. | 2024-12-06 | Wafer placement(Loading) error on the load port1 (Y offset) | 24.65
| JSR Corp. | 2024-12-06 | Wafer placement(Loading) error on the load port1 (X offset) | 19.45
| JSR Corp. | 2024-12-06 | Torque Profile (50mm, 20um/s) | 0.56485
| JSR Corp. | 2024-12-06 | Torque Profile (50mm, 200um/s) | 0.52331
| JSR Corp. | 2024-12-06 | StrokeYMin | -154762
| JSR Corp. | 2024-12-06 | StrokeYMax | 150542.6
| JSR Corp. | 2024-12-06 | StrokeY | 305304.7
| JSR Corp. | 2024-12-06 | StrokeXMin | -205359
| JSR Corp. | 2024-12-06 | StrokeXMax | 203379.8
| JSR Corp. | 2024-12-06 | StrokeX | 408738.6
| JSR Corp. | 2024-12-06 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.48077
| JSR Corp. | 2024-12-06 | Roll motion to X+ full scan | 10.4
| JSR Corp. | 2024-12-06 | RMS of ZeroScan | 0.04
| JSR Corp. | 2024-12-06 | Position repeatability(10repeats, Position Based) | 0.62
| JSR Corp. | 2024-12-06 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5.27e-05
| JSR Corp. | 2024-12-06 | OpticalAxisShiftPerFocusY | 0.0003
| JSR Corp. | 2024-12-06 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000445
| JSR Corp. | 2024-12-06 | OpticalAxisShiftPerFocusX | -0.0018
| JSR Corp. | 2024-12-06 | OPM Repeatability @5mm | 0.554
| JSR Corp. | 2024-12-06 | OPM Repeatability @1mm | 0.385
| JSR Corp. | 2024-12-06 | OPM Repeatability @10mm | 1.158
| JSR Corp. | 2024-12-06 | NCM driving% with AC160TS cantilever | 4.38
| JSR Corp. | 2024-12-06 | Min. Vibration Transmissbility (~100Hz) | 0.031
| JSR Corp. | 2024-12-06 | Max OPM @5mm | 10.232
| JSR Corp. | 2024-12-06 | Max OPM @1mm | 5.573
| JSR Corp. | 2024-12-06 | Max OPM @10mm | 20.374
| JSR Corp. | 2024-12-06 | LowerLimit | 2376.7
| JSR Corp. | 2024-12-06 | LowerLimit | 4396
| JSR Corp. | 2024-12-06 | Lateral(C-D) sensitivity | 19.1
| JSR Corp. | 2024-12-06 | Lateral(C-D) noise | 0.24665
| JSR Corp. | 2024-12-06 | Image offset after moving from -Y to +Y (Y scan) | 35.8887
| JSR Corp. | 2024-12-06 | Image offset after moving from -Y to +Y (X scan) | 21.9727
| JSR Corp. | 2024-12-06 | Image offset after moving from -X to +X (Y scan) | 23.4375
| JSR Corp. | 2024-12-06 | Image offset after moving from -X to +X (X scan) | 22.7051
| JSR Corp. | 2024-12-06 | Image offset after moving from +Y to -Y (Y scan) | 52.7344
| JSR Corp. | 2024-12-06 | Image offset after moving from +Y to -Y (X scan) | 25.6348
| JSR Corp. | 2024-12-06 | Image offset after moving from +X to -X (Y scan) | 16.8457
| JSR Corp. | 2024-12-06 | Image offset after moving from +X to -X (X scan) | 5.8594
| JSR Corp. | 2024-12-06 | Full range of Z Height during Z Scan sweep | 18.08
| JSR Corp. | 2024-12-06 | Full range of Z Detector during Z Scan sweep | 13.44
| JSR Corp. | 2024-12-06 | Full range of Y2 Detector Raw during Y Scan sweep | 14.71
| JSR Corp. | 2024-12-06 | Full range of Y2 Detector during Y Scan sweep | 206
| JSR Corp. | 2024-12-06 | Full range of Y1 Detector Raw during Y Scan sweep | 15.26
| JSR Corp. | 2024-12-06 | Full range of Y1 Detector during Y Scan sweep | 213.5
| JSR Corp. | 2024-12-06 | Full range of Y Detector during Y Scan sweep | 213.5
| JSR Corp. | 2024-12-06 | Full range of X2 Detector Raw during X Scan sweep | 15.9
| JSR Corp. | 2024-12-06 | Full range of X2 Detector during X Scan sweep | 223.3
| JSR Corp. | 2024-12-06 | Full range of X1 Detector Raw during X Scan sweep | 14.79
| JSR Corp. | 2024-12-06 | Full range of X1 Detector during X Scan sweep | 206.5
| JSR Corp. | 2024-12-06 | Full range of X Detector during X Scan sweep | 223.26
| JSR Corp. | 2024-12-06 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| JSR Corp. | 2024-12-06 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.046396
| JSR Corp. | 2024-12-06 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25611
| JSR Corp. | 2024-12-06 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.34836
| JSR Corp. | 2024-12-06 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.063093
| JSR Corp. | 2024-12-06 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14816
| JSR Corp. | 2024-12-06 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.81681
| JSR Corp. | 2024-12-06 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.59722
| JSR Corp. | 2024-12-06 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.10798
| JSR Corp. | 2024-12-05 | Roughness Repeatability 80x80um(1σ) | 0.002667
| JSR Corp. | 2024-12-05 | Roughness Repeatability 80x80um(RSD%) | 1.61
| JSR Corp. | 2024-12-05 | Roughness Repeatability 1x1um(1σ) | 0.001
| JSR Corp. | 2024-12-05 | Roughness Repeatability 1x1um(RSD%) | 0.895
| NVIDIA (Chuan Zhang) | 2024-12-02 | Z Scanner position at -3μm | 0.0023
| NVIDIA (Chuan Zhang) | 2024-12-02 | Z Scanner position at +3μm | 0.0012
| NVIDIA (Chuan Zhang) | 2024-12-02 | Z Linearity by Z Scanner position
| NVIDIA (Chuan Zhang) | 2024-12-02 | Z Linearity by 3 different sample height | 0.994
| NVIDIA (Chuan Zhang) | 2024-12-02 | Z Height noise | 0.027
| NVIDIA (Chuan Zhang) | 2024-12-02 | Y Scan Stroke | 118.8
| NVIDIA (Chuan Zhang) | 2024-12-02 | Y Detector noise (XY Servo on) | 0.31
| NVIDIA (Chuan Zhang) | 2024-12-02 | Y Detector noise (XY Servo off) | 0.32
| NVIDIA (Chuan Zhang) | 2024-12-02 | Y axis OPM | 0.69
| NVIDIA (Chuan Zhang) | 2024-12-02 | X Scan Stroke | 115.5
| NVIDIA (Chuan Zhang) | 2024-12-02 | X Detector noise (XY Servo on) | 0.28
| NVIDIA (Chuan Zhang) | 2024-12-02 | X Detector noise (XY Servo off) | 0.27
| NVIDIA (Chuan Zhang) | 2024-12-02 | X axis OPM | 1.25
| NVIDIA (Chuan Zhang) | 2024-12-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.115
| NVIDIA (Chuan Zhang) | 2024-12-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.006
| NVIDIA (Chuan Zhang) | 2024-12-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 5.923
| NVIDIA (Chuan Zhang) | 2024-12-02 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.173
| NVIDIA (Chuan Zhang) | 2024-12-02 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.256
| NVIDIA (Chuan Zhang) | 2024-12-02 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 8.72
| NVIDIA (Chuan Zhang) | 2024-12-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| NVIDIA (Chuan Zhang) | 2024-12-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| NVIDIA (Chuan Zhang) | 2024-12-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| NVIDIA (Chuan Zhang) | 2024-12-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| NVIDIA (Chuan Zhang) | 2024-12-02 | NCM driving% with AC160TS cantilever | 1.39
| NVIDIA (Chuan Zhang) | 2024-12-02 | NCM Amplitude noise | 10.98
| NVIDIA (Chuan Zhang) | 2024-12-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.024
| NVIDIA (Chuan Zhang) | 2024-12-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.419
| NVIDIA (Chuan Zhang) | 2024-12-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 172.2
| NVIDIA (Chuan Zhang) | 2024-12-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 175.5
| NVIDIA (Chuan Zhang) | 2024-12-02 | Lateral(C-D) sensitivity | 7.5
| NVIDIA (Chuan Zhang) | 2024-12-02 | Lateral(C-D) noise | 0.156
| NVIDIA (Chuan Zhang) | 2024-12-02 | Full range of Z Height during Z Scan sweep | 17.1
| NVIDIA (Chuan Zhang) | 2024-12-02 | Full range of Z Detector during Z Scan sweep | 13.5
| NVIDIA (Chuan Zhang) | 2024-12-02 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| NVIDIA (Chuan Zhang) | 2024-12-02 | Full range of Y Detector during Y Scan sweep | 118.2
| NVIDIA (Chuan Zhang) | 2024-12-02 | Full range of X Detector Raw during X Scan sweep | 13.2
| NVIDIA (Chuan Zhang) | 2024-12-02 | Full range of X Detector during X Scan sweep | 115.6
| NVIDIA (Chuan Zhang) | 2024-12-02 | A-B noise | 6.41
| Crystaland Co.,Ltd. | 2024-11-25 | Z Scanner position at -3μm | 0.012462
| Crystaland Co.,Ltd. | 2024-11-25 | Z Scanner position at +3μm | 0.004592
| Crystaland Co.,Ltd. | 2024-11-25 | Z Linearity by Z Scanner position
| Crystaland Co.,Ltd. | 2024-11-25 | Z Linearity by 3 different sample height | 1.0115
| Crystaland Co.,Ltd. | 2024-11-25 | Z Height noise | 0.028
| Crystaland Co.,Ltd. | 2024-11-25 | Y Scan Stroke | 137.5
| Crystaland Co.,Ltd. | 2024-11-25 | -Y Direction software Limit | -76931.9
| Crystaland Co.,Ltd. | 2024-11-25 | Y Detector noise (XY Servo on) | 0.282
| Crystaland Co.,Ltd. | 2024-11-25 | Y Detector noise (XY Servo off) | 0.278
| Crystaland Co.,Ltd. | 2024-11-25 | Y axis OPM | 0.533
| Crystaland Co.,Ltd. | 2024-11-25 | X Scan Stroke | 135.8
| Crystaland Co.,Ltd. | 2024-11-25 | -X Direction software Limit | -76485.6
| Crystaland Co.,Ltd. | 2024-11-25 | X Detector noise (XY Servo on) | 0.29
| Crystaland Co.,Ltd. | 2024-11-25 | X Detector noise (XY Servo off) | 0.294
| Crystaland Co.,Ltd. | 2024-11-25 | X axis OPM | 0.636
| Crystaland Co.,Ltd. | 2024-11-25 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.23
| Crystaland Co.,Ltd. | 2024-11-25 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.645
| Crystaland Co.,Ltd. | 2024-11-25 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.576
| Crystaland Co.,Ltd. | 2024-11-25 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.007
| Crystaland Co.,Ltd. | 2024-11-25 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.311
| Crystaland Co.,Ltd. | 2024-11-25 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.606
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.25
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.23
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.23
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.23
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.222
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| Crystaland Co.,Ltd. | 2024-11-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Crystaland Co.,Ltd. | 2024-11-25 | NCM driving% with AC160TS cantilever | 0.576
| Crystaland Co.,Ltd. | 2024-11-25 | NCM Amplitude noise | 15.29
| Crystaland Co.,Ltd. | 2024-11-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.887
| Crystaland Co.,Ltd. | 2024-11-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.404
| Crystaland Co.,Ltd. | 2024-11-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 196.1
| Crystaland Co.,Ltd. | 2024-11-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 197
| Crystaland Co.,Ltd. | 2024-11-25 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Crystaland Co.,Ltd. | 2024-11-25 | Lateral(C-D) sensitivity | 23.2
| Crystaland Co.,Ltd. | 2024-11-25 | Lateral(C-D) noise | 0.121
| Crystaland Co.,Ltd. | 2024-11-25 | Full range of Z Height during Z Scan sweep | 17.1
| Crystaland Co.,Ltd. | 2024-11-25 | Full range of Z Detector during Z Scan sweep | 14.6
| Crystaland Co.,Ltd. | 2024-11-25 | Full range of Y Detector Raw during Y Scan sweep | 14
| Crystaland Co.,Ltd. | 2024-11-25 | Full range of Y Detector during Y Scan sweep | 137.5
| Crystaland Co.,Ltd. | 2024-11-25 | Full range of X Detector Raw during X Scan sweep | 13.8
| Crystaland Co.,Ltd. | 2024-11-25 | Full range of X Detector during X Scan sweep | 135.8
| Crystaland Co.,Ltd. | 2024-11-25 | A-B noise | 13.24
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Z Scanner position at -3μm | 0.004059
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Z Scanner position at +3μm | 0.014506
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Z linearity by sample height | 1.003
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Z Height noise | 0.032
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Y Scan Stroke | 55
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Y Detector noise (XY Servo on) | 0.28
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Y Detector noise (XY Servo off) | 0.272
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Y axis OPM | 0.876
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | X Scan Stroke | 55
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | X Detector noise (XY Servo on) | 0.317
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | X Detector noise (XY Servo off) | 0.276
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | X axis OPM | 0.385
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.059
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.424
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.117
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.145
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.636
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.997
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | NCM driving% with AC160TS cantilever | 1.628
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.018
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.967
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.6
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.1
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Lateral(C-D) sensitivity | 22.2
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Lateral(C-D) noise | 0.113
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Full range of Z Height during Z Scan sweep | 17.3
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Full range of Z Detector during Z Scan sweep | 12.8
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Full range of Y Detector during Y Scan sweep | 70.7
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Full range of X Detector Raw during X Scan sweep | 13.6
| Charles University Prague, (CU Prague. Rejhon) | 2024-11-21 | Full range of X Detector during X Scan sweep | 69.3
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Z Scanner position at -3μm | 0.002099
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Z Scanner position at +3μm | 0.003436
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Z linearity by sample height | 1.0052
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Z Height noise | 0.029
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y2 Detector noise (XY Servo on) | 0.324
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y2 Detector noise (XY Servo off) | 0.317
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y1 Detector noise (XY Servo on) | 0.325
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y1 Detector noise (XY Servo off) | 0.322
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y Scan Stroke | 179
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y Detector noise (XY Servo on) | 0.234
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y Detector noise (XY Servo off) | 0.228
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Y axis OPM | 1.16
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X2 Detector noise (XY Servo on) | 0.341
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X2 Detector noise (XY Servo off) | 0.337
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X1 Detector noise (XY Servo on) | 0.381
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X1 Detector noise (XY Servo off) | 0.378
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X Scan Stroke | 177.6
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X Detector noise (XY Servo on) | 0.27
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X Detector noise (XY Servo off) | 0.264
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | X axis OPM | 1.4052
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.092
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.053
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.407
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.191
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.204
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.065
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.36
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.38
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.35
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.34
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | NCM Amplitude noise | 17.198
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.93
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.566
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 244.9
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 232.4
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 240.9
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 239.2
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Maximum Y Scan range (Full Stroke) | 178.4
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Maximum X Scan range (Full Stroke) | 178.4
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Lateral(C-D) sensitivity | 20.9
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Lateral(C-D) noise | 0.167
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Full range of Z Height during Z Scan sweep | 17.4
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Full range of Z Detector during Z Scan sweep | 13
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Full range of Y2 Detector raw | 15
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Full range of Y1 Detector raw | 15
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Full range of X2 Detector raw | 14.7
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | Full range of X1 Detector raw | 15
| 전남대학교_고분자융합소재공학부 | 2024-11-20 | A-B noise | 9.009
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Z Scanner position at -3μm | 0.004252
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Z Scanner position at +3μm | 0.008426
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Z linearity by sample height | 1.0135
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Z Height noise | 0.04
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y2 Detector noise (XY Servo on) | 0.245
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y2 Detector noise (XY Servo off) | 0.246
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y1 Detector noise (XY Servo on) | 0.233
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y1 Detector noise (XY Servo off) | 0.236
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y Scan Stroke | 177.3
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y Detector noise (XY Servo on) | 0.171
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y Detector noise (XY Servo off) | 0.173
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Y axis OPM | 0.9608
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X2 Detector noise (XY Servo on) | 0.249
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X2 Detector noise (XY Servo off) | 0.247
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X1 Detector noise (XY Servo on) | 0.244
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X1 Detector noise (XY Servo off) | 0.24
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X Scan Stroke | 178.1
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X Detector noise (XY Servo on) | 0.183
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X Detector noise (XY Servo off) | 0.181
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | X axis OPM | 0.8908
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.062
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.034
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.238
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.195
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.244
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.081
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.53
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.41
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.47
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.41
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.38
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.38
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | NCM Amplitude noise | 12.606
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.326
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.211
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 257
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 251.2
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 249.7
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 253.8
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Maximum Y Scan range (Full Stroke) | 175
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Maximum X Scan range (Full Stroke) | 173
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.09
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Lateral(C-D) sensitivity | 23.5
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Lateral(C-D) noise | 0.169
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Full range of Z Height during Z Scan sweep | 18.6
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Full range of Z Detector during Z Scan sweep | 14.2
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Full range of Y2 Detector raw | 13.6
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Full range of Y1 Detector raw | 13.9
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Full range of X2 Detector raw | 13.5
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | Full range of X1 Detector raw | 13.9
| HuaHong Semiconductor Manufacturing Wuxi (HHSW) | 2024-11-19 | A-B noise | 6.395
| BYD Chengdu | 2024-11-18 | Z Linearity by Z scanner position | 0.81
| BYD Chengdu | 2024-11-18 | Z Height noise | 0.025
| BYD Chengdu | 2024-11-18 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.8
| BYD Chengdu | 2024-11-18 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.9
| BYD Chengdu | 2024-11-18 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12.2
| BYD Chengdu | 2024-11-18 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.3
| BYD Chengdu | 2024-11-18 | Z Drift | 5.1
| BYD Chengdu | 2024-11-18 | Z Detector offset | 81.9
| BYD Chengdu | 2024-11-18 | Y2 Detector noise (XY Servo on) | 0.47
| BYD Chengdu | 2024-11-18 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 253.583
| BYD Chengdu | 2024-11-18 | Y1 Detector noise (XY Servo on) | 0.44
| BYD Chengdu | 2024-11-18 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 246.857
| BYD Chengdu | 2024-11-18 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 178.426
| BYD Chengdu | 2024-11-18 | Y Ringing(Rpv) | 0.38
| BYD Chengdu | 2024-11-18 | Y Non-linearity error | 2.53
| BYD Chengdu | 2024-11-18 | Y Detector noise (XY Servo on) | 0.33
| BYD Chengdu | 2024-11-18 | Y axis, Repeatability @1st position | 0.03
| BYD Chengdu | 2024-11-18 | Y axis, Repeatability @ 4th position | 0.03
| BYD Chengdu | 2024-11-18 | Y axis, Repeatability @ 3rd position | 0.03
| BYD Chengdu | 2024-11-18 | Y axis, Repeatability @ 2nd position | 0.03
| BYD Chengdu | 2024-11-18 | Y axis, Accuracy @1st position | 0.19
| BYD Chengdu | 2024-11-18 | Y axis, Accuracy @ 4th position | 0.23
| BYD Chengdu | 2024-11-18 | Y axis, Accuracy @ 3rd position | 0.18
| BYD Chengdu | 2024-11-18 | Y axis, Accuracy @ 2nd position | 0.18
| BYD Chengdu | 2024-11-18 | Y axis OPM | 1.29
| BYD Chengdu | 2024-11-18 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.4072
| BYD Chengdu | 2024-11-18 | XY Stage positioning Y offset (LLC Translation) | 1.6602
| BYD Chengdu | 2024-11-18 | XY Stage positioning Y offset (Global Align) | 3.3984
| BYD Chengdu | 2024-11-18 | XY Stage positioning Y 1σ (LLC Translation) | 0.3081
| BYD Chengdu | 2024-11-18 | XY Stage positioning X offset (Vision Pattern Matching) | 3.3643
| BYD Chengdu | 2024-11-18 | XY Stage positioning X offset (LLC Translation) | 1.4551
| BYD Chengdu | 2024-11-18 | XY Stage positioning X offset (Global Align) | 4.7559
| BYD Chengdu | 2024-11-18 | XY Stage positioning X 1σ (LLC Translation) | 0.2511
| BYD Chengdu | 2024-11-18 | XY Orthogonality | 90.02
| BYD Chengdu | 2024-11-18 | X2 Detector noise (XY Servo on) | 0.48
| BYD Chengdu | 2024-11-18 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 267.023
| BYD Chengdu | 2024-11-18 | X1 Detector noise (XY Servo on) | 0.46
| BYD Chengdu | 2024-11-18 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 251.381
| BYD Chengdu | 2024-11-18 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 176.076
| BYD Chengdu | 2024-11-18 | X Ringing(Rpv) | 0.08
| BYD Chengdu | 2024-11-18 | X Non-linearity error | 2.69
| BYD Chengdu | 2024-11-18 | X Detector noise (XY Servo on) | 0.35
| BYD Chengdu | 2024-11-18 | X axis, Repeatability @1st position | 0.05
| BYD Chengdu | 2024-11-18 | X axis, Repeatability @ 4th position | 0.04
| BYD Chengdu | 2024-11-18 | X axis, Repeatability @ 3rd position | 0.04
| BYD Chengdu | 2024-11-18 | X axis, Repeatability @ 2nd position | 0.06
| BYD Chengdu | 2024-11-18 | X axis, Accuracy @1st position | 0.22
| BYD Chengdu | 2024-11-18 | X axis, Accuracy @ 4th position | 0.17
| BYD Chengdu | 2024-11-18 | X axis, Accuracy @ 3rd position | 0.28
| BYD Chengdu | 2024-11-18 | X axis, Accuracy @ 2nd position | 0.29
| BYD Chengdu | 2024-11-18 | X axis OPM | 0.87
| BYD Chengdu | 2024-11-18 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.5
| BYD Chengdu | 2024-11-18 | Wafer placement(Loading) error on the load port2 (X offset) | 16.5
| BYD Chengdu | 2024-11-18 | Wafer placement(Loading) error on the load port1 (Y offset) | 28
| BYD Chengdu | 2024-11-18 | Wafer placement(Loading) error on the load port1 (X offset) | 15.8
| BYD Chengdu | 2024-11-18 | Torque Profile (50mm, 20um/s) | 0.404
| BYD Chengdu | 2024-11-18 | Torque Profile (50mm, 200um/s) | 0.366
| BYD Chengdu | 2024-11-18 | StrokeYMin | -152800
| BYD Chengdu | 2024-11-18 | StrokeYMax | 150000
| BYD Chengdu | 2024-11-18 | StrokeY | 302800
| BYD Chengdu | 2024-11-18 | StrokeXMin | -202841
| BYD Chengdu | 2024-11-18 | StrokeXMax | 202675
| BYD Chengdu | 2024-11-18 | StrokeX | 405516
| BYD Chengdu | 2024-11-18 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.8
| BYD Chengdu | 2024-11-18 | Roughness Repeatability 80x80um(1σ) | 0.001
| BYD Chengdu | 2024-11-18 | Roughness Repeatability 80x80um(RSD%) | 0.16
| BYD Chengdu | 2024-11-18 | Roughness Repeatability 1x1um(1σ) | 0.027
| BYD Chengdu | 2024-11-18 | Roughness Repeatability 1x1um(RSD%) | 14.389
| BYD Chengdu | 2024-11-18 | Roll motion to X+ full scan | 28.5
| BYD Chengdu | 2024-11-18 | RMS of ZeroScan | 0.072
| BYD Chengdu | 2024-11-18 | Position repeatability(10repeats, Position Based) | 0.93
| BYD Chengdu | 2024-11-18 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00013
| BYD Chengdu | 2024-11-18 | OpticalAxisShiftPerFocusY | 0.0021
| BYD Chengdu | 2024-11-18 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00022
| BYD Chengdu | 2024-11-18 | OpticalAxisShiftPerFocusX | -0.0086
| BYD Chengdu | 2024-11-18 | OPM Repeatability @5mm | 1.35
| BYD Chengdu | 2024-11-18 | OPM Repeatability @25mm | 2.11
| BYD Chengdu | 2024-11-18 | OPM Repeatability @1mm | 0.98
| BYD Chengdu | 2024-11-18 | OPM Repeatability @10mm | 1.2
| BYD Chengdu | 2024-11-18 | NCM driving% with AC160TS cantilever | 16.7
| BYD Chengdu | 2024-11-18 | Min. Vibration Transmissbility (~100Hz) | 0.075
| BYD Chengdu | 2024-11-18 | Max OPM @5mm | 16.7
| BYD Chengdu | 2024-11-18 | Max OPM @25mm | 65.15
| BYD Chengdu | 2024-11-18 | Max OPM @1mm | 10.15
| BYD Chengdu | 2024-11-18 | Max OPM @10mm | 16.09
| BYD Chengdu | 2024-11-18 | LowerLimit | 2265
| BYD Chengdu | 2024-11-18 | LowerLimit | 4313
| BYD Chengdu | 2024-11-18 | Lateral(C-D) sensitivity | 15.3
| BYD Chengdu | 2024-11-18 | Lateral(C-D) noise | 0.184
| BYD Chengdu | 2024-11-18 | Image offset after moving from -Y to +Y (Y scan) | 65.19
| BYD Chengdu | 2024-11-18 | Image offset after moving from -Y to +Y (X scan) | 9.52
| BYD Chengdu | 2024-11-18 | Image offset after moving from -X to +X (Y scan) | 33.69
| BYD Chengdu | 2024-11-18 | Image offset after moving from -X to +X (X scan) | 81.3
| BYD Chengdu | 2024-11-18 | Image offset after moving from +Y to -Y (Y scan) | 20.51
| BYD Chengdu | 2024-11-18 | Image offset after moving from +Y to -Y (X scan) | 7.32
| BYD Chengdu | 2024-11-18 | Image offset after moving from +X to -X (Y scan) | 10.99
| BYD Chengdu | 2024-11-18 | Image offset after moving from +X to -X (X scan) | 46.14
| BYD Chengdu | 2024-11-18 | Full range of Z Height during Z Scan sweep | 18.2
| BYD Chengdu | 2024-11-18 | Full range of Z Detector during Z Scan sweep | 14.7
| BYD Chengdu | 2024-11-18 | Full range of Y2 Detector Raw during Y Scan sweep | 14.55
| BYD Chengdu | 2024-11-18 | Full range of Y2 Detector during Y Scan sweep | 184.5
| BYD Chengdu | 2024-11-18 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| BYD Chengdu | 2024-11-18 | Full range of Y1 Detector during Y Scan sweep | 180.4
| BYD Chengdu | 2024-11-18 | Full range of Y Detector during Y Scan sweep | 182.5
| BYD Chengdu | 2024-11-18 | Full range of X2 Detector Raw during X Scan sweep | 14.29
| BYD Chengdu | 2024-11-18 | Full range of X2 Detector during X Scan sweep | 190.7
| BYD Chengdu | 2024-11-18 | Full range of X1 Detector Raw during X Scan sweep | 14.42
| BYD Chengdu | 2024-11-18 | Full range of X1 Detector during X Scan sweep | 181.3
| BYD Chengdu | 2024-11-18 | Full range of X Detector during X Scan sweep | 186
| BYD Chengdu | 2024-11-18 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| BYD Chengdu | 2024-11-18 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| BYD Chengdu | 2024-11-18 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| BYD Chengdu | 2024-11-18 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.2
| BYD Chengdu | 2024-11-18 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| BYD Chengdu | 2024-11-18 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| BYD Chengdu | 2024-11-18 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| BYD Chengdu | 2024-11-18 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.18
| BYD Chengdu | 2024-11-18 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Z Scanner position at -3μm | 0.003475
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Z Scanner position at +3μm | 0.015378
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Z linearity by sample height | 1.008
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Z Height noise | 0.03
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y2 Detector noise (XY Servo on) | 0.43
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y2 Detector noise (XY Servo off) | 0.41
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y1 Detector noise (XY Servo on) | 0.33
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y1 Detector noise (XY Servo off) | 0.31
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y Scan Stroke | 195.7
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | -Y Direction software Limit | -149379
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y Detector noise (XY Servo on) | 0.29
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y Detector noise (XY Servo off) | 0.27
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Y axis OPM | 2.52
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X2 Detector noise (XY Servo on) | 0.46
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X2 Detector noise (XY Servo off) | 0.44
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X1 Detector noise (XY Servo on) | 0.4
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X1 Detector noise (XY Servo off) | 0.39
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X Scan Stroke | 189.8
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | -X Direction software Limit | -149514
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X Detector noise (XY Servo on) | 0.32
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X Detector noise (XY Servo off) | 0.3
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | X axis OPM | 3.17
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.218
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.297
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.556
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.212
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.197
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 8.881
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.43
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.41
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.39
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.41
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.38
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | NCM driving% with AC160TS cantilever | 0.421
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.949
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.446
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 263.5
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 265.5
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 255.9
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 243
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Lateral(C-D) sensitivity | 15.9
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Lateral(C-D) noise | 0.148
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Z Height during Z Scan sweep | 17.2
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Z Detector during Z Scan sweep | 12.7
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.8
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Y2 Detector during Y Scan sweep | 190.7
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.8
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Y1 Detector during Y Scan sweep | 194.9
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of Y Detector during Y Scan sweep | 195.4
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of X2 Detector Raw during X Scan sweep | 15
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of X2 Detector during X Scan sweep | 189
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of X1 Detector Raw during X Scan sweep | 15.3
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of X1 Detector during X Scan sweep | 189.3
| Politehnica University of Bucharest (UPB, CSSNT) | 2024-11-14 | Full range of X Detector during X Scan sweep | 189.9
| 충남대학교_기계공학부 | 2024-11-13 | Z Scanner position at -3μm | 0.008149
| 충남대학교_기계공학부 | 2024-11-13 | Z Scanner position at +3μm | 0.003371
| 충남대학교_기계공학부 | 2024-11-13 | Z linearity by sample height | 1.0068
| 충남대학교_기계공학부 | 2024-11-13 | Z Height noise | 0.028
| 충남대학교_기계공학부 | 2024-11-13 | Y Scan Stroke | 55
| 충남대학교_기계공학부 | 2024-11-13 | Y Detector noise (XY Servo on) | 0.3991
| 충남대학교_기계공학부 | 2024-11-13 | Y Detector noise (XY Servo off) | 0.42
| 충남대학교_기계공학부 | 2024-11-13 | Y axis OPM | 0.3456
| 충남대학교_기계공학부 | 2024-11-13 | XY orthogonality | 89.98
| 충남대학교_기계공학부 | 2024-11-13 | X Scan Stroke | 55
| 충남대학교_기계공학부 | 2024-11-13 | X Detector noise (XY Servo on) | 0.389
| 충남대학교_기계공학부 | 2024-11-13 | X Detector noise (XY Servo off) | 0.421
| 충남대학교_기계공학부 | 2024-11-13 | X axis OPM | 0.4146
| 충남대학교_기계공학부 | 2024-11-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.102
| 충남대학교_기계공학부 | 2024-11-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.422
| 충남대학교_기계공학부 | 2024-11-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.341
| 충남대학교_기계공학부 | 2024-11-13 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.062
| 충남대학교_기계공학부 | 2024-11-13 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.545
| 충남대학교_기계공학부 | 2024-11-13 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.691
| 충남대학교_기계공학부 | 2024-11-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 충남대학교_기계공학부 | 2024-11-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 충남대학교_기계공학부 | 2024-11-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| 충남대학교_기계공학부 | 2024-11-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 충남대학교_기계공학부 | 2024-11-13 | NCM Amplitude noise | 14.27
| 충남대학교_기계공학부 | 2024-11-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.972
| 충남대학교_기계공학부 | 2024-11-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.89
| 충남대학교_기계공학부 | 2024-11-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100
| 충남대학교_기계공학부 | 2024-11-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.4
| 충남대학교_기계공학부 | 2024-11-13 | Lateral(C-D) sensitivity | 50.848
| 충남대학교_기계공학부 | 2024-11-13 | Lateral(C-D) noise | 0.053
| 충남대학교_기계공학부 | 2024-11-13 | Full range of Z Height during Z Scan sweep | 18.5
| 충남대학교_기계공학부 | 2024-11-13 | Full range of Z Detector during Z Scan sweep | 13.8
| 충남대학교_기계공학부 | 2024-11-13 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 충남대학교_기계공학부 | 2024-11-13 | Full range of Y Detector during Y Scan sweep | 69.5
| 충남대학교_기계공학부 | 2024-11-13 | Full range of X Detector Raw during X Scan sweep | 14.6
| 충남대학교_기계공학부 | 2024-11-13 | Full range of X Detector during X Scan sweep | 71.4
| 충남대학교_기계공학부 | 2024-11-13 | A-B noise | 8.623
| Indian Institute of Technology-HYD | 2024-11-08 | Z Scanner position at -3μm | 0.003905
| Istanbul Technology University (ITU) | 2024-11-08 | Z Scanner position at -3μm | 0.001143
| Indian Institute of Technology-HYD | 2024-11-08 | Z Scanner position at +3μm | 0.007705
| Istanbul Technology University (ITU) | 2024-11-08 | Z Scanner position at +3μm | 0.010475
| Indian Institute of Technology-HYD | 2024-11-08 | Z Linearity by Z Scanner position
| Istanbul Technology University (ITU) | 2024-11-08 | Z linearity by sample height | 1.0035
| Indian Institute of Technology-HYD | 2024-11-08 | Z Linearity by 3 different sample height | 0.99159
| Istanbul Technology University (ITU) | 2024-11-08 | Z Height noise | 0.03
| Indian Institute of Technology-HYD | 2024-11-08 | Z Height noise | 0.029
| Istanbul Technology University (ITU) | 2024-11-08 | Y Scan Stroke | 55
| Indian Institute of Technology-HYD | 2024-11-08 | Y Scan Stroke | 55
| Istanbul Technology University (ITU) | 2024-11-08 | Y Detector noise (XY Servo on) | 0.249
| Indian Institute of Technology-HYD | 2024-11-08 | Y Detector noise (XY Servo on) | 0.249
| Istanbul Technology University (ITU) | 2024-11-08 | Y Detector noise (XY Servo off) | 0.243
| Indian Institute of Technology-HYD | 2024-11-08 | Y Detector noise (XY Servo off) | 0.242
| Istanbul Technology University (ITU) | 2024-11-08 | Y axis OPM | 0.579
| Indian Institute of Technology-HYD | 2024-11-08 | Y axis OPM | 0.9938
| Istanbul Technology University (ITU) | 2024-11-08 | X Scan Stroke | 55
| Indian Institute of Technology-HYD | 2024-11-08 | X Scan Stroke | 55
| Istanbul Technology University (ITU) | 2024-11-08 | X Detector noise (XY Servo on) | 0.25
| Indian Institute of Technology-HYD | 2024-11-08 | X Detector noise (XY Servo on) | 0.244
| Istanbul Technology University (ITU) | 2024-11-08 | X Detector noise (XY Servo off) | 0.246
| Indian Institute of Technology-HYD | 2024-11-08 | X Detector noise (XY Servo off) | 0.237
| Istanbul Technology University (ITU) | 2024-11-08 | X axis OPM | 0.832
| Indian Institute of Technology-HYD | 2024-11-08 | X axis OPM | 0.7884
| Istanbul Technology University (ITU) | 2024-11-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 1.771
| Indian Institute of Technology-HYD | 2024-11-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.563
| Istanbul Technology University (ITU) | 2024-11-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.471
| Indian Institute of Technology-HYD | 2024-11-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.626
| Istanbul Technology University (ITU) | 2024-11-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 12.603
| Indian Institute of Technology-HYD | 2024-11-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.145
| Istanbul Technology University (ITU) | 2024-11-08 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.169
| Indian Institute of Technology-HYD | 2024-11-08 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.232
| Istanbul Technology University (ITU) | 2024-11-08 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.417
| Indian Institute of Technology-HYD | 2024-11-08 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.642
| Istanbul Technology University (ITU) | 2024-11-08 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.272
| Indian Institute of Technology-HYD | 2024-11-08 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.709
| Istanbul Technology University (ITU) | 2024-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Indian Institute of Technology-HYD | 2024-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Istanbul Technology University (ITU) | 2024-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Indian Institute of Technology-HYD | 2024-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Istanbul Technology University (ITU) | 2024-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Indian Institute of Technology-HYD | 2024-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Istanbul Technology University (ITU) | 2024-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Indian Institute of Technology-HYD | 2024-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Istanbul Technology University (ITU) | 2024-11-08 | NCM driving% with AC160TS cantilever | 2.248
| Indian Institute of Technology-HYD | 2024-11-08 | NCM driving% with AC160TS cantilever | 1.517
| Indian Institute of Technology-HYD | 2024-11-08 | NCM Amplitude noise | 11.223
| Istanbul Technology University (ITU) | 2024-11-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.349
| Indian Institute of Technology-HYD | 2024-11-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.845
| Istanbul Technology University (ITU) | 2024-11-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.261
| Indian Institute of Technology-HYD | 2024-11-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.127
| Istanbul Technology University (ITU) | 2024-11-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 92.3
| Indian Institute of Technology-HYD | 2024-11-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.9
| Istanbul Technology University (ITU) | 2024-11-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 90.3
| Indian Institute of Technology-HYD | 2024-11-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 94.9
| Istanbul Technology University (ITU) | 2024-11-08 | Lateral(C-D) sensitivity | 30.7
| Indian Institute of Technology-HYD | 2024-11-08 | Lateral(C-D) sensitivity | 61.1
| Istanbul Technology University (ITU) | 2024-11-08 | Lateral(C-D) noise | 0.069
| Indian Institute of Technology-HYD | 2024-11-08 | Lateral(C-D) noise | 0.059
| Istanbul Technology University (ITU) | 2024-11-08 | Full range of Z Height during Z Scan sweep | 16
| Indian Institute of Technology-HYD | 2024-11-08 | Full range of Z Height during Z Scan sweep | 17.3
| Istanbul Technology University (ITU) | 2024-11-08 | Full range of Z Detector during Z Scan sweep | 12.2
| Indian Institute of Technology-HYD | 2024-11-08 | Full range of Z Detector during Z Scan sweep | 13.2
| Istanbul Technology University (ITU) | 2024-11-08 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| Indian Institute of Technology-HYD | 2024-11-08 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Istanbul Technology University (ITU) | 2024-11-08 | Full range of Y Detector during Y Scan sweep | 69.4
| Indian Institute of Technology-HYD | 2024-11-08 | Full range of Y Detector during Y Scan sweep | 70.4
| Istanbul Technology University (ITU) | 2024-11-08 | Full range of X Detector Raw during X Scan sweep | 15.3
| Indian Institute of Technology-HYD | 2024-11-08 | Full range of X Detector Raw during X Scan sweep | 14.6
| Istanbul Technology University (ITU) | 2024-11-08 | Full range of X Detector during X Scan sweep | 69.2
| Indian Institute of Technology-HYD | 2024-11-08 | Full range of X Detector during X Scan sweep | 69.4
| Indian Institute of Technology-HYD | 2024-11-08 | A-B noise | 9.17
| Zhejiang Sci-Tech University | 2024-11-04 | Z Scanner position at -3μm | 0.001713
| Zhejiang Sci-Tech University | 2024-11-04 | Z Scanner position at +3μm | 0.008894
| Zhejiang Sci-Tech University | 2024-11-04 | Z linearity by sample height | 1.0048
| Zhejiang Sci-Tech University | 2024-11-04 | Z Height noise | 0.028
| Zhejiang Sci-Tech University | 2024-11-04 | Y Scan Stroke | 55
| Zhejiang Sci-Tech University | 2024-11-04 | Y Detector noise (XY Servo on) | 0.261
| Zhejiang Sci-Tech University | 2024-11-04 | Y Detector noise (XY Servo off) | 0.246
| Zhejiang Sci-Tech University | 2024-11-04 | Y axis OPM | 0.6026
| Zhejiang Sci-Tech University | 2024-11-04 | X Scan Stroke | 55
| Zhejiang Sci-Tech University | 2024-11-04 | X Detector noise (XY Servo on) | 0.245
| Zhejiang Sci-Tech University | 2024-11-04 | X Detector noise (XY Servo off) | 0.239
| Zhejiang Sci-Tech University | 2024-11-04 | X axis OPM | 0.8804
| Zhejiang Sci-Tech University | 2024-11-04 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.009
| Zhejiang Sci-Tech University | 2024-11-04 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.281
| Zhejiang Sci-Tech University | 2024-11-04 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.311
| Zhejiang Sci-Tech University | 2024-11-04 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.197
| Zhejiang Sci-Tech University | 2024-11-04 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.28
| Zhejiang Sci-Tech University | 2024-11-04 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 11.103
| Zhejiang Sci-Tech University | 2024-11-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Zhejiang Sci-Tech University | 2024-11-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Zhejiang Sci-Tech University | 2024-11-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Zhejiang Sci-Tech University | 2024-11-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Zhejiang Sci-Tech University | 2024-11-04 | NCM driving% with AC160TS cantilever | 0.822
| Zhejiang Sci-Tech University | 2024-11-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.42
| Zhejiang Sci-Tech University | 2024-11-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.049
| Zhejiang Sci-Tech University | 2024-11-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.3
| Zhejiang Sci-Tech University | 2024-11-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.1
| Zhejiang Sci-Tech University | 2024-11-04 | Lateral(C-D) sensitivity | 20.5
| Zhejiang Sci-Tech University | 2024-11-04 | Lateral(C-D) noise | 0.11
| Zhejiang Sci-Tech University | 2024-11-04 | Full range of Z Height during Z Scan sweep | 17.9
| Zhejiang Sci-Tech University | 2024-11-04 | Full range of Z Detector during Z Scan sweep | 13.9
| Zhejiang Sci-Tech University | 2024-11-04 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Zhejiang Sci-Tech University | 2024-11-04 | Full range of Y Detector during Y Scan sweep | 68.3
| Zhejiang Sci-Tech University | 2024-11-04 | Full range of X Detector Raw during X Scan sweep | 14.4
| Zhejiang Sci-Tech University | 2024-11-04 | Full range of X Detector during X Scan sweep | 68.4
| Hauhong Wuxi #5 | 2024-11-01 | Z Linearity by Z scanner position | 0.14
| Dongguan Guangmao Technology | 2024-11-01 | Z Linearity by Z scanner position | 0.1
| Hauhong Wuxi #5 | 2024-11-01 | Z Height noise | 0.024
| Dongguan Guangmao Technology | 2024-11-01 | Z Height noise | 0.026
| Hauhong Wuxi #5 | 2024-11-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.5
| Dongguan Guangmao Technology | 2024-11-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.5
| Hauhong Wuxi #5 | 2024-11-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.7
| Dongguan Guangmao Technology | 2024-11-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.5
| Hauhong Wuxi #5 | 2024-11-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| Dongguan Guangmao Technology | 2024-11-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12
| Hauhong Wuxi #5 | 2024-11-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.2
| Dongguan Guangmao Technology | 2024-11-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.8
| Hauhong Wuxi #5 | 2024-11-01 | Z Drift | 4.9
| Dongguan Guangmao Technology | 2024-11-01 | Z Drift | 4.9
| Hauhong Wuxi #5 | 2024-11-01 | Z Detector offset | 39.4
| Dongguan Guangmao Technology | 2024-11-01 | Z Detector offset | 37.9
| Hauhong Wuxi #5 | 2024-11-01 | Y2 Detector noise (XY Servo on) | 0.33
| Dongguan Guangmao Technology | 2024-11-01 | Y2 Detector noise (XY Servo on) | 0.41
| Hauhong Wuxi #5 | 2024-11-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 266.4
| Dongguan Guangmao Technology | 2024-11-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 250.3
| Hauhong Wuxi #5 | 2024-11-01 | Y1 Detector noise (XY Servo on) | 0.32
| Dongguan Guangmao Technology | 2024-11-01 | Y1 Detector noise (XY Servo on) | 0.38
| Hauhong Wuxi #5 | 2024-11-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 263.7
| Dongguan Guangmao Technology | 2024-11-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 254.2
| Hauhong Wuxi #5 | 2024-11-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 164.9
| Dongguan Guangmao Technology | 2024-11-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 175.7
| Hauhong Wuxi #5 | 2024-11-01 | Y Ringing(Rpv) | 0.1
| Dongguan Guangmao Technology | 2024-11-01 | Y Ringing(Rpv) | 0.09
| Hauhong Wuxi #5 | 2024-11-01 | Y Non-linearity error | 2.37
| Dongguan Guangmao Technology | 2024-11-01 | Y Non-linearity error | 2.18
| Hauhong Wuxi #5 | 2024-11-01 | Y Detector noise (XY Servo on) | 0.25
| Dongguan Guangmao Technology | 2024-11-01 | Y Detector noise (XY Servo on) | 0.29
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Repeatability @1st position | 0.03
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Repeatability @1st position | 0.05
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Repeatability @ 4th position | 0.03
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Repeatability @ 4th position | 0.04
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Repeatability @ 3rd position | 0.03
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Repeatability @ 3rd position | 0.05
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Repeatability @ 2nd position | 0.04
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Repeatability @ 2nd position | 0.05
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Accuracy @1st position | 0.2
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Accuracy @1st position | 0.26
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Accuracy @ 4th position | 0.24
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Accuracy @ 4th position | 0.47
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Accuracy @ 3rd position | 0.2
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Accuracy @ 3rd position | 0.29
| Hauhong Wuxi #5 | 2024-11-01 | Y axis, Accuracy @ 2nd position | 0.18
| Dongguan Guangmao Technology | 2024-11-01 | Y axis, Accuracy @ 2nd position | 0.24
| Hauhong Wuxi #5 | 2024-11-01 | Y axis OPM | 1.79
| Dongguan Guangmao Technology | 2024-11-01 | Y axis OPM | 1.76
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.6943
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.8066
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y offset (LLC Translation) | 3.1165
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y offset (LLC Translation) | 2.8027
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y offset (In-Die Align) | 1.5869
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y offset (In-Die Align) | 1.6211
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y offset (Global Align) | 2.6416
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y offset (Global Align) | 3.7354
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3824
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.37879
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y 1σ (LLC Translation) | 0.521
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y 1σ (LLC Translation) | 0.5165
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y 1σ (In-Die Align) | 0.27554
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y 1σ (In-Die Align) | 0.33252
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning Y 1σ (Global Align) | 0.68613
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning Y 1σ (Global Align) | 0.77527
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X offset (Vision Pattern Matching) | 1.1084
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X offset (Vision Pattern Matching) | 1.958
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X offset (LLC Translation) | 2.3779
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X offset (LLC Translation) | 2.2168
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X offset (In-Die Align) | 1.9434
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X offset (In-Die Align) | 1.0742
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X offset (Global Align) | 1.4551
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X offset (Global Align) | 3.3496
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.22382
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.41069
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X 1σ (LLC Translation) | 0.4294
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X 1σ (LLC Translation) | 0.4156
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X 1σ (In-Die Align) | 0.41068
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X 1σ (In-Die Align) | 0.21561
| Hauhong Wuxi #5 | 2024-11-01 | XY Stage positioning X 1σ (Global Align) | 0.35932
| Dongguan Guangmao Technology | 2024-11-01 | XY Stage positioning X 1σ (Global Align) | 0.73759
| Hauhong Wuxi #5 | 2024-11-01 | XY Orthogonality | 90.01
| Dongguan Guangmao Technology | 2024-11-01 | XY Orthogonality | 89.98
| Hauhong Wuxi #5 | 2024-11-01 | X2 Detector noise (XY Servo on) | 0.35
| Dongguan Guangmao Technology | 2024-11-01 | X2 Detector noise (XY Servo on) | 0.35
| Hauhong Wuxi #5 | 2024-11-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 246.7
| Dongguan Guangmao Technology | 2024-11-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 245.2
| Hauhong Wuxi #5 | 2024-11-01 | X1 Detector noise (XY Servo on) | 0.36
| Dongguan Guangmao Technology | 2024-11-01 | X1 Detector noise (XY Servo on) | 0.4
| Hauhong Wuxi #5 | 2024-11-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 252.8
| Dongguan Guangmao Technology | 2024-11-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 274.4
| Hauhong Wuxi #5 | 2024-11-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 170
| Dongguan Guangmao Technology | 2024-11-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 171.3
| Hauhong Wuxi #5 | 2024-11-01 | X Ringing(Rpv) | 0.08
| Dongguan Guangmao Technology | 2024-11-01 | X Ringing(Rpv) | 0.15
| Hauhong Wuxi #5 | 2024-11-01 | X Non-linearity error | 2.63
| Dongguan Guangmao Technology | 2024-11-01 | X Non-linearity error | 2.43
| Hauhong Wuxi #5 | 2024-11-01 | X Detector noise (XY Servo on) | 0.27
| Dongguan Guangmao Technology | 2024-11-01 | X Detector noise (XY Servo on) | 0.28
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Repeatability @1st position | 0.03
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Repeatability @1st position | 0.03
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Repeatability @ 4th position | 0.03
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Repeatability @ 4th position | 0.05
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Repeatability @ 3rd position | 0.05
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Repeatability @ 3rd position | 0.03
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Repeatability @ 2nd position | 0.05
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Repeatability @ 2nd position | 0.04
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Accuracy @1st position | 0.15
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Accuracy @1st position | 0.25
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Accuracy @ 4th position | 0.15
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Accuracy @ 4th position | 0.25
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Accuracy @ 3rd position | 0.26
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Accuracy @ 3rd position | 0.36
| Hauhong Wuxi #5 | 2024-11-01 | X axis, Accuracy @ 2nd position | 0.24
| Dongguan Guangmao Technology | 2024-11-01 | X axis, Accuracy @ 2nd position | 0.2
| Hauhong Wuxi #5 | 2024-11-01 | X axis OPM | 0.53
| Dongguan Guangmao Technology | 2024-11-01 | X axis OPM | 0.36
| Hauhong Wuxi #5 | 2024-11-01 | Wafer placement(Loading) error on the load port2 (Y offset) | 18.2
| Dongguan Guangmao Technology | 2024-11-01 | Wafer placement(Loading) error on the load port2 (Y offset) | 18.2
| Hauhong Wuxi #5 | 2024-11-01 | Wafer placement(Loading) error on the load port2 (X offset) | 23.8
| Dongguan Guangmao Technology | 2024-11-01 | Wafer placement(Loading) error on the load port2 (X offset) | 23.8
| Hauhong Wuxi #5 | 2024-11-01 | Wafer placement(Loading) error on the load port1 (Y offset) | 50.7
| Dongguan Guangmao Technology | 2024-11-01 | Wafer placement(Loading) error on the load port1 (Y offset) | 50.7
| Hauhong Wuxi #5 | 2024-11-01 | Wafer placement(Loading) error on the load port1 (X offset) | 48.9
| Dongguan Guangmao Technology | 2024-11-01 | Wafer placement(Loading) error on the load port1 (X offset) | 48.9
| Hauhong Wuxi #5 | 2024-11-01 | Torque Profile (50mm, 20um/s) | 0.781
| Dongguan Guangmao Technology | 2024-11-01 | Torque Profile (50mm, 20um/s) | 0.842
| Hauhong Wuxi #5 | 2024-11-01 | Torque Profile (50mm, 200um/s) | 0.474
| Dongguan Guangmao Technology | 2024-11-01 | Torque Profile (50mm, 200um/s) | 0.666
| Hauhong Wuxi #5 | 2024-11-01 | StrokeYMin | -153450
| Dongguan Guangmao Technology | 2024-11-01 | StrokeYMin | -152981
| Hauhong Wuxi #5 | 2024-11-01 | StrokeYMax | 150439.8
| Dongguan Guangmao Technology | 2024-11-01 | StrokeYMax | 150053
| Hauhong Wuxi #5 | 2024-11-01 | StrokeY | 303889.3
| Dongguan Guangmao Technology | 2024-11-01 | StrokeY | 303034
| Hauhong Wuxi #5 | 2024-11-01 | StrokeXMin | -204117
| Dongguan Guangmao Technology | 2024-11-01 | StrokeXMin | -203673
| Hauhong Wuxi #5 | 2024-11-01 | StrokeXMax | 203707.6
| Dongguan Guangmao Technology | 2024-11-01 | StrokeXMax | 203073
| Hauhong Wuxi #5 | 2024-11-01 | StrokeX | 407824.5
| Dongguan Guangmao Technology | 2024-11-01 | StrokeX | 406746
| Hauhong Wuxi #5 | 2024-11-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.4
| Dongguan Guangmao Technology | 2024-11-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.4
| Dongguan Guangmao Technology | 2024-11-01 | Roughness Repeatability 80x80um(1σ) | 0.02
| Dongguan Guangmao Technology | 2024-11-01 | Roughness Repeatability 80x80um(RSD%) | 8.98
| Dongguan Guangmao Technology | 2024-11-01 | Roughness Repeatability 1x1um(1σ) | 0.013333
| Dongguan Guangmao Technology | 2024-11-01 | Roughness Repeatability 1x1um(RSD%) | 10.16
| Hauhong Wuxi #5 | 2024-11-01 | Roll motion to X+ full scan | 6.1
| Dongguan Guangmao Technology | 2024-11-01 | Roll motion to X+ full scan | 0.7
| Hauhong Wuxi #5 | 2024-11-01 | RMS of ZeroScan | 0.045
| Dongguan Guangmao Technology | 2024-11-01 | RMS of ZeroScan | 0.049
| Hauhong Wuxi #5 | 2024-11-01 | Position repeatability(10repeats, Position Based) | 0.28
| Dongguan Guangmao Technology | 2024-11-01 | Position repeatability(10repeats, Position Based) | 0.3
| Hauhong Wuxi #5 | 2024-11-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Dongguan Guangmao Technology | 2024-11-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00012
| Hauhong Wuxi #5 | 2024-11-01 | OpticalAxisShiftPerFocusY | -0.0011
| Dongguan Guangmao Technology | 2024-11-01 | OpticalAxisShiftPerFocusY | 0.0007
| Hauhong Wuxi #5 | 2024-11-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| Dongguan Guangmao Technology | 2024-11-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| Hauhong Wuxi #5 | 2024-11-01 | OpticalAxisShiftPerFocusX | -0.0053
| Dongguan Guangmao Technology | 2024-11-01 | OpticalAxisShiftPerFocusX | -0.0003
| Hauhong Wuxi #5 | 2024-11-01 | OPM Repeatability @5mm | 0.76
| Dongguan Guangmao Technology | 2024-11-01 | OPM Repeatability @5mm | 0.3
| Hauhong Wuxi #5 | 2024-11-01 | OPM Repeatability @25mm | 1.6
| Dongguan Guangmao Technology | 2024-11-01 | OPM Repeatability @25mm | 1.36
| Hauhong Wuxi #5 | 2024-11-01 | OPM Repeatability @1mm | 0.32
| Dongguan Guangmao Technology | 2024-11-01 | OPM Repeatability @1mm | 0.27
| Hauhong Wuxi #5 | 2024-11-01 | OPM Repeatability @10mm | 0.98
| Dongguan Guangmao Technology | 2024-11-01 | OPM Repeatability @10mm | 0.61
| Hauhong Wuxi #5 | 2024-11-01 | NCM driving% with AC160TS cantilever | 2.13
| Dongguan Guangmao Technology | 2024-11-01 | NCM driving% with AC160TS cantilever | 3
| Hauhong Wuxi #5 | 2024-11-01 | Min. Vibration Transmissbility (~100Hz) | 0.036
| Dongguan Guangmao Technology | 2024-11-01 | Min. Vibration Transmissbility (~100Hz) | 0.069
| Hauhong Wuxi #5 | 2024-11-01 | Max OPM @5mm | 17.17
| Dongguan Guangmao Technology | 2024-11-01 | Max OPM @5mm | 11.79
| Hauhong Wuxi #5 | 2024-11-01 | Max OPM @25mm | 48.14
| Dongguan Guangmao Technology | 2024-11-01 | Max OPM @25mm | 79.69
| Hauhong Wuxi #5 | 2024-11-01 | Max OPM @1mm | 2.42
| Dongguan Guangmao Technology | 2024-11-01 | Max OPM @1mm | 3.51
| Hauhong Wuxi #5 | 2024-11-01 | Max OPM @10mm | 17.83
| Dongguan Guangmao Technology | 2024-11-01 | Max OPM @10mm | 16.26
| Hauhong Wuxi #5 | 2024-11-01 | LowerLimit | 1679.4
| Hauhong Wuxi #5 | 2024-11-01 | LowerLimit | 4758
| Dongguan Guangmao Technology | 2024-11-01 | LowerLimit | 2332
| Dongguan Guangmao Technology | 2024-11-01 | LowerLimit | 5229
| Hauhong Wuxi #5 | 2024-11-01 | Lateral(C-D) sensitivity | 32.6
| Dongguan Guangmao Technology | 2024-11-01 | Lateral(C-D) sensitivity | 38.8
| Hauhong Wuxi #5 | 2024-11-01 | Lateral(C-D) noise | 0.203
| Dongguan Guangmao Technology | 2024-11-01 | Lateral(C-D) noise | 0.244
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from -Y to +Y (Y scan) | 98.14
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from -Y to +Y (Y scan) | 77.64
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from -Y to +Y (X scan) | 8.79
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from -Y to +Y (X scan) | 34.42
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from -X to +X (Y scan) | 44.47
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from -X to +X (Y scan) | 47.61
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from -X to +X (X scan) | 50.54
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from -X to +X (X scan) | 51.27
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from +Y to -Y (Y scan) | 71.04
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from +Y to -Y (Y scan) | 77.64
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from +Y to -Y (X scan) | 15.38
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from +Y to -Y (X scan) | 30.76
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from +X to -X (Y scan) | 19.78
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from +X to -X (Y scan) | 17.58
| Hauhong Wuxi #5 | 2024-11-01 | Image offset after moving from +X to -X (X scan) | 62.99
| Dongguan Guangmao Technology | 2024-11-01 | Image offset after moving from +X to -X (X scan) | 38.09
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Z Height during Z Scan sweep | 18.5
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Z Height during Z Scan sweep | 18.6
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Z Detector during Z Scan sweep | 15
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Z Detector during Z Scan sweep | 13.4
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.93
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.46
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Y2 Detector during Y Scan sweep | 202
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Y2 Detector during Y Scan sweep | 202.5
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.69
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.31
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Y1 Detector during Y Scan sweep | 198.8
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Y1 Detector during Y Scan sweep | 193.2
| Hauhong Wuxi #5 | 2024-11-01 | Full range of Y Detector during Y Scan sweep | 198.8
| Dongguan Guangmao Technology | 2024-11-01 | Full range of Y Detector during Y Scan sweep | 202.5
| Hauhong Wuxi #5 | 2024-11-01 | Full range of X2 Detector Raw during X Scan sweep | 14.83
| Dongguan Guangmao Technology | 2024-11-01 | Full range of X2 Detector Raw during X Scan sweep | 14.48
| Hauhong Wuxi #5 | 2024-11-01 | Full range of X2 Detector during X Scan sweep | 181.7
| Dongguan Guangmao Technology | 2024-11-01 | Full range of X2 Detector during X Scan sweep | 202.9
| Hauhong Wuxi #5 | 2024-11-01 | Full range of X1 Detector Raw during X Scan sweep | 15.01
| Dongguan Guangmao Technology | 2024-11-01 | Full range of X1 Detector Raw during X Scan sweep | 14.46
| Hauhong Wuxi #5 | 2024-11-01 | Full range of X1 Detector during X Scan sweep | 197.7
| Dongguan Guangmao Technology | 2024-11-01 | Full range of X1 Detector during X Scan sweep | 195.4
| Hauhong Wuxi #5 | 2024-11-01 | Full range of X Detector during X Scan sweep | 181.7
| Dongguan Guangmao Technology | 2024-11-01 | Full range of X Detector during X Scan sweep | 202.9
| Hauhong Wuxi #5 | 2024-11-01 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Dongguan Guangmao Technology | 2024-11-01 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Hauhong Wuxi #5 | 2024-11-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Dongguan Guangmao Technology | 2024-11-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Hauhong Wuxi #5 | 2024-11-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.62
| Dongguan Guangmao Technology | 2024-11-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.62
| Hauhong Wuxi #5 | 2024-11-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.46
| Dongguan Guangmao Technology | 2024-11-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.8
| Hauhong Wuxi #5 | 2024-11-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Dongguan Guangmao Technology | 2024-11-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| Hauhong Wuxi #5 | 2024-11-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Dongguan Guangmao Technology | 2024-11-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Hauhong Wuxi #5 | 2024-11-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Dongguan Guangmao Technology | 2024-11-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| Hauhong Wuxi #5 | 2024-11-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.44
| Dongguan Guangmao Technology | 2024-11-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| Hauhong Wuxi #5 | 2024-11-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Dongguan Guangmao Technology | 2024-11-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| 강원대학교_물리학과 | 2024-10-30 | Z Scanner position at -3μm | 0.000959
| 강원대학교_물리학과 | 2024-10-30 | Z Scanner position at +3μm | 0.001076
| 강원대학교_물리학과 | 2024-10-30 | Z linearity by sample height | 1.0037
| 강원대학교_물리학과 | 2024-10-30 | Z Height noise | 0.016
| 강원대학교_물리학과 | 2024-10-30 | Y Scan Stroke | 108.6
| 강원대학교_물리학과 | 2024-10-30 | Y Detector noise (XY Servo on) | 0.254
| 강원대학교_물리학과 | 2024-10-30 | Y Detector noise (XY Servo off) | 0.257
| 강원대학교_물리학과 | 2024-10-30 | Y axis OPM | 0.6122
| 강원대학교_물리학과 | 2024-10-30 | XY orthogonality | 90.01
| 강원대학교_물리학과 | 2024-10-30 | X Scan Stroke | 106.6
| 강원대학교_물리학과 | 2024-10-30 | X Detector noise (XY Servo on) | 0.255
| 강원대학교_물리학과 | 2024-10-30 | X Detector noise (XY Servo off) | 0.247
| 강원대학교_물리학과 | 2024-10-30 | X axis OPM | 1.3676
| 강원대학교_물리학과 | 2024-10-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.017
| 강원대학교_물리학과 | 2024-10-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.759
| 강원대학교_물리학과 | 2024-10-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.75
| 강원대학교_물리학과 | 2024-10-30 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.19
| 강원대학교_물리학과 | 2024-10-30 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.724
| 강원대학교_물리학과 | 2024-10-30 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.93
| 강원대학교_물리학과 | 2024-10-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 강원대학교_물리학과 | 2024-10-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 강원대학교_물리학과 | 2024-10-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 강원대학교_물리학과 | 2024-10-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 강원대학교_물리학과 | 2024-10-30 | NCM driving% with AC160TS cantilever | 2.101
| 강원대학교_물리학과 | 2024-10-30 | NCM Amplitude noise | 5.788
| 강원대학교_물리학과 | 2024-10-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.505
| 강원대학교_물리학과 | 2024-10-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.011
| 강원대학교_물리학과 | 2024-10-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 147.7
| 강원대학교_물리학과 | 2024-10-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 150.3
| 강원대학교_물리학과 | 2024-10-30 | Lateral(C-D) sensitivity | 6.4
| 강원대학교_물리학과 | 2024-10-30 | Lateral(C-D) noise | 0.173
| 강원대학교_물리학과 | 2024-10-30 | Full range of Z Height during Z Scan sweep | 16.7
| 강원대학교_물리학과 | 2024-10-30 | Full range of Z Detector during Z Scan sweep | 15.9
| 강원대학교_물리학과 | 2024-10-30 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| 강원대학교_물리학과 | 2024-10-30 | Full range of Y Detector during Y Scan sweep | 108.4
| 강원대학교_물리학과 | 2024-10-30 | Full range of X Detector Raw during X Scan sweep | 14.1
| 강원대학교_물리학과 | 2024-10-30 | Full range of X Detector during X Scan sweep | 106.3
| 강원대학교_물리학과 | 2024-10-30 | A-B noise | 15.643
| 한양대학교_화학과 | 2024-10-28 | Z Scanner position at -3μm | 0.004071
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Z Scanner position at -3μm | 0.001948
| Nanyang Technological University_CEE | 2024-10-28 | Z Scanner position at -3μm | 0.002271
| 한양대학교_화학과 | 2024-10-28 | Z Scanner position at +3μm | 0.01708
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Z Scanner position at +3μm | 0.003071
| Nanyang Technological University_CEE | 2024-10-28 | Z Scanner position at +3μm | 0.007922
| 한양대학교_화학과 | 2024-10-28 | Z linearity by sample height | 1.0041
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Z linearity by sample height | 0.99731
| Nanyang Technological University_CEE | 2024-10-28 | Z linearity by sample height | 0.99669
| 한양대학교_화학과 | 2024-10-28 | Z Height noise | 0.033
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Z Height noise | 0.03
| Nanyang Technological University_CEE | 2024-10-28 | Z Height noise | 0.037
| 한양대학교_화학과 | 2024-10-28 | Y Scan Stroke | 55
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Y Scan Stroke | 55
| Nanyang Technological University_CEE | 2024-10-28 | Y Scan Stroke | 55
| 한양대학교_화학과 | 2024-10-28 | Y Detector noise (XY Servo on) | 0.274
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Y Detector noise (XY Servo on) | 0.13
| Nanyang Technological University_CEE | 2024-10-28 | Y Detector noise (XY Servo on) | 0.256
| 한양대학교_화학과 | 2024-10-28 | Y Detector noise (XY Servo off) | 0.267
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Y Detector noise (XY Servo off) | 0.11
| Nanyang Technological University_CEE | 2024-10-28 | Y Detector noise (XY Servo off) | 0.247
| 한양대학교_화학과 | 2024-10-28 | Y axis OPM | 0.6098
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Y axis OPM | 0.3449
| Nanyang Technological University_CEE | 2024-10-28 | Y axis OPM | 1.6728
| 한양대학교_화학과 | 2024-10-28 | X Scan Stroke | 55
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | X Scan Stroke | 55
| Nanyang Technological University_CEE | 2024-10-28 | X Scan Stroke | 55
| 한양대학교_화학과 | 2024-10-28 | X Detector noise (XY Servo on) | 0.271
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | X Detector noise (XY Servo on) | 0.23
| Nanyang Technological University_CEE | 2024-10-28 | X Detector noise (XY Servo on) | 0.259
| 한양대학교_화학과 | 2024-10-28 | X Detector noise (XY Servo off) | 0.261
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | X Detector noise (XY Servo off) | 0.22
| Nanyang Technological University_CEE | 2024-10-28 | X Detector noise (XY Servo off) | 0.248
| 한양대학교_화학과 | 2024-10-28 | X axis OPM | 0.7742
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | X axis OPM | 0.5663
| Nanyang Technological University_CEE | 2024-10-28 | X axis OPM | 1.5484
| 한양대학교_화학과 | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.023
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.016
| Nanyang Technological University_CEE | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.022
| 한양대학교_화학과 | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.715
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.975
| Nanyang Technological University_CEE | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.72
| 한양대학교_화학과 | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.022
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.436
| Nanyang Technological University_CEE | 2024-10-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 5.948
| 한양대학교_화학과 | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.194
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.227
| Nanyang Technological University_CEE | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.134
| 한양대학교_화학과 | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.505
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.412
| Nanyang Technological University_CEE | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.372
| 한양대학교_화학과 | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.067
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.437
| Nanyang Technological University_CEE | 2024-10-28 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.087
| 한양대학교_화학과 | 2024-10-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Nanyang Technological University_CEE | 2024-10-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 한양대학교_화학과 | 2024-10-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| Nanyang Technological University_CEE | 2024-10-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 한양대학교_화학과 | 2024-10-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| Nanyang Technological University_CEE | 2024-10-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 한양대학교_화학과 | 2024-10-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.13
| Nanyang Technological University_CEE | 2024-10-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| 한양대학교_화학과 | 2024-10-28 | NCM driving% with AC160TS cantilever | 0.309
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | NCM driving% with AC160TS cantilever | 0.432
| Nanyang Technological University_CEE | 2024-10-28 | NCM driving% with AC160TS cantilever | 0.428
| 한양대학교_화학과 | 2024-10-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.142
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.619
| Nanyang Technological University_CEE | 2024-10-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.98
| 한양대학교_화학과 | 2024-10-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.17
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.599
| Nanyang Technological University_CEE | 2024-10-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.147
| 한양대학교_화학과 | 2024-10-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 112.9
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 111.1
| Nanyang Technological University_CEE | 2024-10-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.7
| 한양대학교_화학과 | 2024-10-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 111.8
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 107.3
| Nanyang Technological University_CEE | 2024-10-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.9
| 한양대학교_화학과 | 2024-10-28 | Lateral(C-D) sensitivity | 19.7
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Lateral(C-D) sensitivity | 24.2
| Nanyang Technological University_CEE | 2024-10-28 | Lateral(C-D) sensitivity | 39.7
| 한양대학교_화학과 | 2024-10-28 | Lateral(C-D) noise | 0.06
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Lateral(C-D) noise | 0.097
| Nanyang Technological University_CEE | 2024-10-28 | Lateral(C-D) noise | 0.1
| 한양대학교_화학과 | 2024-10-28 | Full range of Z Height during Z Scan sweep | 17.5
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Full range of Z Height during Z Scan sweep | 17.2
| Nanyang Technological University_CEE | 2024-10-28 | Full range of Z Height during Z Scan sweep | 19.3
| 한양대학교_화학과 | 2024-10-28 | Full range of Z Detector during Z Scan sweep | 12.9
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Full range of Z Detector during Z Scan sweep | 12.5
| Nanyang Technological University_CEE | 2024-10-28 | Full range of Z Detector during Z Scan sweep | 13.3
| 한양대학교_화학과 | 2024-10-28 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Nanyang Technological University_CEE | 2024-10-28 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| 한양대학교_화학과 | 2024-10-28 | Full range of Y Detector during Y Scan sweep | 74.2
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Full range of Y Detector during Y Scan sweep | 77
| Nanyang Technological University_CEE | 2024-10-28 | Full range of Y Detector during Y Scan sweep | 75
| 한양대학교_화학과 | 2024-10-28 | Full range of X Detector Raw during X Scan sweep | 13.1
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Full range of X Detector Raw during X Scan sweep | 13.9
| Nanyang Technological University_CEE | 2024-10-28 | Full range of X Detector Raw during X Scan sweep | 14.1
| 한양대학교_화학과 | 2024-10-28 | Full range of X Detector during X Scan sweep | 73.1
| Park Systems Singapore(PSS) SID No. S10014-280624 | 2024-10-28 | Full range of X Detector during X Scan sweep | 74.8
| Nanyang Technological University_CEE | 2024-10-28 | Full range of X Detector during X Scan sweep | 72
| Huawei Technologies | 2024-10-26 | Roughness Repeatability 80x80um(1σ) | 0.001333
| Huawei Technologies | 2024-10-26 | Roughness Repeatability 80x80um(RSD%) | 0.54
| Huawei Technologies | 2024-10-26 | Roughness Repeatability 1x1um(1σ) | 0.000667
| Huawei Technologies | 2024-10-26 | Roughness Repeatability 1x1um(RSD%) | 1.79
| Hsingchen Technology | 2024-10-25 | Z Linearity by Z scanner position | 0.25
| Hsingchen Technology | 2024-10-25 | Z Height noise | 0.027
| Hsingchen Technology | 2024-10-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.9
| Hsingchen Technology | 2024-10-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.8
| Hsingchen Technology | 2024-10-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| Hsingchen Technology | 2024-10-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.7
| Hsingchen Technology | 2024-10-25 | Z Drift | 8
| Hsingchen Technology | 2024-10-25 | Z Detector offset | 74.6
| Hsingchen Technology | 2024-10-25 | Y2 Detector noise (XY Servo on) | 0.461
| Hsingchen Technology | 2024-10-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 267.1
| Hsingchen Technology | 2024-10-25 | Y1 Detector noise (XY Servo on) | 0.474
| Hsingchen Technology | 2024-10-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 251
| Hsingchen Technology | 2024-10-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 173.3
| Hsingchen Technology | 2024-10-25 | Y Ringing(Rpv) | 0.134
| Hsingchen Technology | 2024-10-25 | Y Non-linearity error | 3.15
| Hsingchen Technology | 2024-10-25 | Y Detector noise (XY Servo on) | 0.335
| Hsingchen Technology | 2024-10-25 | Y axis, Repeatability @1st position | 0.06
| Hsingchen Technology | 2024-10-25 | Y axis, Repeatability @ 4th position | 0.05
| Hsingchen Technology | 2024-10-25 | Y axis, Repeatability @ 3rd position | 0.04
| Hsingchen Technology | 2024-10-25 | Y axis, Repeatability @ 2nd position | 0.06
| Hsingchen Technology | 2024-10-25 | Y axis, Accuracy @1st position | 0.24
| Hsingchen Technology | 2024-10-25 | Y axis, Accuracy @ 4th position | 0.25
| Hsingchen Technology | 2024-10-25 | Y axis, Accuracy @ 3rd position | 0.19
| Hsingchen Technology | 2024-10-25 | Y axis, Accuracy @ 2nd position | 0.3
| Hsingchen Technology | 2024-10-25 | Y axis OPM | 1.39
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.5869
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y offset (LLC Translation) | 2.1143
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y offset (In-Die Align) | 1.2305
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y offset (Global Align) | 3.999
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.27274
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y 1σ (LLC Translation) | 0.4149
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y 1σ (In-Die Align) | 0.28149
| Hsingchen Technology | 2024-10-25 | XY Stage positioning Y 1σ (Global Align) | 0.96871
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X offset (Vision Pattern Matching) | 2.9785
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X offset (LLC Translation) | 2.1777
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X offset (In-Die Align) | 2.4561
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X offset (Global Align) | 2.6416
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.59173
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X 1σ (LLC Translation) | 0.439
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X 1σ (In-Die Align) | 0.48095
| Hsingchen Technology | 2024-10-25 | XY Stage positioning X 1σ (Global Align) | 0.54516
| Hsingchen Technology | 2024-10-25 | XY Orthogonality | 89.9
| Hsingchen Technology | 2024-10-25 | X2 Detector noise (XY Servo on) | 0.514
| Hsingchen Technology | 2024-10-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 279.2
| Hsingchen Technology | 2024-10-25 | X1 Detector noise (XY Servo on) | 0.413
| Hsingchen Technology | 2024-10-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 268.9
| Hsingchen Technology | 2024-10-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 177
| Hsingchen Technology | 2024-10-25 | X Ringing(Rpv) | 0.093
| Hsingchen Technology | 2024-10-25 | X Non-linearity error | 2.75
| Hsingchen Technology | 2024-10-25 | X Detector noise (XY Servo on) | 0.359
| Hsingchen Technology | 2024-10-25 | X axis, Repeatability @1st position | 0.08
| Hsingchen Technology | 2024-10-25 | X axis, Repeatability @ 4th position | 0.07
| Hsingchen Technology | 2024-10-25 | X axis, Repeatability @ 3rd position | 0.05
| Hsingchen Technology | 2024-10-25 | X axis, Repeatability @ 2nd position | 0.06
| Hsingchen Technology | 2024-10-25 | X axis, Accuracy @1st position | 0.43
| Hsingchen Technology | 2024-10-25 | X axis, Accuracy @ 4th position | 0.32
| Hsingchen Technology | 2024-10-25 | X axis, Accuracy @ 3rd position | 0.3
| Hsingchen Technology | 2024-10-25 | X axis, Accuracy @ 2nd position | 0.31
| Hsingchen Technology | 2024-10-25 | X axis OPM | 1.17
| Hsingchen Technology | 2024-10-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 24.5
| Hsingchen Technology | 2024-10-25 | Wafer placement(Loading) error on the load port2 (X offset) | 35.1
| Hsingchen Technology | 2024-10-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 45.9
| Hsingchen Technology | 2024-10-25 | Wafer placement(Loading) error on the load port1 (X offset) | 30.9
| Hsingchen Technology | 2024-10-25 | Torque Profile (50mm, 20um/s) | 0.793
| Hsingchen Technology | 2024-10-25 | Torque Profile (50mm, 200um/s) | 0.609
| Hsingchen Technology | 2024-10-25 | StrokeYMin | -152495
| Hsingchen Technology | 2024-10-25 | StrokeYMax | 151575.8
| Hsingchen Technology | 2024-10-25 | StrokeY | 304071.2
| Hsingchen Technology | 2024-10-25 | StrokeXMin | -201190
| Hsingchen Technology | 2024-10-25 | StrokeXMax | 204135.5
| Hsingchen Technology | 2024-10-25 | StrokeX | 405325
| Hsingchen Technology | 2024-10-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.69
| Hsingchen Technology | 2024-10-25 | Roughness Repeatability 80x80um(1σ) | 0.003
| Hsingchen Technology | 2024-10-25 | Roughness Repeatability 80x80um(RSD%) | 2.08
| Hsingchen Technology | 2024-10-25 | Roughness Repeatability 1x1um(1σ) | 0.000667
| Hsingchen Technology | 2024-10-25 | Roughness Repeatability 1x1um(RSD%) | 1.24
| Hsingchen Technology | 2024-10-25 | Roll motion to X+ full scan | 11.7
| Hsingchen Technology | 2024-10-25 | RMS of ZeroScan | 0.041
| Hsingchen Technology | 2024-10-25 | Position repeatability(10repeats, Position Based) | 0.3
| Hsingchen Technology | 2024-10-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00016
| Hsingchen Technology | 2024-10-25 | OpticalAxisShiftPerFocusY | 0.0007
| Hsingchen Technology | 2024-10-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00013
| Hsingchen Technology | 2024-10-25 | OpticalAxisShiftPerFocusX | -0.0022
| Hsingchen Technology | 2024-10-25 | OPM Repeatability @5mm | 0.86
| Hsingchen Technology | 2024-10-25 | OPM Repeatability @25mm | 2.07
| Hsingchen Technology | 2024-10-25 | OPM Repeatability @1mm | 0.52
| Hsingchen Technology | 2024-10-25 | OPM Repeatability @10mm | 1.44
| Hsingchen Technology | 2024-10-25 | NCM driving% with AC160TS cantilever | 2.69
| Hsingchen Technology | 2024-10-25 | Min. Vibration Transmissbility (~100Hz) | 0.042
| Hsingchen Technology | 2024-10-25 | Max OPM @5mm | 25.15
| Hsingchen Technology | 2024-10-25 | Max OPM @25mm | 84.95
| Hsingchen Technology | 2024-10-25 | Max OPM @1mm | 6.88
| Hsingchen Technology | 2024-10-25 | Max OPM @10mm | 25.51
| Hsingchen Technology | 2024-10-25 | LowerLimit | 2363
| Hsingchen Technology | 2024-10-25 | LowerLimit | 4755
| Hsingchen Technology | 2024-10-25 | Lateral(C-D) sensitivity | 44
| Hsingchen Technology | 2024-10-25 | Lateral(C-D) noise | 0.229
| Hsingchen Technology | 2024-10-25 | Image offset after moving from -Y to +Y (Y scan) | 61.52
| Hsingchen Technology | 2024-10-25 | Image offset after moving from -Y to +Y (X scan) | 45.41
| Hsingchen Technology | 2024-10-25 | Image offset after moving from -X to +X (Y scan) | 93.02
| Hsingchen Technology | 2024-10-25 | Image offset after moving from -X to +X (X scan) | 25.63
| Hsingchen Technology | 2024-10-25 | Image offset after moving from +Y to -Y (Y scan) | 82.76
| Hsingchen Technology | 2024-10-25 | Image offset after moving from +Y to -Y (X scan) | 39.55
| Hsingchen Technology | 2024-10-25 | Image offset after moving from +X to -X (Y scan) | 67.38
| Hsingchen Technology | 2024-10-25 | Image offset after moving from +X to -X (X scan) | 56.4
| Hsingchen Technology | 2024-10-25 | Full range of Z Height during Z Scan sweep | 16
| Hsingchen Technology | 2024-10-25 | Full range of Z Detector during Z Scan sweep | 12.33
| Hsingchen Technology | 2024-10-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.47
| Hsingchen Technology | 2024-10-25 | Full range of Y2 Detector during Y Scan sweep | 229.04
| Hsingchen Technology | 2024-10-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.7
| Hsingchen Technology | 2024-10-25 | Full range of Y1 Detector during Y Scan sweep | 225.14
| Hsingchen Technology | 2024-10-25 | Full range of Y Detector during Y Scan sweep | 227.09
| Hsingchen Technology | 2024-10-25 | Full range of X2 Detector Raw during X Scan sweep | 14.48
| Hsingchen Technology | 2024-10-25 | Full range of X2 Detector during X Scan sweep | 224.25
| Hsingchen Technology | 2024-10-25 | Full range of X1 Detector Raw during X Scan sweep | 14.51
| Hsingchen Technology | 2024-10-25 | Full range of X1 Detector during X Scan sweep | 226.1
| Hsingchen Technology | 2024-10-25 | Full range of X Detector during X Scan sweep | 225.17
| Hsingchen Technology | 2024-10-25 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Hsingchen Technology | 2024-10-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Hsingchen Technology | 2024-10-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24
| Hsingchen Technology | 2024-10-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.49
| Hsingchen Technology | 2024-10-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Hsingchen Technology | 2024-10-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Hsingchen Technology | 2024-10-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.34
| Hsingchen Technology | 2024-10-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.38
| Hsingchen Technology | 2024-10-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| NUS CA2DM | 2024-10-24 | Z Scanner position at -3μm | 0.002323
| NUS CA2DM | 2024-10-24 | Z Scanner position at +3μm | 0.00385
| NUS CA2DM | 2024-10-24 | Z linearity by sample height | 1.0001
| NUS CA2DM | 2024-10-24 | Z Height noise | 0.027
| NUS CA2DM | 2024-10-24 | Y2 Detector noise (XY Servo on) | 0.268
| NUS CA2DM | 2024-10-24 | Y2 Detector noise (XY Servo off) | 0.271
| NUS CA2DM | 2024-10-24 | Y1 Detector noise (XY Servo on) | 0.249
| NUS CA2DM | 2024-10-24 | Y1 Detector noise (XY Servo off) | 0.248
| NUS CA2DM | 2024-10-24 | Y Scan Stroke | 173.0215
| NUS CA2DM | 2024-10-24 | Y Detector noise (XY Servo on) | 0.192
| NUS CA2DM | 2024-10-24 | Y Detector noise (XY Servo off) | 0.191
| NUS CA2DM | 2024-10-24 | Y axis OPM | 0.4898
| NUS CA2DM | 2024-10-24 | X2 Detector noise (XY Servo on) | 0.257
| NUS CA2DM | 2024-10-24 | X2 Detector noise (XY Servo off) | 0.247
| NUS CA2DM | 2024-10-24 | X1 Detector noise (XY Servo on) | 0.343
| NUS CA2DM | 2024-10-24 | X1 Detector noise (XY Servo off) | 0.34
| NUS CA2DM | 2024-10-24 | X Scan Stroke | 171.465
| NUS CA2DM | 2024-10-24 | X Detector noise (XY Servo on) | 0.226
| NUS CA2DM | 2024-10-24 | X Detector noise (XY Servo off) | 0.221
| NUS CA2DM | 2024-10-24 | X axis OPM | 0.8988
| NUS CA2DM | 2024-10-24 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.018
| NUS CA2DM | 2024-10-24 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.867
| NUS CA2DM | 2024-10-24 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 5.026
| NUS CA2DM | 2024-10-24 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.198
| NUS CA2DM | 2024-10-24 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.531
| NUS CA2DM | 2024-10-24 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.975
| NUS CA2DM | 2024-10-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| NUS CA2DM | 2024-10-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| NUS CA2DM | 2024-10-24 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.32
| NUS CA2DM | 2024-10-24 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.31
| NUS CA2DM | 2024-10-24 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| NUS CA2DM | 2024-10-24 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.3
| NUS CA2DM | 2024-10-24 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| NUS CA2DM | 2024-10-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| NUS CA2DM | 2024-10-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| NUS CA2DM | 2024-10-24 | NCM Amplitude noise | 11.552
| NUS CA2DM | 2024-10-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.951
| NUS CA2DM | 2024-10-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.043
| NUS CA2DM | 2024-10-24 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 240.3
| NUS CA2DM | 2024-10-24 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.7
| NUS CA2DM | 2024-10-24 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.9
| NUS CA2DM | 2024-10-24 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 239.1
| NUS CA2DM | 2024-10-24 | Maximum Y Scan range (Full Stroke) | 173
| NUS CA2DM | 2024-10-24 | Maximum X Scan range (Full Stroke) | 171.5
| NUS CA2DM | 2024-10-24 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| NUS CA2DM | 2024-10-24 | Lateral(C-D) sensitivity | 19.1
| NUS CA2DM | 2024-10-24 | Lateral(C-D) noise | 0.136
| NUS CA2DM | 2024-10-24 | Full range of Z Height during Z Scan sweep | 17.4
| NUS CA2DM | 2024-10-24 | Full range of Z Detector during Z Scan sweep | 13.9
| NUS CA2DM | 2024-10-24 | Full range of Y2 Detector raw | 14.4
| NUS CA2DM | 2024-10-24 | Full range of Y1 Detector raw | 14.2
| NUS CA2DM | 2024-10-24 | Full range of X2 Detector raw | 14.2
| NUS CA2DM | 2024-10-24 | Full range of X1 Detector raw | 14.2
| NUS CA2DM | 2024-10-24 | A-B noise | 6.839
| Arizona State University (Umberto Celano) | 2024-10-23 | Z Scanner position at -3μm | 0.008191
| Black Semiconductor GmbH | 2024-10-23 | Z Scanner position at -3μm | 0.000544
| Shenzhen Technology University | 2024-10-23 | Z Scanner position at -3μm | 0.002461
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Z Scanner position at -3μm | 0.002969
| Arizona State University (Umberto Celano) | 2024-10-23 | Z Scanner position at +3μm | 0.013653
| Black Semiconductor GmbH | 2024-10-23 | Z Scanner position at +3μm | 0.004541
| Shenzhen Technology University | 2024-10-23 | Z Scanner position at +3μm | 0.007174
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Z Scanner position at +3μm | 0.004789
| Arizona State University (Umberto Celano) | 2024-10-23 | Z Linearity by Z Scanner position
| Black Semiconductor GmbH | 2024-10-23 | Z linearity by sample height | 1.0052
| Shenzhen Technology University | 2024-10-23 | Z linearity by sample height | 1.01
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Z linearity by sample height | 0.99035
| Arizona State University (Umberto Celano) | 2024-10-23 | Z Linearity by 3 different sample height | 0.99954
| Arizona State University (Umberto Celano) | 2024-10-23 | Z Height noise | 0.028
| Black Semiconductor GmbH | 2024-10-23 | Z Height noise | 0.029
| Shenzhen Technology University | 2024-10-23 | Z Height noise | 0.027
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Z Height noise | 0.019
| Black Semiconductor GmbH | 2024-10-23 | Y2 Detector noise (XY Servo on) | 0.23
| Shenzhen Technology University | 2024-10-23 | Y2 Detector noise (XY Servo on) | 0.265
| Black Semiconductor GmbH | 2024-10-23 | Y2 Detector noise (XY Servo off) | 0.233
| Shenzhen Technology University | 2024-10-23 | Y2 Detector noise (XY Servo off) | 0.253
| Black Semiconductor GmbH | 2024-10-23 | Y1 Detector noise (XY Servo on) | 0.222
| Shenzhen Technology University | 2024-10-23 | Y1 Detector noise (XY Servo on) | 0.232
| Black Semiconductor GmbH | 2024-10-23 | Y1 Detector noise (XY Servo off) | 0.218
| Shenzhen Technology University | 2024-10-23 | Y1 Detector noise (XY Servo off) | 0.223
| Arizona State University (Umberto Celano) | 2024-10-23 | Y Scan Stroke | 119.6
| Black Semiconductor GmbH | 2024-10-23 | Y Scan Stroke | 168.2083
| Shenzhen Technology University | 2024-10-23 | Y Scan Stroke | 178.4
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Y Scan Stroke | 110
| Arizona State University (Umberto Celano) | 2024-10-23 | Y Detector noise (XY Servo on) | 0.28
| Black Semiconductor GmbH | 2024-10-23 | Y Detector noise (XY Servo on) | 0.163
| Shenzhen Technology University | 2024-10-23 | Y Detector noise (XY Servo on) | 0.179
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Y Detector noise (XY Servo on) | 0.27
| Arizona State University (Umberto Celano) | 2024-10-23 | Y Detector noise (XY Servo off) | 0.268
| Black Semiconductor GmbH | 2024-10-23 | Y Detector noise (XY Servo off) | 0.164
| Shenzhen Technology University | 2024-10-23 | Y Detector noise (XY Servo off) | 0.17
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Y Detector noise (XY Servo off) | 0.258
| Arizona State University (Umberto Celano) | 2024-10-23 | Y axis OPM | 1.4834
| Black Semiconductor GmbH | 2024-10-23 | Y axis OPM | 1.1838
| Shenzhen Technology University | 2024-10-23 | Y axis OPM | 0.866
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Y axis OPM | 0.93
| 금오공과대학교_공동실험실습관 | 2024-10-23 | XY orthogonality | 90.12
| Black Semiconductor GmbH | 2024-10-23 | X2 Detector noise (XY Servo on) | 0.236
| Shenzhen Technology University | 2024-10-23 | X2 Detector noise (XY Servo on) | 0.239
| Black Semiconductor GmbH | 2024-10-23 | X2 Detector noise (XY Servo off) | 0.233
| Shenzhen Technology University | 2024-10-23 | X2 Detector noise (XY Servo off) | 0.231
| Black Semiconductor GmbH | 2024-10-23 | X1 Detector noise (XY Servo on) | 0.231
| Shenzhen Technology University | 2024-10-23 | X1 Detector noise (XY Servo on) | 0.258
| Black Semiconductor GmbH | 2024-10-23 | X1 Detector noise (XY Servo off) | 0.225
| Shenzhen Technology University | 2024-10-23 | X1 Detector noise (XY Servo off) | 0.246
| Arizona State University (Umberto Celano) | 2024-10-23 | X Scan Stroke | 116.5
| Black Semiconductor GmbH | 2024-10-23 | X Scan Stroke | 176.368
| Shenzhen Technology University | 2024-10-23 | X Scan Stroke | 167.2
| 금오공과대학교_공동실험실습관 | 2024-10-23 | X Scan Stroke | 109.5
| Arizona State University (Umberto Celano) | 2024-10-23 | X Detector noise (XY Servo on) | 0.262
| Black Semiconductor GmbH | 2024-10-23 | X Detector noise (XY Servo on) | 0.173
| Shenzhen Technology University | 2024-10-23 | X Detector noise (XY Servo on) | 0.184
| 금오공과대학교_공동실험실습관 | 2024-10-23 | X Detector noise (XY Servo on) | 0.257
| Arizona State University (Umberto Celano) | 2024-10-23 | X Detector noise (XY Servo off) | 0.256
| Black Semiconductor GmbH | 2024-10-23 | X Detector noise (XY Servo off) | 0.168
| Shenzhen Technology University | 2024-10-23 | X Detector noise (XY Servo off) | 0.176
| 금오공과대학교_공동실험실습관 | 2024-10-23 | X Detector noise (XY Servo off) | 0.248
| Arizona State University (Umberto Celano) | 2024-10-23 | X axis OPM | 1.552
| Black Semiconductor GmbH | 2024-10-23 | X axis OPM | 0.5428
| Shenzhen Technology University | 2024-10-23 | X axis OPM | 1.278
| 금오공과대학교_공동실험실습관 | 2024-10-23 | X axis OPM | 1.54
| Arizona State University (Umberto Celano) | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.648
| Black Semiconductor GmbH | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.092
| Shenzhen Technology University | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.16
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.017
| Arizona State University (Umberto Celano) | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 4.841
| Black Semiconductor GmbH | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.084
| Shenzhen Technology University | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.269
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.001
| Arizona State University (Umberto Celano) | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.77
| Black Semiconductor GmbH | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 5.005
| Shenzhen Technology University | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 10.123
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.369
| Arizona State University (Umberto Celano) | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.173
| Black Semiconductor GmbH | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.164
| Shenzhen Technology University | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.097
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.144
| Arizona State University (Umberto Celano) | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.428
| Black Semiconductor GmbH | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.522
| Shenzhen Technology University | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 1.978
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.68
| Arizona State University (Umberto Celano) | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.421
| Black Semiconductor GmbH | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.281
| Shenzhen Technology University | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.839
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.23
| Arizona State University (Umberto Celano) | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| 금오공과대학교_공동실험실습관 | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Arizona State University (Umberto Celano) | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| 금오공과대학교_공동실험실습관 | 2024-10-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.34
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.36
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.36
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| Shenzhen Technology University | 2024-10-23 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Arizona State University (Umberto Celano) | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Shenzhen Technology University | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 금오공과대학교_공동실험실습관 | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Arizona State University (Umberto Celano) | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| Black Semiconductor GmbH | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Shenzhen Technology University | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| 금오공과대학교_공동실험실습관 | 2024-10-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Arizona State University (Umberto Celano) | 2024-10-23 | NCM driving% with AC160TS cantilever | 0.419
| 금오공과대학교_공동실험실습관 | 2024-10-23 | NCM driving% with AC160TS cantilever | 4.96
| Arizona State University (Umberto Celano) | 2024-10-23 | NCM Amplitude noise | 13.444
| Black Semiconductor GmbH | 2024-10-23 | NCM Amplitude noise | 15.862
| Shenzhen Technology University | 2024-10-23 | NCM Amplitude noise | 12.577
| 금오공과대학교_공동실험실습관 | 2024-10-23 | NCM Amplitude noise | 14.59
| Arizona State University (Umberto Celano) | 2024-10-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.246
| Black Semiconductor GmbH | 2024-10-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.963
| Shenzhen Technology University | 2024-10-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.066
| 금오공과대학교_공동실험실습관 | 2024-10-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.738
| Arizona State University (Umberto Celano) | 2024-10-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.905
| Black Semiconductor GmbH | 2024-10-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 30.255
| Shenzhen Technology University | 2024-10-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.071
| 금오공과대학교_공동실험실습관 | 2024-10-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.432
| Black Semiconductor GmbH | 2024-10-23 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 236.9
| Shenzhen Technology University | 2024-10-23 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 245
| Black Semiconductor GmbH | 2024-10-23 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 228.7
| Shenzhen Technology University | 2024-10-23 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.8
| Arizona State University (Umberto Celano) | 2024-10-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 172.1
| 금오공과대학교_공동실험실습관 | 2024-10-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.5
| Black Semiconductor GmbH | 2024-10-23 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243
| Shenzhen Technology University | 2024-10-23 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 236.1
| Black Semiconductor GmbH | 2024-10-23 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 240.2
| Shenzhen Technology University | 2024-10-23 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 228.4
| Arizona State University (Umberto Celano) | 2024-10-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 166.5
| 금오공과대학교_공동실험실습관 | 2024-10-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156
| Black Semiconductor GmbH | 2024-10-23 | Maximum Y Scan range (Full Stroke) | 168.2
| Shenzhen Technology University | 2024-10-23 | Maximum Y Scan range (Full Stroke) | 178
| Black Semiconductor GmbH | 2024-10-23 | Maximum X Scan range (Full Stroke) | 176.5
| Shenzhen Technology University | 2024-10-23 | Maximum X Scan range (Full Stroke) | 167.2
| Black Semiconductor GmbH | 2024-10-23 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| Shenzhen Technology University | 2024-10-23 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Arizona State University (Umberto Celano) | 2024-10-23 | Lateral(C-D) sensitivity | 7.6
| Black Semiconductor GmbH | 2024-10-23 | Lateral(C-D) sensitivity | 16.4
| Shenzhen Technology University | 2024-10-23 | Lateral(C-D) sensitivity | 16
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Lateral(C-D) sensitivity | 10.3
| Arizona State University (Umberto Celano) | 2024-10-23 | Lateral(C-D) noise | 0.13
| Black Semiconductor GmbH | 2024-10-23 | Lateral(C-D) noise | 0.135
| Shenzhen Technology University | 2024-10-23 | Lateral(C-D) noise | 0.109
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Lateral(C-D) noise | 0.12
| Arizona State University (Umberto Celano) | 2024-10-23 | Full range of Z Height during Z Scan sweep | 17.9
| Black Semiconductor GmbH | 2024-10-23 | Full range of Z Height during Z Scan sweep | 16.9
| Shenzhen Technology University | 2024-10-23 | Full range of Z Height during Z Scan sweep | 17.6
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Full range of Z Height during Z Scan sweep | 17
| Arizona State University (Umberto Celano) | 2024-10-23 | Full range of Z Detector during Z Scan sweep | 13.3
| Black Semiconductor GmbH | 2024-10-23 | Full range of Z Detector during Z Scan sweep | 12.4
| Shenzhen Technology University | 2024-10-23 | Full range of Z Detector during Z Scan sweep | 13.5
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Full range of Z Detector during Z Scan sweep | 13.9
| Black Semiconductor GmbH | 2024-10-23 | Full range of Y2 Detector raw | 14.5
| Shenzhen Technology University | 2024-10-23 | Full range of Y2 Detector raw | 14.7
| Black Semiconductor GmbH | 2024-10-23 | Full range of Y1 Detector raw | 14.4
| Shenzhen Technology University | 2024-10-23 | Full range of Y1 Detector raw | 14.4
| Arizona State University (Umberto Celano) | 2024-10-23 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| Arizona State University (Umberto Celano) | 2024-10-23 | Full range of Y Detector during Y Scan sweep | 119.4
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Full range of Y Detector during Y Scan sweep | 110.2
| Black Semiconductor GmbH | 2024-10-23 | Full range of X2 Detector raw | 14.6
| Shenzhen Technology University | 2024-10-23 | Full range of X2 Detector raw | 14.3
| Black Semiconductor GmbH | 2024-10-23 | Full range of X1 Detector raw | 14.6
| Shenzhen Technology University | 2024-10-23 | Full range of X1 Detector raw | 14.4
| Arizona State University (Umberto Celano) | 2024-10-23 | Full range of X Detector Raw during X Scan sweep | 14
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Full range of X Detector Raw during X Scan sweep | 14
| Arizona State University (Umberto Celano) | 2024-10-23 | Full range of X Detector during X Scan sweep | 116.5
| 금오공과대학교_공동실험실습관 | 2024-10-23 | Full range of X Detector during X Scan sweep | 108.4
| Arizona State University (Umberto Celano) | 2024-10-23 | A-B noise | 7.728
| Black Semiconductor GmbH | 2024-10-23 | A-B noise | 7.19
| Shenzhen Technology University | 2024-10-23 | A-B noise | 7.253
| 금오공과대학교_공동실험실습관 | 2024-10-23 | A-B noise | 17.08
| 고려기연 | 2024-10-22 | Z Scanner position at -3μm | 0.001674
| 고려기연 | 2024-10-22 | Z Scanner position at +3μm | 0.014207
| 고려기연 | 2024-10-22 | Z linearity by sample height | 1.001
| 고려기연 | 2024-10-22 | Z Height noise | 0.026
| 고려기연 | 2024-10-22 | Y Scan Stroke | 55
| 고려기연 | 2024-10-22 | Y Detector noise (XY Servo on) | 0.268
| 고려기연 | 2024-10-22 | Y Detector noise (XY Servo off) | 0.26
| 고려기연 | 2024-10-22 | Y axis OPM | 0.68
| 고려기연 | 2024-10-22 | X Scan Stroke | 55
| 고려기연 | 2024-10-22 | X Detector noise (XY Servo on) | 0.263
| 고려기연 | 2024-10-22 | X Detector noise (XY Servo off) | 0.255
| 고려기연 | 2024-10-22 | X axis OPM | 0.41
| 고려기연 | 2024-10-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| 고려기연 | 2024-10-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 고려기연 | 2024-10-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 고려기연 | 2024-10-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 고려기연 | 2024-10-22 | NCM driving% with AC160TS cantilever | 1.251
| 고려기연 | 2024-10-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.832
| 고려기연 | 2024-10-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.856
| 고려기연 | 2024-10-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.7
| 고려기연 | 2024-10-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104
| 고려기연 | 2024-10-22 | Lateral(C-D) sensitivity | 23.6
| 고려기연 | 2024-10-22 | Lateral(C-D) noise | 0.07
| 고려기연 | 2024-10-22 | Full range of Z Height during Z Scan sweep | 17.5
| 고려기연 | 2024-10-22 | Full range of Z Detector during Z Scan sweep | 14.1
| 고려기연 | 2024-10-22 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| 고려기연 | 2024-10-22 | Full range of Y Detector during Y Scan sweep | 75.1
| 고려기연 | 2024-10-22 | Full range of X Detector Raw during X Scan sweep | 14.6
| 고려기연 | 2024-10-22 | Full range of X Detector during X Scan sweep | 75.8
| Kanto Chemical | 2024-10-18 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.208
| Kanto Chemical | 2024-10-18 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.881
| Kanto Chemical | 2024-10-18 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 13.217
| Kanto Chemical | 2024-10-18 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.156
| Kanto Chemical | 2024-10-18 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 1.967
| Kanto Chemical | 2024-10-18 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 13.217
| 안동대학교_공동실험실습관 | 2024-10-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.097
| 안동대학교_공동실험실습관 | 2024-10-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.027
| 안동대학교_공동실험실습관 | 2024-10-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.016
| 안동대학교_공동실험실습관 | 2024-10-16 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.139
| 안동대학교_공동실험실습관 | 2024-10-16 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.182
| 안동대학교_공동실험실습관 | 2024-10-16 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.535
| SJ Semi Wuxi | 2024-10-15 | Z Linearity by Z scanner position | 0.1
| SJ Semi Wuxi | 2024-10-15 | Z Height noise | 0.032
| SJ Semi Wuxi | 2024-10-15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.3
| SJ Semi Wuxi | 2024-10-15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.3
| SJ Semi Wuxi | 2024-10-15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12.1
| SJ Semi Wuxi | 2024-10-15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.1
| SJ Semi Wuxi | 2024-10-15 | Z Drift | 7.5
| SJ Semi Wuxi | 2024-10-15 | Z Detector offset | 40.8
| SJ Semi Wuxi | 2024-10-15 | Y2 Detector noise (XY Servo on) | 0.33
| SJ Semi Wuxi | 2024-10-15 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 248.9
| SJ Semi Wuxi | 2024-10-15 | Y1 Detector noise (XY Servo on) | 0.32
| SJ Semi Wuxi | 2024-10-15 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 246.6
| SJ Semi Wuxi | 2024-10-15 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 173
| SJ Semi Wuxi | 2024-10-15 | Y Ringing(Rpv) | 0.1
| SJ Semi Wuxi | 2024-10-15 | Y Non-linearity error | 3.35
| SJ Semi Wuxi | 2024-10-15 | Y Detector noise (XY Servo on) | 0.25
| SJ Semi Wuxi | 2024-10-15 | Y axis, Repeatability @1st position | 0.03
| SJ Semi Wuxi | 2024-10-15 | Y axis, Repeatability @ 4th position | 0.04
| SJ Semi Wuxi | 2024-10-15 | Y axis, Repeatability @ 3rd position | 0.04
| SJ Semi Wuxi | 2024-10-15 | Y axis, Repeatability @ 2nd position | 0.03
| SJ Semi Wuxi | 2024-10-15 | Y axis, Accuracy @1st position | 0.34
| SJ Semi Wuxi | 2024-10-15 | Y axis, Accuracy @ 4th position | 0.37
| SJ Semi Wuxi | 2024-10-15 | Y axis, Accuracy @ 3rd position | 0.39
| SJ Semi Wuxi | 2024-10-15 | Y axis, Accuracy @ 2nd position | 0.32
| SJ Semi Wuxi | 2024-10-15 | Y axis OPM | 1.79
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.6943
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y offset (LLC Translation) | 3.1165
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y offset (In-Die Align) | 1.5869
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y offset (Global Align) | 2.6416
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.3824
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y 1σ (LLC Translation) | 0.521
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y 1σ (In-Die Align) | 0.27554
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning Y 1σ (Global Align) | 0.68613
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X offset (Vision Pattern Matching) | 1.1084
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X offset (LLC Translation) | 2.3779
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X offset (In-Die Align) | 1.9434
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X offset (Global Align) | 1.4551
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.22382
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X 1σ (LLC Translation) | 0.4294
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X 1σ (In-Die Align) | 0.41068
| SJ Semi Wuxi | 2024-10-15 | XY Stage positioning X 1σ (Global Align) | 0.35932
| SJ Semi Wuxi | 2024-10-15 | XY Orthogonality | 90
| SJ Semi Wuxi | 2024-10-15 | X2 Detector noise (XY Servo on) | 0.35
| SJ Semi Wuxi | 2024-10-15 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 259.2
| SJ Semi Wuxi | 2024-10-15 | X1 Detector noise (XY Servo on) | 0.36
| SJ Semi Wuxi | 2024-10-15 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 264.7
| SJ Semi Wuxi | 2024-10-15 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 165.7
| SJ Semi Wuxi | 2024-10-15 | X Ringing(Rpv) | 0.08
| SJ Semi Wuxi | 2024-10-15 | X Non-linearity error | 2.69
| SJ Semi Wuxi | 2024-10-15 | X Detector noise (XY Servo on) | 0.27
| SJ Semi Wuxi | 2024-10-15 | X axis, Repeatability @1st position | 0.04
| SJ Semi Wuxi | 2024-10-15 | X axis, Repeatability @ 4th position | 0.05
| SJ Semi Wuxi | 2024-10-15 | X axis, Repeatability @ 3rd position | 0.06
| SJ Semi Wuxi | 2024-10-15 | X axis, Repeatability @ 2nd position | 0.05
| SJ Semi Wuxi | 2024-10-15 | X axis, Accuracy @1st position | 0.18
| SJ Semi Wuxi | 2024-10-15 | X axis, Accuracy @ 4th position | 0.25
| SJ Semi Wuxi | 2024-10-15 | X axis, Accuracy @ 3rd position | 0.28
| SJ Semi Wuxi | 2024-10-15 | X axis, Accuracy @ 2nd position | 0.31
| SJ Semi Wuxi | 2024-10-15 | X axis OPM | 0.53
| SJ Semi Wuxi | 2024-10-15 | Wafer placement(Loading) error on the load port2 (Y offset) | 16.3
| SJ Semi Wuxi | 2024-10-15 | Wafer placement(Loading) error on the load port2 (X offset) | 55.2
| SJ Semi Wuxi | 2024-10-15 | Wafer placement(Loading) error on the load port1 (Y offset) | 45.5
| SJ Semi Wuxi | 2024-10-15 | Wafer placement(Loading) error on the load port1 (X offset) | 72.5
| SJ Semi Wuxi | 2024-10-15 | Torque Profile (50mm, 20um/s) | 0.781
| SJ Semi Wuxi | 2024-10-15 | Torque Profile (50mm, 200um/s) | 0.474
| SJ Semi Wuxi | 2024-10-15 | StrokeYMin | -151933
| SJ Semi Wuxi | 2024-10-15 | StrokeYMax | 154282
| SJ Semi Wuxi | 2024-10-15 | StrokeY | 306215.3
| SJ Semi Wuxi | 2024-10-15 | StrokeXMin | -203805
| SJ Semi Wuxi | 2024-10-15 | StrokeXMax | 202745.2
| SJ Semi Wuxi | 2024-10-15 | StrokeX | 406550.6
| SJ Semi Wuxi | 2024-10-15 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.29
| SJ Semi Wuxi | 2024-10-15 | Roughness Repeatability 80x80um(1σ) | 0.001667
| SJ Semi Wuxi | 2024-10-15 | Roughness Repeatability 80x80um(RSD%) | 1.1
| SJ Semi Wuxi | 2024-10-15 | Roughness Repeatability 1x1um(1σ) | 0.000333
| SJ Semi Wuxi | 2024-10-15 | Roughness Repeatability 1x1um(RSD%) | 0.352
| SJ Semi Wuxi | 2024-10-15 | Roll motion to X+ full scan | 6.1
| SJ Semi Wuxi | 2024-10-15 | RMS of ZeroScan | 0.048
| SJ Semi Wuxi | 2024-10-15 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| SJ Semi Wuxi | 2024-10-15 | OpticalAxisShiftPerFocusY | -0.001
| SJ Semi Wuxi | 2024-10-15 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00017
| SJ Semi Wuxi | 2024-10-15 | OpticalAxisShiftPerFocusX | -0.0025
| SJ Semi Wuxi | 2024-10-15 | OPM Repeatability @5mm | 1.11
| SJ Semi Wuxi | 2024-10-15 | OPM Repeatability @25mm | 3.84
| SJ Semi Wuxi | 2024-10-15 | OPM Repeatability @1mm | 0.42
| SJ Semi Wuxi | 2024-10-15 | OPM Repeatability @10mm | 0.99
| SJ Semi Wuxi | 2024-10-15 | NCM driving% with AC160TS cantilever | 4.4
| SJ Semi Wuxi | 2024-10-15 | Min. Vibration Transmissbility (~100Hz) | 0.041
| SJ Semi Wuxi | 2024-10-15 | Max OPM @5mm | 24.15
| SJ Semi Wuxi | 2024-10-15 | Max OPM @25mm | 100.06
| SJ Semi Wuxi | 2024-10-15 | Max OPM @1mm | 5.81
| SJ Semi Wuxi | 2024-10-15 | Max OPM @10mm | 39.61
| SJ Semi Wuxi | 2024-10-15 | LowerLimit | 2143.19
| SJ Semi Wuxi | 2024-10-15 | LowerLimit | 4999
| SJ Semi Wuxi | 2024-10-15 | Lateral(C-D) sensitivity | 32.6
| SJ Semi Wuxi | 2024-10-15 | Lateral(C-D) noise | 0.203
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from -Y to +Y (Y scan) | 47.61
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from -Y to +Y (X scan) | 21.97
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from -X to +X (Y scan) | 79.1
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from -X to +X (X scan) | 54.2
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from +Y to -Y (Y scan) | 30.76
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from +Y to -Y (X scan) | 60.79
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from +X to -X (Y scan) | 29.3
| SJ Semi Wuxi | 2024-10-15 | Image offset after moving from +X to -X (X scan) | 45.41
| SJ Semi Wuxi | 2024-10-15 | Full range of Z Height during Z Scan sweep | 17.9
| SJ Semi Wuxi | 2024-10-15 | Full range of Z Detector during Z Scan sweep | 14.1
| SJ Semi Wuxi | 2024-10-15 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| SJ Semi Wuxi | 2024-10-15 | Full range of Y2 Detector during Y Scan sweep | 217.4
| SJ Semi Wuxi | 2024-10-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14.14
| SJ Semi Wuxi | 2024-10-15 | Full range of Y1 Detector during Y Scan sweep | 213.5
| SJ Semi Wuxi | 2024-10-15 | Full range of Y Detector during Y Scan sweep | 213.5
| SJ Semi Wuxi | 2024-10-15 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| SJ Semi Wuxi | 2024-10-15 | Full range of X2 Detector during X Scan sweep | 217
| SJ Semi Wuxi | 2024-10-15 | Full range of X1 Detector Raw during X Scan sweep | 14.02
| SJ Semi Wuxi | 2024-10-15 | Full range of X1 Detector during X Scan sweep | 216.6
| SJ Semi Wuxi | 2024-10-15 | Full range of X Detector during X Scan sweep | 217
| SJ Semi Wuxi | 2024-10-15 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| SJ Semi Wuxi | 2024-10-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| SJ Semi Wuxi | 2024-10-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.44
| SJ Semi Wuxi | 2024-10-15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.46
| SJ Semi Wuxi | 2024-10-15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| SJ Semi Wuxi | 2024-10-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| SJ Semi Wuxi | 2024-10-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.32
| SJ Semi Wuxi | 2024-10-15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| SJ Semi Wuxi | 2024-10-15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| 고려기연 | 2024-10-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.093
| 고려기연 | 2024-10-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.477
| 고려기연 | 2024-10-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.407
| 고려기연 | 2024-10-14 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.075
| 고려기연 | 2024-10-14 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 1.894
| 고려기연 | 2024-10-14 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.301
| Genomics | 2024-10-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.164
| Genomics | 2024-10-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.714
| Genomics | 2024-10-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.017
| Genomics | 2024-10-10 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.198
| Genomics | 2024-10-10 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.269
| Genomics | 2024-10-10 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.818
| Quanzhou San'an Semiconductor(2) | 2024-10-10 | Roughness Repeatability 80x80um(1σ) | 0.03
| Quanzhou San'an Semiconductor(2) | 2024-10-10 | Roughness Repeatability 80x80um(RSD%) | 14.71
| Quanzhou San'an Semiconductor(2) | 2024-10-10 | Roughness Repeatability 1x1um(1σ) | 0.006667
| Quanzhou San'an Semiconductor(2) | 2024-10-10 | Roughness Repeatability 1x1um(RSD%) | 4.03
| heguang photomask | 2024-10-07 | Z Linearity by Z scanner position | 0.21
| heguang photomask | 2024-10-07 | Z Linearity by 4 different sample height | 1.003
| heguang photomask | 2024-10-07 | Z Height noise | 0.038
| heguang photomask | 2024-10-07 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.7
| heguang photomask | 2024-10-07 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.6
| heguang photomask | 2024-10-07 | Z Drift | 15.6
| heguang photomask | 2024-10-07 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 63.1
| heguang photomask | 2024-10-07 | Y Ringing(Rpv) | 0.66
| heguang photomask | 2024-10-07 | Y detector noise (XY Servo on) | 0.15
| heguang photomask | 2024-10-07 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 87.16
| heguang photomask | 2024-10-07 | Y axis, Repeatability @1st position | 0.03
| heguang photomask | 2024-10-07 | Y axis, Repeatability @ 4th position | 0.02
| heguang photomask | 2024-10-07 | Y axis, Repeatability @ 3rd position | 0.03
| heguang photomask | 2024-10-07 | Y axis, Repeatability @ 2nd position | 0.02
| heguang photomask | 2024-10-07 | Y axis, Accuracy @1st position | 0.14
| heguang photomask | 2024-10-07 | Y axis, Accuracy @ 4th position | 0.13
| heguang photomask | 2024-10-07 | Y axis, Accuracy @ 3rd position | 0.13
| heguang photomask | 2024-10-07 | Y axis, Accuracy @ 2nd position | 0.13
| heguang photomask | 2024-10-07 | Y axis OPM | 0.91
| heguang photomask | 2024-10-07 | XY Orthogonality | 89.8
| heguang photomask | 2024-10-07 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 62.2
| heguang photomask | 2024-10-07 | X Ringing(Rpv) | 0.3
| heguang photomask | 2024-10-07 | X detector noise (XY Servo on) | 0.18
| heguang photomask | 2024-10-07 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 83.81
| heguang photomask | 2024-10-07 | X axis, Repeatability @1st position | 0.02
| heguang photomask | 2024-10-07 | X axis, Repeatability @ 4th position | 0.02
| heguang photomask | 2024-10-07 | X axis, Repeatability @ 3rd position | 0.02
| heguang photomask | 2024-10-07 | X axis, Repeatability @ 2nd position | 0.02
| heguang photomask | 2024-10-07 | X axis, Accuracy @1st position | 0.11
| heguang photomask | 2024-10-07 | X axis, Accuracy @ 4th position | 0.13
| heguang photomask | 2024-10-07 | X axis, Accuracy @ 3rd position | 0.11
| heguang photomask | 2024-10-07 | X axis, Accuracy @ 2nd position | 0.12
| heguang photomask | 2024-10-07 | X axis OPM | 2.1
| heguang photomask | 2024-10-07 | Scan offset accuracy (Precision) | 3
| heguang photomask | 2024-10-07 | RMS of ZeroScan | 0.047
| heguang photomask | 2024-10-07 | RMS error after Stage Calibration | 1.1
| heguang photomask | 2024-10-07 | Min. Vibration Transmissbility (~100Hz) | 0.069
| heguang photomask | 2024-10-07 | Mask Roughness Repeatability (10Repeats, σ ) | 0.004
| heguang photomask | 2024-10-07 | Mask Roughness Repeatability (10Repeats, RSD%) | 1.3
| heguang photomask | 2024-10-07 | Mask exchange time | 4'54"
| heguang photomask | 2024-10-07 | Image offset after moving from -Y to +Y (Y scan) | 58.59
| heguang photomask | 2024-10-07 | Image offset after moving from -Y to +Y (X scan) | 53.71
| heguang photomask | 2024-10-07 | Image offset after moving from -X to +X (Y scan) | 29.3
| heguang photomask | 2024-10-07 | Image offset after moving from -X to +X (X scan) | 19.53
| heguang photomask | 2024-10-07 | Image offset after moving from +Y to -Y (Y scan) | 34.18
| heguang photomask | 2024-10-07 | Image offset after moving from +Y to -Y (X scan) | 19.53
| heguang photomask | 2024-10-07 | Image offset after moving from +X to -X (Y scan) | 58.59
| heguang photomask | 2024-10-07 | Image offset after moving from +X to -X (X scan) | 19.53
| heguang photomask | 2024-10-07 | Imag offset after Lock stage (Y scan) | 53.71
| heguang photomask | 2024-10-07 | Imag offset after Lock stage (X scan) | 63.48
| heguang photomask | 2024-10-07 | Imag offset after EFEM loading (Y scan) | 97.66
| heguang photomask | 2024-10-07 | Imag offset after EFEM loading (X scan) | 53.71
| heguang photomask | 2024-10-07 | Full range of Z Height during Z Scan sweep | 19.1
| heguang photomask | 2024-10-07 | Full range of Z Detector during Z Scan sweep | 15.4
| heguang photomask | 2024-10-07 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| heguang photomask | 2024-10-07 | Full range of Y Detector during Y Scan sweep | 65.8
| heguang photomask | 2024-10-07 | Full range of X Detector Raw during X Scan sweep | 14.4
| heguang photomask | 2024-10-07 | Full range of X Detector during X Scan sweep | 63.5
| heguang photomask | 2024-10-07 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.505
| heguang photomask | 2024-10-07 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.505
| heguang photomask | 2024-10-07 | Aligned stage accuracy (Y offset-Point Align3) | 3.1
| heguang photomask | 2024-10-07 | Aligned stage accuracy (Y offset-Point Align2) | 2.41
| heguang photomask | 2024-10-07 | Aligned stage accuracy (Y offset-Point Align1) | 1.3
| heguang photomask | 2024-10-07 | Aligned stage accuracy (Y offset) | -0.74
| heguang photomask | 2024-10-07 | Aligned stage accuracy (X offset-Point Align3) | 1.49
| heguang photomask | 2024-10-07 | Aligned stage accuracy (X offset-Point Align2) | 0.65
| heguang photomask | 2024-10-07 | Aligned stage accuracy (X offset-Point Align1) | 0.58
| heguang photomask | 2024-10-07 | Aligned stage accuracy (X offset) | 0.8
| University of the District of Columbia | 2024-10-04 | Z Scanner position at -3μm | 0.00215
| University of the District of Columbia | 2024-10-04 | Z Scanner position at +3μm | 0.006432
| University of the District of Columbia | 2024-10-04 | Z linearity by sample height | 0.99975
| University of the District of Columbia | 2024-10-04 | Z Height noise | 0.027
| University of the District of Columbia | 2024-10-04 | Y Scan Stroke | 111.8
| University of the District of Columbia | 2024-10-04 | Y Detector noise (XY Servo on) | 0.264
| University of the District of Columbia | 2024-10-04 | Y Detector noise (XY Servo off) | 0.261
| University of the District of Columbia | 2024-10-04 | Y axis OPM | 0.6764
| University of the District of Columbia | 2024-10-04 | X Scan Stroke | 114.3
| University of the District of Columbia | 2024-10-04 | X Detector noise (XY Servo on) | 0.263
| University of the District of Columbia | 2024-10-04 | X Detector noise (XY Servo off) | 0.26
| University of the District of Columbia | 2024-10-04 | X axis OPM | 0.5688
| University of the District of Columbia | 2024-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| University of the District of Columbia | 2024-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| University of the District of Columbia | 2024-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| University of the District of Columbia | 2024-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| University of the District of Columbia | 2024-10-04 | NCM driving% with AC160TS cantilever | 0.746
| University of the District of Columbia | 2024-10-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.404
| University of the District of Columbia | 2024-10-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.25
| University of the District of Columbia | 2024-10-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 162.3
| University of the District of Columbia | 2024-10-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.3
| University of the District of Columbia | 2024-10-04 | Lateral(C-D) sensitivity | 15.4
| University of the District of Columbia | 2024-10-04 | Lateral(C-D) noise | 0.119
| University of the District of Columbia | 2024-10-04 | Full range of Z Height during Z Scan sweep | 18.6
| University of the District of Columbia | 2024-10-04 | Full range of Z Detector during Z Scan sweep | 14.8
| University of the District of Columbia | 2024-10-04 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| University of the District of Columbia | 2024-10-04 | Full range of Y Detector during Y Scan sweep | 111.7
| University of the District of Columbia | 2024-10-04 | Full range of X Detector Raw during X Scan sweep | 14.3
| University of the District of Columbia | 2024-10-04 | Full range of X Detector during X Scan sweep | 114.2
| Technical University of Dresden (Niese) | 2024-10-02 | Z Scanner position at -3μm | 0.001848
| Technical University of Dresden (Niese) | 2024-10-02 | Z Scanner position at +3μm | 0.00152
| Technical University of Dresden (Niese) | 2024-10-02 | Z linearity by sample height | 0.97631
| Technical University of Dresden (Niese) | 2024-10-02 | Z Height noise | 0.033
| Technical University of Dresden (Niese) | 2024-10-02 | Y Scan Stroke | 122.7
| Technical University of Dresden (Niese) | 2024-10-02 | Y Detector noise (XY Servo on) | 0.425
| Technical University of Dresden (Niese) | 2024-10-02 | Y Detector noise (XY Servo off) | 0.45
| Technical University of Dresden (Niese) | 2024-10-02 | Y axis OPM | 1.41
| Technical University of Dresden (Niese) | 2024-10-02 | XY orthogonality | 90.49
| Technical University of Dresden (Niese) | 2024-10-02 | X Scan Stroke | 126.8
| Technical University of Dresden (Niese) | 2024-10-02 | X Detector noise (XY Servo on) | 0.448
| Technical University of Dresden (Niese) | 2024-10-02 | X Detector noise (XY Servo off) | 0.483
| Technical University of Dresden (Niese) | 2024-10-02 | X axis OPM | 1.165
| Technical University of Dresden (Niese) | 2024-10-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.01
| Stanford University | 2024-10-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.011
| Technical University of Dresden (Niese) | 2024-10-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.52
| Stanford University | 2024-10-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.308
| Technical University of Dresden (Niese) | 2024-10-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 4.03
| Stanford University | 2024-10-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.086
| Technical University of Dresden (Niese) | 2024-10-02 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.11
| Stanford University | 2024-10-02 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.24
| Technical University of Dresden (Niese) | 2024-10-02 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.18
| Stanford University | 2024-10-02 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.457
| Technical University of Dresden (Niese) | 2024-10-02 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.88
| Stanford University | 2024-10-02 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.732
| Technical University of Dresden (Niese) | 2024-10-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Technical University of Dresden (Niese) | 2024-10-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| Technical University of Dresden (Niese) | 2024-10-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Technical University of Dresden (Niese) | 2024-10-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| Technical University of Dresden (Niese) | 2024-10-02 | NCM Amplitude noise | 16.922
| Technical University of Dresden (Niese) | 2024-10-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.717
| Technical University of Dresden (Niese) | 2024-10-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.375
| Technical University of Dresden (Niese) | 2024-10-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.9
| Technical University of Dresden (Niese) | 2024-10-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 176.6
| Technical University of Dresden (Niese) | 2024-10-02 | Lateral(C-D) sensitivity | 21
| Technical University of Dresden (Niese) | 2024-10-02 | Lateral(C-D) noise | 0.106
| Technical University of Dresden (Niese) | 2024-10-02 | Full range of Z Height during Z Scan sweep | 17.5
| Technical University of Dresden (Niese) | 2024-10-02 | Full range of Z Detector during Z Scan sweep | 13.3
| Technical University of Dresden (Niese) | 2024-10-02 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| Technical University of Dresden (Niese) | 2024-10-02 | Full range of Y Detector during Y Scan sweep | 122.6
| Technical University of Dresden (Niese) | 2024-10-02 | Full range of X Detector Raw during X Scan sweep | 14.3
| Technical University of Dresden (Niese) | 2024-10-02 | Full range of X Detector during X Scan sweep | 125.9
| Technical University of Dresden (Niese) | 2024-10-02 | A-B noise | 6.411
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Z Scanner position at -3μm | 0.003748
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Z Scanner position at +3μm | 0.017646
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Z linearity by sample height | 0.99342
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Z Height noise | 0.034
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Y Scan Stroke | 55
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Y Detector noise (XY Servo on) | 0.394
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Y Detector noise (XY Servo off) | 0.422
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Y axis OPM | 0.6894
| Nagoya University(Prof.Taniyama) | 2024-09-30 | XY orthogonality | 89.98
| Nagoya University(Prof.Taniyama) | 2024-09-30 | X Scan Stroke | 55
| Nagoya University(Prof.Taniyama) | 2024-09-30 | X Detector noise (XY Servo on) | 0.383
| Nagoya University(Prof.Taniyama) | 2024-09-30 | X Detector noise (XY Servo off) | 0.409
| Nagoya University(Prof.Taniyama) | 2024-09-30 | X axis OPM | 1.525
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.023
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.245
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 5.219
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.216
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.221
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.083
| Nagoya University(Prof.Taniyama) | 2024-09-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Nagoya University(Prof.Taniyama) | 2024-09-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| Nagoya University(Prof.Taniyama) | 2024-09-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Nagoya University(Prof.Taniyama) | 2024-09-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| Nagoya University(Prof.Taniyama) | 2024-09-30 | NCM Amplitude noise | 10.329
| Nagoya University(Prof.Taniyama) | 2024-09-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.926
| Nagoya University(Prof.Taniyama) | 2024-09-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.459
| Nagoya University(Prof.Taniyama) | 2024-09-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 94.2
| Nagoya University(Prof.Taniyama) | 2024-09-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 89.8
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Lateral(C-D) sensitivity | 20.3
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Lateral(C-D) noise | 0.052
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Full range of Z Height during Z Scan sweep | 17.5
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Full range of Z Detector during Z Scan sweep | 14.3
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Full range of Y Detector during Y Scan sweep | 68.5
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Full range of X Detector Raw during X Scan sweep | 15
| Nagoya University(Prof.Taniyama) | 2024-09-30 | Full range of X Detector during X Scan sweep | 67.4
| Nagoya University(Prof.Taniyama) | 2024-09-30 | A-B noise | 4.764
| ITB Indonesia | 2024-09-27 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.087
| ITB Indonesia | 2024-09-27 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.02
| ITB Indonesia | 2024-09-27 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.222
| ITB Indonesia | 2024-09-27 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.197
| ITB Indonesia | 2024-09-27 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.462
| ITB Indonesia | 2024-09-27 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 11.404
| DGIST | 2024-09-26 | Z Scanner position at -3μm | 0.010716
| DGIST | 2024-09-26 | Z Scanner position at +3μm | 0.019011
| DGIST | 2024-09-26 | Z linearity by sample height | 1.0052
| DGIST | 2024-09-26 | Z Height noise | 0.028
| DGIST | 2024-09-26 | Y Scan Stroke | 55
| DGIST | 2024-09-26 | Y Detector noise (XY Servo on) | 0.268
| DGIST | 2024-09-26 | Y Detector noise (XY Servo off) | 0.259
| DGIST | 2024-09-26 | Y axis OPM | 0.5992
| DGIST | 2024-09-26 | X Scan Stroke | 55
| DGIST | 2024-09-26 | X Detector noise (XY Servo on) | 0.267
| DGIST | 2024-09-26 | X Detector noise (XY Servo off) | 0.26
| DGIST | 2024-09-26 | X axis OPM | 0.398
| DGIST | 2024-09-26 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.045
| DGIST | 2024-09-26 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.363
| DGIST | 2024-09-26 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.991
| DGIST | 2024-09-26 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.225
| DGIST | 2024-09-26 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.078
| DGIST | 2024-09-26 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.065
| DGIST | 2024-09-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| DGIST | 2024-09-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| DGIST | 2024-09-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| DGIST | 2024-09-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| DGIST | 2024-09-26 | NCM driving% with AC160TS cantilever | 0.51
| DGIST | 2024-09-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.274
| DGIST | 2024-09-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.42
| DGIST | 2024-09-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.9
| DGIST | 2024-09-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103
| DGIST | 2024-09-26 | Lateral(C-D) sensitivity | 20.6
| DGIST | 2024-09-26 | Lateral(C-D) noise | 0.096
| DGIST | 2024-09-26 | Full range of Z Height during Z Scan sweep | 16.2
| DGIST | 2024-09-26 | Full range of Z Detector during Z Scan sweep | 12.2
| DGIST | 2024-09-26 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| DGIST | 2024-09-26 | Full range of Y Detector during Y Scan sweep | 74.1
| DGIST | 2024-09-26 | Full range of X Detector Raw during X Scan sweep | 14.2
| DGIST | 2024-09-26 | Full range of X Detector during X Scan sweep | 72.9
| TSMC F23(Kumamoto) | 2024-09-26 | Z Linearity by Z scanner position | 0.34
| TSMC F23(Kumamoto) | 2024-09-26 | Z Height noise | 0.026
| TSMC F23(Kumamoto) | 2024-09-26 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.4
| TSMC F23(Kumamoto) | 2024-09-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.3
| TSMC F23(Kumamoto) | 2024-09-26 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5
| TSMC F23(Kumamoto) | 2024-09-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.5
| TSMC F23(Kumamoto) | 2024-09-26 | Z Drift | 6.49
| TSMC F23(Kumamoto) | 2024-09-26 | Z Detector offset | 103.1
| TSMC F23(Kumamoto) | 2024-09-26 | Y2 Detector noise (XY Servo on) | 0.44
| TSMC F23(Kumamoto) | 2024-09-26 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 256
| TSMC F23(Kumamoto) | 2024-09-26 | Y1 Detector noise (XY Servo on) | 0.33
| TSMC F23(Kumamoto) | 2024-09-26 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 257.6
| TSMC F23(Kumamoto) | 2024-09-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 171.9
| TSMC F23(Kumamoto) | 2024-09-26 | Y Ringing(Rpv) | 0.07
| TSMC F23(Kumamoto) | 2024-09-26 | Y Non-linearity error | 1.35
| TSMC F23(Kumamoto) | 2024-09-26 | Y Detector noise (XY Servo on) | 0.29
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Repeatability @1st position | 0.03
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Repeatability @ 4th position | 0.04
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Repeatability @ 3rd position | 0.03
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Repeatability @ 2nd position | 0.05
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Accuracy @1st position | 0.28
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Accuracy @ 4th position | 0.33
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Accuracy @ 3rd position | 0.25
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis, Accuracy @ 2nd position | 0.38
| TSMC F23(Kumamoto) | 2024-09-26 | Y axis OPM | 0.57
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.7871
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y offset (LLC Translation) | 2.4023
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y offset (In-Die Align) | 1.5918
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y offset (Global Align) | 11.3672
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.40874
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y 1σ (LLC Translation) | 0.4758
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y 1σ (In-Die Align) | 0.34577
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning Y 1σ (Global Align) | 1.1871
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X offset (Vision Pattern Matching) | 1.7871
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X offset (LLC Translation) | 0.8447
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X offset (In-Die Align) | 0.835
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X offset (Global Align) | 2.5928
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.35265
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X 1σ (LLC Translation) | 0.2079
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X 1σ (In-Die Align) | 0.17219
| TSMC F23(Kumamoto) | 2024-09-26 | XY Stage positioning X 1σ (Global Align) | 0.57695
| TSMC F23(Kumamoto) | 2024-09-26 | XY Orthogonality | 89.98
| TSMC F23(Kumamoto) | 2024-09-26 | X2 Detector noise (XY Servo on) | 0.5
| TSMC F23(Kumamoto) | 2024-09-26 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 273.4
| TSMC F23(Kumamoto) | 2024-09-26 | X1 Detector noise (XY Servo on) | 0.41
| TSMC F23(Kumamoto) | 2024-09-26 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.4
| TSMC F23(Kumamoto) | 2024-09-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 174.2
| TSMC F23(Kumamoto) | 2024-09-26 | X Ringing(Rpv) | 0.15
| TSMC F23(Kumamoto) | 2024-09-26 | X Non-linearity error | 2.28
| TSMC F23(Kumamoto) | 2024-09-26 | X Detector noise (XY Servo on) | 0.34
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Repeatability @1st position | 0.02
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Repeatability @ 4th position | 0.03
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Repeatability @ 3rd position | 0.02
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Repeatability @ 2nd position | 0.03
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Accuracy @1st position | 0.12
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Accuracy @ 4th position | 0.15
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Accuracy @ 3rd position | 0.19
| TSMC F23(Kumamoto) | 2024-09-26 | X axis, Accuracy @ 2nd position | 0.24
| TSMC F23(Kumamoto) | 2024-09-26 | X axis OPM | 1.53
| TSMC F23(Kumamoto) | 2024-09-26 | Wafer placement(Loading) error on the load port2 (Y offset) | 23.76
| TSMC F23(Kumamoto) | 2024-09-26 | Wafer placement(Loading) error on the load port2 (X offset) | 27.35
| TSMC F23(Kumamoto) | 2024-09-26 | Wafer placement(Loading) error on the load port1 (Y offset) | 19.46
| TSMC F23(Kumamoto) | 2024-09-26 | Wafer placement(Loading) error on the load port1 (X offset) | 32.4
| TSMC F23(Kumamoto) | 2024-09-26 | Torque Profile (50mm, 20um/s) | 0.579
| TSMC F23(Kumamoto) | 2024-09-26 | Torque Profile (50mm, 200um/s) | 0.455
| TSMC F23(Kumamoto) | 2024-09-26 | StrokeYMin | -150564
| TSMC F23(Kumamoto) | 2024-09-26 | StrokeYMax | 152541.9
| TSMC F23(Kumamoto) | 2024-09-26 | StrokeY | 303105.9
| TSMC F23(Kumamoto) | 2024-09-26 | StrokeXMin | -203298
| TSMC F23(Kumamoto) | 2024-09-26 | StrokeXMax | 201165.2
| TSMC F23(Kumamoto) | 2024-09-26 | StrokeX | 404463.2
| TSMC F23(Kumamoto) | 2024-09-26 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.64
| TSMC F23(Kumamoto) | 2024-09-26 | Roll motion to X+ full scan | 8.1
| TSMC F23(Kumamoto) | 2024-09-26 | RMS of ZeroScan | 0.052
| TSMC F23(Kumamoto) | 2024-09-26 | Position repeatability(10repeats, Position Based) | 0.54
| TSMC F23(Kumamoto) | 2024-09-26 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| TSMC F23(Kumamoto) | 2024-09-26 | OpticalAxisShiftPerFocusY | 0.0006
| TSMC F23(Kumamoto) | 2024-09-26 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| TSMC F23(Kumamoto) | 2024-09-26 | OpticalAxisShiftPerFocusX | 0.002
| TSMC F23(Kumamoto) | 2024-09-26 | OPM Repeatability @5mm | 0.76
| TSMC F23(Kumamoto) | 2024-09-26 | OPM Repeatability @25mm | 2.17
| TSMC F23(Kumamoto) | 2024-09-26 | OPM Repeatability @1mm | 0.63
| TSMC F23(Kumamoto) | 2024-09-26 | OPM Repeatability @10mm | 1.31
| TSMC F23(Kumamoto) | 2024-09-26 | NCM driving% with AC160TS cantilever | 11.14
| TSMC F23(Kumamoto) | 2024-09-26 | Min. Vibration Transmissbility (~100Hz) | 0.032
| TSMC F23(Kumamoto) | 2024-09-26 | Max OPM @5mm | 11.25
| TSMC F23(Kumamoto) | 2024-09-26 | Max OPM @25mm | 128.75
| TSMC F23(Kumamoto) | 2024-09-26 | Max OPM @1mm | 7.55
| TSMC F23(Kumamoto) | 2024-09-26 | Max OPM @10mm | 29.31
| TSMC F23(Kumamoto) | 2024-09-26 | LowerLimit | 2313.4
| TSMC F23(Kumamoto) | 2024-09-26 | LowerLimit | 4069
| TSMC F23(Kumamoto) | 2024-09-26 | Lateral(C-D) sensitivity | 5.1
| TSMC F23(Kumamoto) | 2024-09-26 | Lateral(C-D) noise | 0.243
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from -Y to +Y (Y scan) | 57.13
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from -Y to +Y (X scan) | 12.45
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from -X to +X (Y scan) | 28.56
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from -X to +X (X scan) | 71.78
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from +Y to -Y (Y scan) | 17.58
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from +Y to -Y (X scan) | 9.52
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from +X to -X (Y scan) | 19.78
| TSMC F23(Kumamoto) | 2024-09-26 | Image offset after moving from +X to -X (X scan) | 69.58
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Z Height during Z Scan sweep | 17
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Z Detector during Z Scan sweep | 12.9
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.61
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Y2 Detector during Y Scan sweep | 186.9
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.85
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Y1 Detector during Y Scan sweep | 191.2
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of Y Detector during Y Scan sweep | 191.2
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of X2 Detector Raw during X Scan sweep | 14.05
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of X2 Detector during X Scan sweep | 191.7
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of X1 Detector Raw during X Scan sweep | 14.46
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of X1 Detector during X Scan sweep | 192
| TSMC F23(Kumamoto) | 2024-09-26 | Full range of X Detector during X Scan sweep | 192
| TSMC F23(Kumamoto) | 2024-09-26 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC F23(Kumamoto) | 2024-09-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| TSMC F23(Kumamoto) | 2024-09-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| TSMC F23(Kumamoto) | 2024-09-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| TSMC F23(Kumamoto) | 2024-09-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| TSMC F23(Kumamoto) | 2024-09-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| TSMC F23(Kumamoto) | 2024-09-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.36
| TSMC F23(Kumamoto) | 2024-09-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22
| TSMC F23(Kumamoto) | 2024-09-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| Qorvo #3 | 2024-09-25 | Roughness Repeatability 80x80um(1σ) | 0.023333
| Qorvo #3 | 2024-09-25 | Roughness Repeatability 80x80um(RSD%) | 6.78
| Qorvo #3 | 2024-09-25 | Roughness Repeatability 1x1um(1σ) | 0.002
| Qorvo #3 | 2024-09-25 | Roughness Repeatability 1x1um(RSD%) | 2.057
| Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.067
| Wuhan Epiflag | 2024-09-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.017
| Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.147
| Wuhan Epiflag | 2024-09-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.518
| Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.399
| Wuhan Epiflag | 2024-09-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.657
| Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.194
| Wuhan Epiflag | 2024-09-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.168
| Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.208
| Wuhan Epiflag | 2024-09-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.542
| Institute of Condensed Matter Chemistry and Technologies for Energy (ICMATE) | 2024-09-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.054
| Wuhan Epiflag | 2024-09-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.009
| Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.051
| Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.74
| Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.681
| Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.068
| Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.129
| Xiamen Silan Advanced Compound Semiconductor | 2024-09-20 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.603
| Winbond | 2024-09-20 | Z Linearity by Z scanner position | 1.35
| Winbond | 2024-09-20 | Z Height noise | 0.027
| Winbond | 2024-09-20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.1
| Winbond | 2024-09-20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.4
| Winbond | 2024-09-20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12.9
| Winbond | 2024-09-20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.4
| Winbond | 2024-09-20 | Z Drift | 7.2
| Winbond | 2024-09-20 | Z Detector offset | 30.3
| Winbond | 2024-09-20 | Y2 Detector noise (XY Servo on) | 0.55
| Winbond | 2024-09-20 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 229.2
| Winbond | 2024-09-20 | Y1 Detector noise (XY Servo on) | 0.53
| Winbond | 2024-09-20 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 251.9
| Winbond | 2024-09-20 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.3
| Winbond | 2024-09-20 | Y Ringing(Rpv) | 0.12
| Winbond | 2024-09-20 | Y Non-linearity error | 2.74
| Winbond | 2024-09-20 | Y Detector noise (XY Servo on) | 0.4
| Winbond | 2024-09-20 | Y axis, Repeatability @1st position | 0.07
| Winbond | 2024-09-20 | Y axis, Repeatability @ 4th position | 0.04
| Winbond | 2024-09-20 | Y axis, Repeatability @ 3rd position | 0.05
| Winbond | 2024-09-20 | Y axis, Repeatability @ 2nd position | 0.05
| Winbond | 2024-09-20 | Y axis, Accuracy @1st position | 1.28
| Winbond | 2024-09-20 | Y axis, Accuracy @ 4th position | 1.57
| Winbond | 2024-09-20 | Y axis, Accuracy @ 3rd position | 1.23
| Winbond | 2024-09-20 | Y axis, Accuracy @ 2nd position | 1.03
| Winbond | 2024-09-20 | Y axis OPM | 0.34
| Winbond | 2024-09-20 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.8457
| Winbond | 2024-09-20 | XY Stage positioning Y offset (LLC Translation) | 2.2949
| Winbond | 2024-09-20 | XY Stage positioning Y offset (Global Align) | 4.3066
| Winbond | 2024-09-20 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.57211
| Winbond | 2024-09-20 | XY Stage positioning Y 1σ (LLC Translation) | 0.4482
| Winbond | 2024-09-20 | XY Stage positioning Y 1σ (Global Align) | 1.1772
| Winbond | 2024-09-20 | XY Stage positioning X offset (Vision Pattern Matching) | 1.8604
| Winbond | 2024-09-20 | XY Stage positioning X offset (LLC Translation) | 2.3047
| Winbond | 2024-09-20 | XY Stage positioning X offset (Global Align) | 3.3398
| Winbond | 2024-09-20 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.53423
| Winbond | 2024-09-20 | XY Stage positioning X 1σ (LLC Translation) | 0.432
| Winbond | 2024-09-20 | XY Stage positioning X 1σ (Global Align) | 0.69487
| Winbond | 2024-09-20 | XY Orthogonality | 90.1
| Winbond | 2024-09-20 | X2 Detector noise (XY Servo on) | 0.48
| Winbond | 2024-09-20 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 253.4
| Winbond | 2024-09-20 | X1 Detector noise (XY Servo on) | 0.56
| Winbond | 2024-09-20 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 246.2
| Winbond | 2024-09-20 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 171.2
| Winbond | 2024-09-20 | X Ringing(Rpv) | 0.08
| Winbond | 2024-09-20 | X Non-linearity error | 2.72
| Winbond | 2024-09-20 | X Detector noise (XY Servo on) | 0.39
| Winbond | 2024-09-20 | X axis, Repeatability @1st position | 0.06
| Winbond | 2024-09-20 | X axis, Repeatability @ 4th position | 0.04
| Winbond | 2024-09-20 | X axis, Repeatability @ 3rd position | 0.09
| Winbond | 2024-09-20 | X axis, Repeatability @ 2nd position | 0.04
| Winbond | 2024-09-20 | X axis, Accuracy @1st position | 0.35
| Winbond | 2024-09-20 | X axis, Accuracy @ 4th position | 0.2
| Winbond | 2024-09-20 | X axis, Accuracy @ 3rd position | 0.51
| Winbond | 2024-09-20 | X axis, Accuracy @ 2nd position | 0.19
| Winbond | 2024-09-20 | X axis OPM | 0.46
| Winbond | 2024-09-20 | Wafer placement(Loading) error on the load port2 (Y offset) | 30.1
| Winbond | 2024-09-20 | Wafer placement(Loading) error on the load port2 (X offset) | 26.6
| Winbond | 2024-09-20 | Wafer placement(Loading) error on the load port1 (Y offset) | 27.4
| Winbond | 2024-09-20 | Wafer placement(Loading) error on the load port1 (X offset) | 53.9
| Winbond | 2024-09-20 | Torque Profile (50mm, 20um/s) | 0.913
| Winbond | 2024-09-20 | Torque Profile (50mm, 200um/s) | 0.647
| Winbond | 2024-09-20 | StrokeYMin | -152507
| Winbond | 2024-09-20 | StrokeYMax | 150008.7
| Winbond | 2024-09-20 | StrokeY | 302516
| Winbond | 2024-09-20 | StrokeXMin | -200010
| Winbond | 2024-09-20 | StrokeXMax | 202207.2
| Winbond | 2024-09-20 | StrokeX | 402217.3
| Winbond | 2024-09-20 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.96
| Winbond | 2024-09-20 | Roughness Repeatability 80x80um(1σ) | 0.003333
| Winbond | 2024-09-20 | Roughness Repeatability 80x80um(RSD%) | 1.55
| Winbond | 2024-09-20 | Roughness Repeatability 1x1um(1σ) | 0.001
| Winbond | 2024-09-20 | Roughness Repeatability 1x1um(RSD%) | 1.007
| Winbond | 2024-09-20 | Roll motion to X+ full scan | 42.2
| Winbond | 2024-09-20 | RMS of ZeroScan | 0.029
| Winbond | 2024-09-20 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00025
| Winbond | 2024-09-20 | OpticalAxisShiftPerFocusY | -0.0006
| Winbond | 2024-09-20 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00066
| Winbond | 2024-09-20 | OpticalAxisShiftPerFocusX | -0.0058
| Winbond | 2024-09-20 | OPM Repeatability @5mm | 0.58
| Winbond | 2024-09-20 | OPM Repeatability @25mm | 10.32
| Winbond | 2024-09-20 | OPM Repeatability @1mm | 0.64
| Winbond | 2024-09-20 | OPM Repeatability @10mm | 0.96
| Winbond | 2024-09-20 | NCM driving% with AC160TS cantilever | 6.14
| Winbond | 2024-09-20 | Min. Vibration Transmissbility (~100Hz) | 0.033
| Winbond | 2024-09-20 | Max OPM @5mm | 19.71
| Winbond | 2024-09-20 | Max OPM @25mm | 91.88
| Winbond | 2024-09-20 | Max OPM @1mm | 9.48
| Winbond | 2024-09-20 | Max OPM @10mm | 25.25
| Winbond | 2024-09-20 | LowerLimit | 2076
| Winbond | 2024-09-20 | LowerLimit | 4631
| Winbond | 2024-09-20 | Lateral(C-D) sensitivity | 83.2
| Winbond | 2024-09-20 | Lateral(C-D) noise | 0.208
| Winbond | 2024-09-20 | Image offset after moving from -Y to +Y (Y scan) | 46.14
| Winbond | 2024-09-20 | Image offset after moving from -Y to +Y (X scan) | 33.69
| Winbond | 2024-09-20 | Image offset after moving from -X to +X (Y scan) | 110.6
| Winbond | 2024-09-20 | Image offset after moving from -X to +X (X scan) | 27.1
| Winbond | 2024-09-20 | Image offset after moving from +Y to -Y (Y scan) | 79.1
| Winbond | 2024-09-20 | Image offset after moving from +Y to -Y (X scan) | 46.14
| Winbond | 2024-09-20 | Image offset after moving from +X to -X (Y scan) | 13.18
| Winbond | 2024-09-20 | Image offset after moving from +X to -X (X scan) | 9.52
| Winbond | 2024-09-20 | Full range of Z Height during Z Scan sweep | 18.7
| Winbond | 2024-09-20 | Full range of Z Detector during Z Scan sweep | 14
| Winbond | 2024-09-20 | Full range of Y2 Detector Raw during Y Scan sweep | 13.72
| Winbond | 2024-09-20 | Full range of Y2 Detector during Y Scan sweep | 202.2
| Winbond | 2024-09-20 | Full range of Y1 Detector Raw during Y Scan sweep | 13.35
| Winbond | 2024-09-20 | Full range of Y1 Detector during Y Scan sweep | 214.7
| Winbond | 2024-09-20 | Full range of Y Detector during Y Scan sweep | 208.4
| Winbond | 2024-09-20 | Full range of X2 Detector Raw during X Scan sweep | 13.32
| Winbond | 2024-09-20 | Full range of X2 Detector during X Scan sweep | 206.1
| Winbond | 2024-09-20 | Full range of X1 Detector Raw during X Scan sweep | 13.38
| Winbond | 2024-09-20 | Full range of X1 Detector during X Scan sweep | 203
| Winbond | 2024-09-20 | Full range of X Detector during X Scan sweep | 204.6
| Winbond | 2024-09-20 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Winbond | 2024-09-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Winbond | 2024-09-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| Winbond | 2024-09-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.33
| Winbond | 2024-09-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Winbond | 2024-09-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Winbond | 2024-09-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| Winbond | 2024-09-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| Winbond | 2024-09-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Z Scanner position at -3μm | 0.00041
| Harvard University (Nocera) | 2024-09-19 | Z Scanner position at -3μm | 0.005133
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Z Scanner position at +3μm | 0.012058
| Harvard University (Nocera) | 2024-09-19 | Z Scanner position at +3μm | 0.014509
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Z linearity by sample height | 1.0038
| Harvard University (Nocera) | 2024-09-19 | Z linearity by sample height | 1.0065
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Z Height noise | 0.028
| Harvard University (Nocera) | 2024-09-19 | Z Height noise | 0.028
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Y Scan Stroke | 55
| Harvard University (Nocera) | 2024-09-19 | Y Scan Stroke | 113.3
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Y Detector noise (XY Servo on) | 0.279
| Harvard University (Nocera) | 2024-09-19 | Y Detector noise (XY Servo on) | 0.322
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Y Detector noise (XY Servo off) | 0.254
| Harvard University (Nocera) | 2024-09-19 | Y Detector noise (XY Servo off) | 0.335
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Y axis OPM | 0.4558
| Harvard University (Nocera) | 2024-09-19 | Y axis OPM | 0.60375
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | X Scan Stroke | 55
| Harvard University (Nocera) | 2024-09-19 | X Scan Stroke | 110.5
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | X Detector noise (XY Servo on) | 0.28
| Harvard University (Nocera) | 2024-09-19 | X Detector noise (XY Servo on) | 0.285
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | X Detector noise (XY Servo off) | 0.252
| Harvard University (Nocera) | 2024-09-19 | X Detector noise (XY Servo off) | 0.287
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | X axis OPM | 0.6284
| Harvard University (Nocera) | 2024-09-19 | X axis OPM | 0.8077
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.04
| Harvard University (Nocera) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.15
| Rolex (Dudon) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.021
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.916
| Harvard University (Nocera) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.22
| Rolex (Dudon) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.222
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.338
| Harvard University (Nocera) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.67
| Rolex (Dudon) | 2024-09-19 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.361
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.227
| Harvard University (Nocera) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.09
| Rolex (Dudon) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.104
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.392
| Harvard University (Nocera) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 1.86
| Rolex (Dudon) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.38
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.686
| Harvard University (Nocera) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.56
| Rolex (Dudon) | 2024-09-19 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 11.185
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Harvard University (Nocera) | 2024-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Harvard University (Nocera) | 2024-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Harvard University (Nocera) | 2024-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| Harvard University (Nocera) | 2024-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | NCM driving% with AC160TS cantilever | 0.629
| Harvard University (Nocera) | 2024-09-19 | NCM driving% with AC160TS cantilever | 0.401
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.56
| Harvard University (Nocera) | 2024-09-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.093
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.725
| Harvard University (Nocera) | 2024-09-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.912
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 109
| Harvard University (Nocera) | 2024-09-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 162.7
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 109.2
| Harvard University (Nocera) | 2024-09-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Lateral(C-D) sensitivity | 4.1
| Harvard University (Nocera) | 2024-09-19 | Lateral(C-D) sensitivity | 23.9
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Lateral(C-D) noise | 0.085
| Harvard University (Nocera) | 2024-09-19 | Lateral(C-D) noise | 0.087
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Full range of Z Height during Z Scan sweep | 18.5
| Harvard University (Nocera) | 2024-09-19 | Full range of Z Height during Z Scan sweep | 17.6
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Full range of Z Detector during Z Scan sweep | 13.3
| Harvard University (Nocera) | 2024-09-19 | Full range of Z Detector during Z Scan sweep | 13.1
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Harvard University (Nocera) | 2024-09-19 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Full range of Y Detector during Y Scan sweep | 75.7
| Harvard University (Nocera) | 2024-09-19 | Full range of Y Detector during Y Scan sweep | 113.3
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Full range of X Detector Raw during X Scan sweep | 14
| Harvard University (Nocera) | 2024-09-19 | Full range of X Detector Raw during X Scan sweep | 13.9
| National University of Singapore_IFIM(Institute for Functional Intelligent Materials) | 2024-09-19 | Full range of X Detector during X Scan sweep | 76.5
| Harvard University (Nocera) | 2024-09-19 | Full range of X Detector during X Scan sweep | 110.5
| INN Demokritos, (El Sachat) | 2024-09-13 | Z Scanner position at -3μm | 0.000103
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Z Scanner position at -3μm | 0.003397
| INN Demokritos, (El Sachat) | 2024-09-13 | Z Scanner position at +3μm | 0.000446
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Z Scanner position at +3μm | 0.007015
| INN Demokritos, (El Sachat) | 2024-09-13 | Z Linearity by Z Scanner position
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Z linearity by sample height | 0.99907
| INN Demokritos, (El Sachat) | 2024-09-13 | Z Linearity by 3 different sample height | 1.0128
| INN Demokritos, (El Sachat) | 2024-09-13 | Z Height noise | 0.0264
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Z Height noise | 0.026
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y2 Detector noise (XY Servo on) | 0.242
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y2 Detector noise (XY Servo off) | 0.244
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y1 Detector noise (XY Servo on) | 0.236
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y1 Detector noise (XY Servo off) | 0.236
| INN Demokritos, (El Sachat) | 2024-09-13 | Y Scan Stroke | 55
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y Scan Stroke | 196.1
| INN Demokritos, (El Sachat) | 2024-09-13 | Y Detector noise (XY Servo on) | 0.2565
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y Detector noise (XY Servo on) | 0.178
| INN Demokritos, (El Sachat) | 2024-09-13 | Y Detector noise (XY Servo off) | 0.2473
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y Detector noise (XY Servo off) | 0.178
| INN Demokritos, (El Sachat) | 2024-09-13 | Y axis OPM | 1.0846
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Y axis OPM | 0.9732
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X2 Detector noise (XY Servo on) | 0.219
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X2 Detector noise (XY Servo off) | 0.213
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X1 Detector noise (XY Servo on) | 0.235
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X1 Detector noise (XY Servo off) | 0.236
| INN Demokritos, (El Sachat) | 2024-09-13 | X Scan Stroke | 55
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X Scan Stroke | 199.3
| INN Demokritos, (El Sachat) | 2024-09-13 | X Detector noise (XY Servo on) | 0.2825
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X Detector noise (XY Servo on) | 0.171
| INN Demokritos, (El Sachat) | 2024-09-13 | X Detector noise (XY Servo off) | 0.2891
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X Detector noise (XY Servo off) | 0.169
| INN Demokritos, (El Sachat) | 2024-09-13 | X axis OPM | 0.8456
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | X axis OPM | 1.6936
| INN Demokritos, (El Sachat) | 2024-09-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.059
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.096
| INN Demokritos, (El Sachat) | 2024-09-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 4.845
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.273
| INN Demokritos, (El Sachat) | 2024-09-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 11.167
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.101
| INN Demokritos, (El Sachat) | 2024-09-13 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.205
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.185
| INN Demokritos, (El Sachat) | 2024-09-13 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.243
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 1.919
| INN Demokritos, (El Sachat) | 2024-09-13 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.582
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.198
| INN Demokritos, (El Sachat) | 2024-09-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.207
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| INN Demokritos, (El Sachat) | 2024-09-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.274
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.36
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.36
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.35
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.35
| INN Demokritos, (El Sachat) | 2024-09-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.114
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| INN Demokritos, (El Sachat) | 2024-09-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.215
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| INN Demokritos, (El Sachat) | 2024-09-13 | NCM driving% with AC160TS cantilever | 3.503
| INN Demokritos, (El Sachat) | 2024-09-13 | NCM Amplitude noise | 13.7175
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | NCM Amplitude noise | 15.103
| INN Demokritos, (El Sachat) | 2024-09-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.491
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.482
| INN Demokritos, (El Sachat) | 2024-09-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.81
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.831
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 272.1
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 259
| INN Demokritos, (El Sachat) | 2024-09-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 233
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 266.5
| INN Demokritos, (El Sachat) | 2024-09-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Maximum Y Scan range (Full Stroke) | 187.7
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Maximum X Scan range (Full Stroke) | 182.4
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| INN Demokritos, (El Sachat) | 2024-09-13 | Lateral(C-D) sensitivity | 3.9
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Lateral(C-D) sensitivity | 4.9
| INN Demokritos, (El Sachat) | 2024-09-13 | Lateral(C-D) noise | 0.07
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Lateral(C-D) noise | 0.161
| INN Demokritos, (El Sachat) | 2024-09-13 | Full range of Z Height during Z Scan sweep | 17.1
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Full range of Z Height during Z Scan sweep | 16.3
| INN Demokritos, (El Sachat) | 2024-09-13 | Full range of Z Detector during Z Scan sweep | 13.8
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Full range of Z Detector during Z Scan sweep | 12.5
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Full range of Y2 Detector raw | 14
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Full range of Y1 Detector raw | 14.2
| INN Demokritos, (El Sachat) | 2024-09-13 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| INN Demokritos, (El Sachat) | 2024-09-13 | Full range of Y Detector during Y Scan sweep | 71.9
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Full range of X2 Detector raw | 15.2
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | Full range of X1 Detector raw | 14.1
| INN Demokritos, (El Sachat) | 2024-09-13 | Full range of X Detector Raw during X Scan sweep | 14.2
| INN Demokritos, (El Sachat) | 2024-09-13 | Full range of X Detector during X Scan sweep | 71
| INN Demokritos, (El Sachat) | 2024-09-13 | A-B noise | 7.869
| University of Cambridge (Ferrari), (PSE) | 2024-09-13 | A-B noise | 7.796
| Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.041
| Zhejiang University_Physics | 2024-09-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.153
| Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 5.587
| Zhejiang University_Physics | 2024-09-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.829
| Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 13.333
| Zhejiang University_Physics | 2024-09-10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.924
| Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.192
| Zhejiang University_Physics | 2024-09-10 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.227
| Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.46
| Zhejiang University_Physics | 2024-09-10 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.127
| Suzhou Institute of Nano-Tech and Nano-Bionics(SINANO) | 2024-09-10 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.401
| Zhejiang University_Physics | 2024-09-10 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.9
| Sant'Anna Pisa (Poliziani) | 2024-09-09 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.051
| Alba Lulia University (SchaeferSE) | 2024-09-09 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.043
| Sant'Anna Pisa (Poliziani) | 2024-09-09 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 4.118
| Alba Lulia University (SchaeferSE) | 2024-09-09 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 4.447
| Sant'Anna Pisa (Poliziani) | 2024-09-09 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.251
| Alba Lulia University (SchaeferSE) | 2024-09-09 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 10.553
| Sant'Anna Pisa (Poliziani) | 2024-09-09 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.188
| Alba Lulia University (SchaeferSE) | 2024-09-09 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.216
| Sant'Anna Pisa (Poliziani) | 2024-09-09 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.445
| Alba Lulia University (SchaeferSE) | 2024-09-09 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.208
| Sant'Anna Pisa (Poliziani) | 2024-09-09 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.952
| Alba Lulia University (SchaeferSE) | 2024-09-09 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.73
| San'an-ST Microelectronics | 2024-09-06 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.063
| San'an-ST Microelectronics | 2024-09-06 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.402
| San'an-ST Microelectronics | 2024-09-06 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.099
| San'an-ST Microelectronics | 2024-09-06 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.203
| San'an-ST Microelectronics | 2024-09-06 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.364
| San'an-ST Microelectronics | 2024-09-06 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.442
| Hirosaki University | 2024-09-03 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.008
| Hirosaki University | 2024-09-03 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.88
| Hirosaki University | 2024-09-03 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.172
| Hirosaki University | 2024-09-03 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.158
| Hirosaki University | 2024-09-03 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.062
| Hirosaki University | 2024-09-03 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.156
| TSMC AP3 RD | 2024-09-03 | Z Linearity by Z scanner position | 1.72
| TSMC AP3 RD | 2024-09-03 | Z Height noise | 0.026
| TSMC AP3 RD | 2024-09-03 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.1
| TSMC AP3 RD | 2024-09-03 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.5
| TSMC AP3 RD | 2024-09-03 | Z Drift | 15.3
| TSMC AP3 RD | 2024-09-03 | Z Detector offset | 63.1
| TSMC AP3 RD | 2024-09-03 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 107.8
| TSMC AP3 RD | 2024-09-03 | Y Ringing(Rpv) | 0.21
| TSMC AP3 RD | 2024-09-03 | Y detector noise (XY Servo on) | 0.26
| TSMC AP3 RD | 2024-09-03 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 154.136
| TSMC AP3 RD | 2024-09-03 | Y axis, Repeatability @1st position | 0.07
| TSMC AP3 RD | 2024-09-03 | Y axis, Repeatability @ 4th position | 0.08
| TSMC AP3 RD | 2024-09-03 | Y axis, Repeatability @ 3rd position | 0.07
| TSMC AP3 RD | 2024-09-03 | Y axis, Repeatability @ 2nd position | 0.09
| TSMC AP3 RD | 2024-09-03 | Y axis, Accuracy @1st position | 0.38
| TSMC AP3 RD | 2024-09-03 | Y axis, Accuracy @ 4th position | 0.52
| TSMC AP3 RD | 2024-09-03 | Y axis, Accuracy @ 3rd position | 0.36
| TSMC AP3 RD | 2024-09-03 | Y axis, Accuracy @ 2nd position | 0.54
| TSMC AP3 RD | 2024-09-03 | Y axis OPM | 0.76
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y offset (Vision Pattern Matching) | 0.9961
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y offset (LLC Translation) | 1.709
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y offset (In-Die Align) | 1.9873
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y offset (Global Align) | 3.1348
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y 1σ (Vision Pattern Matching) | 0.19311
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y 1σ (In-Die Align) | 0.39905
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning Y 1σ (Global Align) | 0.81277
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X offset (Vision Pattern Matching) | 1.85
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X offset (LLC Translation) | 1.43
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X offset (In-Die Align) | 1.4795
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X offset (Global Align) | 2.2852
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X 1σ (Vision Pattern Matching) | 0.39498
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X 1σ (In-Die Align) | 0.28169
| TSMC AP3 RD | 2024-09-03 | XY Stage positioning X 1σ (Global Align) | 0.43894
| TSMC AP3 RD | 2024-09-03 | XY Orthogonality | 90.1
| TSMC AP3 RD | 2024-09-03 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 106.4
| TSMC AP3 RD | 2024-09-03 | X Ringing(Rpv) | 0.27
| TSMC AP3 RD | 2024-09-03 | X detector noise (XY Servo on) | 0.17
| TSMC AP3 RD | 2024-09-03 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 150.908
| TSMC AP3 RD | 2024-09-03 | X axis, Repeatability @1st position | 0.04
| TSMC AP3 RD | 2024-09-03 | X axis, Repeatability @ 4th position | 0.03
| TSMC AP3 RD | 2024-09-03 | X axis, Repeatability @ 3rd position | 0.04
| TSMC AP3 RD | 2024-09-03 | X axis, Repeatability @ 2nd position | 0.04
| TSMC AP3 RD | 2024-09-03 | X axis, Accuracy @1st position | 0.18
| TSMC AP3 RD | 2024-09-03 | X axis, Accuracy @ 4th position | 0.18
| TSMC AP3 RD | 2024-09-03 | X axis, Accuracy @ 3rd position | 0.21
| TSMC AP3 RD | 2024-09-03 | X axis, Accuracy @ 2nd position | 0.19
| TSMC AP3 RD | 2024-09-03 | X axis OPM | 1.63
| TSMC AP3 RD | 2024-09-03 | StrokeYMin | -153439
| TSMC AP3 RD | 2024-09-03 | StrokeYMax | 150634
| TSMC AP3 RD | 2024-09-03 | StrokeY | 304073
| TSMC AP3 RD | 2024-09-03 | StrokeXMin | -226644
| TSMC AP3 RD | 2024-09-03 | StrokeXMax | 225195
| TSMC AP3 RD | 2024-09-03 | StrokeX | 451839
| TSMC AP3 RD | 2024-09-03 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.82
| TSMC AP3 RD | 2024-09-03 | RMS of ZeroScan | 0.084
| TSMC AP3 RD | 2024-09-03 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4e-05
| TSMC AP3 RD | 2024-09-03 | OpticalAxisShiftPerFocusY | -0.0029
| TSMC AP3 RD | 2024-09-03 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0
| TSMC AP3 RD | 2024-09-03 | OpticalAxisShiftPerFocusX | 0.0013
| TSMC AP3 RD | 2024-09-03 | NCM driving% with AC160TS cantilever | 15.42
| TSMC AP3 RD | 2024-09-03 | Min. Vibration Transmissbility (~100Hz) | 0.173
| TSMC AP3 RD | 2024-09-03 | LowerLimit | 1944
| TSMC AP3 RD | 2024-09-03 | LowerLimit | 6093
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from -Y to +Y (Y scan) | 30.76
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from -Y to +Y (X scan) | 46.14
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from -X to +X (Y scan) | 111.33
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from -X to +X (X scan) | 24.9
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from +Y to -Y (Y scan) | 13.18
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from +Y to -Y (X scan) | 71.78
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from +X to -X (Y scan) | 57.13
| TSMC AP3 RD | 2024-09-03 | Image offset after moving from +X to -X (X scan) | 24.17
| TSMC AP3 RD | 2024-09-03 | Full range of Z Height during Z Scan sweep | 18.9
| TSMC AP3 RD | 2024-09-03 | Full range of Z Height during Z Scan sweep | 18.9
| TSMC AP3 RD | 2024-09-03 | Full range of Z Detector during Z Scan sweep | 14.5
| TSMC AP3 RD | 2024-09-03 | Full range of Z Detector during Z Scan sweep | 14.5
| TSMC AP3 RD | 2024-09-03 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| TSMC AP3 RD | 2024-09-03 | Full range of Y Detector during Y Scan sweep | 116.9
| TSMC AP3 RD | 2024-09-03 | Full range of X Detector Raw during X Scan sweep | 14.3
| TSMC AP3 RD | 2024-09-03 | Full range of X Detector during X Scan sweep | 109.5
| TSMC AP3 RD | 2024-09-03 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 101.01
| TSMC AP3 RD | 2024-09-03 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 101.01
| TSMC AP3 RD | 2024-09-03 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| TSMC AP3 RD | 2024-09-03 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.34
| Polimi, Italy | 2024-09-02 | Z Scanner position at -3μm | 0.004032
| 고려대학교 융합에너지공학과 | 2024-09-02 | Z Scanner position at -3μm | 0.0006
| Polimi, Italy | 2024-09-02 | Z Scanner position at +3μm | 0.004577
| 고려대학교 융합에너지공학과 | 2024-09-02 | Z Scanner position at +3μm | 0.0052
| 고려대학교 융합에너지공학과 | 2024-09-02 | Z linearity by sample height | 1.01
| Polimi, Italy | 2024-09-02 | Z Linearity by 3 different sample height | 1
| Polimi, Italy | 2024-09-02 | Z Height noise | 0.027
| 고려대학교 융합에너지공학과 | 2024-09-02 | Z Height noise | 0.026
| Polimi, Italy | 2024-09-02 | Y Scan Stroke | 116.5
| 고려대학교 융합에너지공학과 | 2024-09-02 | Y Scan Stroke | 75.6
| Polimi, Italy | 2024-09-02 | Y Detector noise (XY Servo on) | 0.28
| 고려대학교 융합에너지공학과 | 2024-09-02 | Y Detector noise (XY Servo on) | 0.26
| 고려대학교 융합에너지공학과 | 2024-09-02 | Y Detector noise (XY Servo off) | 0.25
| Polimi, Italy | 2024-09-02 | Y Detector noise (XY Servo off) | 0.274
| Polimi, Italy | 2024-09-02 | Y axis OPM | 0.8564
| 고려대학교 융합에너지공학과 | 2024-09-02 | Y axis OPM | 0.44
| Polimi, Italy | 2024-09-02 | X Scan Stroke | 120.3
| 고려대학교 융합에너지공학과 | 2024-09-02 | X Scan Stroke | 77.6
| Polimi, Italy | 2024-09-02 | X Detector noise (XY Servo on) | 0.279
| 고려대학교 융합에너지공학과 | 2024-09-02 | X Detector noise (XY Servo on) | 0.26
| 고려대학교 융합에너지공학과 | 2024-09-02 | X Detector noise (XY Servo off) | 0.25
| Polimi, Italy | 2024-09-02 | X Detector noise (XY Servo off) | 0.279
| Polimi, Italy | 2024-09-02 | X axis OPM | 0.6856
| 고려대학교 융합에너지공학과 | 2024-09-02 | X axis OPM | 0.59
| Polimi, Italy | 2024-09-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 고려대학교 융합에너지공학과 | 2024-09-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Polimi, Italy | 2024-09-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 고려대학교 융합에너지공학과 | 2024-09-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Polimi, Italy | 2024-09-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 고려대학교 융합에너지공학과 | 2024-09-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| Polimi, Italy | 2024-09-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 고려대학교 융합에너지공학과 | 2024-09-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Polimi, Italy | 2024-09-02 | NCM driving% with AC160TS cantilever | 0.412
| 고려대학교 융합에너지공학과 | 2024-09-02 | NCM driving% with AC160TS cantilever | 0.11
| Polimi, Italy | 2024-09-02 | NCM Amplitude noise | 10.514
| Polimi, Italy | 2024-09-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.337
| 고려대학교 융합에너지공학과 | 2024-09-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.845
| Polimi, Italy | 2024-09-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.236
| 고려대학교 융합에너지공학과 | 2024-09-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.306
| Polimi, Italy | 2024-09-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 170.3
| 고려대학교 융합에너지공학과 | 2024-09-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.5
| Polimi, Italy | 2024-09-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167
| 고려대학교 융합에너지공학과 | 2024-09-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.4
| Polimi, Italy | 2024-09-02 | Lateral(C-D) sensitivity | 15.2
| 고려대학교 융합에너지공학과 | 2024-09-02 | Lateral(C-D) sensitivity | 18.5
| Polimi, Italy | 2024-09-02 | Lateral(C-D) noise | 0.1
| 고려대학교 융합에너지공학과 | 2024-09-02 | Lateral(C-D) noise | 0.097
| Polimi, Italy | 2024-09-02 | Full range of Z Height during Z Scan sweep | 18.5
| 고려대학교 융합에너지공학과 | 2024-09-02 | Full range of Z Height during Z Scan sweep | 17.3
| Polimi, Italy | 2024-09-02 | Full range of Z Detector during Z Scan sweep | 14.6
| 고려대학교 융합에너지공학과 | 2024-09-02 | Full range of Z Detector during Z Scan sweep | 13.2
| Polimi, Italy | 2024-09-02 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 고려대학교 융합에너지공학과 | 2024-09-02 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Polimi, Italy | 2024-09-02 | Full range of Y Detector during Y Scan sweep | 118.2
| 고려대학교 융합에너지공학과 | 2024-09-02 | Full range of Y Detector during Y Scan sweep | 76.7
| Polimi, Italy | 2024-09-02 | Full range of X Detector Raw during X Scan sweep | 14.3
| 고려대학교 융합에너지공학과 | 2024-09-02 | Full range of X Detector Raw during X Scan sweep | 14.8
| Polimi, Italy | 2024-09-02 | Full range of X Detector during X Scan sweep | 119.2
| 고려대학교 융합에너지공학과 | 2024-09-02 | Full range of X Detector during X Scan sweep | 77.2
| Polimi, Italy | 2024-09-02 | A-B noise | 6.003
| Shenzhen Technology University | 2024-08-30 | Z Scanner position at -3μm | 0.002461
| Shenzhen Technology University | 2024-08-30 | Z Scanner position at +3μm | 0.007174
| Shenzhen Technology University | 2024-08-30 | Z linearity by sample height | 1.0101
| Shenzhen Technology University | 2024-08-30 | Z Height noise | 0.027
| Shenzhen Technology University | 2024-08-30 | Y2 Detector noise (XY Servo on) | 0.265
| Shenzhen Technology University | 2024-08-30 | Y2 Detector noise (XY Servo off) | 0.253
| Shenzhen Technology University | 2024-08-30 | Y1 Detector noise (XY Servo on) | 0.232
| Shenzhen Technology University | 2024-08-30 | Y1 Detector noise (XY Servo off) | 0.223
| Shenzhen Technology University | 2024-08-30 | Y Scan Stroke | 178.4
| Shenzhen Technology University | 2024-08-30 | Y Detector noise (XY Servo on) | 0.179
| Shenzhen Technology University | 2024-08-30 | Y Detector noise (XY Servo off) | 0.17
| Shenzhen Technology University | 2024-08-30 | Y axis OPM | 0.866
| Shenzhen Technology University | 2024-08-30 | X2 Detector noise (XY Servo on) | 0.239
| Shenzhen Technology University | 2024-08-30 | X2 Detector noise (XY Servo off) | 0.231
| Shenzhen Technology University | 2024-08-30 | X1 Detector noise (XY Servo on) | 0.258
| Shenzhen Technology University | 2024-08-30 | X1 Detector noise (XY Servo off) | 0.246
| Shenzhen Technology University | 2024-08-30 | X Scan Stroke | 167.2
| Shenzhen Technology University | 2024-08-30 | X Detector noise (XY Servo on) | 0.184
| Shenzhen Technology University | 2024-08-30 | X Detector noise (XY Servo off) | 0.176
| Shenzhen Technology University | 2024-08-30 | X axis OPM | 1.278
| Hunan San'an Semiconductor | 2024-08-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.05
| Hunan San'an Semiconductor | 2024-08-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.157
| Hunan San'an Semiconductor | 2024-08-30 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.074
| Hunan San'an Semiconductor | 2024-08-30 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.199
| Hunan San'an Semiconductor | 2024-08-30 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.319
| Hunan San'an Semiconductor | 2024-08-30 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.757
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| Shenzhen Technology University | 2024-08-30 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Shenzhen Technology University | 2024-08-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Shenzhen Technology University | 2024-08-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Shenzhen Technology University | 2024-08-30 | NCM Amplitude noise | 12.577
| Shenzhen Technology University | 2024-08-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.066
| Shenzhen Technology University | 2024-08-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.071
| Shenzhen Technology University | 2024-08-30 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 245
| Shenzhen Technology University | 2024-08-30 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.8
| Shenzhen Technology University | 2024-08-30 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 236.1
| Shenzhen Technology University | 2024-08-30 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 228.4
| Shenzhen Technology University | 2024-08-30 | Maximum Y Scan range (Full Stroke) | 178
| Shenzhen Technology University | 2024-08-30 | Maximum X Scan range (Full Stroke) | 167.2
| Shenzhen Technology University | 2024-08-30 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Shenzhen Technology University | 2024-08-30 | Lateral(C-D) sensitivity | 16
| Shenzhen Technology University | 2024-08-30 | Lateral(C-D) noise | 0.109
| Shenzhen Technology University | 2024-08-30 | Full range of Z Height during Z Scan sweep | 17.6
| Shenzhen Technology University | 2024-08-30 | Full range of Z Detector during Z Scan sweep | 13.5
| Shenzhen Technology University | 2024-08-30 | Full range of Y2 Detector raw | 14.7
| Shenzhen Technology University | 2024-08-30 | Full range of Y1 Detector raw | 14.4
| Shenzhen Technology University | 2024-08-30 | Full range of X2 Detector raw | 14.3
| Shenzhen Technology University | 2024-08-30 | Full range of X1 Detector raw | 14.4
| Shenzhen Technology University | 2024-08-30 | A-B noise | 7.253
| Samsung LED | 2024-08-30 | Z Linearity by Z scanner position | 1.4
| Samsung LED | 2024-08-30 | Z Height noise | 0.036
| Samsung LED | 2024-08-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 29.5
| Samsung LED | 2024-08-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 29.4
| Samsung LED | 2024-08-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.8
| Samsung LED | 2024-08-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.9
| Samsung LED | 2024-08-30 | Z Drift | 13.7
| Samsung LED | 2024-08-30 | Z Detector offset | 21.2
| Samsung LED | 2024-08-30 | Y2 Detector noise (XY Servo on) | 0.44
| Samsung LED | 2024-08-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 247.6
| Samsung LED | 2024-08-30 | Y1 Detector noise (XY Servo on) | 0.41
| Samsung LED | 2024-08-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 241.6
| Samsung LED | 2024-08-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 154.3
| Samsung LED | 2024-08-30 | Y Ringing(Rpv) | 0.18
| Samsung LED | 2024-08-30 | Y Non-linearity error | 2.86
| Samsung LED | 2024-08-30 | Y Detector noise (XY Servo on) | 0.38
| Samsung LED | 2024-08-30 | Y axis, Repeatability @1st position | 0.02
| Samsung LED | 2024-08-30 | Y axis, Repeatability @ 4th position | 0.03
| Samsung LED | 2024-08-30 | Y axis, Repeatability @ 3rd position | 0.05
| Samsung LED | 2024-08-30 | Y axis, Repeatability @ 2nd position | 0.03
| Samsung LED | 2024-08-30 | Y axis, Accuracy @1st position | 0.28
| Samsung LED | 2024-08-30 | Y axis, Accuracy @ 4th position | 0.37
| Samsung LED | 2024-08-30 | Y axis, Accuracy @ 3rd position | 0.43
| Samsung LED | 2024-08-30 | Y axis, Accuracy @ 2nd position | 0.27
| Samsung LED | 2024-08-30 | Y axis OPM | 0.81
| Samsung LED | 2024-08-30 | XY Orthogonality | 89.9
| Samsung LED | 2024-08-30 | X2 Detector noise (XY Servo on) | 0.29
| Samsung LED | 2024-08-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 247
| Samsung LED | 2024-08-30 | X1 Detector noise (XY Servo on) | 0.31
| Samsung LED | 2024-08-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 246.9
| Samsung LED | 2024-08-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 152.4
| Samsung LED | 2024-08-30 | X Ringing(Rpv) | 0.1
| Samsung LED | 2024-08-30 | X Non-linearity error | 3.42
| Samsung LED | 2024-08-30 | X Detector noise (XY Servo on) | 0.24
| Samsung LED | 2024-08-30 | X axis, Repeatability @1st position | 0.05
| Samsung LED | 2024-08-30 | X axis, Repeatability @ 4th position | 0.08
| Samsung LED | 2024-08-30 | X axis, Repeatability @ 3rd position | 0.06
| Samsung LED | 2024-08-30 | X axis, Repeatability @ 2nd position | 0.06
| Samsung LED | 2024-08-30 | X axis, Accuracy @1st position | 0.23
| Samsung LED | 2024-08-30 | X axis, Accuracy @ 4th position | 0.39
| Samsung LED | 2024-08-30 | X axis, Accuracy @ 3rd position | 0.32
| Samsung LED | 2024-08-30 | X axis, Accuracy @ 2nd position | 0.33
| Samsung LED | 2024-08-30 | X axis OPM | 0.12
| Samsung LED | 2024-08-30 | Wafer placement(Loading) error on the load port2 (Y offset) | 36.61
| Samsung LED | 2024-08-30 | Wafer placement(Loading) error on the load port2 (X offset) | 33.56
| Samsung LED | 2024-08-30 | Wafer placement(Loading) error on the load port1 (Y offset) | 21.66
| Samsung LED | 2024-08-30 | Wafer placement(Loading) error on the load port1 (X offset) | 28.8
| Samsung LED | 2024-08-30 | Torque Profile (50mm, 20um/s) | 0.601
| Samsung LED | 2024-08-30 | Torque Profile (50mm, 200um/s) | 0.547
| Samsung LED | 2024-08-30 | StrokeYMin | -152658
| Samsung LED | 2024-08-30 | StrokeYMax | 150997
| Samsung LED | 2024-08-30 | StrokeY | 303655
| Samsung LED | 2024-08-30 | StrokeXMin | -205723
| Samsung LED | 2024-08-30 | StrokeXMax | 203632
| Samsung LED | 2024-08-30 | StrokeX | 409355
| Samsung LED | 2024-08-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.14
| Samsung LED | 2024-08-30 | Roll motion to X+ full scan | 21.9
| Samsung LED | 2024-08-30 | RMS of ZeroScan | 0.074
| Samsung LED | 2024-08-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Samsung LED | 2024-08-30 | OpticalAxisShiftPerFocusY | -0.002
| Samsung LED | 2024-08-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00033
| Samsung LED | 2024-08-30 | OpticalAxisShiftPerFocusX | -0.0039
| Samsung LED | 2024-08-30 | OPM Repeatability @5mm | 0.84
| Samsung LED | 2024-08-30 | OPM Repeatability @25mm | 6.05
| Samsung LED | 2024-08-30 | OPM Repeatability @1mm | 0.67
| Samsung LED | 2024-08-30 | OPM Repeatability @10mm | 3.02
| Samsung LED | 2024-08-30 | NCM driving% with AC160TS cantilever | 4.9
| Samsung LED | 2024-08-30 | Min. Vibration Transmissbility (~100Hz) | 0.048
| Samsung LED | 2024-08-30 | Max OPM @5mm | 34.91
| Samsung LED | 2024-08-30 | Max OPM @25mm | 185.91
| Samsung LED | 2024-08-30 | Max OPM @1mm | 10.51
| Samsung LED | 2024-08-30 | Max OPM @10mm | 76.98
| Samsung LED | 2024-08-30 | LowerLimit | 2085
| Samsung LED | 2024-08-30 | LowerLimit | 4973
| Samsung LED | 2024-08-30 | Lateral(C-D) sensitivity | 124.4
| Samsung LED | 2024-08-30 | Lateral(C-D) noise | 0.194
| Samsung LED | 2024-08-30 | Image offset after moving from -Y to +Y (Y scan) | 71.78
| Samsung LED | 2024-08-30 | Image offset after moving from -Y to +Y (X scan) | 27.1
| Samsung LED | 2024-08-30 | Image offset after moving from -X to +X (Y scan) | 45.41
| Samsung LED | 2024-08-30 | Image offset after moving from -X to +X (X scan) | 33.69
| Samsung LED | 2024-08-30 | Image offset after moving from +Y to -Y (Y scan) | 27.1
| Samsung LED | 2024-08-30 | Image offset after moving from +Y to -Y (X scan) | 13.18
| Samsung LED | 2024-08-30 | Image offset after moving from +X to -X (Y scan) | 37.35
| Samsung LED | 2024-08-30 | Image offset after moving from +X to -X (X scan) | 26.37
| Samsung LED | 2024-08-30 | Full range of Z Height during Z Scan sweep | 17.2
| Samsung LED | 2024-08-30 | Full range of Z Detector during Z Scan sweep | 11.6
| Samsung LED | 2024-08-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.08
| Samsung LED | 2024-08-30 | Full range of Y2 Detector during Y Scan sweep | 175.3
| Samsung LED | 2024-08-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.29
| Samsung LED | 2024-08-30 | Full range of Y1 Detector during Y Scan sweep | 169.6
| Samsung LED | 2024-08-30 | Full range of Y Detector during Y Scan sweep | 172.4
| Samsung LED | 2024-08-30 | Full range of X2 Detector Raw during X Scan sweep | 13.98
| Samsung LED | 2024-08-30 | Full range of X2 Detector during X Scan sweep | 175.9
| Samsung LED | 2024-08-30 | Full range of X1 Detector Raw during X Scan sweep | 14.09
| Samsung LED | 2024-08-30 | Full range of X1 Detector during X Scan sweep | 176.1
| Samsung LED | 2024-08-30 | Full range of X Detector during X Scan sweep | 176
| Samsung LED | 2024-08-30 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Samsung LED | 2024-08-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Samsung LED | 2024-08-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| Samsung LED | 2024-08-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.78
| Samsung LED | 2024-08-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| Samsung LED | 2024-08-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Samsung LED | 2024-08-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| Samsung LED | 2024-08-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.36
| Samsung LED | 2024-08-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| Samsung LED | 2024-08-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.2
| Samsung LED | 2024-08-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 1.7
| Samsung LED | 2024-08-30 | %GRR w/ roughness of non-patterened wafer | 5.4
| University of Bologna | 2024-08-29 | Z Scanner position at -3μm | 0.009915
| University of Bologna | 2024-08-29 | Z Scanner position at +3μm | 0.010789
| University of Bologna | 2024-08-29 | Z linearity by sample height | 1.0103
| University of Bologna | 2024-08-29 | Z Height noise | 0.028
| University of Bologna | 2024-08-29 | Y Scan Stroke | 190.7
| University of Bologna | 2024-08-29 | Y Detector noise (XY Servo on) | 0.371
| University of Bologna | 2024-08-29 | Y Detector noise (XY Servo off) | 0.372
| University of Bologna | 2024-08-29 | Y axis OPM | 1.8764
| University of Bologna | 2024-08-29 | X Scan Stroke | 186.9
| University of Bologna | 2024-08-29 | X Detector noise (XY Servo on) | 0.362
| University of Bologna | 2024-08-29 | X Detector noise (XY Servo off) | 0.354
| University of Bologna | 2024-08-29 | X axis OPM | 0.8574
| University of Bologna | 2024-08-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.51
| University of Bologna | 2024-08-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.88
| University of Bologna | 2024-08-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.49
| University of Bologna | 2024-08-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.87
| University of Bologna | 2024-08-29 | NCM Amplitude noise | 14.676
| University of Bologna | 2024-08-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.426
| University of Bologna | 2024-08-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.747
| University of Bologna | 2024-08-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 237.4
| University of Bologna | 2024-08-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 235.5
| University of Bologna | 2024-08-29 | Lateral(C-D) sensitivity | 17.2
| University of Bologna | 2024-08-29 | Lateral(C-D) noise | 0.144
| University of Bologna | 2024-08-29 | Full range of Z Height during Z Scan sweep | 16.6
| University of Bologna | 2024-08-29 | Full range of Z Detector during Z Scan sweep | 11.9
| University of Bologna | 2024-08-29 | Full range of Y Detector Raw during Y Scan sweep | 15.9
| University of Bologna | 2024-08-29 | Full range of Y Detector during Y Scan sweep | 190
| University of Bologna | 2024-08-29 | Full range of X Detector Raw during X Scan sweep | 15.9
| University of Bologna | 2024-08-29 | Full range of X Detector during X Scan sweep | 187
| University of Bologna | 2024-08-29 | A-B noise | 9.953
| 주성엔지니어링 | 2024-08-28 | Z Scanner position at -3μm | 0.0068
| Hirosaki University | 2024-08-28 | Z Scanner position at -3μm | 0.000183
| 주성엔지니어링 | 2024-08-28 | Z Scanner position at +3μm | 0.0014
| Hirosaki University | 2024-08-28 | Z Scanner position at +3μm | 0.003358
| Hirosaki University | 2024-08-28 | Z linearity by sample height | 1.0012
| 주성엔지니어링 | 2024-08-28 | Z linearity by sample height | 1.006
| Hirosaki University | 2024-08-28 | Z Height noise | 0.03
| 주성엔지니어링 | 2024-08-28 | Z Height noise | 0.02
| Hirosaki University | 2024-08-28 | Y Scan Stroke | 55
| 주성엔지니어링 | 2024-08-28 | Y Scan Stroke | 113.3
| Hirosaki University | 2024-08-28 | Y Detector noise (XY Servo on) | 0.396
| 주성엔지니어링 | 2024-08-28 | Y Detector noise (XY Servo on) | 0.25
| Hirosaki University | 2024-08-28 | Y Detector noise (XY Servo off) | 0.422
| 주성엔지니어링 | 2024-08-28 | Y Detector noise (XY Servo off) | 0.24
| Hirosaki University | 2024-08-28 | Y axis OPM | 0.6796
| 주성엔지니어링 | 2024-08-28 | Y axis OPM | 1.36
| Hirosaki University | 2024-08-28 | XY orthogonality | 90.15
| 주성엔지니어링 | 2024-08-28 | XY orthogonality | 90.09
| Hirosaki University | 2024-08-28 | X Scan Stroke | 55
| 주성엔지니어링 | 2024-08-28 | X Scan Stroke | 108.9
| Hirosaki University | 2024-08-28 | X Detector noise (XY Servo on) | 0.388
| 주성엔지니어링 | 2024-08-28 | X Detector noise (XY Servo on) | 0.25
| Hirosaki University | 2024-08-28 | X Detector noise (XY Servo off) | 0.415
| 주성엔지니어링 | 2024-08-28 | X Detector noise (XY Servo off) | 0.24
| Hirosaki University | 2024-08-28 | X axis OPM | 0.9574
| 주성엔지니어링 | 2024-08-28 | X axis OPM | 1.12
| 주성엔지니어링 | 2024-08-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.037
| 주성엔지니어링 | 2024-08-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.243
| 주성엔지니어링 | 2024-08-28 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 7.19
| 주성엔지니어링 | 2024-08-28 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.149
| 주성엔지니어링 | 2024-08-28 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.575
| 주성엔지니어링 | 2024-08-28 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.785
| Hirosaki University | 2024-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 주성엔지니어링 | 2024-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Hirosaki University | 2024-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| 주성엔지니어링 | 2024-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Hirosaki University | 2024-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 주성엔지니어링 | 2024-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Hirosaki University | 2024-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 주성엔지니어링 | 2024-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 주성엔지니어링 | 2024-08-28 | NCM driving% with AC160TS cantilever | 2.26
| Hirosaki University | 2024-08-28 | NCM Amplitude noise | 8.397
| 주성엔지니어링 | 2024-08-28 | NCM Amplitude noise | 5.2
| Hirosaki University | 2024-08-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.169
| 주성엔지니어링 | 2024-08-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.814
| Hirosaki University | 2024-08-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.411
| 주성엔지니어링 | 2024-08-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 20.791
| Hirosaki University | 2024-08-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 95.8
| 주성엔지니어링 | 2024-08-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 154.1
| Hirosaki University | 2024-08-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.8
| 주성엔지니어링 | 2024-08-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.4
| Hirosaki University | 2024-08-28 | Lateral(C-D) sensitivity | 18.2
| 주성엔지니어링 | 2024-08-28 | Lateral(C-D) sensitivity | 7.8
| Hirosaki University | 2024-08-28 | Lateral(C-D) noise | 0.046
| 주성엔지니어링 | 2024-08-28 | Lateral(C-D) noise | 0.12
| Hirosaki University | 2024-08-28 | Full range of Z Height during Z Scan sweep | 18.1
| 주성엔지니어링 | 2024-08-28 | Full range of Z Height during Z Scan sweep | 15.4
| Hirosaki University | 2024-08-28 | Full range of Z Detector during Z Scan sweep | 13.2
| 주성엔지니어링 | 2024-08-28 | Full range of Z Detector during Z Scan sweep | 14.8
| Hirosaki University | 2024-08-28 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| 주성엔지니어링 | 2024-08-28 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Hirosaki University | 2024-08-28 | Full range of Y Detector during Y Scan sweep | 68.3
| 주성엔지니어링 | 2024-08-28 | Full range of Y Detector during Y Scan sweep | 113.1
| Hirosaki University | 2024-08-28 | Full range of X Detector Raw during X Scan sweep | 13.7
| 주성엔지니어링 | 2024-08-28 | Full range of X Detector Raw during X Scan sweep | 14.1
| Hirosaki University | 2024-08-28 | Full range of X Detector during X Scan sweep | 65.5
| 주성엔지니어링 | 2024-08-28 | Full range of X Detector during X Scan sweep | 108.9
| Hirosaki University | 2024-08-28 | A-B noise | 6.325
| 주성엔지니어링 | 2024-08-28 | A-B noise | 10.13
| Dupont US | 2024-08-28 | Z Linearity by Z scanner position | 1
| TSMC F20 EBO | 2024-08-28 | Z Linearity by Z scanner position | 1.3
| Hauhong Wuxi #4 | 2024-08-28 | Z Linearity by Z scanner position | 0.6
| Dupont US | 2024-08-28 | Z Linearity by 4 different sample height | 1.003
| TSMC F20 EBO | 2024-08-28 | Z Linearity by 4 different sample height | 1.003
| Dupont US | 2024-08-28 | Z Height noise | 0.023
| TSMC F20 EBO | 2024-08-28 | Z Height noise | 0.028
| Hauhong Wuxi #4 | 2024-08-28 | Z Height noise | 0.025
| Dupont US | 2024-08-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.7
| Hauhong Wuxi #4 | 2024-08-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.5
| Dupont US | 2024-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.7
| TSMC F20 EBO | 2024-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.88
| Hauhong Wuxi #4 | 2024-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.6
| Dupont US | 2024-08-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 12
| Hauhong Wuxi #4 | 2024-08-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6
| Dupont US | 2024-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.8
| TSMC F20 EBO | 2024-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.55
| Hauhong Wuxi #4 | 2024-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6
| Dupont US | 2024-08-28 | Z Drift | 12.7
| TSMC F20 EBO | 2024-08-28 | Z Drift | 17.85
| Hauhong Wuxi #4 | 2024-08-28 | Z Drift | 11
| Dupont US | 2024-08-28 | Z Detector offset | 168.2
| Hauhong Wuxi #4 | 2024-08-28 | Z Detector offset | 49.8
| Dupont US | 2024-08-28 | Y2 Detector noise (XY Servo on) | 0.45
| Hauhong Wuxi #4 | 2024-08-28 | Y2 Detector noise (XY Servo on) | 0.42
| Dupont US | 2024-08-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 246.6
| Hauhong Wuxi #4 | 2024-08-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 265.4
| Dupont US | 2024-08-28 | Y1 Detector noise (XY Servo on) | 0.45
| Hauhong Wuxi #4 | 2024-08-28 | Y1 Detector noise (XY Servo on) | 0.38
| Dupont US | 2024-08-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 219.9
| Hauhong Wuxi #4 | 2024-08-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 261.5
| Dupont US | 2024-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 159.1
| TSMC F20 EBO | 2024-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 61.6
| Hauhong Wuxi #4 | 2024-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 164.9
| Dupont US | 2024-08-28 | Y Ringing(Rpv) | 0.26
| TSMC F20 EBO | 2024-08-28 | Y Ringing(Rpv) | 0.162
| Hauhong Wuxi #4 | 2024-08-28 | Y Ringing(Rpv) | 1.29
| Hauhong Wuxi #4 | 2024-08-28 | Y Non-linearity error | 3.3
| Dupont US | 2024-08-28 | Y Detector noise (XY Servo on) | 0.36
| TSMC F20 EBO | 2024-08-28 | Y detector noise (XY Servo on) | 0.171
| Hauhong Wuxi #4 | 2024-08-28 | Y Detector noise (XY Servo on) | 0.3
| TSMC F20 EBO | 2024-08-28 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 88.74
| Dupont US | 2024-08-28 | Y axis, Repeatability @1st position | 0.04
| TSMC F20 EBO | 2024-08-28 | Y axis, Repeatability @1st position | 0.04
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Repeatability @1st position | 0.04
| Dupont US | 2024-08-28 | Y axis, Repeatability @ 4th position | 0.03
| TSMC F20 EBO | 2024-08-28 | Y axis, Repeatability @ 4th position | 0.03
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Repeatability @ 4th position | 0.05
| Dupont US | 2024-08-28 | Y axis, Repeatability @ 3rd position | 0.02
| TSMC F20 EBO | 2024-08-28 | Y axis, Repeatability @ 3rd position | 0.03
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Repeatability @ 3rd position | 0.04
| Dupont US | 2024-08-28 | Y axis, Repeatability @ 2nd position | 0.04
| TSMC F20 EBO | 2024-08-28 | Y axis, Repeatability @ 2nd position | 0.04
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Repeatability @ 2nd position | 0.06
| Dupont US | 2024-08-28 | Y axis, Accuracy @1st position | 0.31
| TSMC F20 EBO | 2024-08-28 | Y axis, Accuracy @1st position | 0.23
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Accuracy @1st position | 0.51
| Dupont US | 2024-08-28 | Y axis, Accuracy @ 4th position | 0.24
| TSMC F20 EBO | 2024-08-28 | Y axis, Accuracy @ 4th position | 0.23
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Accuracy @ 4th position | 0.5
| Dupont US | 2024-08-28 | Y axis, Accuracy @ 3rd position | 0.21
| TSMC F20 EBO | 2024-08-28 | Y axis, Accuracy @ 3rd position | 0.23
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Accuracy @ 3rd position | 0.37
| Dupont US | 2024-08-28 | Y axis, Accuracy @ 2nd position | 0.25
| TSMC F20 EBO | 2024-08-28 | Y axis, Accuracy @ 2nd position | 0.25
| Hauhong Wuxi #4 | 2024-08-28 | Y axis, Accuracy @ 2nd position | 0.39
| Dupont US | 2024-08-28 | Y axis OPM | 0.96
| TSMC F20 EBO | 2024-08-28 | Y axis OPM | 1.11
| Hauhong Wuxi #4 | 2024-08-28 | Y axis OPM | 1.39
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.5527
| Dupont US | 2024-08-28 | XY Stage positioning Y offset (Vision Pattern Matching)
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning Y offset (LLC Translation) | 2.6855
| Dupont US | 2024-08-28 | XY Stage positioning Y offset (LLC Translation) | 2.4219
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning Y offset (In-Die Align) | 2.0459
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning Y offset (Global Align) | 3.4961
| Dupont US | 2024-08-28 | XY Stage positioning Y offset (Global Align) | 2.5391
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning Y 1σ (In-Die Align) | 0.36602
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning X offset (Vision Pattern Matching) | 2.6465
| Dupont US | 2024-08-28 | XY Stage positioning X offset (Vision Pattern Matching)
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning X offset (LLC Translation) | 2.9004
| Dupont US | 2024-08-28 | XY Stage positioning X offset (LLC Translation) | 2.7441
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning X offset (In-Die Align) | 1.7236
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning X offset (Global Align) | 2.5928
| Dupont US | 2024-08-28 | XY Stage positioning X offset (Global Align) | 2.5391
| Hauhong Wuxi #4 | 2024-08-28 | XY Stage positioning X 1σ (In-Die Align) | 0.41436
| Dupont US | 2024-08-28 | XY Orthogonality | 89.98
| TSMC F20 EBO | 2024-08-28 | XY Orthogonality | 89.9
| Hauhong Wuxi #4 | 2024-08-28 | XY Orthogonality | 89.86
| Dupont US | 2024-08-28 | X2 Detector noise (XY Servo on) | 0.45
| Hauhong Wuxi #4 | 2024-08-28 | X2 Detector noise (XY Servo on) | 0.37
| Dupont US | 2024-08-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 237.4
| Hauhong Wuxi #4 | 2024-08-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 240.5
| Dupont US | 2024-08-28 | X1 Detector noise (XY Servo on) | 0.47
| Hauhong Wuxi #4 | 2024-08-28 | X1 Detector noise (XY Servo on) | 0.37
| Dupont US | 2024-08-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 258.6
| Hauhong Wuxi #4 | 2024-08-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 252.9
| Dupont US | 2024-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 170
| TSMC F20 EBO | 2024-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 60.02
| Hauhong Wuxi #4 | 2024-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 170
| Dupont US | 2024-08-28 | X Ringing(Rpv) | 0.15
| TSMC F20 EBO | 2024-08-28 | X Ringing(Rpv) | 0.161
| Hauhong Wuxi #4 | 2024-08-28 | X Ringing(Rpv) | 0.57
| Hauhong Wuxi #4 | 2024-08-28 | X Non-linearity error | 3.07
| Dupont US | 2024-08-28 | X Detector noise (XY Servo on) | 0.35
| TSMC F20 EBO | 2024-08-28 | X detector noise (XY Servo on) | 0.17
| Hauhong Wuxi #4 | 2024-08-28 | X Detector noise (XY Servo on) | 0.28
| TSMC F20 EBO | 2024-08-28 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 85.14
| Dupont US | 2024-08-28 | X axis, Repeatability @1st position | 0.04
| TSMC F20 EBO | 2024-08-28 | X axis, Repeatability @1st position | 0.03
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Repeatability @1st position | 0.06
| Dupont US | 2024-08-28 | X axis, Repeatability @ 4th position | 0.06
| TSMC F20 EBO | 2024-08-28 | X axis, Repeatability @ 4th position | 0.02
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Repeatability @ 4th position | 0.04
| Dupont US | 2024-08-28 | X axis, Repeatability @ 3rd position | 0.04
| TSMC F20 EBO | 2024-08-28 | X axis, Repeatability @ 3rd position | 0.03
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Repeatability @ 3rd position | 0.04
| Dupont US | 2024-08-28 | X axis, Repeatability @ 2nd position | 0.04
| TSMC F20 EBO | 2024-08-28 | X axis, Repeatability @ 2nd position | 0.03
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Repeatability @ 2nd position | 0.06
| Dupont US | 2024-08-28 | X axis, Accuracy @1st position | 0.2
| TSMC F20 EBO | 2024-08-28 | X axis, Accuracy @1st position | 0.13
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Accuracy @1st position | 0.35
| Dupont US | 2024-08-28 | X axis, Accuracy @ 4th position | 0.26
| TSMC F20 EBO | 2024-08-28 | X axis, Accuracy @ 4th position | 0.14
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Accuracy @ 4th position | 0.23
| Dupont US | 2024-08-28 | X axis, Accuracy @ 3rd position | 0.22
| TSMC F20 EBO | 2024-08-28 | X axis, Accuracy @ 3rd position | 0.15
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Accuracy @ 3rd position | 0.26
| Dupont US | 2024-08-28 | X axis, Accuracy @ 2nd position | 0.21
| TSMC F20 EBO | 2024-08-28 | X axis, Accuracy @ 2nd position | 0.16
| Hauhong Wuxi #4 | 2024-08-28 | X axis, Accuracy @ 2nd position | 0.28
| Dupont US | 2024-08-28 | X axis OPM | 1.04
| TSMC F20 EBO | 2024-08-28 | X axis OPM | 1.28
| Hauhong Wuxi #4 | 2024-08-28 | X axis OPM | 1.17
| Dupont US | 2024-08-28 | Wafer placement(Loading) error on the load port2 (Y offset) | 22.4
| Hauhong Wuxi #4 | 2024-08-28 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.5
| Dupont US | 2024-08-28 | Wafer placement(Loading) error on the load port2 (X offset) | 34.3
| Hauhong Wuxi #4 | 2024-08-28 | Wafer placement(Loading) error on the load port2 (X offset) | 36.1
| Dupont US | 2024-08-28 | Wafer placement(Loading) error on the load port1 (Y offset) | 17.9
| Hauhong Wuxi #4 | 2024-08-28 | Wafer placement(Loading) error on the load port1 (Y offset) | 25.1
| Dupont US | 2024-08-28 | Wafer placement(Loading) error on the load port1 (X offset) | 34.7
| Hauhong Wuxi #4 | 2024-08-28 | Wafer placement(Loading) error on the load port1 (X offset) | 75.8
| Hauhong Wuxi #4 | 2024-08-28 | Torque Profile (50mm, 20um/s) | 0.964
| Hauhong Wuxi #4 | 2024-08-28 | Torque Profile (50mm, 200um/s) | 0.838
| Dupont US | 2024-08-28 | StrokeYMin | -152427
| Hauhong Wuxi #4 | 2024-08-28 | StrokeYMin | -153685
| Dupont US | 2024-08-28 | StrokeYMax | 150028
| Hauhong Wuxi #4 | 2024-08-28 | StrokeYMax | 150107
| Dupont US | 2024-08-28 | StrokeY | 302455
| Hauhong Wuxi #4 | 2024-08-28 | StrokeY | 303792
| Dupont US | 2024-08-28 | StrokeXMin | -203644
| Hauhong Wuxi #4 | 2024-08-28 | StrokeXMin | -201334
| Dupont US | 2024-08-28 | StrokeXMax | 203118
| Hauhong Wuxi #4 | 2024-08-28 | StrokeXMax | 203206
| Dupont US | 2024-08-28 | StrokeX | 406762
| Hauhong Wuxi #4 | 2024-08-28 | StrokeX | 404540
| Dupont US | 2024-08-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7
| Hauhong Wuxi #4 | 2024-08-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7
| TSMC F20 EBO | 2024-08-28 | Scan offset accuracy (Precision) | 1.426
| Dupont US | 2024-08-28 | Roll motion to X+ full scan | 16.9
| Hauhong Wuxi #4 | 2024-08-28 | Roll motion to X+ full scan | 22.3
| Dupont US | 2024-08-28 | RMS of ZeroScan | 0.054
| TSMC F20 EBO | 2024-08-28 | RMS of ZeroScan | 0.037
| Hauhong Wuxi #4 | 2024-08-28 | RMS of ZeroScan | 0.058
| TSMC F20 EBO | 2024-08-28 | RMS error after Stage Calibration | 1.06
| Dupont US | 2024-08-28 | Position repeatability(10repeats, Position Based) | 0.38
| Hauhong Wuxi #4 | 2024-08-28 | Position repeatability(10repeats, Position Based) | 0.31
| Dupont US | 2024-08-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00027
| Hauhong Wuxi #4 | 2024-08-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00021
| Dupont US | 2024-08-28 | OpticalAxisShiftPerFocusY | -0.0004
| Hauhong Wuxi #4 | 2024-08-28 | OpticalAxisShiftPerFocusY | -0.0006
| Dupont US | 2024-08-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00033
| Hauhong Wuxi #4 | 2024-08-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.0003
| Dupont US | 2024-08-28 | OpticalAxisShiftPerFocusX | -0.002
| Hauhong Wuxi #4 | 2024-08-28 | OpticalAxisShiftPerFocusX | -0.0023
| Dupont US | 2024-08-28 | OPM Repeatability @5mm | 0.88
| Hauhong Wuxi #4 | 2024-08-28 | OPM Repeatability @5mm | 0.93
| Dupont US | 2024-08-28 | OPM Repeatability @25mm | 3.48
| Hauhong Wuxi #4 | 2024-08-28 | OPM Repeatability @25mm | 4.1
| Dupont US | 2024-08-28 | OPM Repeatability @1mm | 0.77
| Hauhong Wuxi #4 | 2024-08-28 | OPM Repeatability @1mm | 0.61
| Dupont US | 2024-08-28 | OPM Repeatability @10mm | 1.41
| Hauhong Wuxi #4 | 2024-08-28 | OPM Repeatability @10mm | 1.49
| Dupont US | 2024-08-28 | NCM driving% with AC160TS cantilever | 28
| Hauhong Wuxi #4 | 2024-08-28 | NCM driving% with AC160TS cantilever | 5.6
| Dupont US | 2024-08-28 | Min. Vibration Transmissbility (~100Hz) | 0.029
| TSMC F20 EBO | 2024-08-28 | Min. Vibration Transmissbility (~100Hz) | 0.032
| Hauhong Wuxi #4 | 2024-08-28 | Min. Vibration Transmissbility (~100Hz) | 0.047
| Dupont US | 2024-08-28 | Max OPM @5mm | 34.86
| Hauhong Wuxi #4 | 2024-08-28 | Max OPM @5mm | 24.22
| Dupont US | 2024-08-28 | Max OPM @25mm | 123.71
| Hauhong Wuxi #4 | 2024-08-28 | Max OPM @25mm | 105.27
| Dupont US | 2024-08-28 | Max OPM @1mm | 6.4
| Hauhong Wuxi #4 | 2024-08-28 | Max OPM @1mm | 5.53
| Dupont US | 2024-08-28 | Max OPM @10mm | 51.09
| Hauhong Wuxi #4 | 2024-08-28 | Max OPM @10mm | 28.6
| TSMC F20 EBO | 2024-08-28 | Mask Roughness Repeatability (10Repeats, σ ) | 0.0035
| TSMC F20 EBO | 2024-08-28 | Mask Roughness Repeatability (10Repeats, RSD%) | 2.6
| TSMC F20 EBO | 2024-08-28 | Mask exchange time | 2.97
| Dupont US | 2024-08-28 | LowerLimit | 2367
| Dupont US | 2024-08-28 | LowerLimit | 5000
| Hauhong Wuxi #4 | 2024-08-28 | LowerLimit | 2390
| Hauhong Wuxi #4 | 2024-08-28 | LowerLimit | 4946
| Dupont US | 2024-08-28 | Lateral(C-D) sensitivity | 17.3
| Hauhong Wuxi #4 | 2024-08-28 | Lateral(C-D) sensitivity | 17.3
| Dupont US | 2024-08-28 | Lateral(C-D) noise | 0.2
| Hauhong Wuxi #4 | 2024-08-28 | Lateral(C-D) noise | 0.2
| Dupont US | 2024-08-28 | Image offset after moving from -Y to +Y (Y scan) | 41.02
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from -Y to +Y (Y scan) | 19.53
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from -Y to +Y (Y scan) | 76.9
| Dupont US | 2024-08-28 | Image offset after moving from -Y to +Y (X scan) | 40.28
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from -Y to +Y (X scan) | 19.53
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from -Y to +Y (X scan) | 21.97
| Dupont US | 2024-08-28 | Image offset after moving from -X to +X (Y scan) | 18.31
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from -X to +X (Y scan) | 24.41
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from -X to +X (Y scan) | 46.88
| Dupont US | 2024-08-28 | Image offset after moving from -X to +X (X scan) | 93.02
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from -X to +X (X scan) | 29.29
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from -X to +X (X scan) | 20.51
| Dupont US | 2024-08-28 | Image offset after moving from +Y to -Y (Y scan) | 48.34
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from +Y to -Y (Y scan) | 48.82
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from +Y to -Y (Y scan) | 50.54
| Dupont US | 2024-08-28 | Image offset after moving from +Y to -Y (X scan) | 16.11
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from +Y to -Y (X scan) | 24.41
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from +Y to -Y (X scan) | 15.38
| Dupont US | 2024-08-28 | Image offset after moving from +X to -X (Y scan) | 13.18
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from +X to -X (Y scan) | 68.35
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from +X to -X (Y scan) | 25.63
| Dupont US | 2024-08-28 | Image offset after moving from +X to -X (X scan) | 43.95
| TSMC F20 EBO | 2024-08-28 | Image offset after moving from +X to -X (X scan) | 14.64
| Hauhong Wuxi #4 | 2024-08-28 | Image offset after moving from +X to -X (X scan) | 17.58
| TSMC F20 EBO | 2024-08-28 | Imag offset after Lock stage (Y scan) | 97.656
| TSMC F20 EBO | 2024-08-28 | Imag offset after Lock stage (X scan) | 78.125
| TSMC F20 EBO | 2024-08-28 | Imag offset after EFEM loading (Y scan) | 156.25
| TSMC F20 EBO | 2024-08-28 | Imag offset after EFEM loading (X scan) | 185.54
| Dupont US | 2024-08-28 | Full range of Z Height during Z Scan sweep | 17.4
| TSMC F20 EBO | 2024-08-28 | Full range of Z Height during Z Scan sweep | 18.8
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Z Height during Z Scan sweep | 17.3
| Dupont US | 2024-08-28 | Full range of Z Detector during Z Scan sweep | 14.1
| TSMC F20 EBO | 2024-08-28 | Full range of Z Detector during Z Scan sweep | 14.1
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Z Detector during Z Scan sweep | 13.9
| Dupont US | 2024-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.41
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Dupont US | 2024-08-28 | Full range of Y2 Detector during Y Scan sweep | 188.3
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Y2 Detector during Y Scan sweep | 198.9
| Dupont US | 2024-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.75
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.55
| Dupont US | 2024-08-28 | Full range of Y1 Detector during Y Scan sweep | 173.4
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Y1 Detector during Y Scan sweep | 196.5
| TSMC F20 EBO | 2024-08-28 | Full range of Y Detector Raw during Y Scan sweep | 14.98
| Dupont US | 2024-08-28 | Full range of Y Detector during Y Scan sweep | 188.3
| TSMC F20 EBO | 2024-08-28 | Full range of Y Detector during Y Scan sweep | 66.45
| Hauhong Wuxi #4 | 2024-08-28 | Full range of Y Detector during Y Scan sweep | 198.9
| Dupont US | 2024-08-28 | Full range of X2 Detector Raw during X Scan sweep | 14.29
| Hauhong Wuxi #4 | 2024-08-28 | Full range of X2 Detector Raw during X Scan sweep | 14.12
| Dupont US | 2024-08-28 | Full range of X2 Detector during X Scan sweep | 180.9
| Hauhong Wuxi #4 | 2024-08-28 | Full range of X2 Detector during X Scan sweep | 197.4
| Dupont US | 2024-08-28 | Full range of X1 Detector Raw during X Scan sweep | 13.98
| Hauhong Wuxi #4 | 2024-08-28 | Full range of X1 Detector Raw during X Scan sweep | 14.25
| Dupont US | 2024-08-28 | Full range of X1 Detector during X Scan sweep | 193.3
| Hauhong Wuxi #4 | 2024-08-28 | Full range of X1 Detector during X Scan sweep | 192.7
| TSMC F20 EBO | 2024-08-28 | Full range of X Detector Raw during X Scan sweep | 14.87
| Dupont US | 2024-08-28 | Full range of X Detector during X Scan sweep | 180.9
| TSMC F20 EBO | 2024-08-28 | Full range of X Detector during X Scan sweep | 63.31
| Hauhong Wuxi #4 | 2024-08-28 | Full range of X Detector during X Scan sweep | 195.1
| TSMC F20 EBO | 2024-08-28 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.505
| TSMC F20 EBO | 2024-08-28 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.505
| Dupont US | 2024-08-28 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Hauhong Wuxi #4 | 2024-08-28 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (Y offset-Point Align3) | 11.59
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (Y offset-Point Align2) | 11.04
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (Y offset-Point Align1) | 8.34
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (Y offset) | 0.71
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (X offset-Point Align3) | 1.08
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (X offset-Point Align2) | 4.3
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (X offset-Point Align1) | 3.39
| TSMC F20 EBO | 2024-08-28 | Aligned stage accuracy (X offset) | 0.84
| Dupont US | 2024-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Hauhong Wuxi #4 | 2024-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.16
| Dupont US | 2024-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.73
| Hauhong Wuxi #4 | 2024-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.9
| Dupont US | 2024-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.44
| Hauhong Wuxi #4 | 2024-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.83
| Dupont US | 2024-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Hauhong Wuxi #4 | 2024-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| Dupont US | 2024-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| Hauhong Wuxi #4 | 2024-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Dupont US | 2024-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.68
| Hauhong Wuxi #4 | 2024-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| Dupont US | 2024-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.57
| Hauhong Wuxi #4 | 2024-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.48
| Dupont US | 2024-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| Hauhong Wuxi #4 | 2024-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| App Team Demo | 2024-08-26 | Z Scanner position at -3μm | 0.004031
| App Team Demo | 2024-08-26 | Z Scanner position at +3μm | 0.009694
| App Team Demo | 2024-08-26 | Z linearity by sample height | 0.99886
| App Team Demo | 2024-08-26 | Z Height noise | 0.029
| App Team Demo | 2024-08-26 | Y Scan Stroke | 55
| App Team Demo | 2024-08-26 | Y Detector noise (XY Servo on) | 0.397
| App Team Demo | 2024-08-26 | Y Detector noise (XY Servo off) | 0.423
| App Team Demo | 2024-08-26 | Y axis OPM | 0.7574
| App Team Demo | 2024-08-26 | XY orthogonality | 89.74
| App Team Demo | 2024-08-26 | X Scan Stroke | 55
| App Team Demo | 2024-08-26 | X Detector noise (XY Servo on) | 0.42
| App Team Demo | 2024-08-26 | X Detector noise (XY Servo off) | 0.438
| App Team Demo | 2024-08-26 | X axis OPM | 1.3486
| App Team Demo | 2024-08-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| App Team Demo | 2024-08-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| App Team Demo | 2024-08-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| App Team Demo | 2024-08-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| App Team Demo | 2024-08-26 | NCM Amplitude noise | 14.245
| App Team Demo | 2024-08-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.381
| App Team Demo | 2024-08-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.722
| App Team Demo | 2024-08-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.7
| App Team Demo | 2024-08-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.5
| App Team Demo | 2024-08-26 | Lateral(C-D) sensitivity | 21.3
| App Team Demo | 2024-08-26 | Lateral(C-D) noise | 0.045
| App Team Demo | 2024-08-26 | Full range of Z Height during Z Scan sweep | 18.2
| App Team Demo | 2024-08-26 | Full range of Z Detector during Z Scan sweep | 13.7
| App Team Demo | 2024-08-26 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| App Team Demo | 2024-08-26 | Full range of Y Detector during Y Scan sweep | 65.9
| App Team Demo | 2024-08-26 | Full range of X Detector Raw during X Scan sweep | 13.7
| App Team Demo | 2024-08-26 | Full range of X Detector during X Scan sweep | 69.7
| App Team Demo | 2024-08-26 | A-B noise | 7.144
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Z Scanner position at -3μm | 0.0065
| 서강대학교 물리학과 | 2024-08-23 | Z Scanner position at -3μm | 0.0032
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Z Scanner position at +3μm | 0.0021
| 서강대학교 물리학과 | 2024-08-23 | Z Scanner position at +3μm | 0.0044
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Z linearity by sample height | 1.008
| 서강대학교 물리학과 | 2024-08-23 | Z linearity by sample height | 1.018
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Z Height noise | 0.018
| 서강대학교 물리학과 | 2024-08-23 | Z Height noise | 0.029
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Y Scan Stroke | 110.3
| 서강대학교 물리학과 | 2024-08-23 | Y Scan Stroke | 55
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Y Detector noise (XY Servo on) | 0.26
| 서강대학교 물리학과 | 2024-08-23 | Y Detector noise (XY Servo on) | 0.38
| 서강대학교 물리학과 | 2024-08-23 | Y Detector noise (XY Servo off) | 0.42
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Y Detector noise (XY Servo off) | 0.25
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Y axis OPM | 0.94
| 서강대학교 물리학과 | 2024-08-23 | Y axis OPM | 0.29
| University of Arkansas (Hugh Churchill) | 2024-08-23 | XY orthogonality | 90.15
| University of Arkansas (Hugh Churchill) | 2024-08-23 | X Scan Stroke | 109.4
| 서강대학교 물리학과 | 2024-08-23 | X Scan Stroke | 55
| University of Arkansas (Hugh Churchill) | 2024-08-23 | X Detector noise (XY Servo on) | 0.26
| 서강대학교 물리학과 | 2024-08-23 | X Detector noise (XY Servo on) | 0.38
| 서강대학교 물리학과 | 2024-08-23 | X Detector noise (XY Servo off) | 0.42
| University of Arkansas (Hugh Churchill) | 2024-08-23 | X Detector noise (XY Servo off) | 0.25
| University of Arkansas (Hugh Churchill) | 2024-08-23 | X axis OPM | 1.09
| 서강대학교 물리학과 | 2024-08-23 | X axis OPM | 1.05
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.885
| 서강대학교_물리학과 | 2024-08-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.012
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 20.928
| 서강대학교_물리학과 | 2024-08-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.756
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 45.4
| 서강대학교_물리학과 | 2024-08-23 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.108
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.23
| 서강대학교_물리학과 | 2024-08-23 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.206
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.673
| 서강대학교_물리학과 | 2024-08-23 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.542
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 11.076
| 서강대학교_물리학과 | 2024-08-23 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.338
| University of Arkansas (Hugh Churchill) | 2024-08-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 서강대학교 물리학과 | 2024-08-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| University of Arkansas (Hugh Churchill) | 2024-08-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 서강대학교 물리학과 | 2024-08-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| University of Arkansas (Hugh Churchill) | 2024-08-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 서강대학교 물리학과 | 2024-08-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| University of Arkansas (Hugh Churchill) | 2024-08-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 서강대학교 물리학과 | 2024-08-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| University of Arkansas (Hugh Churchill) | 2024-08-23 | NCM driving% with AC160TS cantilever | 4.26
| 서강대학교 물리학과 | 2024-08-23 | NCM driving% with AC160TS cantilever | 0.771
| University of Arkansas (Hugh Churchill) | 2024-08-23 | NCM Amplitude noise | 11.86
| University of Arkansas (Hugh Churchill) | 2024-08-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.727
| 서강대학교 물리학과 | 2024-08-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.386
| University of Arkansas (Hugh Churchill) | 2024-08-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.57
| 서강대학교 물리학과 | 2024-08-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.116
| University of Arkansas (Hugh Churchill) | 2024-08-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 151.6
| 서강대학교 물리학과 | 2024-08-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.2
| University of Arkansas (Hugh Churchill) | 2024-08-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.6
| 서강대학교 물리학과 | 2024-08-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Lateral(C-D) sensitivity | 11.2
| 서강대학교 물리학과 | 2024-08-23 | Lateral(C-D) sensitivity | 19.2
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Lateral(C-D) noise | 0.102
| 서강대학교 물리학과 | 2024-08-23 | Lateral(C-D) noise | 0.037
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Full range of Z Height during Z Scan sweep | 16.8
| 서강대학교 물리학과 | 2024-08-23 | Full range of Z Height during Z Scan sweep | 18.2
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Full range of Z Detector during Z Scan sweep | 14.2
| 서강대학교 물리학과 | 2024-08-23 | Full range of Z Detector during Z Scan sweep | 12.9
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 서강대학교 물리학과 | 2024-08-23 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Full range of Y Detector during Y Scan sweep | 110.3
| 서강대학교 물리학과 | 2024-08-23 | Full range of Y Detector during Y Scan sweep | 69.1
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Full range of X Detector Raw during X Scan sweep | 14.2
| 서강대학교 물리학과 | 2024-08-23 | Full range of X Detector Raw during X Scan sweep | 13.6
| University of Arkansas (Hugh Churchill) | 2024-08-23 | Full range of X Detector during X Scan sweep | 109.5
| 서강대학교 물리학과 | 2024-08-23 | Full range of X Detector during X Scan sweep | 67.2
| University of Arkansas (Hugh Churchill) | 2024-08-23 | A-B noise | 18.5
| University of Nevada(Ben King) | 2024-08-21 | Z Scanner position at -3μm | 0.0075
| University of Nevada(Ben King) | 2024-08-21 | Z Scanner position at +3μm | 0.007
| University of Nevada(Ben King) | 2024-08-21 | Z linearity by sample height | 1.008
| University of Nevada(Ben King) | 2024-08-21 | Z Height noise | 0.027
| University of Nevada(Ben King) | 2024-08-21 | Y Scan Stroke | 55
| University of Nevada(Ben King) | 2024-08-21 | Y Detector noise (XY Servo on) | 0.43
| University of Nevada(Ben King) | 2024-08-21 | Y Detector noise (XY Servo off) | 0.45
| University of Nevada(Ben King) | 2024-08-21 | Y axis OPM | 0.83
| University of Nevada(Ben King) | 2024-08-21 | X Scan Stroke | 55
| University of Nevada(Ben King) | 2024-08-21 | X Detector noise (XY Servo on) | 0.45
| University of Nevada(Ben King) | 2024-08-21 | X Detector noise (XY Servo off) | 0.46
| University of Nevada(Ben King) | 2024-08-21 | X axis OPM | 1.03
| University of Nevada(Ben King) | 2024-08-21 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.046
| University of Nevada(Ben King) | 2024-08-21 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 1.612
| University of Nevada(Ben King) | 2024-08-21 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 9.772
| University of Nevada(Ben King) | 2024-08-21 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.172
| University of Nevada(Ben King) | 2024-08-21 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.21
| University of Nevada(Ben King) | 2024-08-21 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.827
| University of Nevada(Ben King) | 2024-08-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| University of Nevada(Ben King) | 2024-08-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| University of Nevada(Ben King) | 2024-08-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| University of Nevada(Ben King) | 2024-08-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| University of Nevada(Ben King) | 2024-08-21 | NCM driving% with AC160TS cantilever | 2.133
| University of Nevada(Ben King) | 2024-08-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.556
| University of Nevada(Ben King) | 2024-08-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.198
| University of Nevada(Ben King) | 2024-08-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 89.9
| University of Nevada(Ben King) | 2024-08-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 92.8
| University of Nevada(Ben King) | 2024-08-21 | Lateral(C-D) sensitivity | 16.2
| University of Nevada(Ben King) | 2024-08-21 | Lateral(C-D) noise | 0.04
| University of Nevada(Ben King) | 2024-08-21 | Full range of Z Height during Z Scan sweep | 16.8
| University of Nevada(Ben King) | 2024-08-21 | Full range of Z Detector during Z Scan sweep | 12.8
| University of Nevada(Ben King) | 2024-08-21 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| University of Nevada(Ben King) | 2024-08-21 | Full range of Y Detector during Y Scan sweep | 66.1
| University of Nevada(Ben King) | 2024-08-21 | Full range of X Detector Raw during X Scan sweep | 14.9
| University of Nevada(Ben King) | 2024-08-21 | Full range of X Detector during X Scan sweep | 69.2
| Shandong University | 2024-08-20 | Z Scanner position at -3μm | 0.004927
| 한양대학교 화학과 | 2024-08-20 | Z Scanner position at -3μm | 0.004071
| Alba Iulia University, Romania | 2024-08-20 | Z Scanner position at -3μm | 0.000565
| Shandong University | 2024-08-20 | Z Scanner position at +3μm | 0.001702
| 한양대학교 화학과 | 2024-08-20 | Z Scanner position at +3μm | 0.01708
| Alba Iulia University, Romania | 2024-08-20 | Z Scanner position at +3μm | 0.005447
| Alba Iulia University, Romania | 2024-08-20 | Z linearity by sample height | 1.0152
| Shandong University | 2024-08-20 | Z Linearity by 3 different sample height | 1.0123
| 한양대학교 화학과 | 2024-08-20 | Z Linearity by 3 different sample height | 1.0041
| Shandong University | 2024-08-20 | Z Height noise | 0.032
| 한양대학교 화학과 | 2024-08-20 | Z Height noise | 0.033
| Alba Iulia University, Romania | 2024-08-20 | Z Height noise | 0.029
| Alba Iulia University, Romania | 2024-08-20 | Y2 Detector noise (XY Servo on) | 0.231
| Alba Iulia University, Romania | 2024-08-20 | Y2 Detector noise (XY Servo off) | 0.225
| Alba Iulia University, Romania | 2024-08-20 | Y1 Detector noise (XY Servo on) | 0.22
| Alba Iulia University, Romania | 2024-08-20 | Y1 Detector noise (XY Servo off) | 0.22
| Shandong University | 2024-08-20 | Y Scan Stroke | 110.7
| 한양대학교 화학과 | 2024-08-20 | Y Scan Stroke | 55
| Alba Iulia University, Romania | 2024-08-20 | Y Scan Stroke | 177.6
| Shandong University | 2024-08-20 | Y Detector noise (XY Servo on) | 0.261
| 한양대학교 화학과 | 2024-08-20 | Y Detector noise (XY Servo on) | 0.274
| Alba Iulia University, Romania | 2024-08-20 | Y Detector noise (XY Servo on) | 0.167
| Alba Iulia University, Romania | 2024-08-20 | Y Detector noise (XY Servo off) | 0.163
| Shandong University | 2024-08-20 | Y Detector noise (XY Servo off) | 0.257
| 한양대학교 화학과 | 2024-08-20 | Y Detector noise (XY Servo off) | 0.267
| Shandong University | 2024-08-20 | Y axis OPM | 1.2244
| 한양대학교 화학과 | 2024-08-20 | Y axis OPM | 0.6098
| Alba Iulia University, Romania | 2024-08-20 | Y axis OPM | 1.268
| Alba Iulia University, Romania | 2024-08-20 | X2 Detector noise (XY Servo on) | 0.242
| Alba Iulia University, Romania | 2024-08-20 | X2 Detector noise (XY Servo off) | 0.246
| Alba Iulia University, Romania | 2024-08-20 | X1 Detector noise (XY Servo on) | 0.229
| Alba Iulia University, Romania | 2024-08-20 | X1 Detector noise (XY Servo off) | 0.226
| Shandong University | 2024-08-20 | X Scan Stroke | 109.6
| 한양대학교 화학과 | 2024-08-20 | X Scan Stroke | 55
| Alba Iulia University, Romania | 2024-08-20 | X Scan Stroke | 174.1
| Shandong University | 2024-08-20 | X Detector noise (XY Servo on) | 0.257
| 한양대학교 화학과 | 2024-08-20 | X Detector noise (XY Servo on) | 0.271
| Alba Iulia University, Romania | 2024-08-20 | X Detector noise (XY Servo on) | 0.178
| Alba Iulia University, Romania | 2024-08-20 | X Detector noise (XY Servo off) | 0.175
| Shandong University | 2024-08-20 | X Detector noise (XY Servo off) | 0.25
| 한양대학교 화학과 | 2024-08-20 | X Detector noise (XY Servo off) | 0.261
| Shandong University | 2024-08-20 | X axis OPM | 1.1914
| 한양대학교 화학과 | 2024-08-20 | X axis OPM | 0.7742
| Alba Iulia University, Romania | 2024-08-20 | X axis OPM | 1.193
| Shandong University | 2024-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 한양대학교 화학과 | 2024-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| Shandong University | 2024-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| 한양대학교 화학과 | 2024-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.41
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.39
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.4
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.38
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.39
| Shandong University | 2024-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 한양대학교 화학과 | 2024-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Shandong University | 2024-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 한양대학교 화학과 | 2024-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Alba Iulia University, Romania | 2024-08-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Shandong University | 2024-08-20 | NCM driving% with AC160TS cantilever | 0.362
| 한양대학교 화학과 | 2024-08-20 | NCM driving% with AC160TS cantilever | 0.309
| Shandong University | 2024-08-20 | NCM Amplitude noise | 13.426
| 한양대학교 화학과 | 2024-08-20 | NCM Amplitude noise | 15.074
| Alba Iulia University, Romania | 2024-08-20 | NCM Amplitude noise | 17.616
| Shandong University | 2024-08-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.004
| 한양대학교 화학과 | 2024-08-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.142
| Alba Iulia University, Romania | 2024-08-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.44
| Shandong University | 2024-08-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.77
| 한양대학교 화학과 | 2024-08-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.17
| Alba Iulia University, Romania | 2024-08-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.827
| Alba Iulia University, Romania | 2024-08-20 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 248.1
| Alba Iulia University, Romania | 2024-08-20 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 241.8
| Shandong University | 2024-08-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 161.5
| 한양대학교 화학과 | 2024-08-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 112.9
| Alba Iulia University, Romania | 2024-08-20 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 249.4
| Alba Iulia University, Romania | 2024-08-20 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 241
| Shandong University | 2024-08-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 158
| 한양대학교 화학과 | 2024-08-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 111.8
| Alba Iulia University, Romania | 2024-08-20 | Maximum Y Scan range (Full Stroke) | 178.1
| Alba Iulia University, Romania | 2024-08-20 | Maximum X Scan range (Full Stroke) | 178
| Alba Iulia University, Romania | 2024-08-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Shandong University | 2024-08-20 | Lateral(C-D) sensitivity | 18.5
| 한양대학교 화학과 | 2024-08-20 | Lateral(C-D) sensitivity | 19.7
| Alba Iulia University, Romania | 2024-08-20 | Lateral(C-D) sensitivity | 17
| Shandong University | 2024-08-20 | Lateral(C-D) noise | 0.095
| 한양대학교 화학과 | 2024-08-20 | Lateral(C-D) noise | 0.06
| Alba Iulia University, Romania | 2024-08-20 | Lateral(C-D) noise | 0.142
| Shandong University | 2024-08-20 | Full range of Z Height during Z Scan sweep | 17.8
| 한양대학교 화학과 | 2024-08-20 | Full range of Z Height during Z Scan sweep | 17.5
| Alba Iulia University, Romania | 2024-08-20 | Full range of Z Height during Z Scan sweep | 18.6
| Shandong University | 2024-08-20 | Full range of Z Detector during Z Scan sweep | 12.8
| 한양대학교 화학과 | 2024-08-20 | Full range of Z Detector during Z Scan sweep | 12.9
| Alba Iulia University, Romania | 2024-08-20 | Full range of Z Detector during Z Scan sweep | 13.3
| Alba Iulia University, Romania | 2024-08-20 | Full range of Y2 Detector raw | 14.5
| Alba Iulia University, Romania | 2024-08-20 | Full range of Y1 Detector raw | 14.4
| Shandong University | 2024-08-20 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| 한양대학교 화학과 | 2024-08-20 | Full range of Y Detector Raw during Y Scan sweep | 13.1
| Shandong University | 2024-08-20 | Full range of Y Detector during Y Scan sweep | 110.7
| 한양대학교 화학과 | 2024-08-20 | Full range of Y Detector during Y Scan sweep | 74.2
| Alba Iulia University, Romania | 2024-08-20 | Full range of X2 Detector raw | 14.4
| Alba Iulia University, Romania | 2024-08-20 | Full range of X1 Detector raw | 14.7
| Shandong University | 2024-08-20 | Full range of X Detector Raw during X Scan sweep | 13.9
| 한양대학교 화학과 | 2024-08-20 | Full range of X Detector Raw during X Scan sweep | 13.1
| Shandong University | 2024-08-20 | Full range of X Detector during X Scan sweep | 109.6
| 한양대학교 화학과 | 2024-08-20 | Full range of X Detector during X Scan sweep | 73.1
| Shandong University | 2024-08-20 | A-B noise | 11.187
| 한양대학교 화학과 | 2024-08-20 | A-B noise | 11.306
| Alba Iulia University, Romania | 2024-08-20 | A-B noise | 7.106
| 서울대학교 화학생물공학부 | 2024-08-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.064
| 서울대학교 화학생물공학부 | 2024-08-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 3.089
| 서울대학교 화학생물공학부 | 2024-08-16 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 6.797
| 서울대학교 화학생물공학부 | 2024-08-16 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.203
| 서울대학교 화학생물공학부 | 2024-08-16 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.224
| 서울대학교 화학생물공학부 | 2024-08-16 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.57
| Tokyo Metropolitan University | 2024-08-14 | Z Scanner position at -3μm | 0.0169
| Tokyo Metropolitan University | 2024-08-14 | Z Scanner position at +3μm | 0.0151
| Tokyo Metropolitan University | 2024-08-14 | Z linearity by sample height | 1.019
| Tokyo Metropolitan University | 2024-08-14 | Z Height noise | 0.03
| Tokyo Metropolitan University | 2024-08-14 | Y Scan Stroke | 55
| Tokyo Metropolitan University | 2024-08-14 | Y Detector noise (XY Servo on) | 0.41
| Tokyo Metropolitan University | 2024-08-14 | Y Detector noise (XY Servo off) | 0.41
| Tokyo Metropolitan University | 2024-08-14 | Y axis OPM | 0.49
| Tokyo Metropolitan University | 2024-08-14 | X Scan Stroke | 55
| Tokyo Metropolitan University | 2024-08-14 | X Detector noise (XY Servo on) | 0.43
| Tokyo Metropolitan University | 2024-08-14 | X Detector noise (XY Servo off) | 0.43
| Tokyo Metropolitan University | 2024-08-14 | X axis OPM | 0.86
| Tokyo Metropolitan University | 2024-08-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.284
| Tokyo Metropolitan University | 2024-08-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 5.186
| Tokyo Metropolitan University | 2024-08-14 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.677
| Tokyo Metropolitan University | 2024-08-14 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.232
| Tokyo Metropolitan University | 2024-08-14 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.208
| Tokyo Metropolitan University | 2024-08-14 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 9.698
| Tokyo Metropolitan University | 2024-08-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Tokyo Metropolitan University | 2024-08-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Tokyo Metropolitan University | 2024-08-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| Tokyo Metropolitan University | 2024-08-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Tokyo Metropolitan University | 2024-08-14 | NCM driving% with AC160TS cantilever | 1.7
| Tokyo Metropolitan University | 2024-08-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.804
| Tokyo Metropolitan University | 2024-08-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.067
| Tokyo Metropolitan University | 2024-08-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 90.6
| Tokyo Metropolitan University | 2024-08-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.6
| Tokyo Metropolitan University | 2024-08-14 | Lateral(C-D) sensitivity | 16.8
| Tokyo Metropolitan University | 2024-08-14 | Lateral(C-D) noise | 0.057
| Tokyo Metropolitan University | 2024-08-14 | Full range of Z Height during Z Scan sweep | 16.9
| Tokyo Metropolitan University | 2024-08-14 | Full range of Z Detector during Z Scan sweep | 12.1
| Tokyo Metropolitan University | 2024-08-14 | Full range of Y Detector Raw during Y Scan sweep | 15.8
| Tokyo Metropolitan University | 2024-08-14 | Full range of Y Detector during Y Scan sweep | 71.6
| Tokyo Metropolitan University | 2024-08-14 | Full range of X Detector Raw during X Scan sweep | 14.8
| Tokyo Metropolitan University | 2024-08-14 | Full range of X Detector during X Scan sweep | 72.2
| Global Wafers America | 2024-08-12 | Z Linearity by Z scanner position | 0.88
| Global Wafers America | 2024-08-12 | Z Linearity by 4 different sample height | 0.999
| Global Wafers America | 2024-08-12 | Z Height noise | 0.03
| Global Wafers America | 2024-08-12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7
| Global Wafers America | 2024-08-12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7
| Global Wafers America | 2024-08-12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| Global Wafers America | 2024-08-12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.7
| Global Wafers America | 2024-08-12 | Z Drift | 5.16
| Global Wafers America | 2024-08-12 | Z Detector offset | 114.4
| Global Wafers America | 2024-08-12 | Y2 Detector noise (XY Servo on) | 0.36
| Global Wafers America | 2024-08-12 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 247.8
| Global Wafers America | 2024-08-12 | Y1 Detector noise (XY Servo on) | 0.28
| Global Wafers America | 2024-08-12 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 251.6
| Global Wafers America | 2024-08-12 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 174
| Global Wafers America | 2024-08-12 | Y Ringing(Rpv) | 0.24
| Global Wafers America | 2024-08-12 | Y Detector noise (XY Servo on) | 0.24
| Global Wafers America | 2024-08-12 | Y axis, Repeatability @1st position | 0.05
| Global Wafers America | 2024-08-12 | Y axis, Repeatability @ 4th position | 0.06
| Global Wafers America | 2024-08-12 | Y axis, Repeatability @ 3rd position | 0.05
| Global Wafers America | 2024-08-12 | Y axis, Repeatability @ 2nd position | 0.05
| Global Wafers America | 2024-08-12 | Y axis, Accuracy @1st position | 0.29
| Global Wafers America | 2024-08-12 | Y axis, Accuracy @ 4th position | 0.44
| Global Wafers America | 2024-08-12 | Y axis, Accuracy @ 3rd position | 0.35
| Global Wafers America | 2024-08-12 | Y axis, Accuracy @ 2nd position | 0.29
| Global Wafers America | 2024-08-12 | Y axis OPM | 0.43
| Global Wafers America | 2024-08-12 | XY Orthogonality | 90.17
| Global Wafers America | 2024-08-12 | X2 Detector noise (XY Servo on) | 0.29
| Global Wafers America | 2024-08-12 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 235.7
| Global Wafers America | 2024-08-12 | X1 Detector noise (XY Servo on) | 0.28
| Global Wafers America | 2024-08-12 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 241.8
| Global Wafers America | 2024-08-12 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 162.2
| Global Wafers America | 2024-08-12 | X Ringing(Rpv) | 0.08
| Global Wafers America | 2024-08-12 | X Detector noise (XY Servo on) | 0.22
| Global Wafers America | 2024-08-12 | X axis, Repeatability @1st position | 0.03
| Global Wafers America | 2024-08-12 | X axis, Repeatability @ 4th position | 0.06
| Global Wafers America | 2024-08-12 | X axis, Repeatability @ 3rd position | 0.04
| Global Wafers America | 2024-08-12 | X axis, Repeatability @ 2nd position | 0.05
| Global Wafers America | 2024-08-12 | X axis, Accuracy @1st position | 0.18
| Global Wafers America | 2024-08-12 | X axis, Accuracy @ 4th position | 0.31
| Global Wafers America | 2024-08-12 | X axis, Accuracy @ 3rd position | 0.2
| Global Wafers America | 2024-08-12 | X axis, Accuracy @ 2nd position | 0.25
| Global Wafers America | 2024-08-12 | X axis OPM | 0.33
| Global Wafers America | 2024-08-12 | Wafer placement(Loading) error on the load port2 (Y offset) | 73.27
| Global Wafers America | 2024-08-12 | Wafer placement(Loading) error on the load port2 (X offset) | 57.7
| Global Wafers America | 2024-08-12 | Wafer placement(Loading) error on the load port1 (Y offset) | 39.06
| Global Wafers America | 2024-08-12 | Wafer placement(Loading) error on the load port1 (X offset) | 65
| Global Wafers America | 2024-08-12 | StrokeYMin | -152602
| Global Wafers America | 2024-08-12 | StrokeYMax | 150485.9
| Global Wafers America | 2024-08-12 | StrokeY | 303087.6
| Global Wafers America | 2024-08-12 | StrokeXMin | -203502
| Global Wafers America | 2024-08-12 | StrokeXMax | 203859.5
| Global Wafers America | 2024-08-12 | StrokeX | 407361.8
| Global Wafers America | 2024-08-12 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.96
| Global Wafers America | 2024-08-12 | Roll motion to X+ full scan | 13.4
| Global Wafers America | 2024-08-12 | RMS of ZeroScan | 0.041
| Global Wafers America | 2024-08-12 | Position repeatability(10repeats, Position Based) | 0.92
| Global Wafers America | 2024-08-12 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8e-05
| Global Wafers America | 2024-08-12 | OpticalAxisShiftPerFocusY | 0.0019
| Global Wafers America | 2024-08-12 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| Global Wafers America | 2024-08-12 | OpticalAxisShiftPerFocusX | -0.0026
| Global Wafers America | 2024-08-12 | OPM Repeatability @5mm | 0.383
| Global Wafers America | 2024-08-12 | OPM Repeatability @25mm | 2.998
| Global Wafers America | 2024-08-12 | OPM Repeatability @1mm | 0.503
| Global Wafers America | 2024-08-12 | OPM Repeatability @10mm | 1.272
| Global Wafers America | 2024-08-12 | NCM driving% with AC160TS cantilever | 10.18
| Global Wafers America | 2024-08-12 | Min. Vibration Transmissbility (~100Hz) | 0.035
| Global Wafers America | 2024-08-12 | Max OPM @5mm | 15.589
| Global Wafers America | 2024-08-12 | Max OPM @25mm | 52.497
| Global Wafers America | 2024-08-12 | Max OPM @1mm | 7.667
| Global Wafers America | 2024-08-12 | Max OPM @10mm | 16.334
| Global Wafers America | 2024-08-12 | LowerLimit | 1400
| Global Wafers America | 2024-08-12 | LowerLimit | 5497
| Global Wafers America | 2024-08-12 | Lateral(C-D) sensitivity | 11.9
| Global Wafers America | 2024-08-12 | Lateral(C-D) noise | 0.239
| Global Wafers America | 2024-08-12 | Image offset after moving from -Y to +Y (Y scan) | 74.71
| Global Wafers America | 2024-08-12 | Image offset after moving from -Y to +Y (X scan) | 37.35
| Global Wafers America | 2024-08-12 | Image offset after moving from -X to +X (Y scan) | 31.49
| Global Wafers America | 2024-08-12 | Image offset after moving from -X to +X (X scan) | 64.45
| Global Wafers America | 2024-08-12 | Image offset after moving from +Y to -Y (Y scan) | 20.51
| Global Wafers America | 2024-08-12 | Image offset after moving from +Y to -Y (X scan) | 11.72
| Global Wafers America | 2024-08-12 | Image offset after moving from +X to -X (Y scan) | 44.68
| Global Wafers America | 2024-08-12 | Image offset after moving from +X to -X (X scan) | 101.07
| Global Wafers America | 2024-08-12 | Full range of Z Height during Z Scan sweep | 17
| Global Wafers America | 2024-08-12 | Full range of Z Detector during Z Scan sweep | 13.02
| Global Wafers America | 2024-08-12 | Full range of Y2 Detector Raw during Y Scan sweep | 14.33
| Global Wafers America | 2024-08-12 | Full range of Y2 Detector during Y Scan sweep | 196.7
| Global Wafers America | 2024-08-12 | Full range of Y1 Detector Raw during Y Scan sweep | 14.25
| Global Wafers America | 2024-08-12 | Full range of Y1 Detector during Y Scan sweep | 195.9
| Global Wafers America | 2024-08-12 | Full range of Y Detector during Y Scan sweep | 196.7
| Global Wafers America | 2024-08-12 | Full range of X2 Detector Raw during X Scan sweep | 14.19
| Global Wafers America | 2024-08-12 | Full range of X2 Detector during X Scan sweep | 188
| Global Wafers America | 2024-08-12 | Full range of X1 Detector Raw during X Scan sweep | 14.11
| Global Wafers America | 2024-08-12 | Full range of X1 Detector during X Scan sweep | 194.9
| Global Wafers America | 2024-08-12 | Full range of X Detector during X Scan sweep | 188
| Global Wafers America | 2024-08-12 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Global Wafers America | 2024-08-12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Global Wafers America | 2024-08-12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| Global Wafers America | 2024-08-12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.48
| Global Wafers America | 2024-08-12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Global Wafers America | 2024-08-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Global Wafers America | 2024-08-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.19
| Global Wafers America | 2024-08-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.52
| Global Wafers America | 2024-08-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 |  | 103495
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 |  | -101836
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 |  | 101687
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 |  | -102740
| 서울대학교 화학생물공학부 | 2024-08-09 | Z Scanner position at -3μm | 0.010073
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Z Scanner position at -3μm | 0.0005
| 서울대학교 화학생물공학부 | 2024-08-09 | Z Scanner position at +3μm | 0.011049
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Z Scanner position at +3μm | 0.0007
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Z linearity by sample height | 1.005
| 서울대학교 화학생물공학부 | 2024-08-09 | Z Linearity by 3 different sample height | 1.0084
| 서울대학교 화학생물공학부 | 2024-08-09 | Z Height noise | 0.032
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Z Height noise | 0.027
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y2 Detector noise (XY Servo on) | 0.28
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y2 Detector noise (XY Servo off) | 0.25
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y1 Detector noise (XY Servo on) | 0.26
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y1 Detector noise (XY Servo off) | 0.25
| 서울대학교 화학생물공학부 | 2024-08-09 | Y Scan Stroke | 55
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y Scan Stroke | 185.9
| 서울대학교 화학생물공학부 | 2024-08-09 | Y Detector noise (XY Servo on) | 0.216
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y Detector noise (XY Servo on) | 0.19
| 서울대학교 화학생물공학부 | 2024-08-09 | Y Detector noise (XY Servo off) | 0.208
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y Detector noise (XY Servo off) | 0.18
| 서울대학교 화학생물공학부 | 2024-08-09 | Y axis OPM | 0.548
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Y axis OPM | 0.77
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X2 Detector noise (XY Servo on) | 0.31
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X2 Detector noise (XY Servo off) | 0.28
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X1 Detector noise (XY Servo on) | 0.27
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X1 Detector noise (XY Servo off) | 0.26
| 서울대학교 화학생물공학부 | 2024-08-09 | X Scan Stroke | 55
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X Scan Stroke | 180.1
| 서울대학교 화학생물공학부 | 2024-08-09 | X Detector noise (XY Servo on) | 0.228
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X Detector noise (XY Servo on) | 0.23
| 서울대학교 화학생물공학부 | 2024-08-09 | X Detector noise (XY Servo off) | 0.221
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X Detector noise (XY Servo off) | 0.21
| 서울대학교 화학생물공학부 | 2024-08-09 | X axis OPM | 1.136
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | X axis OPM | 0.52
| 서울대학교 화학생물공학부 | 2024-08-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| 서울대학교 화학생물공학부 | 2024-08-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.36
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.32
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| 서울대학교 화학생물공학부 | 2024-08-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| 서울대학교 화학생물공학부 | 2024-08-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| 서울대학교 화학생물공학부 | 2024-08-09 | NCM driving% with AC160TS cantilever | 0.666
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | NCM driving% with AC160TS cantilever | 0.68
| 서울대학교 화학생물공학부 | 2024-08-09 | NCM Amplitude noise | 15.513
| 서울대학교 화학생물공학부 | 2024-08-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.001
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.323
| 서울대학교 화학생물공학부 | 2024-08-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.932
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.739
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 268.5
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 260.9
| 서울대학교 화학생물공학부 | 2024-08-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.7
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 250
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 257.4
| 서울대학교 화학생물공학부 | 2024-08-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 108.6
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| 서울대학교 화학생물공학부 | 2024-08-09 | Lateral(C-D) sensitivity | 14.3
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Lateral(C-D) sensitivity | 4.2
| 서울대학교 화학생물공학부 | 2024-08-09 | Lateral(C-D) noise | 0.074
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Lateral(C-D) noise | 0.068
| 서울대학교 화학생물공학부 | 2024-08-09 | Full range of Z Height during Z Scan sweep | 17.5
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Z Height during Z Scan sweep | 18.2
| 서울대학교 화학생물공학부 | 2024-08-09 | Full range of Z Detector during Z Scan sweep | 13
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Z Detector during Z Scan sweep | 14.1
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Y2 Detector Raw during Y Scan sweep | 13.9
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Y2 Detector during Y Scan sweep | 186.3
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Y1 Detector during Y Scan sweep | 185.1
| 서울대학교 화학생물공학부 | 2024-08-09 | Full range of Y Detector Raw during Y Scan sweep | 14
| 서울대학교 화학생물공학부 | 2024-08-09 | Full range of Y Detector during Y Scan sweep | 71.1
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of Y Detector during Y Scan sweep | 185.7
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of X2 Detector during X Scan sweep | 179.3
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of X1 Detector during X Scan sweep | 180.2
| 서울대학교 화학생물공학부 | 2024-08-09 | Full range of X Detector Raw during X Scan sweep | 13.8
| 서울대학교 화학생물공학부 | 2024-08-09 | Full range of X Detector during X Scan sweep | 75.1
| Suzhou Xiangcheng Industrial Technology Research Institute | 2024-08-09 | Full range of X Detector during X Scan sweep | 180.6
| 서울대학교 화학생물공학부 | 2024-08-09 | A-B noise | 6.487
| University of Toronto(De La Barrera) | 2024-08-08 | Z Scanner position at -6μm | 0.0007
| University of Toronto(De La Barrera) | 2024-08-08 | Z Scanner position at -3μm | 0.0127
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Z Scanner position at -3μm | 0.0006
| University of Limerick (PSE) | 2024-08-08 | Z Scanner position at -3μm | 0.0005
| University of Toronto(De La Barrera) | 2024-08-08 | Z Scanner position at +6μm | 0.0031
| University of Toronto(De La Barrera) | 2024-08-08 | Z Scanner position at +3μm | 0.0156
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Z Scanner position at +3μm | 0.0047
| University of Limerick (PSE) | 2024-08-08 | Z Scanner position at +3μm | 0
| University of Toronto(De La Barrera) | 2024-08-08 | Z linearity by sample height | 1.011
| University of Toronto(De La Barrera) | 2024-08-08 | Z linearity by sample height | 1.014
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Z linearity by sample height | 0.997
| University of Limerick (PSE) | 2024-08-08 | Z linearity by sample height | 1.003
| University of Toronto(De La Barrera) | 2024-08-08 | Z Height noise | 0.028
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Z Height noise | 0.027
| University of Limerick (PSE) | 2024-08-08 | Z Height noise | 0.025
| University of Toronto(De La Barrera) | 2024-08-08 | Y Scan Stroke | 125.7
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Y Scan Stroke | 55
| University of Limerick (PSE) | 2024-08-08 | Y Scan Stroke | 55
| University of Toronto(De La Barrera) | 2024-08-08 | Y Detector noise (XY Servo on) | 0.38
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Y Detector noise (XY Servo on) | 0.33
| University of Limerick (PSE) | 2024-08-08 | Y Detector noise (XY Servo on) | 0.25
| University of Toronto(De La Barrera) | 2024-08-08 | Y Detector noise (XY Servo off) | 0.4
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Y Detector noise (XY Servo off) | 0.37
| University of Limerick (PSE) | 2024-08-08 | Y Detector noise (XY Servo off) | 0.25
| University of Toronto(De La Barrera) | 2024-08-08 | Y axis OPM | 1.68
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Y axis OPM | 1.49
| University of Limerick (PSE) | 2024-08-08 | Y axis OPM | 0.49
| University of Toronto(De La Barrera) | 2024-08-08 | X Scan Stroke | 132.1
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | X Scan Stroke | 55
| University of Limerick (PSE) | 2024-08-08 | X Scan Stroke | 55
| University of Toronto(De La Barrera) | 2024-08-08 | X Detector noise (XY Servo on) | 0.37
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | X Detector noise (XY Servo on) | 0.34
| University of Limerick (PSE) | 2024-08-08 | X Detector noise (XY Servo on) | 0.25
| University of Toronto(De La Barrera) | 2024-08-08 | X Detector noise (XY Servo off) | 0.39
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | X Detector noise (XY Servo off) | 0.37
| University of Limerick (PSE) | 2024-08-08 | X Detector noise (XY Servo off) | 0.24
| University of Toronto(De La Barrera) | 2024-08-08 | X axis OPM | 1.53
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | X axis OPM | 1.62
| University of Limerick (PSE) | 2024-08-08 | X axis OPM | 0.84
| 목포해양대학교 | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.165
| University of Toronto(De La Barrera) | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.087
| University of Limerick (PSE) | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.222
| 목포해양대학교 | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 4.819
| University of Toronto(De La Barrera) | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.615
| University of Limerick (PSE) | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 0.988
| 목포해양대학교 | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 10.895
| University of Toronto(De La Barrera) | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 8.772
| University of Limerick (PSE) | 2024-08-08 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 13.601
| 목포해양대학교 | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.196
| University of Toronto(De La Barrera) | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.152
| University of Limerick (PSE) | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.104
| 목포해양대학교 | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.494
| University of Toronto(De La Barrera) | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.173
| University of Limerick (PSE) | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.204
| 목포해양대학교 | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.051
| University of Toronto(De La Barrera) | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.746
| University of Limerick (PSE) | 2024-08-08 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.946
| University of Toronto(De La Barrera) | 2024-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| University of Limerick (PSE) | 2024-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Toronto(De La Barrera) | 2024-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| University of Limerick (PSE) | 2024-08-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| University of Toronto(De La Barrera) | 2024-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| University of Limerick (PSE) | 2024-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of Toronto(De La Barrera) | 2024-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| University of Limerick (PSE) | 2024-08-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| University of Toronto(De La Barrera) | 2024-08-08 | NCM driving% with AC160TS cantilever | 0.862
| University of Toronto(De La Barrera) | 2024-08-08 | NCM driving% with AC160TS cantilever | 1.15
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | NCM driving% with AC160TS cantilever | 0.529
| University of Limerick (PSE) | 2024-08-08 | NCM driving% with AC160TS cantilever | 2.699
| University of Toronto(De La Barrera) | 2024-08-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.632
| University of Toronto(De La Barrera) | 2024-08-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 26.06
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.988
| University of Limerick (PSE) | 2024-08-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.165
| University of Toronto(De La Barrera) | 2024-08-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.182
| University of Toronto(De La Barrera) | 2024-08-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 56.916
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.201
| University of Limerick (PSE) | 2024-08-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.193
| University of Toronto(De La Barrera) | 2024-08-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 175.3
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 95.3
| University of Limerick (PSE) | 2024-08-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 110.2
| University of Toronto(De La Barrera) | 2024-08-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 173.1
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.6
| University of Limerick (PSE) | 2024-08-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.2
| University of Toronto(De La Barrera) | 2024-08-08 | Lateral(C-D) sensitivity | 11.7
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Lateral(C-D) sensitivity | 23.5
| University of Limerick (PSE) | 2024-08-08 | Lateral(C-D) sensitivity | 12
| University of Toronto(De La Barrera) | 2024-08-08 | Lateral(C-D) noise | 0.075
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Lateral(C-D) noise | 0.052
| University of Limerick (PSE) | 2024-08-08 | Lateral(C-D) noise | 0.065
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of Z Height during Z Scan sweep | 16.1
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of Z Height during Z Scan sweep | 43.4
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Full range of Z Height during Z Scan sweep | 16.1
| University of Limerick (PSE) | 2024-08-08 | Full range of Z Height during Z Scan sweep | 16.2
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of Z Detector during Z Scan sweep | 11.8
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of Z Detector during Z Scan sweep | 15.3
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Full range of Z Detector during Z Scan sweep | 12.3
| University of Limerick (PSE) | 2024-08-08 | Full range of Z Detector during Z Scan sweep | 14
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of Y Detector Raw during Y Scan sweep | 14
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| University of Limerick (PSE) | 2024-08-08 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of Y Detector during Y Scan sweep | 122.5
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Full range of Y Detector during Y Scan sweep | 68.6
| University of Limerick (PSE) | 2024-08-08 | Full range of Y Detector during Y Scan sweep | 75
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of X Detector Raw during X Scan sweep | 15
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Full range of X Detector Raw during X Scan sweep | 13.9
| University of Limerick (PSE) | 2024-08-08 | Full range of X Detector Raw during X Scan sweep | 14.2
| University of Toronto(De La Barrera) | 2024-08-08 | Full range of X Detector during X Scan sweep | 129.2
| Optotransmitter-Umweltzschutz Technology (OUT) | 2024-08-08 | Full range of X Detector during X Scan sweep | 68.1
| University of Limerick (PSE) | 2024-08-08 | Full range of X Detector during X Scan sweep | 75.3
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Z Scanner position at -3μm | 0.0048
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Z Scanner position at +3μm | 0.01
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Z linearity by sample height | 0.992
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Z Height noise | 0.019
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Y Scan Stroke | 113.1
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Y Detector noise (XY Servo on) | 0.25
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Y Detector noise (XY Servo off) | 0.24
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Y axis OPM | 1.13
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | XY orthogonality | 90
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | X Scan Stroke | 112
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | X Detector noise (XY Servo on) | 0.26
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | X Detector noise (XY Servo off) | 0.24
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | X axis OPM | 1.28
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | NCM driving% with AC160TS cantilever | 1.867
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | NCM Amplitude noise | 11.85
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.792
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.974
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160.8
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Lateral(C-D) sensitivity | 9.2
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Lateral(C-D) noise | 0.106
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Full range of Z Height during Z Scan sweep | 17.6
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Full range of Z Detector during Z Scan sweep | 16
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Full range of Y Detector during Y Scan sweep | 113.1
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Full range of X Detector Raw during X Scan sweep | 13.9
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | Full range of X Detector during X Scan sweep | 112
| Indian Institute of Technology Madras, Chennai (IIT Madras) | 2024-08-06 | A-B noise | 16.34
| University of Limerick | 2024-08-05 | Z Scanner position at -3μm | 0.000504
| University of Limerick | 2024-08-05 | Z Scanner position at +3μm | 0
| University of Limerick | 2024-08-05 | Z Linearity by 3 different sample height | 1.0027
| University of Limerick | 2024-08-05 | Z Height noise | 0.025
| University of Limerick | 2024-08-05 | Y Scan Stroke | 55
| University of Limerick | 2024-08-05 | Y Detector noise (XY Servo on) | 0.251
| University of Limerick | 2024-08-05 | Y Detector noise (XY Servo off) | 0.25
| University of Limerick | 2024-08-05 | Y axis OPM | 0.486
| University of Limerick | 2024-08-05 | X Scan Stroke | 55
| University of Limerick | 2024-08-05 | X Detector noise (XY Servo on) | 0.245
| University of Limerick | 2024-08-05 | X Detector noise (XY Servo off) | 0.24
| University of Limerick | 2024-08-05 | X axis OPM | 0.8424
| University of Limerick | 2024-08-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Limerick | 2024-08-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| University of Limerick | 2024-08-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of Limerick | 2024-08-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| University of Limerick | 2024-08-05 | NCM driving% with AC160TS cantilever | 2.699
| University of Limerick | 2024-08-05 | NCM Amplitude noise | 9.841
| University of Limerick | 2024-08-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.165
| University of Limerick | 2024-08-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.193
| University of Limerick | 2024-08-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 110.2
| University of Limerick | 2024-08-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.2
| University of Limerick | 2024-08-05 | Lateral(C-D) sensitivity | 12
| University of Limerick | 2024-08-05 | Lateral(C-D) noise | 0.065
| University of Limerick | 2024-08-05 | Full range of Z Height during Z Scan sweep | 16.2
| University of Limerick | 2024-08-05 | Full range of Z Detector during Z Scan sweep | 14
| University of Limerick | 2024-08-05 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| University of Limerick | 2024-08-05 | Full range of Y Detector during Y Scan sweep | 75
| University of Limerick | 2024-08-05 | Full range of X Detector Raw during X Scan sweep | 14.2
| University of Limerick | 2024-08-05 | Full range of X Detector during X Scan sweep | 75.3
| University of Limerick | 2024-08-05 | A-B noise | 13.337
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Z Scanner position at -3μm | 0.004581
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Z Scanner position at +3μm | 0.011118
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Z linearity by sample height | 1.084
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Z Height noise | 0.029
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y2 Detector noise (XY Servo on) | 0.306
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y2 Detector noise (XY Servo off) | 0.266
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y1 Detector noise (XY Servo on) | 0.302
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y1 Detector noise (XY Servo off) | 0.281
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y Scan Stroke | 169.7
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y Detector noise (XY Servo on) | 0.222
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y Detector noise (XY Servo off) | 0.194
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Y axis OPM | 0.8446
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X2 Detector noise (XY Servo on) | 0.364
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X2 Detector noise (XY Servo off) | 0.34
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X1 Detector noise (XY Servo on) | 0.311
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X1 Detector noise (XY Servo off) | 0.287
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X Scan Stroke | 158.8
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X Detector noise (XY Servo on) | 0.25
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X Detector noise (XY Servo off) | 0.227
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | X axis OPM | 1.3762
| Osaka University | 2024-08-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC | 0.078
| Osaka University | 2024-08-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC | 2.067
| Osaka University | 2024-08-02 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC | 13.098
| Osaka University | 2024-08-02 | Spring constant calibration by Sader Method with 3XC-NA-500DC | 0.181
| Osaka University | 2024-08-02 | Spring constant calibration by Sader Method with 3XC-NA-240AC | 2.316
| Osaka University | 2024-08-02 | Spring constant calibration by Sader Method with 3XC-NA-200AC | 10.032
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.33
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.28
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.3
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | NCM Amplitude noise | 12.374
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.075
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.501
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 230.1
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 233.4
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 216.8
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 224.3
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Maximum Y Scan range (Full Stroke) | 169.7
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Maximum X Scan range (Full Stroke) | 158.7
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Lateral(C-D) sensitivity | 17.7
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Lateral(C-D) noise | 0.073
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Full range of Z Height during Z Scan sweep | 17.4
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Full range of Z Detector during Z Scan sweep | 12.6
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Full range of Y2 Detector raw | 14.6
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Full range of Y1 Detector raw | 14.7
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Full range of X2 Detector raw | 14.3
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | Full range of X1 Detector raw | 14.5
| Jiangsu Hi-print Technology Co., Ltd | 2024-08-02 | A-B noise | 7.361
| Sincere International Trade Co., Ltd | 2024-07-31 | Z Scanner position at -3μm | 0.004198
| Sincere International Trade Co., Ltd | 2024-07-31 | Z Scanner position at +3μm | 0.00628
| Sincere International Trade Co., Ltd | 2024-07-31 | Z Linearity by Z Scanner position
| Sincere International Trade Co., Ltd | 2024-07-31 | Z Linearity by 3 different sample height | 1.0011
| Sincere International Trade Co., Ltd | 2024-07-31 | Z Height noise | 0.0264
| Sincere International Trade Co., Ltd | 2024-07-31 | Y Scan Stroke | 55
| Sincere International Trade Co., Ltd | 2024-07-31 | Y Detector noise (XY Servo on) | 0.2586
| Sincere International Trade Co., Ltd | 2024-07-31 | Y Detector noise (XY Servo off) | 0.2524
| Sincere International Trade Co., Ltd | 2024-07-31 | Y axis OPM | 0.8566
| Sincere International Trade Co., Ltd | 2024-07-31 | X Scan Stroke | 55
| Sincere International Trade Co., Ltd | 2024-07-31 | X Detector noise (XY Servo on) | 0.257
| Sincere International Trade Co., Ltd | 2024-07-31 | X Detector noise (XY Servo off) | 0.2499
| Sincere International Trade Co., Ltd | 2024-07-31 | X axis OPM | 1.1062
| Sincere International Trade Co., Ltd | 2024-07-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Sincere International Trade Co., Ltd | 2024-07-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Sincere International Trade Co., Ltd | 2024-07-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.147
| Sincere International Trade Co., Ltd | 2024-07-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.333
| Sincere International Trade Co., Ltd | 2024-07-31 | NCM driving% with AC160TS cantilever | 0.13
| Sincere International Trade Co., Ltd | 2024-07-31 | NCM Amplitude noise | 8.5969
| Sincere International Trade Co., Ltd | 2024-07-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.128
| Sincere International Trade Co., Ltd | 2024-07-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.141
| Sincere International Trade Co., Ltd | 2024-07-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 108.6
| Sincere International Trade Co., Ltd | 2024-07-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.3
| Sincere International Trade Co., Ltd | 2024-07-31 | Lateral(C-D) sensitivity | 4.7
| Sincere International Trade Co., Ltd | 2024-07-31 | Lateral(C-D) noise | 0.088
| Sincere International Trade Co., Ltd | 2024-07-31 | Full range of Z Height during Z Scan sweep | 17.6
| Sincere International Trade Co., Ltd | 2024-07-31 | Full range of Z Detector during Z Scan sweep | 13.5
| Sincere International Trade Co., Ltd | 2024-07-31 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| Sincere International Trade Co., Ltd | 2024-07-31 | Full range of Y Detector during Y Scan sweep | 75
| Sincere International Trade Co., Ltd | 2024-07-31 | Full range of X Detector Raw during X Scan sweep | 13.7
| Sincere International Trade Co., Ltd | 2024-07-31 | Full range of X Detector during X Scan sweep | 73.1
| Sincere International Trade Co., Ltd | 2024-07-31 | A-B noise | 15.6368
| ATC(Demoroom) | 2024-07-30 | Z Linearity by Z scanner position | 0.6
| ATC(Demoroom) | 2024-07-30 | Z Height noise | 0.024
| ATC(Demoroom) | 2024-07-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.8
| ATC(Demoroom) | 2024-07-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.9
| ATC(Demoroom) | 2024-07-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.6
| ATC(Demoroom) | 2024-07-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.5
| ATC(Demoroom) | 2024-07-30 | Z Drift | 17.6
| ATC(Demoroom) | 2024-07-30 | Z Detector offset | 60.4
| ATC(Demoroom) | 2024-07-30 | Y2 Detector noise (XY Servo on) | 0.78
| ATC(Demoroom) | 2024-07-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 278.2
| ATC(Demoroom) | 2024-07-30 | Y1 Detector noise (XY Servo on) | 0.78
| ATC(Demoroom) | 2024-07-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 278.7
| ATC(Demoroom) | 2024-07-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 177.6
| ATC(Demoroom) | 2024-07-30 | Y Ringing(Rpv) | 0.19
| ATC(Demoroom) | 2024-07-30 | Y Non-linearity error | 3.55
| ATC(Demoroom) | 2024-07-30 | Y Detector noise (XY Servo on) | 0.75
| ATC(Demoroom) | 2024-07-30 | Y axis, Repeatability @1st position | 0.02
| ATC(Demoroom) | 2024-07-30 | Y axis, Repeatability @ 4th position | 0.02
| ATC(Demoroom) | 2024-07-30 | Y axis, Repeatability @ 3rd position | 0.03
| ATC(Demoroom) | 2024-07-30 | Y axis, Repeatability @ 2nd position | 0.03
| ATC(Demoroom) | 2024-07-30 | Y axis, Accuracy @1st position | 0.11
| ATC(Demoroom) | 2024-07-30 | Y axis, Accuracy @ 4th position | 0.12
| ATC(Demoroom) | 2024-07-30 | Y axis, Accuracy @ 3rd position | 0.13
| ATC(Demoroom) | 2024-07-30 | Y axis, Accuracy @ 2nd position | 0.16
| ATC(Demoroom) | 2024-07-30 | Y axis OPM | 1.22
| ATC(Demoroom) | 2024-07-30 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.42
| ATC(Demoroom) | 2024-07-30 | XY Stage positioning Y offset (LLC Translation) | 1.91
| ATC(Demoroom) | 2024-07-30 | XY Stage positioning Y offset (Global Align) | 3.22
| ATC(Demoroom) | 2024-07-30 | XY Stage positioning X offset (Vision Pattern Matching) | 2.34
| ATC(Demoroom) | 2024-07-30 | XY Stage positioning X offset (LLC Translation) | 1.87
| ATC(Demoroom) | 2024-07-30 | XY Stage positioning X offset (Global Align) | 3.27
| ATC(Demoroom) | 2024-07-30 | XY Orthogonality | 90.1
| ATC(Demoroom) | 2024-07-30 | X2 Detector noise (XY Servo on) | 0.4
| ATC(Demoroom) | 2024-07-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 275.7
| ATC(Demoroom) | 2024-07-30 | X1 Detector noise (XY Servo on) | 0.36
| ATC(Demoroom) | 2024-07-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 276.4
| ATC(Demoroom) | 2024-07-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 167.1
| ATC(Demoroom) | 2024-07-30 | X Ringing(Rpv) | 0.11
| ATC(Demoroom) | 2024-07-30 | X Non-linearity error | 2.8
| ATC(Demoroom) | 2024-07-30 | X Detector noise (XY Servo on) | 0.31
| ATC(Demoroom) | 2024-07-30 | X axis, Repeatability @1st position | 0.03
| ATC(Demoroom) | 2024-07-30 | X axis, Repeatability @ 4th position | 0.07
| ATC(Demoroom) | 2024-07-30 | X axis, Repeatability @ 3rd position | 0.04
| ATC(Demoroom) | 2024-07-30 | X axis, Repeatability @ 2nd position | 0.08
| ATC(Demoroom) | 2024-07-30 | X axis, Accuracy @1st position | 0.16
| ATC(Demoroom) | 2024-07-30 | X axis, Accuracy @ 4th position | 0.35
| ATC(Demoroom) | 2024-07-30 | X axis, Accuracy @ 3rd position | 0.2
| ATC(Demoroom) | 2024-07-30 | X axis, Accuracy @ 2nd position | 0.34
| ATC(Demoroom) | 2024-07-30 | X axis OPM | 2.4
| ATC(Demoroom) | 2024-07-30 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.2
| ATC(Demoroom) | 2024-07-30 | Wafer placement(Loading) error on the load port2 (X offset) | 25.8
| ATC(Demoroom) | 2024-07-30 | Wafer placement(Loading) error on the load port1 (Y offset) | 31
| ATC(Demoroom) | 2024-07-30 | Wafer placement(Loading) error on the load port1 (X offset) | 24.3
| ATC(Demoroom) | 2024-07-30 | Torque Profile (50mm, 20um/s) | 0.531
| ATC(Demoroom) | 2024-07-30 | Torque Profile (50mm, 200um/s) | 0.518
| ATC(Demoroom) | 2024-07-30 | StrokeYMin | -152742
| ATC(Demoroom) | 2024-07-30 | StrokeYMax | 150359
| ATC(Demoroom) | 2024-07-30 | StrokeY | 303101
| ATC(Demoroom) | 2024-07-30 | StrokeXMin | -204014
| ATC(Demoroom) | 2024-07-30 | StrokeXMax | 203781
| ATC(Demoroom) | 2024-07-30 | StrokeX | 407795
| ATC(Demoroom) | 2024-07-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.74
| ATC(Demoroom) | 2024-07-30 | Roll motion to X+ full scan | 3.8
| ATC(Demoroom) | 2024-07-30 | RMS of ZeroScan | 0.052
| ATC(Demoroom) | 2024-07-30 | Position repeatability(10repeats, Position Based) | 0.25
| ATC(Demoroom) | 2024-07-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| ATC(Demoroom) | 2024-07-30 | OpticalAxisShiftPerFocusY | -0.0025
| ATC(Demoroom) | 2024-07-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 3e-05
| ATC(Demoroom) | 2024-07-30 | OpticalAxisShiftPerFocusX | -0.0023
| ATC(Demoroom) | 2024-07-30 | OPM Repeatability @5mm | 0.68
| ATC(Demoroom) | 2024-07-30 | OPM Repeatability @25mm | 1.43
| ATC(Demoroom) | 2024-07-30 | OPM Repeatability @1mm | 0.5
| ATC(Demoroom) | 2024-07-30 | OPM Repeatability @10mm | 1.06
| ATC(Demoroom) | 2024-07-30 | NCM driving% with AC160TS cantilever | 15.7
| ATC(Demoroom) | 2024-07-30 | Min. Vibration Transmissbility (~100Hz) | 0.017
| ATC(Demoroom) | 2024-07-30 | Max OPM @5mm | 17.47
| ATC(Demoroom) | 2024-07-30 | Max OPM @25mm | 82.83
| ATC(Demoroom) | 2024-07-30 | Max OPM @1mm | 8.49
| ATC(Demoroom) | 2024-07-30 | Max OPM @10mm | 21.23
| ATC(Demoroom) | 2024-07-30 | LowerLimit | 2117.25
| ATC(Demoroom) | 2024-07-30 | LowerLimit | 4728
| ATC(Demoroom) | 2024-07-30 | Lateral(C-D) sensitivity | 45.366
| ATC(Demoroom) | 2024-07-30 | Lateral(C-D) noise | 0.213
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from -Y to +Y (Y scan) | 49.07
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from -Y to +Y (X scan) | 25.63
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from -X to +X (Y scan) | 35.89
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from -X to +X (X scan) | 78.37
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from +Y to -Y (Y scan) | 117.19
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from +Y to -Y (X scan) | 19.04
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from +X to -X (Y scan) | 17.58
| ATC(Demoroom) | 2024-07-30 | Image offset after moving from +X to -X (X scan) | 59.33
| ATC(Demoroom) | 2024-07-30 | Full range of Z Height during Z Scan sweep | 17.1
| ATC(Demoroom) | 2024-07-30 | Full range of Z Detector during Z Scan sweep | 13.7
| ATC(Demoroom) | 2024-07-30 | Full range of Y2 Detector Raw during Y Scan sweep | 13.88
| ATC(Demoroom) | 2024-07-30 | Full range of Y2 Detector during Y Scan sweep | 217.4
| ATC(Demoroom) | 2024-07-30 | Full range of Y1 Detector Raw during Y Scan sweep | 13.92
| ATC(Demoroom) | 2024-07-30 | Full range of Y1 Detector during Y Scan sweep | 198.6
| ATC(Demoroom) | 2024-07-30 | Full range of Y Detector during Y Scan sweep | 208
| ATC(Demoroom) | 2024-07-30 | Full range of X2 Detector Raw during X Scan sweep | 13.62
| ATC(Demoroom) | 2024-07-30 | Full range of X2 Detector during X Scan sweep | 207.5
| ATC(Demoroom) | 2024-07-30 | Full range of X1 Detector Raw during X Scan sweep | 13.95
| ATC(Demoroom) | 2024-07-30 | Full range of X1 Detector during X Scan sweep | 197.9
| ATC(Demoroom) | 2024-07-30 | Full range of X Detector during X Scan sweep | 202.7
| ATC(Demoroom) | 2024-07-30 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| ATC(Demoroom) | 2024-07-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| ATC(Demoroom) | 2024-07-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.63
| ATC(Demoroom) | 2024-07-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.41
| ATC(Demoroom) | 2024-07-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| ATC(Demoroom) | 2024-07-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| ATC(Demoroom) | 2024-07-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| ATC(Demoroom) | 2024-07-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.78
| ATC(Demoroom) | 2024-07-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| ATC(Demoroom) | 2024-07-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 11.6
| ATC(Demoroom) | 2024-07-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.8
| ATC(Demoroom) | 2024-07-30 | %GRR w/ roughness of non-patterened wafer | 12.6
| Osaka University | 2024-07-29 | Z Scanner position at -3μm | 0.00483
| Osaka University | 2024-07-29 | Z Scanner position at +3μm | 0.001675
| Osaka University | 2024-07-29 | Z linearity by sample height | 1.0179
| Osaka University | 2024-07-29 | Z Height noise | 0.024
| Osaka University | 2024-07-29 | Y Scan Stroke | 197
| Osaka University | 2024-07-29 | Y Detector noise (XY Servo on) | 0.398
| Osaka University | 2024-07-29 | Y Detector noise (XY Servo off) | 0.391
| Osaka University | 2024-07-29 | Y axis OPM | 0.8282
| Osaka University | 2024-07-29 | X Scan Stroke | 184.5
| Osaka University | 2024-07-29 | X Detector noise (XY Servo on) | 0.379
| Osaka University | 2024-07-29 | X Detector noise (XY Servo off) | 0.369
| Osaka University | 2024-07-29 | X axis OPM | 1.3638
| Osaka University | 2024-07-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.6
| Osaka University | 2024-07-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.81
| Osaka University | 2024-07-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.6
| Osaka University | 2024-07-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.83
| Osaka University | 2024-07-29 | NCM Amplitude noise | 12.321
| Osaka University | 2024-07-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.962
| Osaka University | 2024-07-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.256
| Osaka University | 2024-07-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 247.4
| Osaka University | 2024-07-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 240.1
| Osaka University | 2024-07-29 | Lateral(C-D) sensitivity | 14.8
| Osaka University | 2024-07-29 | Lateral(C-D) noise | 0.136
| Osaka University | 2024-07-29 | Full range of Z Height during Z Scan sweep | 17.3
| Osaka University | 2024-07-29 | Full range of Z Detector during Z Scan sweep | 13.7
| Osaka University | 2024-07-29 | Full range of Y Detector Raw during Y Scan sweep | 15.9
| Osaka University | 2024-07-29 | Full range of Y Detector during Y Scan sweep | 198.1
| Osaka University | 2024-07-29 | Full range of X Detector Raw during X Scan sweep | 15.3
| Osaka University | 2024-07-29 | Full range of X Detector during X Scan sweep | 189
| Osaka University | 2024-07-29 | A-B noise | 11.045
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 |  | 194213.6
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 |  | -110769
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 |  | 100544
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 |  | -103378
| Park Systems Europe GmbH (PSE) | 2024-07-26 |  | 194396.9
| Park Systems Europe GmbH (PSE) | 2024-07-26 |  | -110166
| Park Systems Europe GmbH (PSE) | 2024-07-26 |  | 102984.4
| Park Systems Europe GmbH (PSE) | 2024-07-26 |  | -102269
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Z Scanner position at -3μm | 0.0077
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Z Scanner position at -3μm | 0.0098
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Z Scanner position at +3μm | 0.0064
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Z Scanner position at +3μm | 0.0095
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Z linearity by sample height | 1
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Z linearity by sample height | 1
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Z Height noise | 0.014
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Z Height noise | 0.016
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y2 Detector noise (XY Servo on) | 0.21
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y2 Detector noise (XY Servo on) | 0.23
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y2 Detector noise (XY Servo off) | 0.19
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y2 Detector noise (XY Servo off) | 0.22
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y1 Detector noise (XY servo on) | 0.19
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y1 Detector noise (XY servo on) | 0.23
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y1 Detector noise (XY servo off) | 0.18
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y1 Detector noise (XY servo off) | 0.21
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y Scan Stroke | 167.9
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y Scan Stroke | 178.3
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y Non-linearity error | 0.16
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y Detector noise (XY servo on) | 0.14
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y Detector noise (XY servo on) | 0.17
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y Detector noise (XY servo off) | 0.13
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y Detector noise (XY servo off) | 0.16
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y axis OPM (-52000,-52000) | 1.63
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y axis OPM (-52000,-52000) | 2.13
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y axis OPM (-52000,52000) | 2.34
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y axis OPM (-52000,52000) | 1.9
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y axis OPM (52000,-52000) | 1.63
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y axis OPM (52000,-52000) | 1.93
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y axis OPM (52000,52000) | 1.62
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y axis OPM (52000,52000) | 1.93
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Y axis OPM (0,0) | 0.94
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Y axis OPM (0,0) | 1.72
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | XY Orthogonality (-52000,-52000) | 90.2
| Park Systems Europe GmbH (PSE) | 2024-07-26 | XY Orthogonality (-52000,-52000) | 89.69
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | XY Orthogonality (-52000,52000) | 89.8
| Park Systems Europe GmbH (PSE) | 2024-07-26 | XY Orthogonality (-52000,52000) | 89.97
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | XY Orthogonality (52000,-52000) | 89.9
| Park Systems Europe GmbH (PSE) | 2024-07-26 | XY Orthogonality (52000,-52000) | 89.77
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | XY Orthogonality (52000,52000) | 90.2
| Park Systems Europe GmbH (PSE) | 2024-07-26 | XY Orthogonality (52000,52000) | 90.01
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | XY Orthogonality (0,0) | 90.1
| Park Systems Europe GmbH (PSE) | 2024-07-26 | XY Orthogonality (0,0) | 89.8
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X2 Detector noise (XY Servo on) | 0.22
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X2 Detector noise (XY Servo on) | 0.3
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X2 Detector noise (XY Servo off) | 0.21
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X2 Detector noise (XY Servo off) | 0.27
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X1 Detector noise (XY servo on) | 0.21
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X1 Detector noise (XY servo on) | 0.35
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X1 Detector noise (XY servo off) | 0.2
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X1 Detector noise (XY servo off) | 0.34
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X Scan Stroke | 172
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X Scan Stroke | 177.6
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X Non-linearity error | 0.18
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X Detector noise (XY servo on) | 0.17
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X Detector noise (XY servo on) | 0.25
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X Detector noise (XY servo off) | 0.15
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X Detector noise (XY servo off) | 0.22
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X axis OPM (-52000,-52000) | 2.79
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X axis OPM (-52000,-52000) | 1.49
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X axis OPM (-52000,52000) | 2.21
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X axis OPM (-52000,52000) | 3.86
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X axis OPM (52000,-52000) | 2.79
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X axis OPM (52000,-52000) | 1.53
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X axis OPM (52000,52000) | 0.41
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X axis OPM (52000,52000) | 0.73
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | X axis OPM (0,0) | 1.69
| Park Systems Europe GmbH (PSE) | 2024-07-26 | X axis OPM (0,0) | 0.88
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.17
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.19
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.17
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.2
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.21
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.16
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.16
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.16
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.23
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.15
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.15
| Park Systems Europe GmbH (PSE) | 2024-07-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | NCM driving% with AC160TS cantilever | 0.32
| Park Systems Europe GmbH (PSE) | 2024-07-26 | NCM driving% with AC160TS cantilever | 1.58
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | NCM Amplitude noise | 12.37
| Park Systems Europe GmbH (PSE) | 2024-07-26 | NCM Amplitude noise | 6.97
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.043
| Park Systems Europe GmbH (PSE) | 2024-07-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.774
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.058
| Park Systems Europe GmbH (PSE) | 2024-07-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 20.794
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 246.3
| Park Systems Europe GmbH (PSE) | 2024-07-26 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 264
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 239
| Park Systems Europe GmbH (PSE) | 2024-07-26 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 256.7
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 250.7
| Park Systems Europe GmbH (PSE) | 2024-07-26 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 274.2
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 247.1
| Park Systems Europe GmbH (PSE) | 2024-07-26 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 251.9
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Lateral(C-D) sensitivity | 18.6
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Lateral(C-D) sensitivity | 12.8
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Lateral(C-D) noise | 0.112
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Lateral(C-D) noise | 0.106
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of Z Height during Z Scan sweep | 15.6
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of Z Height during Z Scan sweep | 15.7
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of Y2 Detector Raw during Y Scan sweep | 13.8
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of Y2 Detector Raw during Y Scan sweep | 13.6
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of Y2 Detector during Y Scan sweep | 169.6
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of Y2 Detector during Y Scan sweep | 174.3
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of Y1 Detector Raw during Y Scan sweep | 13.9
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of Y1 Detector Raw during Y Scan sweep | 13.8
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of Y1 Detector during Y Scan sweep | 166
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of Y1 Detector during Y Scan sweep | 172.8
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of Y Detector during Y Scan sweep | 167.7
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of Y Detector during Y Scan sweep | 173.6
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of X2 Detector Raw during X Scan sweep | 13.2
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of X2 Detector during X Scan sweep | 173.3
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of X2 Detector during X Scan sweep | 176
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of X1 Detector Raw during X Scan sweep | 13.7
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of X1 Detector Raw during X Scan sweep | 13.4
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of X1 Detector during X Scan sweep | 168.8
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of X1 Detector during X Scan sweep | 166.6
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | Full range of X Detector during X Scan sweep | 171
| Park Systems Europe GmbH (PSE) | 2024-07-26 | Full range of X Detector during X Scan sweep | 171.3
| Park Systems Japan (PSJ) SID No. J34003-220524 | 2024-07-26 | A-B noise | 11.33
| Park Systems Europe GmbH (PSE) | 2024-07-26 | A-B noise | 9.79
| Henan Academy of Sciences | 2024-07-25 | Z Scanner position at -3μm | 0.0006
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Z Scanner position at -3μm | 0.0015
| Henan Academy of Sciences | 2024-07-25 | Z Scanner position at +3μm | 0.0012
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Z Scanner position at +3μm | 0.0074
| Henan Academy of Sciences | 2024-07-25 | Z linearity by sample height | 1.014
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Z linearity by sample height | 1.013
| Henan Academy of Sciences | 2024-07-25 | Z Height noise | 0.026
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Z Height noise | 0.018
| Henan Academy of Sciences | 2024-07-25 | Y Scan Stroke | 119.3
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Y Scan Stroke | 115.3
| Henan Academy of Sciences | 2024-07-25 | Y Detector noise (XY Servo on) | 0.27
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Y Detector noise (XY Servo on) | 0.26
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Y Detector noise (XY Servo off) | 0.26
| Henan Academy of Sciences | 2024-07-25 | Y Detector noise (XY Servo off) | 0.26
| Henan Academy of Sciences | 2024-07-25 | Y axis OPM | 0.99
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Y axis OPM | 1.49
| Henan Academy of Sciences | 2024-07-25 | X Scan Stroke | 119.7
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | X Scan Stroke | 116.1
| Henan Academy of Sciences | 2024-07-25 | X Detector noise (XY Servo on) | 0.27
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | X Detector noise (XY Servo on) | 0.26
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | X Detector noise (XY Servo off) | 0.25
| Henan Academy of Sciences | 2024-07-25 | X Detector noise (XY Servo off) | 0.27
| Henan Academy of Sciences | 2024-07-25 | X axis OPM | 1.17
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | X axis OPM | 0.83
| Henan Academy of Sciences | 2024-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Henan Academy of Sciences | 2024-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Henan Academy of Sciences | 2024-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Henan Academy of Sciences | 2024-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Henan Academy of Sciences | 2024-07-25 | NCM driving% with AC160TS cantilever | 1.15
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | NCM driving% with AC160TS cantilever | 0.875
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | NCM Amplitude noise | 13.13
| Henan Academy of Sciences | 2024-07-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.588
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.187
| Henan Academy of Sciences | 2024-07-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.418
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.601
| Henan Academy of Sciences | 2024-07-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 152.6
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.4
| Henan Academy of Sciences | 2024-07-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 155.6
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160.8
| Henan Academy of Sciences | 2024-07-25 | Lateral(C-D) sensitivity | 12.9
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Lateral(C-D) sensitivity | 8.3
| Henan Academy of Sciences | 2024-07-25 | Lateral(C-D) noise | 0.083
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Lateral(C-D) noise | 0.124
| Henan Academy of Sciences | 2024-07-25 | Full range of Z Height during Z Scan sweep | 19
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Full range of Z Height during Z Scan sweep | 18
| Henan Academy of Sciences | 2024-07-25 | Full range of Z Detector during Z Scan sweep | 16.2
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Full range of Z Detector during Z Scan sweep | 14.6
| Henan Academy of Sciences | 2024-07-25 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Henan Academy of Sciences | 2024-07-25 | Full range of Y Detector during Y Scan sweep | 118.6
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Full range of Y Detector during Y Scan sweep | 115.3
| Henan Academy of Sciences | 2024-07-25 | Full range of X Detector Raw during X Scan sweep | 15
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Full range of X Detector Raw during X Scan sweep | 14.4
| Henan Academy of Sciences | 2024-07-25 | Full range of X Detector during X Scan sweep | 115.4
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | Full range of X Detector during X Scan sweep | 116.1
| Park Systems Inc. (PSI) SID No. - U28009-180424 | 2024-07-25 | A-B noise | 13.66
| Wuhan Epiflag | 2024-07-22 | Z Scanner position at -3μm | 0.0091
| Wuhan Epiflag | 2024-07-22 | Z Scanner position at +3μm | 0.0084
| Wuhan Epiflag | 2024-07-22 | Z Linearity by Z Scanner position
| Wuhan Epiflag | 2024-07-22 | Z Linearity by 3 different sample height | 1.005
| Wuhan Epiflag | 2024-07-22 | Z Height noise | 0.00033
| Wuhan Epiflag | 2024-07-22 | Y Scan Stroke | 135.3
| Wuhan Epiflag | 2024-07-22 | -Y Direction software Limit | -75580
| Wuhan Epiflag | 2024-07-22 | Y Detector noise (XY Servo on) | 0.308
| Wuhan Epiflag | 2024-07-22 | Y Detector noise (XY Servo off) | 0.296
| Wuhan Epiflag | 2024-07-22 | Y axis OPM | 0.8188
| Wuhan Epiflag | 2024-07-22 | X Scan Stroke | 137.1
| Wuhan Epiflag | 2024-07-22 | -X Direction software Limit | -75590.6
| Wuhan Epiflag | 2024-07-22 | X Detector noise (XY Servo on) | 0.31
| Wuhan Epiflag | 2024-07-22 | X Detector noise (XY Servo off) | 0.297
| Wuhan Epiflag | 2024-07-22 | X axis OPM | 1.1538
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.27
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.28
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.28
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.27
| Wuhan Epiflag | 2024-07-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| Wuhan Epiflag | 2024-07-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Wuhan Epiflag | 2024-07-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.19
| Wuhan Epiflag | 2024-07-22 | NCM driving% with AC160TS cantilever | 0.249
| Wuhan Epiflag | 2024-07-22 | NCM Amplitude noise | 14.359
| Wuhan Epiflag | 2024-07-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.093
| Wuhan Epiflag | 2024-07-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.388
| Wuhan Epiflag | 2024-07-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 196.4
| Wuhan Epiflag | 2024-07-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 191
| Wuhan Epiflag | 2024-07-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Wuhan Epiflag | 2024-07-22 | Lateral(C-D) sensitivity | 17.7
| Wuhan Epiflag | 2024-07-22 | Lateral(C-D) noise | 0.118
| Wuhan Epiflag | 2024-07-22 | Full range of Z Height during Z Scan sweep | 17.4
| Wuhan Epiflag | 2024-07-22 | Full range of Z Detector during Z Scan sweep | 12.3
| Wuhan Epiflag | 2024-07-22 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| Wuhan Epiflag | 2024-07-22 | Full range of Y Detector during Y Scan sweep | 135.3
| Wuhan Epiflag | 2024-07-22 | Full range of X Detector Raw during X Scan sweep | 14.4
| Wuhan Epiflag | 2024-07-22 | Full range of X Detector during X Scan sweep | 137.1
| Wuhan Epiflag | 2024-07-22 | A-B noise | 7.296
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Z Scanner position at -3μm | 0.0103
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Z Scanner position at +3μm | 0.0059
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Z Linearity by 3 different sample height | 0.991
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Z Height noise | 0.03
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Y Scan Stroke | 55
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Y Detector noise (XY Servo on) | 0.246
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Y Detector noise (XY Servo off) | 0.24
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Y axis OPM | 0.8156
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | X Scan Stroke | 55
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | X Detector noise (XY Servo on) | 0.252
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | X Detector noise (XY Servo off) | 0.245
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | X axis OPM | 0.765
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | NCM driving% with AC160TS cantilever | 3.055
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | NCM Amplitude noise | 8.734
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.076
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.274
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.5
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.5
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Lateral(C-D) sensitivity | 17.8
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Lateral(C-D) noise | 0.07
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Full range of Z Height during Z Scan sweep | 17.9
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Full range of Z Detector during Z Scan sweep | 14.2
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Full range of Y Detector during Y Scan sweep | 73.7
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Full range of X Detector Raw during X Scan sweep | 14.5
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | Full range of X Detector during X Scan sweep | 74
| 목포해양대학교 기관시스템공학부 | 2024-07-19 | A-B noise | 13.676
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Z Scanner position at -3μm | 0.0042
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Z Scanner position at +3μm | 0.0058
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Z linearity by sample height | 1.004
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Z Height noise | 0.028
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y2 Detector noise (XY Servo on) | 0.211
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y2 Detector noise (XY Servo off) | 0.207
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y1 Detector noise (XY Servo on) | 0.22
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y1 Detector noise (XY Servo off) | 0.216
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y Scan Stroke | 174.3
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y Detector noise (XY Servo on) | 0.153
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y Detector noise (XY Servo off) | 0.15
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Y axis OPM | 1.21
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X2 Detector noise (XY Servo on) | 0.221
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X2 Detector noise (XY Servo off) | 0.217
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X1 Detector noise (XY Servo on) | 0.237
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X1 Detector noise (XY Servo off) | 0.234
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X Scan Stroke | 173.6
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X Detector noise (XY Servo on) | 0.168
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X Detector noise (XY Servo off) | 0.166
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | X axis OPM | 1.17
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.36
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.37
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | NCM Amplitude noise | 15.26
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.745
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.353
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 230
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 239.6
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 227.4
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 245
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Maximum Y Scan range (Full Stroke) | 174.6
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Maximum X Scan range (Full Stroke) | 173.4
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Lateral(C-D) sensitivity | 19
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Lateral(C-D) noise | 0.107
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Full range of Z Height during Z Scan sweep | 17.2
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Full range of Z Detector during Z Scan sweep | 12.6
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Full range of Y2 Detector raw | 14.9
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Full range of Y1 Detector raw | 14.8
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Full range of X2 Detector raw | 14.7
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | Full range of X1 Detector raw | 14.7
| Hunan San'an Semiconductor Co., Ltd | 2024-07-18 | A-B noise | 7.47
| Bosch | 2024-07-12 | Z Linearity by Z scanner position | 1.7
| Huahong Chengdu | 2024-07-12 | Z Linearity by Z scanner position | 0.7
| Intel Hillsboro#2 | 2024-07-12 | Z Linearity by Z scanner position | 0.2
| Bosch | 2024-07-12 | Z Height noise | 0.035
| Huahong Chengdu | 2024-07-12 | Z Height noise | 0.027
| Intel Hillsboro#2 | 2024-07-12 | Z Height noise | 0.07
| Bosch | 2024-07-12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.8
| Huahong Chengdu | 2024-07-12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7
| Intel Hillsboro#2 | 2024-07-12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 64.0503
| Bosch | 2024-07-12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.8
| Huahong Chengdu | 2024-07-12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7
| Intel Hillsboro#2 | 2024-07-12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 63.9784
| Bosch | 2024-07-12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.9
| Huahong Chengdu | 2024-07-12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.8
| Intel Hillsboro#2 | 2024-07-12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.6509
| Bosch | 2024-07-12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12
| Huahong Chengdu | 2024-07-12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.9
| Intel Hillsboro#2 | 2024-07-12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.5724
| Bosch | 2024-07-12 | Z Drift | 12.4
| Huahong Chengdu | 2024-07-12 | Z Drift | 11
| Intel Hillsboro#2 | 2024-07-12 | Z Drift | 16.4
| Bosch | 2024-07-12 | Z Detector offset | 50.4
| Huahong Chengdu | 2024-07-12 | Z Detector offset | 38.1
| Intel Hillsboro#2 | 2024-07-12 | Z Detector offset | 122.6
| Bosch | 2024-07-12 | Y2 Detector noise (XY Servo on) | 0.26
| Huahong Chengdu | 2024-07-12 | Y2 Detector noise (XY Servo on) | 0.48
| Intel Hillsboro#2 | 2024-07-12 | Y2 Detector noise (XY Servo on) | 0.49
| Bosch | 2024-07-12 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 265.7
| Huahong Chengdu | 2024-07-12 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 284.4
| Intel Hillsboro#2 | 2024-07-12 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 273.4
| Bosch | 2024-07-12 | Y1 Detector noise (XY Servo on) | 0.28
| Huahong Chengdu | 2024-07-12 | Y1 Detector noise (XY Servo on) | 0.46
| Intel Hillsboro#2 | 2024-07-12 | Y1 Detector noise (XY Servo on) | 0.45
| Bosch | 2024-07-12 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 257.4
| Huahong Chengdu | 2024-07-12 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 282.7
| Intel Hillsboro#2 | 2024-07-12 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 269.2
| Bosch | 2024-07-12 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 153
| Huahong Chengdu | 2024-07-12 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 151.4
| Intel Hillsboro#2 | 2024-07-12 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 176
| Bosch | 2024-07-12 | Y Ringing(Rpv) | 0.13
| Huahong Chengdu | 2024-07-12 | Y Ringing(Rpv) | 0.05
| Intel Hillsboro#2 | 2024-07-12 | Y Ringing(Rpv) | 0.87
| Bosch | 2024-07-12 | Y Non-linearity error | 3.61
| Huahong Chengdu | 2024-07-12 | Y Non-linearity error | 3.78
| Intel Hillsboro#2 | 2024-07-12 | Y Non-linearity error | 2.55
| Bosch | 2024-07-12 | Y Detector noise (XY Servo on) | 0.2
| Huahong Chengdu | 2024-07-12 | Y Detector noise (XY Servo on) | 0.42
| Intel Hillsboro#2 | 2024-07-12 | Y Detector noise (XY Servo on) | 0.42
| Bosch | 2024-07-12 | Y axis, Repeatability @1st position | 0.03
| Huahong Chengdu | 2024-07-12 | Y axis, Repeatability @1st position | 0.03
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Repeatability @1st position | 0.02
| Bosch | 2024-07-12 | Y axis, Repeatability @ 4th position | 0.04
| Huahong Chengdu | 2024-07-12 | Y axis, Repeatability @ 4th position | 0.03
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Repeatability @ 4th position | 0.03
| Bosch | 2024-07-12 | Y axis, Repeatability @ 3rd position | 0.04
| Huahong Chengdu | 2024-07-12 | Y axis, Repeatability @ 3rd position | 0.03
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Repeatability @ 3rd position | 0.02
| Bosch | 2024-07-12 | Y axis, Repeatability @ 2nd position | 0.05
| Huahong Chengdu | 2024-07-12 | Y axis, Repeatability @ 2nd position | 0.02
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Repeatability @ 2nd position | 0.03
| Bosch | 2024-07-12 | Y axis, Accuracy @1st position | 0.17
| Huahong Chengdu | 2024-07-12 | Y axis, Accuracy @1st position | 0.36
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Accuracy @1st position | 0.13
| Bosch | 2024-07-12 | Y axis, Accuracy @ 4th position | 0.2
| Huahong Chengdu | 2024-07-12 | Y axis, Accuracy @ 4th position | 0.41
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Accuracy @ 4th position | 0.15
| Bosch | 2024-07-12 | Y axis, Accuracy @ 3rd position | 0.21
| Huahong Chengdu | 2024-07-12 | Y axis, Accuracy @ 3rd position | 0.37
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Accuracy @ 3rd position | 0.12
| Bosch | 2024-07-12 | Y axis, Accuracy @ 2nd position | 0.22
| Huahong Chengdu | 2024-07-12 | Y axis, Accuracy @ 2nd position | 0.28
| Intel Hillsboro#2 | 2024-07-12 | Y axis, Accuracy @ 2nd position | 0.12
| Bosch | 2024-07-12 | Y axis OPM | 0.45
| Huahong Chengdu | 2024-07-12 | Y axis OPM | 0.5
| Intel Hillsboro#2 | 2024-07-12 | Y axis OPM | 1.09
| Huahong Chengdu | 2024-07-12 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.793
| Huahong Chengdu | 2024-07-12 | XY Stage positioning Y offset (LLC Translation) | 1.416
| Huahong Chengdu | 2024-07-12 | XY Stage positioning Y offset (Global Align) | 4.0967
| Huahong Chengdu | 2024-07-12 | XY Stage positioning X offset (Vision Pattern Matching) | 2.0898
| Huahong Chengdu | 2024-07-12 | XY Stage positioning X offset (LLC Translation) | 2.4121
| Huahong Chengdu | 2024-07-12 | XY Stage positioning X offset (Global Align) | 4.9121
| Bosch | 2024-07-12 | XY Orthogonality | 89.78
| Huahong Chengdu | 2024-07-12 | XY Orthogonality | 89.86
| Intel Hillsboro#2 | 2024-07-12 | XY Orthogonality | 90
| Bosch | 2024-07-12 | X2 Detector noise (XY Servo on) | 0.27
| Huahong Chengdu | 2024-07-12 | X2 Detector noise (XY Servo on) | 0.36
| Intel Hillsboro#2 | 2024-07-12 | X2 Detector noise (XY Servo on) | 0.34
| Bosch | 2024-07-12 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 263.6
| Huahong Chengdu | 2024-07-12 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 293.2
| Intel Hillsboro#2 | 2024-07-12 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 280.1
| Bosch | 2024-07-12 | X1 Detector noise (XY Servo on) | 0.26
| Huahong Chengdu | 2024-07-12 | X1 Detector noise (XY Servo on) | 0.38
| Intel Hillsboro#2 | 2024-07-12 | X1 Detector noise (XY Servo on) | 0.33
| Bosch | 2024-07-12 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 258.8
| Huahong Chengdu | 2024-07-12 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 288
| Intel Hillsboro#2 | 2024-07-12 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.6
| Bosch | 2024-07-12 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 155.1
| Huahong Chengdu | 2024-07-12 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 178.5
| Intel Hillsboro#2 | 2024-07-12 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 173.7
| Bosch | 2024-07-12 | X Ringing(Rpv) | 0.13
| Huahong Chengdu | 2024-07-12 | X Ringing(Rpv) | 0.08
| Intel Hillsboro#2 | 2024-07-12 | X Ringing(Rpv) | 0.93
| Bosch | 2024-07-12 | X Non-linearity error | 2.66
| Huahong Chengdu | 2024-07-12 | X Non-linearity error | 2.53
| Intel Hillsboro#2 | 2024-07-12 | X Non-linearity error | 3.28
| Bosch | 2024-07-12 | X Detector noise (XY Servo on) | 0.19
| Huahong Chengdu | 2024-07-12 | X Detector noise (XY Servo on) | 0.3
| Intel Hillsboro#2 | 2024-07-12 | X Detector noise (XY Servo on) | 0.26
| Bosch | 2024-07-12 | X axis, Repeatability @1st position | 0.02
| Huahong Chengdu | 2024-07-12 | X axis, Repeatability @1st position | 0.05
| Intel Hillsboro#2 | 2024-07-12 | X axis, Repeatability @1st position | 0.04
| Bosch | 2024-07-12 | X axis, Repeatability @ 4th position | 0.03
| Huahong Chengdu | 2024-07-12 | X axis, Repeatability @ 4th position | 0.08
| Intel Hillsboro#2 | 2024-07-12 | X axis, Repeatability @ 4th position | 0.03
| Bosch | 2024-07-12 | X axis, Repeatability @ 3rd position | 0.04
| Huahong Chengdu | 2024-07-12 | X axis, Repeatability @ 3rd position | 0.06
| Intel Hillsboro#2 | 2024-07-12 | X axis, Repeatability @ 3rd position | 0.03
| Bosch | 2024-07-12 | X axis, Repeatability @ 2nd position | 0.04
| Huahong Chengdu | 2024-07-12 | X axis, Repeatability @ 2nd position | 0.1
| Intel Hillsboro#2 | 2024-07-12 | X axis, Repeatability @ 2nd position | 0.03
| Bosch | 2024-07-12 | X axis, Accuracy @1st position | 0.12
| Huahong Chengdu | 2024-07-12 | X axis, Accuracy @1st position | 0.35
| Intel Hillsboro#2 | 2024-07-12 | X axis, Accuracy @1st position | 0.15
| Bosch | 2024-07-12 | X axis, Accuracy @ 4th position | 0.14
| Huahong Chengdu | 2024-07-12 | X axis, Accuracy @ 4th position | 0.41
| Intel Hillsboro#2 | 2024-07-12 | X axis, Accuracy @ 4th position | 0.15
| Bosch | 2024-07-12 | X axis, Accuracy @ 3rd position | 0.2
| Huahong Chengdu | 2024-07-12 | X axis, Accuracy @ 3rd position | 0.44
| Intel Hillsboro#2 | 2024-07-12 | X axis, Accuracy @ 3rd position | 0.17
| Bosch | 2024-07-12 | X axis, Accuracy @ 2nd position | 0.18
| Huahong Chengdu | 2024-07-12 | X axis, Accuracy @ 2nd position | 0.44
| Intel Hillsboro#2 | 2024-07-12 | X axis, Accuracy @ 2nd position | 0.14
| Bosch | 2024-07-12 | X axis OPM | 0.56
| Huahong Chengdu | 2024-07-12 | X axis OPM | 0.2
| Intel Hillsboro#2 | 2024-07-12 | X axis OPM | 0.9
| Bosch | 2024-07-12 | Wafer placement(Loading) error on the load port2 (Y offset) | 33.6
| Huahong Chengdu | 2024-07-12 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.5
| Intel Hillsboro#2 | 2024-07-12 | Wafer placement(Loading) error on the load port2 (Y offset) | 16.2
| Bosch | 2024-07-12 | Wafer placement(Loading) error on the load port2 (X offset) | 25.6
| Huahong Chengdu | 2024-07-12 | Wafer placement(Loading) error on the load port2 (X offset) | 36.1
| Intel Hillsboro#2 | 2024-07-12 | Wafer placement(Loading) error on the load port2 (X offset) | 25.5
| Bosch | 2024-07-12 | Wafer placement(Loading) error on the load port1 (Y offset) | 24.3
| Huahong Chengdu | 2024-07-12 | Wafer placement(Loading) error on the load port1 (Y offset) | 25.1
| Intel Hillsboro#2 | 2024-07-12 | Wafer placement(Loading) error on the load port1 (Y offset) | 32.7
| Bosch | 2024-07-12 | Wafer placement(Loading) error on the load port1 (X offset) | 30.5
| Huahong Chengdu | 2024-07-12 | Wafer placement(Loading) error on the load port1 (X offset) | 75.8
| Intel Hillsboro#2 | 2024-07-12 | Wafer placement(Loading) error on the load port1 (X offset) | 35.4
| Huahong Chengdu | 2024-07-12 | Torque Profile (50mm, 20um/s) | 0.756
| Intel Hillsboro#2 | 2024-07-12 | Torque Profile (50mm, 20um/s) | -0.658
| Huahong Chengdu | 2024-07-12 | Torque Profile (50mm, 200um/s) | 0.659
| Intel Hillsboro#2 | 2024-07-12 | Torque Profile (50mm, 200um/s) | -0.577
| Bosch | 2024-07-12 | StrokeYMin | -101895
| Huahong Chengdu | 2024-07-12 | StrokeYMin | -153685
| Intel Hillsboro#2 | 2024-07-12 | StrokeYMin | -152820
| Bosch | 2024-07-12 | StrokeYMax | 104452.5
| Huahong Chengdu | 2024-07-12 | StrokeYMax | 150107
| Intel Hillsboro#2 | 2024-07-12 | StrokeYMax | 150386
| Bosch | 2024-07-12 | StrokeY | 206347.1
| Huahong Chengdu | 2024-07-12 | StrokeY | 303792
| Intel Hillsboro#2 | 2024-07-12 | StrokeY | 303206
| Bosch | 2024-07-12 | StrokeXMin | -152699
| Huahong Chengdu | 2024-07-12 | StrokeXMin | -201334
| Intel Hillsboro#2 | 2024-07-12 | StrokeXMin | -203131
| Bosch | 2024-07-12 | StrokeXMax | 152675.7
| Huahong Chengdu | 2024-07-12 | StrokeXMax | 203206
| Intel Hillsboro#2 | 2024-07-12 | StrokeXMax | 202607
| Bosch | 2024-07-12 | StrokeX | 305374.7
| Huahong Chengdu | 2024-07-12 | StrokeX | 404540
| Intel Hillsboro#2 | 2024-07-12 | StrokeX | 405738
| Bosch | 2024-07-12 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.8
| Huahong Chengdu | 2024-07-12 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.8
| Intel Hillsboro#2 | 2024-07-12 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.68
| Huahong Chengdu | 2024-07-12 | Roll motion to X+ full scan | 9.3
| Intel Hillsboro#2 | 2024-07-12 | Roll motion to X+ full scan | 31.9
| Bosch | 2024-07-12 | RMS of ZeroScan | 0.036
| Huahong Chengdu | 2024-07-12 | RMS of ZeroScan | 0.058
| Intel Hillsboro#2 | 2024-07-12 | RMS of ZeroScan | 0.057
| Huahong Chengdu | 2024-07-12 | Position repeatability(10repeats, Position Based) | 0.27
| Intel Hillsboro#2 | 2024-07-12 | Position repeatability(10repeats, Position Based) | 0.18
| Bosch | 2024-07-12 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00014
| Huahong Chengdu | 2024-07-12 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Intel Hillsboro#2 | 2024-07-12 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4e-05
| Bosch | 2024-07-12 | OpticalAxisShiftPerFocusY | 0.0016
| Huahong Chengdu | 2024-07-12 | OpticalAxisShiftPerFocusY | -0.0015
| Intel Hillsboro#2 | 2024-07-12 | OpticalAxisShiftPerFocusY | -0.0017
| Bosch | 2024-07-12 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 9e-05
| Huahong Chengdu | 2024-07-12 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00047
| Intel Hillsboro#2 | 2024-07-12 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| Bosch | 2024-07-12 | OpticalAxisShiftPerFocusX | -0.0035
| Huahong Chengdu | 2024-07-12 | OpticalAxisShiftPerFocusX | -0.0013
| Intel Hillsboro#2 | 2024-07-12 | OpticalAxisShiftPerFocusX | -0.0058
| Huahong Chengdu | 2024-07-12 | OPM Repeatability @5mm | 0.74
| Intel Hillsboro#2 | 2024-07-12 | OPM Repeatability @5mm | 1.2
| Huahong Chengdu | 2024-07-12 | OPM Repeatability @25mm | 9.88
| Intel Hillsboro#2 | 2024-07-12 | OPM Repeatability @25mm | 1.72
| Huahong Chengdu | 2024-07-12 | OPM Repeatability @1mm | 0.62
| Intel Hillsboro#2 | 2024-07-12 | OPM Repeatability @1mm | 1.07
| Huahong Chengdu | 2024-07-12 | OPM Repeatability @10mm | 1.96
| Intel Hillsboro#2 | 2024-07-12 | OPM Repeatability @10mm | 1.34
| Bosch | 2024-07-12 | NCM driving% with AC160TS cantilever | 6
| Huahong Chengdu | 2024-07-12 | NCM driving% with AC160TS cantilever | 8.1
| Intel Hillsboro#2 | 2024-07-12 | NCM driving% with AC160TS cantilever | 4.7
| Bosch | 2024-07-12 | Min. Vibration Transmissbility (~100Hz) | 0.015
| Huahong Chengdu | 2024-07-12 | Min. Vibration Transmissbility (~100Hz) | 0.047
| Intel Hillsboro#2 | 2024-07-12 | Min. Vibration Transmissbility (~100Hz) | 0.033
| Huahong Chengdu | 2024-07-12 | Max OPM @5mm | 14.02
| Intel Hillsboro#2 | 2024-07-12 | Max OPM @5mm | 29.63
| Huahong Chengdu | 2024-07-12 | Max OPM @25mm | 73.84
| Intel Hillsboro#2 | 2024-07-12 | Max OPM @25mm | 67.72
| Huahong Chengdu | 2024-07-12 | Max OPM @1mm | 5.23
| Intel Hillsboro#2 | 2024-07-12 | Max OPM @1mm | 12.49
| Huahong Chengdu | 2024-07-12 | Max OPM @10mm | 28.39
| Intel Hillsboro#2 | 2024-07-12 | Max OPM @10mm | 38.26
| Bosch | 2024-07-12 | LowerLimit | 2010
| Bosch | 2024-07-12 | LowerLimit | 11650
| Huahong Chengdu | 2024-07-12 | LowerLimit | 2122
| Huahong Chengdu | 2024-07-12 | LowerLimit | 5200
| Intel Hillsboro#2 | 2024-07-12 | LowerLimit | 1276.6
| Intel Hillsboro#2 | 2024-07-12 | LowerLimit | 5200
| Bosch | 2024-07-12 | Lateral(C-D) sensitivity | 31.8
| Huahong Chengdu | 2024-07-12 | Lateral(C-D) sensitivity | 0.659
| Intel Hillsboro#2 | 2024-07-12 | Lateral(C-D) sensitivity | 47.1
| Bosch | 2024-07-12 | Lateral(C-D) noise | 0.151
| Huahong Chengdu | 2024-07-12 | Lateral(C-D) noise | 0.756
| Intel Hillsboro#2 | 2024-07-12 | Lateral(C-D) noise | 0.176
| Bosch | 2024-07-12 | Image offset after moving from -Y to +Y (Y scan) | 62.26
| Huahong Chengdu | 2024-07-12 | Image offset after moving from -Y to +Y (Y scan) | 136.23
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from -Y to +Y (Y scan) | 28.56
| Bosch | 2024-07-12 | Image offset after moving from -Y to +Y (X scan) | 8.79
| Huahong Chengdu | 2024-07-12 | Image offset after moving from -Y to +Y (X scan) | 38.82
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from -Y to +Y (X scan) | 18.31
| Bosch | 2024-07-12 | Image offset after moving from -X to +X (Y scan) | 14.65
| Huahong Chengdu | 2024-07-12 | Image offset after moving from -X to +X (Y scan) | 21.97
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from -X to +X (Y scan) | 40.28
| Bosch | 2024-07-12 | Image offset after moving from -X to +X (X scan) | 46.14
| Huahong Chengdu | 2024-07-12 | Image offset after moving from -X to +X (X scan) | 13.18
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from -X to +X (X scan) | 37.35
| Bosch | 2024-07-12 | Image offset after moving from +Y to -Y (Y scan) | 87.89
| Huahong Chengdu | 2024-07-12 | Image offset after moving from +Y to -Y (Y scan) | 69.58
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from +Y to -Y (Y scan) | 95.94
| Bosch | 2024-07-12 | Image offset after moving from +Y to -Y (X scan) | 13.18
| Huahong Chengdu | 2024-07-12 | Image offset after moving from +Y to -Y (X scan) | 16.85
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from +Y to -Y (X scan) | 24.9
| Bosch | 2024-07-12 | Image offset after moving from +X to -X (Y scan) | 138.43
| Huahong Chengdu | 2024-07-12 | Image offset after moving from +X to -X (Y scan) | 9.52
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from +X to -X (Y scan) | 13.91
| Bosch | 2024-07-12 | Image offset after moving from +X to -X (X scan) | 55.66
| Huahong Chengdu | 2024-07-12 | Image offset after moving from +X to -X (X scan) | 9.52
| Intel Hillsboro#2 | 2024-07-12 | Image offset after moving from +X to -X (X scan) | 76.17
| Bosch | 2024-07-12 | Full range of Z Height during Z Scan sweep | 17.8
| Huahong Chengdu | 2024-07-12 | Full range of Z Height during Z Scan sweep | 17
| Intel Hillsboro#2 | 2024-07-12 | Full range of Z Height during Z Scan sweep | 40.4
| Bosch | 2024-07-12 | Full range of Z Detector during Z Scan sweep | 13.3
| Huahong Chengdu | 2024-07-12 | Full range of Z Detector during Z Scan sweep | 12.7
| Intel Hillsboro#2 | 2024-07-12 | Full range of Z Detector during Z Scan sweep | 13.6
| Bosch | 2024-07-12 | Full range of Y2 Detector Raw during Y Scan sweep | 13.45
| Huahong Chengdu | 2024-07-12 | Full range of Y2 Detector Raw during Y Scan sweep | 13.92
| Intel Hillsboro#2 | 2024-07-12 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| Bosch | 2024-07-12 | Full range of Y2 Detector during Y Scan sweep | 188
| Huahong Chengdu | 2024-07-12 | Full range of Y2 Detector during Y Scan sweep | 193.4
| Intel Hillsboro#2 | 2024-07-12 | Full range of Y2 Detector during Y Scan sweep | 180.6
| Bosch | 2024-07-12 | Full range of Y1 Detector Raw during Y Scan sweep | 13.46
| Huahong Chengdu | 2024-07-12 | Full range of Y1 Detector Raw during Y Scan sweep | 14.01
| Intel Hillsboro#2 | 2024-07-12 | Full range of Y1 Detector Raw during Y Scan sweep | 14.37
| Bosch | 2024-07-12 | Full range of Y1 Detector during Y Scan sweep | 179.9
| Huahong Chengdu | 2024-07-12 | Full range of Y1 Detector during Y Scan sweep | 198.9
| Intel Hillsboro#2 | 2024-07-12 | Full range of Y1 Detector during Y Scan sweep | 185.4
| Bosch | 2024-07-12 | Full range of Y Detector during Y Scan sweep | 179.9
| Huahong Chengdu | 2024-07-12 | Full range of Y Detector during Y Scan sweep | 193.4
| Intel Hillsboro#2 | 2024-07-12 | Full range of Y Detector during Y Scan sweep | 180.6
| Bosch | 2024-07-12 | Full range of X2 Detector Raw during X Scan sweep | 13.32
| Huahong Chengdu | 2024-07-12 | Full range of X2 Detector Raw during X Scan sweep | 13.92
| Intel Hillsboro#2 | 2024-07-12 | Full range of X2 Detector Raw during X Scan sweep | 14.64
| Bosch | 2024-07-12 | Full range of X2 Detector during X Scan sweep | 183.3
| Huahong Chengdu | 2024-07-12 | Full range of X2 Detector during X Scan sweep | 201.9
| Intel Hillsboro#2 | 2024-07-12 | Full range of X2 Detector during X Scan sweep | 191
| Bosch | 2024-07-12 | Full range of X1 Detector Raw during X Scan sweep | 13.66
| Huahong Chengdu | 2024-07-12 | Full range of X1 Detector Raw during X Scan sweep | 13.91
| Intel Hillsboro#2 | 2024-07-12 | Full range of X1 Detector Raw during X Scan sweep | 14.47
| Bosch | 2024-07-12 | Full range of X1 Detector during X Scan sweep | 185.9
| Huahong Chengdu | 2024-07-12 | Full range of X1 Detector during X Scan sweep | 199.3
| Intel Hillsboro#2 | 2024-07-12 | Full range of X1 Detector during X Scan sweep | 195.2
| Bosch | 2024-07-12 | Full range of X Detector during X Scan sweep | 185.9
| Huahong Chengdu | 2024-07-12 | Full range of X Detector during X Scan sweep | 201.9
| Intel Hillsboro#2 | 2024-07-12 | Full range of X Detector during X Scan sweep | 195.2
| Bosch | 2024-07-12 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Huahong Chengdu | 2024-07-12 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Intel Hillsboro#2 | 2024-07-12 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Huahong Chengdu | 2024-07-12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.33
| Intel Hillsboro#2 | 2024-07-12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.16
| Huahong Chengdu | 2024-07-12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| Intel Hillsboro#2 | 2024-07-12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.9
| Huahong Chengdu | 2024-07-12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| Intel Hillsboro#2 | 2024-07-12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| Huahong Chengdu | 2024-07-12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.33
| Intel Hillsboro#2 | 2024-07-12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Bosch | 2024-07-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Huahong Chengdu | 2024-07-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.41
| Intel Hillsboro#2 | 2024-07-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Bosch | 2024-07-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| Huahong Chengdu | 2024-07-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Intel Hillsboro#2 | 2024-07-12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| Bosch | 2024-07-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| Huahong Chengdu | 2024-07-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| Intel Hillsboro#2 | 2024-07-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| Bosch | 2024-07-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Huahong Chengdu | 2024-07-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.49
| Intel Hillsboro#2 | 2024-07-12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Bosch | 2024-07-12 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.6
| Huahong Chengdu | 2024-07-12 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.2
| Bosch | 2024-07-12 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 12.2
| Huahong Chengdu | 2024-07-12 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.5
| Bosch | 2024-07-12 | %GRR w/ roughness of non-patterened wafer | 13.4
| Huahong Chengdu | 2024-07-12 | %GRR w/ roughness of non-patterened wafer | 5
| 엘티씨에이엠㈜ | 2024-07-11 | Z Scanner position at -3μm | 0.0097
| 엘티씨에이엠㈜ | 2024-07-11 | Z Scanner position at +3μm | 0.0164
| 엘티씨에이엠㈜ | 2024-07-11 | Z linearity by sample height | 1.003
| 엘티씨에이엠㈜ | 2024-07-11 | Z Height noise | 0.032
| 엘티씨에이엠㈜ | 2024-07-11 | Y2 Detector noise (XY Servo on) | 0.27
| 엘티씨에이엠㈜ | 2024-07-11 | Y2 Detector noise (XY Servo off) | 0.3
| 엘티씨에이엠㈜ | 2024-07-11 | Y1 Detector noise (XY Servo on) | 0.27
| 엘티씨에이엠㈜ | 2024-07-11 | Y1 Detector noise (XY Servo off) | 0.24
| 엘티씨에이엠㈜ | 2024-07-11 | Y Scan Stroke | 195.2
| 엘티씨에이엠㈜ | 2024-07-11 | -Y Direction software Limit | 150753.5
| 엘티씨에이엠㈜ | 2024-07-11 | Y Detector noise (XY Servo on) | 0.21
| 엘티씨에이엠㈜ | 2024-07-11 | Y Detector noise (XY Servo off) | 0.18
| 엘티씨에이엠㈜ | 2024-07-11 | Y axis OPM | 3.38
| 엘티씨에이엠㈜ | 2024-07-11 | X2 Detector noise (XY Servo on) | 0.32
| 엘티씨에이엠㈜ | 2024-07-11 | X2 Detector noise (XY Servo off) | 0.33
| 엘티씨에이엠㈜ | 2024-07-11 | X1 Detector noise (XY Servo on) | 0.28
| 엘티씨에이엠㈜ | 2024-07-11 | X1 Detector noise (XY Servo off) | 0.3
| 엘티씨에이엠㈜ | 2024-07-11 | X Scan Stroke | 202.9
| 엘티씨에이엠㈜ | 2024-07-11 | -X Direction software Limit | 150259.5
| 엘티씨에이엠㈜ | 2024-07-11 | X Detector noise (XY Servo on) | 0.24
| 엘티씨에이엠㈜ | 2024-07-11 | X Detector noise (XY Servo off) | 0.26
| 엘티씨에이엠㈜ | 2024-07-11 | X axis OPM | 1.69
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.58
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.39
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.41
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.34
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.36
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.34
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| 엘티씨에이엠㈜ | 2024-07-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| 엘티씨에이엠㈜ | 2024-07-11 | NCM driving% with AC160TS cantilever | 0.7
| 엘티씨에이엠㈜ | 2024-07-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.088
| 엘티씨에이엠㈜ | 2024-07-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.146
| 엘티씨에이엠㈜ | 2024-07-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 276.1
| 엘티씨에이엠㈜ | 2024-07-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 271.8
| 엘티씨에이엠㈜ | 2024-07-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 291.6
| 엘티씨에이엠㈜ | 2024-07-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 278.4
| 엘티씨에이엠㈜ | 2024-07-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| 엘티씨에이엠㈜ | 2024-07-11 | Lateral(C-D) sensitivity | 14.5
| 엘티씨에이엠㈜ | 2024-07-11 | Lateral(C-D) noise | 0.164
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Z Height during Z Scan sweep | 17.9
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Z Detector during Z Scan sweep | 13.7
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Y2 Detector during Y Scan sweep | 193.4
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Y1 Detector during Y Scan sweep | 197.2
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of Y Detector during Y Scan sweep | 195.2
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of X2 Detector during X Scan sweep | 200
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of X1 Detector during X Scan sweep | 194.8
| 엘티씨에이엠㈜ | 2024-07-11 | Full range of X Detector during X Scan sweep | 197.4
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Z Scanner position at -3μm | 0.0012
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Z Scanner position at -3μm | 0.0004
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Z Scanner position at +3μm | 0.0007
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Z Scanner position at +3μm | 0.0087
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Z linearity by sample height | 1.013
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Z linearity by sample height | 0.999
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Z Height noise | 0.026
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Z Height noise | 0.029
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y2 Detector noise (XY Servo on) | 0.229
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y2 Detector noise (XY Servo off) | 0.226
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y1 Detector noise (XY Servo on) | 0.303
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y1 Detector noise (XY Servo off) | 0.292
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Y Scan Stroke | 184.9
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y Scan Stroke | 169.439
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Y Detector noise (XY Servo on) | 0.214
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y Detector noise (XY Servo on) | 0.195
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Y Detector noise (XY Servo off) | 0.217
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y Detector noise (XY Servo off) | 0.199
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Y axis OPM | 1.5822
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Y axis OPM | 0.6688
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X2 Detector noise (XY Servo on) | 0.241
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X2 Detector noise (XY Servo off) | 0.224
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X1 Detector noise (XY Servo on) | 0.331
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X1 Detector noise (XY Servo off) | 0.316
| 충남대학교 분석과학기술대학원 | 2024-07-09 | X Scan Stroke | 190
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X Scan Stroke | 153.008
| 충남대학교 분석과학기술대학원 | 2024-07-09 | X Detector noise (XY Servo on) | 0.24
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X Detector noise (XY Servo on) | 0.214
| 충남대학교 분석과학기술대학원 | 2024-07-09 | X Detector noise (XY Servo off) | 0.234
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X Detector noise (XY Servo off) | 0.22
| 충남대학교 분석과학기술대학원 | 2024-07-09 | X axis OPM | 1.127
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | X axis OPM | 1.551
| 충남대학교 분석과학기술대학원 | 2024-07-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.39
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 충남대학교 분석과학기술대학원 | 2024-07-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.26
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.28
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.27
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.25
| 충남대학교 분석과학기술대학원 | 2024-07-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 충남대학교 분석과학기술대학원 | 2024-07-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 충남대학교 분석과학기술대학원 | 2024-07-09 | NCM Amplitude noise | 16.805
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | NCM Amplitude noise | 14.38
| 충남대학교 분석과학기술대학원 | 2024-07-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.252
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.133
| 충남대학교 분석과학기술대학원 | 2024-07-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.458
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.308
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 230.4
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 229.1
| 충남대학교 분석과학기술대학원 | 2024-07-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 253.6
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 211.3
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 207.9
| 충남대학교 분석과학기술대학원 | 2024-07-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 265.5
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Maximum Y Scan range (Full Stroke) | 169.3
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Maximum X Scan range (Full Stroke) | 152.6
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Lateral(C-D) sensitivity | 14.4
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Lateral(C-D) sensitivity | 19.4
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Lateral(C-D) noise | 0.13
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Lateral(C-D) noise | 0.127
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Full range of Z Height during Z Scan sweep | 18
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Full range of Z Height during Z Scan sweep | 17.6
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Full range of Z Detector during Z Scan sweep | 14.2
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Full range of Z Detector during Z Scan sweep | 12.5
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Full range of Y2 Detector raw | 14.7
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Full range of Y1 Detector raw | 14.8
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Full range of Y Detector during Y Scan sweep | 184.1
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Full range of X2 Detector raw | 14.6
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | Full range of X1 Detector raw | 14.6
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Full range of X Detector Raw during X Scan sweep | 14.5
| 충남대학교 분석과학기술대학원 | 2024-07-09 | Full range of X Detector during X Scan sweep | 192.2
| 충남대학교 분석과학기술대학원 | 2024-07-09 | A-B noise | 7.894
| Sanan ST-Microelectronics Co., Ltd | 2024-07-09 | A-B noise | 6.33
| Kyocera SLD Laser, Inc | 2024-06-28 |  | 76810
| Kyocera SLD Laser, Inc | 2024-06-28 |  | -77214
| Kyocera SLD Laser, Inc | 2024-06-28 |  | 76936
| Kyocera SLD Laser, Inc | 2024-06-28 |  | -77047
| 에이앤아이 | 2024-06-28 | Z Scanner position at -3μm | 0.0101
| Kyocera SLD Laser, Inc | 2024-06-28 | Z Scanner position at -3μm | 0.0011
| University of Bath (Prof. Kamal) | 2024-06-28 | Z Scanner position at -3μm | 0.0046
| 에이앤아이 | 2024-06-28 | Z Scanner position at +3μm | 0.0037
| Kyocera SLD Laser, Inc | 2024-06-28 | Z Scanner position at +3μm | 0.0041
| University of Bath (Prof. Kamal) | 2024-06-28 | Z Scanner position at +3μm | 0.0119
| 에이앤아이 | 2024-06-28 | Z linearity by sample height | 1
| Kyocera SLD Laser, Inc | 2024-06-28 | Z linearity by sample height | 1.014
| University of Bath (Prof. Kamal) | 2024-06-28 | Z linearity by sample height | 1.011
| 에이앤아이 | 2024-06-28 | Z Height noise | 0.029
| Kyocera SLD Laser, Inc | 2024-06-28 | Z Height noise | 0.028
| University of Bath (Prof. Kamal) | 2024-06-28 | Z Height noise | 0.022
| 에이앤아이 | 2024-06-28 | Y2 Detector noise (XY Servo on) | 0.26
| Kyocera SLD Laser, Inc | 2024-06-28 | Y2 Detector noise (XY Servo on) | 0.28
| 에이앤아이 | 2024-06-28 | Y2 Detector noise (XY Servo off) | 0.26
| Kyocera SLD Laser, Inc | 2024-06-28 | Y2 Detector noise (XY Servo off) | 0.27
| 에이앤아이 | 2024-06-28 | Y1 Detector noise (XY servo on) | 0.26
| Kyocera SLD Laser, Inc | 2024-06-28 | Y1 Detector noise (XY servo on) | 0.33
| 에이앤아이 | 2024-06-28 | Y1 Detector noise (XY servo off) | 0.26
| Kyocera SLD Laser, Inc | 2024-06-28 | Y1 Detector noise (XY servo off) | 0.32
| 에이앤아이 | 2024-06-28 | Y Scan Stroke | 198.4
| Kyocera SLD Laser, Inc | 2024-06-28 | Y Scan Stroke | 185
| University of Bath (Prof. Kamal) | 2024-06-28 | Y Scan Stroke | 55
| 에이앤아이 | 2024-06-28 | -Y Direction software Limit | -150283
| 에이앤아이 | 2024-06-28 | Y Detector noise (XY servo on) | 0.2
| Kyocera SLD Laser, Inc | 2024-06-28 | Y Detector noise (XY servo on) | 0.23
| University of Bath (Prof. Kamal) | 2024-06-28 | Y Detector noise (XY Servo on) | 0.25
| 에이앤아이 | 2024-06-28 | Y Detector noise (XY servo off) | 0.19
| Kyocera SLD Laser, Inc | 2024-06-28 | Y Detector noise (XY servo off) | 0.22
| University of Bath (Prof. Kamal) | 2024-06-28 | Y Detector noise (XY Servo off) | 0.25
| 에이앤아이 | 2024-06-28 | Y axis OPM | 2.45
| Kyocera SLD Laser, Inc | 2024-06-28 | Y axis OPM | 0.88
| University of Bath (Prof. Kamal) | 2024-06-28 | Y axis OPM | 1.33
| 에이앤아이 | 2024-06-28 | X2 Detector noise (XY Servo on) | 0.27
| Kyocera SLD Laser, Inc | 2024-06-28 | X2 Detector noise (XY Servo on) | 0.26
| 에이앤아이 | 2024-06-28 | X2 Detector noise (XY Servo off) | 0.26
| Kyocera SLD Laser, Inc | 2024-06-28 | X2 Detector noise (XY Servo off) | 0.25
| 에이앤아이 | 2024-06-28 | X1 Detector noise (XY servo on) | 0.28
| Kyocera SLD Laser, Inc | 2024-06-28 | X1 Detector noise (XY servo on) | 0.27
| 에이앤아이 | 2024-06-28 | X1 Detector noise (XY servo off) | 0.27
| Kyocera SLD Laser, Inc | 2024-06-28 | X1 Detector noise (XY servo off) | 0.26
| 에이앤아이 | 2024-06-28 | X Scan Stroke | 194.6
| Kyocera SLD Laser, Inc | 2024-06-28 | X Scan Stroke | 182
| University of Bath (Prof. Kamal) | 2024-06-28 | X Scan Stroke | 55
| 에이앤아이 | 2024-06-28 | -X Direction software Limit | -150440
| 에이앤아이 | 2024-06-28 | X Detector noise (XY servo on) | 0.21
| Kyocera SLD Laser, Inc | 2024-06-28 | X Detector noise (XY servo on) | 0.19
| University of Bath (Prof. Kamal) | 2024-06-28 | X Detector noise (XY Servo on) | 0.25
| 에이앤아이 | 2024-06-28 | X Detector noise (XY servo off) | 0.21
| Kyocera SLD Laser, Inc | 2024-06-28 | X Detector noise (XY servo off) | 0.19
| University of Bath (Prof. Kamal) | 2024-06-28 | X Detector noise (XY Servo off) | 0.24
| 에이앤아이 | 2024-06-28 | X axis OPM | 1.94
| Kyocera SLD Laser, Inc | 2024-06-28 | X axis OPM | 0.83
| University of Bath (Prof. Kamal) | 2024-06-28 | X axis OPM | 0.56
| 에이앤아이 | 2024-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| University of Bath (Prof. Kamal) | 2024-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 에이앤아이 | 2024-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Bath (Prof. Kamal) | 2024-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.33
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.31
| 에이앤아이 | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.41
| 에이앤아이 | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.43
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.32
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.32
| 에이앤아이 | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.41
| 에이앤아이 | 2024-06-28 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.36
| 에이앤아이 | 2024-06-28 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.37
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| 에이앤아이 | 2024-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| University of Bath (Prof. Kamal) | 2024-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| 에이앤아이 | 2024-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Kyocera SLD Laser, Inc | 2024-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| University of Bath (Prof. Kamal) | 2024-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 에이앤아이 | 2024-06-28 | NCM driving% with AC160TS cantilever | 0.62
| Kyocera SLD Laser, Inc | 2024-06-28 | NCM driving% with AC160TS cantilever | 0.496
| University of Bath (Prof. Kamal) | 2024-06-28 | NCM driving% with AC160TS cantilever | 0.199
| 에이앤아이 | 2024-06-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.464
| Kyocera SLD Laser, Inc | 2024-06-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.871
| University of Bath (Prof. Kamal) | 2024-06-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.206
| 에이앤아이 | 2024-06-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.067
| Kyocera SLD Laser, Inc | 2024-06-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.702
| University of Bath (Prof. Kamal) | 2024-06-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.776
| 에이앤아이 | 2024-06-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 276.8
| Kyocera SLD Laser, Inc | 2024-06-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 266.3
| 에이앤아이 | 2024-06-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 284.1
| Kyocera SLD Laser, Inc | 2024-06-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 268.4
| University of Bath (Prof. Kamal) | 2024-06-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 108.7
| 에이앤아이 | 2024-06-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 266.3
| Kyocera SLD Laser, Inc | 2024-06-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 267.6
| 에이앤아이 | 2024-06-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 285.8
| Kyocera SLD Laser, Inc | 2024-06-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 265
| University of Bath (Prof. Kamal) | 2024-06-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.9
| 에이앤아이 | 2024-06-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Kyocera SLD Laser, Inc | 2024-06-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| 에이앤아이 | 2024-06-28 | Lateral(C-D) sensitivity | 16.3
| Kyocera SLD Laser, Inc | 2024-06-28 | Lateral(C-D) sensitivity | 12.4
| University of Bath (Prof. Kamal) | 2024-06-28 | Lateral(C-D) sensitivity | 21.1
| 에이앤아이 | 2024-06-28 | Lateral(C-D) noise | 0.175
| Kyocera SLD Laser, Inc | 2024-06-28 | Lateral(C-D) noise | 0.144
| University of Bath (Prof. Kamal) | 2024-06-28 | Lateral(C-D) noise | 0.103
| 에이앤아이 | 2024-06-28 | Full range of Z Height during Z Scan sweep | 16.7
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Z Height during Z Scan sweep | 17.9
| University of Bath (Prof. Kamal) | 2024-06-28 | Full range of Z Height during Z Scan sweep | 17.9
| 에이앤아이 | 2024-06-28 | Full range of Z Detector during Z Scan sweep | 12.4
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Z Detector during Z Scan sweep | 13.5
| University of Bath (Prof. Kamal) | 2024-06-28 | Full range of Z Detector during Z Scan sweep | 13.4
| 에이앤아이 | 2024-06-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| 에이앤아이 | 2024-06-28 | Full range of Y2 Detector during Y Scan sweep | 193.7
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Y2 Detector during Y Scan sweep | 186.2
| 에이앤아이 | 2024-06-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.3
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| 에이앤아이 | 2024-06-28 | Full range of Y1 Detector during Y Scan sweep | 200
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Y1 Detector during Y Scan sweep | 187.9
| University of Bath (Prof. Kamal) | 2024-06-28 | Full range of Y Detector Raw during Y Scan sweep | 14
| 에이앤아이 | 2024-06-28 | Full range of Y Detector during Y Scan sweep | 196.9
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of Y Detector during Y Scan sweep | 186.8
| University of Bath (Prof. Kamal) | 2024-06-28 | Full range of Y Detector during Y Scan sweep | 76.2
| 에이앤아이 | 2024-06-28 | Full range of X2 Detector Raw during X Scan sweep | 14
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of X2 Detector Raw during X Scan sweep | 13.8
| 에이앤아이 | 2024-06-28 | Full range of X2 Detector during X Scan sweep | 186
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of X2 Detector during X Scan sweep | 184.5
| 에이앤아이 | 2024-06-28 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| 에이앤아이 | 2024-06-28 | Full range of X1 Detector during X Scan sweep | 197.5
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of X1 Detector during X Scan sweep | 183.5
| University of Bath (Prof. Kamal) | 2024-06-28 | Full range of X Detector Raw during X Scan sweep | 14.4
| 에이앤아이 | 2024-06-28 | Full range of X Detector during X Scan sweep | 191.7
| Kyocera SLD Laser, Inc | 2024-06-28 | Full range of X Detector during X Scan sweep | 183.8
| University of Bath (Prof. Kamal) | 2024-06-28 | Full range of X Detector during X Scan sweep | 74.6
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Linearity by Z scanner position | 1.66
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Linearity by 4 different sample height | 0.99898
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Height noise | 0.042
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.8681
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.8443
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.7143
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.6812
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Drift | 3.821
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Z Detector offset | 19.55
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y2 Detector noise (XY Servo on) | 0.279
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 249.1393
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y1 Detector noise (XY Servo on) | 0.322
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 256.058
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 166.955
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y Ringing(Rpv) | 0.147
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y Non-linearity error | 2.94
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y Detector noise (XY Servo on) | 0.223
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Repeatability @1st position | 0.02
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Repeatability @ 4th position | 0.02
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Repeatability @ 3rd position | 0.03
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Repeatability @ 2nd position | 0.02
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Accuracy @1st position | 0.24
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Accuracy @ 4th position | 0.22
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Accuracy @ 3rd position | 0.23
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis, Accuracy @ 2nd position | 0.2
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Y axis OPM | 0.629
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Stage positioning Y offset (Vision Pattern Matching) | 3.42
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Stage positioning Y offset (LLC Translation) | 2.21
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Stage positioning Y offset (Global Align) | 4.7
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Stage positioning X offset (Vision Pattern Matching) | 2.2
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Stage positioning X offset (LLC Translation) | 2.3
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Stage positioning X offset (Global Align) | 3.11
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | XY Orthogonality | 89.9
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X2 Detector noise (XY Servo on) | 0.296
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 258.7886
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X1 Detector noise (XY Servo on) | 0.296
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 257.5073
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 169.89
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X Ringing(Rpv) | 0.127
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X Non-linearity error | 1.96
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X Detector noise (XY Servo on) | 0.229
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Repeatability @1st position | 0.03
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Repeatability @ 4th position | 0.04
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Repeatability @ 3rd position | 0.02
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Repeatability @ 2nd position | 0.06
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Accuracy @1st position | 0.18
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Accuracy @ 4th position | 0.23
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Accuracy @ 3rd position | 0.13
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis, Accuracy @ 2nd position | 0.26
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | X axis OPM | 0.912
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Wafer placement(Loading) error on the load port2 (Y offset) | 33.3
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Wafer placement(Loading) error on the load port2 (X offset) | 16.7
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Wafer placement(Loading) error on the load port1 (Y offset) | 26.67
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Wafer placement(Loading) error on the load port1 (X offset) | 45.1
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Torque Profile (50mm, 20um/s) | 3.324
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Torque Profile (50mm, 200um/s) | 2.059
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeYMin | -153223
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeYMax | 150292
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeY | 0
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeY | 0
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeY | 303515
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeXMin | -203564
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeXMax | 203908
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeX | 0
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeX | 0
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | StrokeX | 407472
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.682
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Roll motion to X+ full scan | 7.4
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | RMS of ZeroScan | 0.069
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Position repeatability(10repeats, Position Based) | 0.5
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OpticalAxisShiftPerFocusY | -0.0005
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00048
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OpticalAxisShiftPerFocusX | -0.0021
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OPM Repeatability @5mm | 0.893
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OPM Repeatability @25mm | 2.192
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OPM Repeatability @1mm | 0.533
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | OPM Repeatability @10mm | 3.51
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | NCM driving% with AC160TS cantilever | 2.03
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Min. Vibration Transmissbility (~100Hz) | 0.026
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Max OPM @5mm | 10.944
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Max OPM @25mm | 30.795
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Max OPM @1mm | 7.83
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Max OPM @10mm | 21.296
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | LowerLimit | 2211
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | LowerLimit | 4760
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Lateral(C-D) sensitivity | 10.6
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Lateral(C-D) noise | 0.245
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from -Y to +Y (Y scan) | 44.6777
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from -Y to +Y (X scan) | 66.6504
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from -X to +X (Y scan) | 35.1563
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from -X to +X (X scan) | 43.9453
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from +Y to -Y (Y scan) | 30.7617
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from +Y to -Y (X scan) | 48.3398
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from +X to -X (Y scan) | 27.0996
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Image offset after moving from +X to -X (X scan) | 42.4805
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of Z Height during Z Scan sweep | 15.6
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of Z Detector during Z Scan sweep | 12.51
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.55
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of Y2 Detector during Y Scan sweep | 179.69
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.17
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of Y1 Detector during Y Scan sweep | 180.74
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of X2 Detector Raw during X Scan sweep | 14.62
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of X2 Detector during X Scan sweep | 187.54
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of X1 Detector Raw during X Scan sweep | 14.45
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Full range of X1 Detector during X Scan sweep | 187.04
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.81
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.82
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.55
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.73
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.2
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.7
| TSMC F20 #3 (Hsinchu) | 2024-06-28 | %GRR w/ roughness of non-patterened wafer | 5
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Z Scanner position at -3μm | 0.0043
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Z Scanner position at +3μm | 0.0124
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Z linearity by sample height | 0.995
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Z Height noise | 0.025
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Y Scan Stroke | 55
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Y Detector noise (XY Servo on) | 0.25
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Y Detector noise (XY Servo off) | 0.25
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Y axis OPM | 0.76
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | X Scan Stroke | 55
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | X Detector noise (XY Servo on) | 0.25
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | X Detector noise (XY Servo off) | 0.25
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | X axis OPM | 0.59
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | NCM driving% with AC160TS cantilever | 0.391
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.457
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.169
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.5
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Lateral(C-D) sensitivity | 19.1
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Lateral(C-D) noise | 0.075
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Full range of Z Height during Z Scan sweep | 16.4
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Full range of Z Detector during Z Scan sweep | 13.6
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Full range of Y Detector during Y Scan sweep | 75.5
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Full range of X Detector Raw during X Scan sweep | 14.7
| 한국폴리텍대학교(안성, 나노측정과) | 2024-06-27 | Full range of X Detector during X Scan sweep | 76.8
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Linearity by Z scanner position | 0.62
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Linearity by Z scanner position | 0.59
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Linearity by 4 different sample height | 1.0034
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Linearity by 4 different sample height | 1
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Height noise | 0.026
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Height noise | 0.034
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.2899
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.1012
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.2141
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.0821
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5949
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.8664
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.5336
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6803
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Drift | 2.832
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Drift | 4.577
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Z Detector offset | 19
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Z Detector offset | 22.512
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y2 Detector noise (XY Servo on) | 0.374
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y2 Detector noise (XY Servo on) | 0.516
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.6341
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 221.5019
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y1 Detector noise (XY Servo on) | 0.382
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y1 Detector noise (XY Servo on) | 0.491
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 262.0146
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 233.0596
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 175.716
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 159.65
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y Ringing(Rpv) | 0.062
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y Ringing(Rpv) | 0.361
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y Non-linearity error | 2.84
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y Non-linearity error | 2.91
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y Detector noise (XY Servo on) | 0.325
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y Detector noise (XY Servo on) | 0.435
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Repeatability @1st position | 0.07
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Repeatability @1st position | 0.07
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Repeatability @ 4th position | 0.06
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Repeatability @ 4th position | 0.08
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Repeatability @ 3rd position | 0.06
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Repeatability @ 3rd position | 0.05
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Repeatability @ 2nd position | 0.05
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Repeatability @ 2nd position | 0.1
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Accuracy @1st position | 0.35
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Accuracy @1st position | 1.43
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Accuracy @ 4th position | 0.31
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Accuracy @ 4th position | 1.8
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Accuracy @ 3rd position | 0.26
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Accuracy @ 3rd position | 1.27
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis, Accuracy @ 2nd position | 0.27
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis, Accuracy @ 2nd position | 1.1
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Y axis OPM | 0.325
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Y axis OPM | 1.712
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Stage positioning Y offset (Vision Pattern Matching) | 3.64
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.06
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Stage positioning Y offset (LLC Translation) | 2.42
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Stage positioning Y offset (LLC Translation) | 1.33
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Stage positioning Y offset (Global Align) | 2.7344
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Stage positioning Y offset (Global Align) | 1.05
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Stage positioning X offset (Vision Pattern Matching) | 3.81
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Stage positioning X offset (Vision Pattern Matching) | 2.09
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Stage positioning X offset (LLC Translation) | 2.2
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Stage positioning X offset (LLC Translation) | 1.48
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Stage positioning X offset (Global Align) | 1.7969
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Stage positioning X offset (Global Align) | 1.58
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | XY Orthogonality | 90.1
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | XY Orthogonality | 90
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X2 Detector noise (XY Servo on) | 0.334
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X2 Detector noise (XY Servo on) | 0.415
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 283.9799
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 256.8832
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X1 Detector noise (XY Servo on) | 0.325
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X1 Detector noise (XY Servo on) | 0.411
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 283.2863
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.2106
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 175.281
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 177.224
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X Ringing(Rpv) | 0.095
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X Ringing(Rpv) | 0.316
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X Non-linearity error | 2.7
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X Non-linearity error | 2.72
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X Detector noise (XY Servo on) | 0.261
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X Detector noise (XY Servo on) | 0.307
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Repeatability @1st position | 0.06
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Repeatability @1st position | 0.04
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Repeatability @ 4th position | 0.06
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Repeatability @ 4th position | 0.06
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Repeatability @ 3rd position | 0.06
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Repeatability @ 3rd position | 0.04
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Repeatability @ 2nd position | 0.08
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Repeatability @ 2nd position | 0.07
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Accuracy @1st position | 0.27
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Accuracy @1st position | 0.23
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Accuracy @ 4th position | 0.26
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Accuracy @ 4th position | 0.33
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Accuracy @ 3rd position | 0.33
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Accuracy @ 3rd position | 0.45
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis, Accuracy @ 2nd position | 0.4
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis, Accuracy @ 2nd position | 0.33
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | X axis OPM | 0.262
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | X axis OPM | 0.544
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Wafer placement(Loading) error on the load port2 (Y offset) | 61.39
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Wafer placement(Loading) error on the load port2 (Y offset) | 19.1
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Wafer placement(Loading) error on the load port2 (X offset) | 42.85
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Wafer placement(Loading) error on the load port2 (X offset) | 21.8
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Wafer placement(Loading) error on the load port1 (Y offset) | 63.49
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Wafer placement(Loading) error on the load port1 (Y offset) | 47.55
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Wafer placement(Loading) error on the load port1 (X offset) | 43.46
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Wafer placement(Loading) error on the load port1 (X offset) | 42.36
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Torque Profile (50mm, 20um/s) | 2.526
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Torque Profile (50mm, 20um/s) | 0.802
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Torque Profile (50mm, 200um/s) | 2.297
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Torque Profile (50mm, 200um/s) | 0.564
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | StrokeYMin | -153987
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | StrokeYMin | -1153405
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | StrokeYMax | 150378.9
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | StrokeYMax | 1150941
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | StrokeY | 304365.9
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | StrokeY | 2304346
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | StrokeXMin | -203915
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | StrokeXMin | -1204117
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | StrokeXMax | 203031.5
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | StrokeXMax | 1205033
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | StrokeX | 406946.4
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | StrokeX | 2409150
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.805
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.707
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Roll motion to X+ full scan | 3.8
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Roll motion to X+ full scan | 0.3
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | RMS of ZeroScan | 0.057
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | RMS of ZeroScan | 0.047
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Position repeatability(10repeats, Position Based) | 0.78
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Position repeatability(10repeats, Position Based) | 0.56
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4e-05
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OpticalAxisShiftPerFocusY | 0.0006
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OpticalAxisShiftPerFocusY | 0.0018
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00015
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OpticalAxisShiftPerFocusX | -0.0052
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OpticalAxisShiftPerFocusX | -0.001
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OPM Repeatability @5mm | 0.677
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OPM Repeatability @5mm | 0.505
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OPM Repeatability @25mm | 1.123
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OPM Repeatability @25mm | 5.814
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OPM Repeatability @1mm | 0.619
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OPM Repeatability @1mm | 0.848
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | OPM Repeatability @10mm | 0.696
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | OPM Repeatability @10mm | 4.768
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | NCM driving% with AC160TS cantilever | 24.72
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | NCM driving% with AC160TS cantilever | 5.16
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Min. Vibration Transmissbility (~100Hz) | 0.027
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Min. Vibration Transmissbility (~100Hz) | 0.025
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Max OPM @5mm | 16.666
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Max OPM @5mm | 17.438
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Max OPM @25mm | 41.216
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Max OPM @25mm | 167.094
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Max OPM @1mm | 7.842
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Max OPM @1mm | 5.677
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Max OPM @10mm | 25.066
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Max OPM @10mm | 79.248
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | LowerLimit | 2223.81
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | LowerLimit | 4689
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | LowerLimit | 2460
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | LowerLimit | 5219
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Lateral(C-D) sensitivity | 17.5
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Lateral(C-D) sensitivity | 16.2
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Lateral(C-D) noise | 0.162
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Lateral(C-D) noise | 0.247
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from -Y to +Y (Y scan) | 32.959
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from -Y to +Y (Y scan) | 31.4941
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from -Y to +Y (X scan) | 45.4102
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from -Y to +Y (X scan) | 19.7754
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from -X to +X (Y scan) | 79.1016
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from -X to +X (Y scan) | 43.9453
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from -X to +X (X scan) | 29.2969
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from -X to +X (X scan) | 98.1445
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from +Y to -Y (Y scan) | 38.8184
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from +Y to -Y (Y scan) | 34.4238
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from +Y to -Y (X scan) | 25.6348
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from +Y to -Y (X scan) | 13.1836
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from +X to -X (Y scan) | 56.3965
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from +X to -X (Y scan) | 25.6348
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Image offset after moving from +X to -X (X scan) | 15.3809
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Image offset after moving from +X to -X (X scan) | 71.7773
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Z Height during Z Scan sweep | 17.3
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Z Height during Z Scan sweep | 0.032
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Z Height during Z Scan sweep | 18.2735
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Z Detector during Z Scan sweep | 13.78
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Z Detector during Z Scan sweep | 12.8
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Y2 Detector Raw during Y Scan sweep | 14.11
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Y2 Detector Raw during Y Scan sweep | 14.09
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Y2 Detector during Y Scan sweep | 213.03
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Y2 Detector during Y Scan sweep | 210.9574
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Y1 Detector Raw during Y Scan sweep | 14.42
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Y1 Detector Raw during Y Scan sweep | 14.06
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Y1 Detector during Y Scan sweep | 206.25
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Y1 Detector during Y Scan sweep | 204.3323
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of Y Detector during Y Scan sweep | 213
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of Y Detector during Y Scan sweep | 210.9574
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of X2 Detector Raw during X Scan sweep | 13.96
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of X2 Detector Raw during X Scan sweep | 14.56
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of X2 Detector during X Scan sweep | 215.59
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of X2 Detector during X Scan sweep | 203.9906
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of X1 Detector Raw during X Scan sweep | 14.06
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of X1 Detector Raw during X Scan sweep | 14.43
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of X1 Detector during X Scan sweep | 221.95
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of X1 Detector during X Scan sweep | 195.2669
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Full range of X Detector during X Scan sweep | 221.95
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Full range of X Detector during X Scan sweep | 204.3321
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.53
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.21
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 1.1
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 3.7
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.3
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 1.5
| TSMC F20 #2 (Hsinchu) | 2024-06-27 | %GRR w/ roughness of non-patterened wafer | 2.6
| Rong Semiconductor (Huaian) Corporation | 2024-06-27 | %GRR w/ roughness of non-patterened wafer | 4
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Linearity by Z scanner position | 0.41
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Height noise | 0.058
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 58.78
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 58.77
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 26.11
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 26.16
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Drift | 5.71
| Qorvo #2(Thomas Russel) | 2024-06-26 | Z Detector offset | 31.8
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y2 Detector noise (XY Servo on) | 0.23
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 261.6
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y1 Detector noise (XY Servo on) | 0.243
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 255.5
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 182.1
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y Ringing(Rpv) | 0.08
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y Non-linearity error | 1.81
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y Detector noise (XY Servo on) | 0.243
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Repeatability @1st position | 0.03
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Repeatability @ 4th position | 0.04
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Repeatability @ 3rd position | 0.04
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Repeatability @ 2nd position | 0.05
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Accuracy @1st position | 0.17
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Accuracy @ 4th position | 0.2
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Accuracy @ 3rd position | 0.21
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis, Accuracy @ 2nd position | 0.22
| Qorvo #2(Thomas Russel) | 2024-06-26 | Y axis OPM | 0.61
| Qorvo #2(Thomas Russel) | 2024-06-26 | XY Orthogonality | 90
| Qorvo #2(Thomas Russel) | 2024-06-26 | X2 Detector noise (XY Servo on) | 0.267
| Qorvo #2(Thomas Russel) | 2024-06-26 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 256.4
| Qorvo #2(Thomas Russel) | 2024-06-26 | X1 Detector noise (XY Servo on) | 0.242
| Qorvo #2(Thomas Russel) | 2024-06-26 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 258.3
| Qorvo #2(Thomas Russel) | 2024-06-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 169.6
| Qorvo #2(Thomas Russel) | 2024-06-26 | X Ringing(Rpv) | 0.27
| Qorvo #2(Thomas Russel) | 2024-06-26 | X Non-linearity error | 2.23
| Qorvo #2(Thomas Russel) | 2024-06-26 | X Detector noise (XY Servo on) | 0.242
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Repeatability @1st position | 0.02
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Repeatability @ 4th position | 0.03
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Repeatability @ 3rd position | 0.04
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Repeatability @ 2nd position | 0.04
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Accuracy @1st position | 0.12
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Accuracy @ 4th position | 0.14
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Accuracy @ 3rd position | 0.2
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis, Accuracy @ 2nd position | 0.18
| Qorvo #2(Thomas Russel) | 2024-06-26 | X axis OPM | 0.69
| Qorvo #2(Thomas Russel) | 2024-06-26 | Wafer placement(Loading) error on the load port2 (Y offset) | 41.27
| Qorvo #2(Thomas Russel) | 2024-06-26 | Wafer placement(Loading) error on the load port2 (X offset) | 47.65
| Qorvo #2(Thomas Russel) | 2024-06-26 | Wafer placement(Loading) error on the load port1 (Y offset) | 29.52
| Qorvo #2(Thomas Russel) | 2024-06-26 | Wafer placement(Loading) error on the load port1 (X offset) | 23.65
| Qorvo #2(Thomas Russel) | 2024-06-26 | Torque Profile (50mm, 20um/s) | 2.275
| Qorvo #2(Thomas Russel) | 2024-06-26 | Torque Profile (50mm, 200um/s) | 2.353
| Qorvo #2(Thomas Russel) | 2024-06-26 | StrokeYMin | -101857
| Qorvo #2(Thomas Russel) | 2024-06-26 | StrokeYMax | 104302
| Qorvo #2(Thomas Russel) | 2024-06-26 | StrokeY | 206159
| Qorvo #2(Thomas Russel) | 2024-06-26 | StrokeXMin | -155896
| Qorvo #2(Thomas Russel) | 2024-06-26 | StrokeXMax | 152427
| Qorvo #2(Thomas Russel) | 2024-06-26 | StrokeX | 308323
| Qorvo #2(Thomas Russel) | 2024-06-26 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.496
| Qorvo #2(Thomas Russel) | 2024-06-26 | Roll motion to X+ full scan | 15.8
| Qorvo #2(Thomas Russel) | 2024-06-26 | RMS of ZeroScan | 0.024
| Qorvo #2(Thomas Russel) | 2024-06-26 | Position repeatability(10repeats, Position Based) | 0.08
| Qorvo #2(Thomas Russel) | 2024-06-26 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00014
| Qorvo #2(Thomas Russel) | 2024-06-26 | OpticalAxisShiftPerFocusY | 0.0016
| Qorvo #2(Thomas Russel) | 2024-06-26 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 9.19e-05
| Qorvo #2(Thomas Russel) | 2024-06-26 | OpticalAxisShiftPerFocusX | -0.0035
| Qorvo #2(Thomas Russel) | 2024-06-26 | OPM Repeatability @5mm | 0.76
| Qorvo #2(Thomas Russel) | 2024-06-26 | OPM Repeatability @25mm | 5.64
| Qorvo #2(Thomas Russel) | 2024-06-26 | OPM Repeatability @1mm | 0.58
| Qorvo #2(Thomas Russel) | 2024-06-26 | OPM Repeatability @10mm | 1.91
| Qorvo #2(Thomas Russel) | 2024-06-26 | NCM driving% with AC160TS cantilever | 12.45
| Qorvo #2(Thomas Russel) | 2024-06-26 | Min. Vibration Transmissbility (~100Hz) | 0.016
| Qorvo #2(Thomas Russel) | 2024-06-26 | Max OPM @5mm | 11.15
| Qorvo #2(Thomas Russel) | 2024-06-26 | Max OPM @25mm | 99.28
| Qorvo #2(Thomas Russel) | 2024-06-26 | Max OPM @1mm | 8.29
| Qorvo #2(Thomas Russel) | 2024-06-26 | Max OPM @10mm | 18.76
| Qorvo #2(Thomas Russel) | 2024-06-26 | LowerLimit | 2315
| Qorvo #2(Thomas Russel) | 2024-06-26 | LowerLimit | 4435
| Qorvo #2(Thomas Russel) | 2024-06-26 | Lateral(C-D) sensitivity | 31.8
| Qorvo #2(Thomas Russel) | 2024-06-26 | Lateral(C-D) noise | 0.151
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from -Y to +Y (Y scan) | 14.64
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from -Y to +Y (X scan) | 8.78
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from -X to +X (Y scan) | 62.25
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from -X to +X (X scan) | 46.14
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from +Y to -Y (Y scan) | 138.42
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from +Y to -Y (X scan) | 13.18
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from +X to -X (Y scan) | 87.89
| Qorvo #2(Thomas Russel) | 2024-06-26 | Image offset after moving from +X to -X (X scan) | 55.66
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Z Height during Z Scan sweep | 41
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Z Detector during Z Scan sweep | 13.9
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Y2 Detector Raw during Y Scan sweep | 13.67
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Y2 Detector during Y Scan sweep | 168.3
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Y1 Detector Raw during Y Scan sweep | 13.87
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Y1 Detector during Y Scan sweep | 187.6
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of Y Detector during Y Scan sweep | 187.6
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of X2 Detector Raw during X Scan sweep | 13.75
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of X2 Detector during X Scan sweep | 183.9
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of X1 Detector during X Scan sweep | 191.5
| Qorvo #2(Thomas Russel) | 2024-06-26 | Full range of X Detector during X Scan sweep | 191.5
| Qorvo #2(Thomas Russel) | 2024-06-26 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.54
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| Qorvo #2(Thomas Russel) | 2024-06-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Qorvo #2(Thomas Russel) | 2024-06-26 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 9.8
| Qorvo #2(Thomas Russel) | 2024-06-26 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.3
| Qorvo #2(Thomas Russel) | 2024-06-26 | %GRR w/ roughness of non-patterened wafer | 10.3
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Z Scanner position at -3μm | 0.0099
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Z Scanner position at -3μm | 0.0013
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Z Scanner position at +3μm | 0.0128
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Z Scanner position at +3μm | 0.0078
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Z linearity by sample height | 0.999
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Z linearity by sample height | 0.995
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Z Height noise | 0.028
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Z Height noise | 0.028
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y2 Detector noise (XY Servo on) | 0.3
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y2 Detector noise (XY Servo off) | 0.29
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y1 Detector noise (XY servo on) | 0.26
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y1 Detector noise (XY servo off) | 0.25
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Y Scan Stroke | 55
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y Scan Stroke | 177.7
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Y Detector noise (XY Servo on) | 0.26
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y Detector noise (XY servo on) | 0.2
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Y Detector noise (XY Servo off) | 0.25
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y Detector noise (XY servo off) | 0.2
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Y axis OPM | 0.85
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Y axis OPM | 0.98
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X2 Detector noise (XY Servo on) | 0.32
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X2 Detector noise (XY Servo off) | 0.31
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X1 Detector noise (XY servo on) | 0.23
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X1 Detector noise (XY servo off) | 0.22
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | X Scan Stroke | 55
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X Scan Stroke | 169.3
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | X Detector noise (XY Servo on) | 0.27
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X Detector noise (XY servo on) | 0.21
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | X Detector noise (XY Servo off) | 0.26
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X Detector noise (XY servo off) | 0.2
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | X axis OPM | 1.13
| Konfoong Materials International Co., Ltd. | 2024-06-25 | X axis OPM | 0.97
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.27
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.27
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.28
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Konfoong Materials International Co., Ltd. | 2024-06-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | NCM driving% with AC160TS cantilever | 1.58
| Konfoong Materials International Co., Ltd. | 2024-06-25 | NCM driving% with AC160TS cantilever | 0.98
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.112
| Konfoong Materials International Co., Ltd. | 2024-06-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.128
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.714
| Konfoong Materials International Co., Ltd. | 2024-06-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.687
| Konfoong Materials International Co., Ltd. | 2024-06-25 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 250.6
| Konfoong Materials International Co., Ltd. | 2024-06-25 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 236.6
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.2
| Konfoong Materials International Co., Ltd. | 2024-06-25 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 233.8
| Konfoong Materials International Co., Ltd. | 2024-06-25 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100.2
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Lateral(C-D) sensitivity | 18.9
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Lateral(C-D) noise | 0.125
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Full range of Z Height during Z Scan sweep | 16
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Z Height during Z Scan sweep | 16.1
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Full range of Z Detector during Z Scan sweep | 12.4
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Z Detector during Z Scan sweep | 11.6
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Y2 Detector during Y Scan sweep | 181.5
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.7
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Y1 Detector during Y Scan sweep | 173.8
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Full range of Y Detector during Y Scan sweep | 76.4
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of Y Detector during Y Scan sweep | 177.7
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of X2 Detector during X Scan sweep | 169.3
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of X1 Detector during X Scan sweep | 169.1
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Full range of X Detector Raw during X Scan sweep | 15.2
| National Research Council, Institute for microelectronics and microsystems (CNR IMM, Giannazzo) (Gembetti) | 2024-06-25 | Full range of X Detector during X Scan sweep | 76.2
| Konfoong Materials International Co., Ltd. | 2024-06-25 | Full range of X Detector during X Scan sweep | 169.2
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Linearity by Z scanner position | 1.28
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Linearity by Z scanner position | 1.28
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Linearity by 4 different sample height | 1.0062
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Linearity by 4 different sample height | 1.0062
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Height noise | 0.028
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Height noise | 0.028
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 23.7127
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 23.7127
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.5446
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.5446
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8279
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8279
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.8767
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.8767
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Drift | 5.6829
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Drift | 5.6829
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Detector offset | 20.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Z Detector offset | 20.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y2 Detector noise (XY Servo on) | 0.379
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y2 Detector noise (XY Servo on) | 0.379
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 254.9199
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 254.9199
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y1 Detector noise (XY Servo on) | 0.36
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y1 Detector noise (XY Servo on) | 0.36
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 271.101
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 271.101
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 175.616
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 175.616
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Ringing(Rpv) | 0.109
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Ringing(Rpv) | 0.109
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Non-linearity error | 2.96
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Non-linearity error | 2.96
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Detector noise (XY Servo on) | 0.282
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y Detector noise (XY Servo on) | 0.282
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @1st position | 0.06
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @1st position | 0.06
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @ 4th position | 0.03
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @ 4th position | 0.03
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @ 3rd position | 0.03
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @ 3rd position | 0.03
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @ 2nd position | 0.04
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Repeatability @ 2nd position | 0.04
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @1st position | 0.44
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @1st position | 0.44
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @ 4th position | 0.49
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @ 4th position | 0.49
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @ 3rd position | 0.37
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @ 3rd position | 0.37
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @ 2nd position | 0.35
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis, Accuracy @ 2nd position | 0.35
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis OPM | 1.422
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Y axis OPM | 1.422
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.33
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Stage positioning Y offset (LLC Translation) | 1.99
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Stage positioning Y offset (Global Align) | 6.28
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Stage positioning X offset (Vision Pattern Matching) | 2.59
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Stage positioning X offset (LLC Translation) | 1.53
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Stage positioning X offset (Global Align) | 3.19
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Orthogonality | 90
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | XY Orthogonality | 90
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X2 Detector noise (XY Servo on) | 0.444
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X2 Detector noise (XY Servo on) | 0.444
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 273.6736
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 273.6736
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X1 Detector noise (XY Servo on) | 0.372
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X1 Detector noise (XY Servo on) | 0.372
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 279.7258
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 279.7258
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 176.763
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 176.763
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Ringing(Rpv) | 0.086
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Ringing(Rpv) | 0.086
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Non-linearity error | 2.12
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Non-linearity error | 2.12
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Detector noise (XY Servo on) | 0.31
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X Detector noise (XY Servo on) | 0.31
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @1st position | 0.04
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @1st position | 0.04
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @ 4th position | 0.03
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @ 4th position | 0.03
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @ 3rd position | 0.06
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @ 3rd position | 0.06
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @ 2nd position | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Repeatability @ 2nd position | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @1st position | 0.22
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @1st position | 0.22
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @ 4th position | 0.17
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @ 4th position | 0.17
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @ 3rd position | 0.29
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @ 3rd position | 0.29
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @ 2nd position | 0.22
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis, Accuracy @ 2nd position | 0.22
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis OPM | 0.755
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | X axis OPM | 0.755
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port2 (Y offset) | 0.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port2 (Y offset) | 0.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port2 (X offset) | 1.25
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port2 (X offset) | 1.25
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port1 (Y offset) | 0.85
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port1 (Y offset) | 0.85
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port1 (X offset) | 1.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Wafer placement(Loading) error on the load port1 (X offset) | 1.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Torque Profile (50mm, 20um/s) | 2.145
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Torque Profile (50mm, 20um/s) | 2.145
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Torque Profile (50mm, 200um/s) | 2.628
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Torque Profile (50mm, 200um/s) | 2.628
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeYMin | -152840
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeYMin | -152840
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeYMax | 152210.5
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeYMax | 152210.5
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeY | 305050
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeY | 305050
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeXMin | -205415
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeXMin | -205415
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeXMax | 203917.6
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeXMax | 203917.6
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeX | 409332.2
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | StrokeX | 409332.2
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.372
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.372
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Roll motion to X+ full scan | 13.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Roll motion to X+ full scan | 13.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | RMS of ZeroScan | 0.039
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | RMS of ZeroScan | 0.039
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Position repeatability(10repeats, Position Based) | 0.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Position repeatability(10repeats, Position Based) | 0.8
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.000213
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.000213
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusY | 0.0029
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusY | 0.0029
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000166
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000166
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusX | -0.0041
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OpticalAxisShiftPerFocusX | -0.0041
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @5mm | 0.576
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @5mm | 0.576
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @25mm | 4.074
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @25mm | 4.074
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @1mm | 0.533
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @1mm | 0.533
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @10mm | 1.273
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | OPM Repeatability @10mm | 1.273
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | NCM driving% with AC160TS cantilever | 40.52
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | NCM driving% with AC160TS cantilever | 40.52
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Min. Vibration Transmissbility (~100Hz) | 0.028
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Min. Vibration Transmissbility (~100Hz) | 0.028
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @5mm | 10.877
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @5mm | 10.877
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @25mm | 79.177
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @25mm | 79.177
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @1mm | 3.93
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @1mm | 3.93
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @10mm | 31.709
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Max OPM @10mm | 31.709
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | LowerLimit | 2000
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | LowerLimit | 4430
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | LowerLimit | 2000
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | LowerLimit | 4430
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Lateral(C-D) sensitivity | 14.4
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Lateral(C-D) sensitivity | 14.4
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Lateral(C-D) noise | 0.216
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Lateral(C-D) noise | 0.216
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -Y to +Y (Y scan) | 24.1699
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -Y to +Y (Y scan) | 24.1699
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -Y to +Y (X scan) | 27.832
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -Y to +Y (X scan) | 27.832
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -X to +X (Y scan) | 17.5781
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -X to +X (Y scan) | 17.5781
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -X to +X (X scan) | 27.0996
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from -X to +X (X scan) | 27.0996
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +Y to -Y (Y scan) | 71.7773
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +Y to -Y (Y scan) | 71.7773
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +Y to -Y (X scan) | 8.7891
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +Y to -Y (X scan) | 8.7891
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +X to -X (Y scan) | 22.7051
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +X to -X (Y scan) | 22.7051
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +X to -X (X scan) | 44.6777
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Image offset after moving from +X to -X (X scan) | 44.6777
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Z Height during Z Scan sweep | 16.4
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Z Height during Z Scan sweep | 16.4
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Z Detector during Z Scan sweep | 13.87
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Z Detector during Z Scan sweep | 13.87
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.76
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.76
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y2 Detector during Y Scan sweep | 211.56
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y2 Detector during Y Scan sweep | 211.56
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y1 Detector during Y Scan sweep | 215.65
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y1 Detector during Y Scan sweep | 215.65
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y Detector during Y Scan sweep | 211.56
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of Y Detector during Y Scan sweep | 211.56
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X2 Detector Raw during X Scan sweep | 14.39
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X2 Detector Raw during X Scan sweep | 14.39
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X2 Detector during X Scan sweep | 216.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X2 Detector during X Scan sweep | 216.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X1 Detector Raw during X Scan sweep | 14.28
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X1 Detector Raw during X Scan sweep | 14.28
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X1 Detector during X Scan sweep | 222.4
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X1 Detector during X Scan sweep | 222.4
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X Detector during X Scan sweep | 216.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Full range of X Detector during X Scan sweep | 216.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 0.6
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 0.6
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.3
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.3
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | %GRR w/ roughness of non-patterened wafer | 4.3
| Guangzhou Zen Semiconductor Corporation | 2024-06-24 | %GRR w/ roughness of non-patterened wafer | 4.3
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Z Scanner position at -3μm | 0.0039
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Z Scanner position at -3μm | 0.0016
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Z Scanner position at +3μm | 0.0143
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Z Scanner position at +3μm | 0.0022
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Z linearity by sample height | 0.996
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Z linearity by sample height | 1.012
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Z Height noise | 0.024
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Z Height noise | 0.03
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y2 Detector noise (XY Servo on) | 0.22
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y2 Detector noise (XY Servo off) | 0.21
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y1 Detector noise (XY Servo on) | 0.27
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y1 Detector noise (XY Servo off) | 0.25
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y Scan Stroke | 183.8
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Y Scan Stroke | 112.6
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y Detector noise (XY Servo on) | 0.18
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Y Detector noise (XY Servo on) | 0.26
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y Detector noise (XY Servo off) | 0.16
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Y Detector noise (XY Servo off) | 0.26
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Y axis OPM | 0.78
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Y axis OPM | 0.45
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X2 Detector noise (XY Servo on) | 0.22
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X2 Detector noise (XY Servo off) | 0.21
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X1 Detector noise (XY Servo on) | 0.25
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X1 Detector noise (XY Servo off) | 0.23
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X Scan Stroke | 166.1
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | X Scan Stroke | 110
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X Detector noise (XY Servo on) | 0.19
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | X Detector noise (XY Servo on) | 0.26
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X Detector noise (XY Servo off) | 0.17
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | X Detector noise (XY Servo off) | 0.26
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | X axis OPM | 1.58
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | X axis OPM | 0.7
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.16
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.16
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.14
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.13
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.13
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | NCM Amplitude noise | 11.1
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | NCM Amplitude noise | 13.24
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.728
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.314
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.78
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.285
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 246.3
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 267.9
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 158.5
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 240.8
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 230.5
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156.9
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Maximum Y Scan range (Full Stroke) | 182
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Maximum X Scan range (Full Stroke) | 164.9
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Lateral(C-D) sensitivity | 11.6
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Lateral(C-D) sensitivity | 15.2
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Lateral(C-D) noise | 0.105
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Lateral(C-D) noise | 0.09
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Full range of Z Height during Z Scan sweep | 17.6
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Full range of Z Height during Z Scan sweep | 18.2
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Full range of Z Detector during Z Scan sweep | 13.4
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Full range of Y2 Detector raw | 14
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Full range of Y1 Detector raw | 14.3
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Full range of Y Detector during Y Scan sweep | 113.8
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Full range of X2 Detector raw | 14
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | Full range of X1 Detector raw | 14.1
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Full range of X Detector Raw during X Scan sweep | 14.3
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | Full range of X Detector during X Scan sweep | 112
| Park Systems Inc. (PSI) SID No. - U34002-300424 | 2024-06-21 | A-B noise | 13.12
| Park Systems Inc. (PSI) SID No. - U10086-170424 | 2024-06-21 | A-B noise | 8.17
| Tsinghua University | 2024-06-20 | Z Scanner position at -3μm | 0
| Tsinghua University | 2024-06-20 | Z Scanner position at +3μm | 0.0022
| Tsinghua University | 2024-06-20 | Z linearity by sample height | 1.012
| Tsinghua University | 2024-06-20 | Z Height noise | 0.03
| Tsinghua University | 2024-06-20 | Y Detector noise (XY Servo on) | 0.29
| Tsinghua University | 2024-06-20 | Y Detector noise (XY Servo off) | 0.29
| Tsinghua University | 2024-06-20 | Y axis OPM | 1.13
| Tsinghua University | 2024-06-20 | X Detector noise (XY Servo on) | 0.26
| Tsinghua University | 2024-06-20 | X Detector noise (XY Servo off) | 0.26
| Tsinghua University | 2024-06-20 | X axis OPM | 0.96
| Tsinghua University | 2024-06-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Tsinghua University | 2024-06-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Tsinghua University | 2024-06-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Tsinghua University | 2024-06-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Tsinghua University | 2024-06-20 | NCM Amplitude noise | 13.74
| Tsinghua University | 2024-06-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.168
| Tsinghua University | 2024-06-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.328
| Tsinghua University | 2024-06-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.4
| Tsinghua University | 2024-06-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 166.5
| Tsinghua University | 2024-06-20 | Lateral(C-D) sensitivity | 13.2
| Tsinghua University | 2024-06-20 | Lateral(C-D) noise | 0.106
| Tsinghua University | 2024-06-20 | Full range of Z Height during Z Scan sweep | 18.4
| Tsinghua University | 2024-06-20 | Full range of Z Detector during Z Scan sweep | 12.9
| Tsinghua University | 2024-06-20 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Tsinghua University | 2024-06-20 | Full range of Y Detector during Y Scan sweep | 120.6
| Tsinghua University | 2024-06-20 | Full range of X Detector Raw during X Scan sweep | 14.7
| Tsinghua University | 2024-06-20 | Full range of X Detector during X Scan sweep | 122
| Tsinghua University | 2024-06-20 | A-B noise | 7.38
| Sanya Research Institute of Hainan University | 2024-06-18 | Z Scanner position at -3μm | 0.0011
| Sanya Research Institute of Hainan University | 2024-06-18 | Z Scanner position at +3μm | 0.0033
| Sanya Research Institute of Hainan University | 2024-06-18 | Z linearity by sample height | 1.005
| Sanya Research Institute of Hainan University | 2024-06-18 | Z Height noise | 0.027
| Sanya Research Institute of Hainan University | 2024-06-18 | Y2 Detector noise (XY Servo on) | 0.21
| Sanya Research Institute of Hainan University | 2024-06-18 | Y2 Detector noise (XY Servo off) | 0.21
| Sanya Research Institute of Hainan University | 2024-06-18 | Y1 Detector noise (XY Servo on) | 0.22
| Sanya Research Institute of Hainan University | 2024-06-18 | Y1 Detector noise (XY Servo off) | 0.22
| Sanya Research Institute of Hainan University | 2024-06-18 | Y Scan Stroke | 165.9
| Sanya Research Institute of Hainan University | 2024-06-18 | Y Detector noise (XY Servo on) | 0.16
| Sanya Research Institute of Hainan University | 2024-06-18 | Y Detector noise (XY Servo off) | 0.15
| Sanya Research Institute of Hainan University | 2024-06-18 | Y axis OPM | 1.67
| Sanya Research Institute of Hainan University | 2024-06-18 | X2 Detector noise (XY Servo on) | 0.23
| Sanya Research Institute of Hainan University | 2024-06-18 | X2 Detector noise (XY Servo off) | 0.22
| Sanya Research Institute of Hainan University | 2024-06-18 | X1 Detector noise (XY Servo on) | 0.23
| Sanya Research Institute of Hainan University | 2024-06-18 | X1 Detector noise (XY Servo off) | 0.22
| Sanya Research Institute of Hainan University | 2024-06-18 | X Scan Stroke | 167.8
| Sanya Research Institute of Hainan University | 2024-06-18 | X Detector noise (XY Servo on) | 0.17
| Sanya Research Institute of Hainan University | 2024-06-18 | X Detector noise (XY Servo off) | 0.16
| Sanya Research Institute of Hainan University | 2024-06-18 | X axis OPM | 1.15
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.34
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.33
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.26
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.27
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Sanya Research Institute of Hainan University | 2024-06-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Sanya Research Institute of Hainan University | 2024-06-18 | NCM Amplitude noise | 15.83
| Sanya Research Institute of Hainan University | 2024-06-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.073
| Sanya Research Institute of Hainan University | 2024-06-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.98
| Sanya Research Institute of Hainan University | 2024-06-18 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 229
| Sanya Research Institute of Hainan University | 2024-06-18 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 236.7
| Sanya Research Institute of Hainan University | 2024-06-18 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 229.5
| Sanya Research Institute of Hainan University | 2024-06-18 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 233.3
| Sanya Research Institute of Hainan University | 2024-06-18 | Maximum Y Scan range (Full Stroke) | 165.8
| Sanya Research Institute of Hainan University | 2024-06-18 | Maximum X Scan range (Full Stroke) | 167.6
| Sanya Research Institute of Hainan University | 2024-06-18 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Sanya Research Institute of Hainan University | 2024-06-18 | Lateral(C-D) sensitivity | 13.9
| Sanya Research Institute of Hainan University | 2024-06-18 | Lateral(C-D) noise | 0.116
| Sanya Research Institute of Hainan University | 2024-06-18 | Full range of Z Height during Z Scan sweep | 17.4
| Sanya Research Institute of Hainan University | 2024-06-18 | Full range of Z Detector during Z Scan sweep | 13.4
| Sanya Research Institute of Hainan University | 2024-06-18 | Full range of Y2 Detector raw | 14
| Sanya Research Institute of Hainan University | 2024-06-18 | Full range of Y1 Detector raw | 14.4
| Sanya Research Institute of Hainan University | 2024-06-18 | Full range of X2 Detector raw | 14.2
| Sanya Research Institute of Hainan University | 2024-06-18 | Full range of X1 Detector raw | 14.7
| Sanya Research Institute of Hainan University | 2024-06-18 | A-B noise | 7.29
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Linearity by Z scanner position | 0.48
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Linearity by 4 different sample height | 1.001
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Height noise | 0.028
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 23.7035
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.679
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4424
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4947
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Drift | 14.09
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Z Detector offset | 113.3
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y2 Detector noise (XY Servo on) | 0.15
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 269.2055
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y1 Detector noise (XY Servo on) | 0.154
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 258.9749
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 173.61
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y Ringing(Rpv) | 0.125
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y Non-linearity error | 3.33
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y Detector noise (XY Servo on) | 0.107
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Repeatability @1st position | 0.02
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Repeatability @ 4th position | 0.03
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Repeatability @ 3rd position | 0.02
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Repeatability @ 2nd position | 0.02
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Accuracy @1st position | 0.19
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Accuracy @ 4th position | 0.21
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Accuracy @ 3rd position | 0.16
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis, Accuracy @ 2nd position | 0.19
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Y axis OPM | 0.416
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.22
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Stage positioning Y offset (LLC Translation) | 9.18
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Stage positioning Y offset (Global Align) | 6.19
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Stage positioning X offset (Vision Pattern Matching) | 2.39
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Stage positioning X offset (LLC Translation) | 2.8
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Stage positioning X offset (Global Align) | 2.65
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | XY Orthogonality | 90
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X2 Detector noise (XY Servo on) | 0.162
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 270.9235
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X1 Detector noise (XY Servo on) | 0.149
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 273.2301
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 174.039
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X Ringing(Rpv) | 0.11
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X Non-linearity error | 3.14
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X Detector noise (XY Servo on) | 0.111
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Repeatability @1st position | 0.04
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Repeatability @ 4th position | 0.04
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Repeatability @ 3rd position | 0.03
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Repeatability @ 2nd position | 0.03
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Accuracy @1st position | 0.3
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Accuracy @ 4th position | 0.19
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Accuracy @ 3rd position | 0.19
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis, Accuracy @ 2nd position | 0.12
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | X axis OPM | 0.616
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Wafer placement(Loading) error on the load port2 (Y offset) | 31.47
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Wafer placement(Loading) error on the load port2 (X offset) | 25.79
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Wafer placement(Loading) error on the load port1 (Y offset) | 28.96
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Wafer placement(Loading) error on the load port1 (X offset) | 41.49
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Torque Profile (50mm, 20um/s) | 2.313
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Torque Profile (50mm, 200um/s) | 1.611
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | StrokeYMin | -151640
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | StrokeYMax | 153525
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | StrokeY | 305164.8
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | StrokeXMin | -204720
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | StrokeXMax | 205132.7
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | StrokeX | 409852.7
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.534
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Roll motion to X+ full scan | 1
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | RMS of ZeroScan | 0.054
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Position repeatability(10repeats, Position Based) | 0.52
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.0002
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OpticalAxisShiftPerFocusY | -0.0016
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00094
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OpticalAxisShiftPerFocusX | -0.0003
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OPM Repeatability @5mm | 1.596
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OPM Repeatability @25mm | 2.182
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OPM Repeatability @1mm | 0.911
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | OPM Repeatability @10mm | 1.578
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | NCM driving% with AC160TS cantilever | 12.97
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Min. Vibration Transmissbility (~100Hz) | 0.023
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Max OPM @5mm | 10.921
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Max OPM @25mm | 76.119
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Max OPM @1mm | 6.117
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Max OPM @10mm | 32.337
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | LowerLimit | 2087.6
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | LowerLimit | 5200
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Lateral(C-D) sensitivity | 17
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Lateral(C-D) noise | 0.187
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from -Y to +Y (Y scan) | 87.15
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from -Y to +Y (X scan) | 19.04
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from -X to +X (Y scan) | 32.222
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from -X to +X (X scan) | 36.62
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from +Y to -Y (Y scan) | 65.91
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from +Y to -Y (X scan) | 19.04
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from +X to -X (Y scan) | 31.49
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Image offset after moving from +X to -X (X scan) | 26.36
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Z Height during Z Scan sweep | 20.9439
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Z Detector during Z Scan sweep | 14.69
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Y2 Detector during Y Scan sweep | 214
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Y1 Detector during Y Scan sweep | 205
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of Y Detector during Y Scan sweep | 205
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of X2 Detector Raw during X Scan sweep | 13.7
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of X2 Detector during X Scan sweep | 211
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of X1 Detector during X Scan sweep | 210
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Full range of X Detector during X Scan sweep | 210
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.57
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.57
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.79
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 0
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.7
| TSMC F20 #1 (Hsinchu) | 2024-06-17 | %GRR w/ roughness of non-patterened wafer | 3.7
| PSCN Guangzhou Demo | 2024-06-14 | Z Scanner position at -3μm | 0.0067
| 서울과학기술대학교 | 2024-06-14 | Z Scanner position at -3μm | 0.003
| PSCN Guangzhou Demo | 2024-06-14 | Z Scanner position at +3μm | 0.0114
| 서울과학기술대학교 | 2024-06-14 | Z Scanner position at +3μm | 0.0178
| PSCN Guangzhou Demo | 2024-06-14 | Z linearity by sample height | 1.019
| 서울과학기술대학교 | 2024-06-14 | Z linearity by sample height | 1.013
| PSCN Guangzhou Demo | 2024-06-14 | Z Height noise | 0.026
| 서울과학기술대학교 | 2024-06-14 | Z Height noise | 0.028
| PSCN Guangzhou Demo | 2024-06-14 | Y2 Detector noise (XY Servo on) | 0.25
| PSCN Guangzhou Demo | 2024-06-14 | Y2 Detector noise (XY Servo off) | 0.25
| PSCN Guangzhou Demo | 2024-06-14 | Y1 Detector noise (XY Servo on) | 0.24
| PSCN Guangzhou Demo | 2024-06-14 | Y1 Detector noise (XY Servo off) | 0.24
| PSCN Guangzhou Demo | 2024-06-14 | Y Scan Stroke | 177.2
| 서울과학기술대학교 | 2024-06-14 | Y Scan Stroke | 55
| PSCN Guangzhou Demo | 2024-06-14 | Y Detector noise (XY Servo on) | 0.17
| 서울과학기술대학교 | 2024-06-14 | Y Detector noise (XY Servo on) | 0.38
| PSCN Guangzhou Demo | 2024-06-14 | Y Detector noise (XY Servo off) | 0.17
| 서울과학기술대학교 | 2024-06-14 | Y Detector noise (XY Servo off) | 0.42
| PSCN Guangzhou Demo | 2024-06-14 | Y axis OPM | 0.66
| 서울과학기술대학교 | 2024-06-14 | Y axis OPM | 0.85
| PSCN Guangzhou Demo | 2024-06-14 | X2 Detector noise (XY Servo on) | 0.31
| PSCN Guangzhou Demo | 2024-06-14 | X2 Detector noise (XY Servo off) | 0.29
| PSCN Guangzhou Demo | 2024-06-14 | X1 Detector noise (XY Servo on) | 0.27
| PSCN Guangzhou Demo | 2024-06-14 | X1 Detector noise (XY Servo off) | 0.26
| PSCN Guangzhou Demo | 2024-06-14 | X Scan Stroke | 179.5
| 서울과학기술대학교 | 2024-06-14 | X Scan Stroke | 55
| PSCN Guangzhou Demo | 2024-06-14 | X Detector noise (XY Servo on) | 0.23
| 서울과학기술대학교 | 2024-06-14 | X Detector noise (XY Servo on) | 0.37
| PSCN Guangzhou Demo | 2024-06-14 | X Detector noise (XY Servo off) | 0.22
| 서울과학기술대학교 | 2024-06-14 | X Detector noise (XY Servo off) | 0.4
| PSCN Guangzhou Demo | 2024-06-14 | X axis OPM | 1.08
| 서울과학기술대학교 | 2024-06-14 | X axis OPM | 0.89
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 서울과학기술대학교 | 2024-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 서울과학기술대학교 | 2024-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.34
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.28
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 서울과학기술대학교 | 2024-06-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| PSCN Guangzhou Demo | 2024-06-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| 서울과학기술대학교 | 2024-06-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| 서울과학기술대학교 | 2024-06-14 | NCM driving% with AC160TS cantilever | 0.223
| PSCN Guangzhou Demo | 2024-06-14 | NCM Amplitude noise | 16.01
| PSCN Guangzhou Demo | 2024-06-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.665
| 서울과학기술대학교 | 2024-06-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.996
| PSCN Guangzhou Demo | 2024-06-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.527
| 서울과학기술대학교 | 2024-06-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.906
| PSCN Guangzhou Demo | 2024-06-14 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 260.1
| PSCN Guangzhou Demo | 2024-06-14 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 250.7
| 서울과학기술대학교 | 2024-06-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102
| PSCN Guangzhou Demo | 2024-06-14 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 256.5
| PSCN Guangzhou Demo | 2024-06-14 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 263.9
| 서울과학기술대학교 | 2024-06-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98
| PSCN Guangzhou Demo | 2024-06-14 | Maximum Y Scan range (Full Stroke) | 177.5
| PSCN Guangzhou Demo | 2024-06-14 | Maximum X Scan range (Full Stroke) | 179.3
| PSCN Guangzhou Demo | 2024-06-14 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| PSCN Guangzhou Demo | 2024-06-14 | Lateral(C-D) sensitivity | 19.4
| PSCN Guangzhou Demo | 2024-06-14 | Lateral(C-D) noise | 0.122
| PSCN Guangzhou Demo | 2024-06-14 | Full range of Z Height during Z Scan sweep | 16.7
| 서울과학기술대학교 | 2024-06-14 | Full range of Z Height during Z Scan sweep | 17.7
| PSCN Guangzhou Demo | 2024-06-14 | Full range of Z Detector during Z Scan sweep | 13.6
| 서울과학기술대학교 | 2024-06-14 | Full range of Z Detector during Z Scan sweep | 14.2
| PSCN Guangzhou Demo | 2024-06-14 | Full range of Y2 Detector raw | 13.7
| PSCN Guangzhou Demo | 2024-06-14 | Full range of Y1 Detector raw | 14.1
| 서울과학기술대학교 | 2024-06-14 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| 서울과학기술대학교 | 2024-06-14 | Full range of Y Detector during Y Scan sweep | 67.9
| PSCN Guangzhou Demo | 2024-06-14 | Full range of X2 Detector raw | 13.9
| PSCN Guangzhou Demo | 2024-06-14 | Full range of X1 Detector raw | 13.7
| 서울과학기술대학교 | 2024-06-14 | Full range of X Detector Raw during X Scan sweep | 14.1
| 서울과학기술대학교 | 2024-06-14 | Full range of X Detector during X Scan sweep | 70.7
| PSCN Guangzhou Demo | 2024-06-14 | A-B noise | 6.25
| Chongqing San'an Semiconductor | 2024-06-11 | Z Scanner position at -3μm | 0.0065
| Chongqing San'an Semiconductor | 2024-06-11 | Z Scanner position at +3μm | 0.006
| Chongqing San'an Semiconductor | 2024-06-11 | Z linearity by sample height | 1.011
| Chongqing San'an Semiconductor | 2024-06-11 | Z Height noise | 0.032
| Chongqing San'an Semiconductor | 2024-06-11 | Y2 Detector noise (XY Servo on) | 0.26
| Chongqing San'an Semiconductor | 2024-06-11 | Y2 Detector noise (XY Servo off) | 0.27
| Chongqing San'an Semiconductor | 2024-06-11 | Y1 Detector noise (XY Servo on) | 0.25
| Chongqing San'an Semiconductor | 2024-06-11 | Y1 Detector noise (XY Servo off) | 0.25
| Chongqing San'an Semiconductor | 2024-06-11 | Y Scan Stroke | 179.7
| Chongqing San'an Semiconductor | 2024-06-11 | Y Detector noise (XY Servo on) | 0.18
| Chongqing San'an Semiconductor | 2024-06-11 | Y Detector noise (XY Servo off) | 0.19
| Chongqing San'an Semiconductor | 2024-06-11 | Y axis OPM | 1.35
| Chongqing San'an Semiconductor | 2024-06-11 | X2 Detector noise (XY Servo on) | 0.24
| Chongqing San'an Semiconductor | 2024-06-11 | X2 Detector noise (XY Servo off) | 0.24
| Chongqing San'an Semiconductor | 2024-06-11 | X1 Detector noise (XY Servo on) | 0.3
| Chongqing San'an Semiconductor | 2024-06-11 | X1 Detector noise (XY Servo off) | 0.3
| Chongqing San'an Semiconductor | 2024-06-11 | X Scan Stroke | 175
| Chongqing San'an Semiconductor | 2024-06-11 | X Detector noise (XY Servo on) | 0.21
| Chongqing San'an Semiconductor | 2024-06-11 | X Detector noise (XY Servo off) | 0.2
| Chongqing San'an Semiconductor | 2024-06-11 | X axis OPM | 0.8
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.31
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.35
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.3
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Chongqing San'an Semiconductor | 2024-06-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Chongqing San'an Semiconductor | 2024-06-11 | NCM Amplitude noise | 12.18
| Chongqing San'an Semiconductor | 2024-06-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.421
| Chongqing San'an Semiconductor | 2024-06-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.784
| Chongqing San'an Semiconductor | 2024-06-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 253.9
| Chongqing San'an Semiconductor | 2024-06-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 250.5
| Chongqing San'an Semiconductor | 2024-06-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 245.8
| Chongqing San'an Semiconductor | 2024-06-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 240.3
| Chongqing San'an Semiconductor | 2024-06-11 | Maximum Y Scan range (Full Stroke) | 175.3
| Chongqing San'an Semiconductor | 2024-06-11 | Maximum X Scan range (Full Stroke) | 175.3
| Chongqing San'an Semiconductor | 2024-06-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Chongqing San'an Semiconductor | 2024-06-11 | Lateral(C-D) sensitivity | 12.4
| Chongqing San'an Semiconductor | 2024-06-11 | Lateral(C-D) noise | 0.145
| Chongqing San'an Semiconductor | 2024-06-11 | Full range of Z Height during Z Scan sweep | 16.5
| Chongqing San'an Semiconductor | 2024-06-11 | Full range of Z Detector during Z Scan sweep | 11.5
| Chongqing San'an Semiconductor | 2024-06-11 | Full range of Y2 Detector raw | 14.1
| Chongqing San'an Semiconductor | 2024-06-11 | Full range of Y1 Detector raw | 14.4
| Chongqing San'an Semiconductor | 2024-06-11 | Full range of X2 Detector raw | 14.4
| Chongqing San'an Semiconductor | 2024-06-11 | Full range of X1 Detector raw | 14.4
| Chongqing San'an Semiconductor | 2024-06-11 | A-B noise | 13.52
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Z Scanner position at -3μm | 0.006
| GRG Test Group | 2024-06-05 | Z Scanner position at -3μm | 0.0041
| University of Texas (Austin) | 2024-06-05 | Z Scanner position at -3μm | 0.002
| GRG Test Group | 2024-06-05 | Z Scanner position at +3μm | 0.0046
| University of Texas (Austin) | 2024-06-05 | Z Scanner position at +3μm | 0.0023
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Z Scanner position at +3μm | 0.0025
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Z linearity by sample height | 0.988
| GRG Test Group | 2024-06-05 | Z linearity by sample height | 1.001
| University of Texas (Austin) | 2024-06-05 | Z linearity by sample height | 1.006
| GRG Test Group | 2024-06-05 | Z Linearity by 3 different sample height | 1.012
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Z Height noise | 0.027
| GRG Test Group | 2024-06-05 | Z Height noise | 0.029
| University of Texas (Austin) | 2024-06-05 | Z Height noise | 0.032
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Y Scan Stroke | 55
| GRG Test Group | 2024-06-05 | Y Scan Stroke | 117.6
| University of Texas (Austin) | 2024-06-05 | Y Scan Stroke | 55
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Y Detector noise (XY Servo on) | 0.43
| GRG Test Group | 2024-06-05 | Y Detector noise (XY Servo on) | 0.27
| University of Texas (Austin) | 2024-06-05 | Y Detector noise (XY Servo on) | 0.26
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Y Detector noise (XY Servo off) | 0.46
| GRG Test Group | 2024-06-05 | Y Detector noise (XY Servo off) | 0.26
| University of Texas (Austin) | 2024-06-05 | Y Detector noise (XY Servo off) | 0.25
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Y axis OPM | 0.87
| GRG Test Group | 2024-06-05 | Y axis OPM | 1.73
| University of Texas (Austin) | 2024-06-05 | Y axis OPM | 1.5
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | X Scan Stroke | 55
| GRG Test Group | 2024-06-05 | X Scan Stroke | 114.6
| University of Texas (Austin) | 2024-06-05 | X Scan Stroke | 55
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | X Detector noise (XY Servo on) | 0.44
| GRG Test Group | 2024-06-05 | X Detector noise (XY Servo on) | 0.26
| University of Texas (Austin) | 2024-06-05 | X Detector noise (XY Servo on) | 0.26
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | X Detector noise (XY Servo off) | 0.48
| GRG Test Group | 2024-06-05 | X Detector noise (XY Servo off) | 0.26
| University of Texas (Austin) | 2024-06-05 | X Detector noise (XY Servo off) | 0.25
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | X axis OPM | 0.7
| GRG Test Group | 2024-06-05 | X axis OPM | 1.66
| University of Texas (Austin) | 2024-06-05 | X axis OPM | 1.34
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| GRG Test Group | 2024-06-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| University of Texas (Austin) | 2024-06-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| GRG Test Group | 2024-06-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| University of Texas (Austin) | 2024-06-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| GRG Test Group | 2024-06-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| University of Texas (Austin) | 2024-06-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| GRG Test Group | 2024-06-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| University of Texas (Austin) | 2024-06-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | NCM Amplitude noise | 15.58
| GRG Test Group | 2024-06-05 | NCM Amplitude noise | 3.42
| University of Texas (Austin) | 2024-06-05 | NCM Amplitude noise | 14.5
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.294
| GRG Test Group | 2024-06-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.318
| GRG Test Group | 2024-06-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 26.015
| University of Texas (Austin) | 2024-06-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.244
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.919
| GRG Test Group | 2024-06-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.043
| GRG Test Group | 2024-06-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 69.152
| University of Texas (Austin) | 2024-06-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.1
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 97
| GRG Test Group | 2024-06-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 167.5
| University of Texas (Austin) | 2024-06-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.9
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.7
| GRG Test Group | 2024-06-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167.4
| University of Texas (Austin) | 2024-06-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.6
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Lateral(C-D) sensitivity | 17.8
| GRG Test Group | 2024-06-05 | Lateral(C-D) sensitivity | 19.3
| University of Texas (Austin) | 2024-06-05 | Lateral(C-D) sensitivity | 58.939
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Lateral(C-D) noise | 0.063
| GRG Test Group | 2024-06-05 | Lateral(C-D) noise | 0.101
| University of Texas (Austin) | 2024-06-05 | Lateral(C-D) noise | 0.022
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Full range of Z Height during Z Scan sweep | 18
| GRG Test Group | 2024-06-05 | Full range of Z Height during Z Scan sweep | 18
| GRG Test Group | 2024-06-05 | Full range of Z Height during Z Scan sweep | 41.6
| University of Texas (Austin) | 2024-06-05 | Full range of Z Height during Z Scan sweep | 18.1
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Full range of Z Detector during Z Scan sweep | 14.5
| GRG Test Group | 2024-06-05 | Full range of Z Detector during Z Scan sweep | 12.8
| GRG Test Group | 2024-06-05 | Full range of Z Detector during Z Scan sweep | 12
| University of Texas (Austin) | 2024-06-05 | Full range of Z Detector during Z Scan sweep | 13.3
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| GRG Test Group | 2024-06-05 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| University of Texas (Austin) | 2024-06-05 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Full range of Y Detector during Y Scan sweep | 71.8
| GRG Test Group | 2024-06-05 | Full range of Y Detector during Y Scan sweep | 117.7
| University of Texas (Austin) | 2024-06-05 | Full range of Y Detector during Y Scan sweep | 75.5
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Full range of X Detector Raw during X Scan sweep | 13.8
| GRG Test Group | 2024-06-05 | Full range of X Detector Raw during X Scan sweep | 13.7
| University of Texas (Austin) | 2024-06-05 | Full range of X Detector Raw during X Scan sweep | 14.6
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | Full range of X Detector during X Scan sweep | 70.1
| GRG Test Group | 2024-06-05 | Full range of X Detector during X Scan sweep | 114.6
| University of Texas (Austin) | 2024-06-05 | Full range of X Detector during X Scan sweep | 75
| GRG Test Group | 2024-06-05 | Approach position at -5μm | 0.0071
| GRG Test Group | 2024-06-05 | Approach position at +5μm | 0.0079
| University of Electronic Science and Technology of China_NX7 | 2024-06-05 | A-B noise | 8.88
| GRG Test Group | 2024-06-05 | A-B noise | 11.35
| University of Texas (Austin) | 2024-06-05 | A-B noise | 9.29
| Siltronic Singapore | 2024-05-31 | Z Linearity by Z scanner position | 1.65
| Siltronic Singapore | 2024-05-31 | Z Linearity by 4 different sample height | 1.0027
| Siltronic Singapore | 2024-05-31 | Z Height noise | 0.024
| Siltronic Singapore | 2024-05-31 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 23.8232
| Siltronic Singapore | 2024-05-31 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.8232
| Siltronic Singapore | 2024-05-31 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2974
| Siltronic Singapore | 2024-05-31 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2436
| Siltronic Singapore | 2024-05-31 | Z Drift | 7.572
| Siltronic Singapore | 2024-05-31 | Z Detector offset | 8.5
| Siltronic Singapore | 2024-05-31 | Y2 Detector noise (XY Servo on) | 0.423
| Siltronic Singapore | 2024-05-31 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 273.784
| Siltronic Singapore | 2024-05-31 | Y1 Detector noise (XY Servo on) | 0.47
| Siltronic Singapore | 2024-05-31 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 287.4114
| Siltronic Singapore | 2024-05-31 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 170.922
| Siltronic Singapore | 2024-05-31 | Y Ringing(Rpv) | 0.132
| Siltronic Singapore | 2024-05-31 | Y Non-linearity error | 1.86
| Siltronic Singapore | 2024-05-31 | Y Detector noise (XY Servo on) | 0.334
| Siltronic Singapore | 2024-05-31 | Y axis, Repeatability @1st position | 0.03
| Siltronic Singapore | 2024-05-31 | Y axis, Repeatability @ 4th position | 0.02
| Siltronic Singapore | 2024-05-31 | Y axis, Repeatability @ 3rd position | 0.03
| Siltronic Singapore | 2024-05-31 | Y axis, Repeatability @ 2nd position | 0.02
| Siltronic Singapore | 2024-05-31 | Y axis, Accuracy @1st position | 0.17
| Siltronic Singapore | 2024-05-31 | Y axis, Accuracy @ 4th position | 0.19
| Siltronic Singapore | 2024-05-31 | Y axis, Accuracy @ 3rd position | 0.2
| Siltronic Singapore | 2024-05-31 | Y axis, Accuracy @ 2nd position | 0.18
| Siltronic Singapore | 2024-05-31 | Y axis OPM | 1.265
| Siltronic Singapore | 2024-05-31 | XY Orthogonality | 89.8
| Siltronic Singapore | 2024-05-31 | X2 Detector noise (XY Servo on) | 0.521
| Siltronic Singapore | 2024-05-31 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 294.2037
| Siltronic Singapore | 2024-05-31 | X1 Detector noise (XY Servo on) | 0.491
| Siltronic Singapore | 2024-05-31 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 273.5288
| Siltronic Singapore | 2024-05-31 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 178.333
| Siltronic Singapore | 2024-05-31 | X Ringing(Rpv) | 0.1
| Siltronic Singapore | 2024-05-31 | X Non-linearity error | 2.41
| Siltronic Singapore | 2024-05-31 | X Detector noise (XY Servo on) | 0.408
| Siltronic Singapore | 2024-05-31 | X axis, Repeatability @1st position | 0.04
| Siltronic Singapore | 2024-05-31 | X axis, Repeatability @ 4th position | 0.03
| Siltronic Singapore | 2024-05-31 | X axis, Repeatability @ 3rd position | 0.03
| Siltronic Singapore | 2024-05-31 | X axis, Repeatability @ 2nd position | 0.03
| Siltronic Singapore | 2024-05-31 | X axis, Accuracy @1st position | 0.22
| Siltronic Singapore | 2024-05-31 | X axis, Accuracy @ 4th position | 0.16
| Siltronic Singapore | 2024-05-31 | X axis, Accuracy @ 3rd position | 0.16
| Siltronic Singapore | 2024-05-31 | X axis, Accuracy @ 2nd position | 0.17
| Siltronic Singapore | 2024-05-31 | X axis OPM | 0.577
| Siltronic Singapore | 2024-05-31 | Wafer placement(Loading) error on the load port2 (Y offset) | 15.22
| Siltronic Singapore | 2024-05-31 | Wafer placement(Loading) error on the load port2 (X offset) | 29.29
| Siltronic Singapore | 2024-05-31 | Wafer placement(Loading) error on the load port1 (Y offset) | 60.37
| Siltronic Singapore | 2024-05-31 | Wafer placement(Loading) error on the load port1 (X offset) | 13.35
| Siltronic Singapore | 2024-05-31 | StrokeYMin | -152549
| Siltronic Singapore | 2024-05-31 | StrokeYMax | 150705.3
| Siltronic Singapore | 2024-05-31 | StrokeY | 303253.8
| Siltronic Singapore | 2024-05-31 | StrokeXMin | -204535
| Siltronic Singapore | 2024-05-31 | StrokeXMax | 203098.9
| Siltronic Singapore | 2024-05-31 | StrokeX | 407634
| Siltronic Singapore | 2024-05-31 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.267
| Siltronic Singapore | 2024-05-31 | RMS of ZeroScan | 0.059
| Siltronic Singapore | 2024-05-31 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| Siltronic Singapore | 2024-05-31 | OpticalAxisShiftPerFocusY | -0.0008
| Siltronic Singapore | 2024-05-31 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00015
| Siltronic Singapore | 2024-05-31 | OpticalAxisShiftPerFocusX | -0.0037
| Siltronic Singapore | 2024-05-31 | NCM driving% with AC160TS cantilever | 9.12
| Siltronic Singapore | 2024-05-31 | Min. Vibration Transmissbility (~100Hz) | 0.021
| Siltronic Singapore | 2024-05-31 | LowerLimit | 1900
| Siltronic Singapore | 2024-05-31 | LowerLimit | 5260
| Siltronic Singapore | 2024-05-31 | Lateral(C-D) sensitivity | 22.7
| Siltronic Singapore | 2024-05-31 | Lateral(C-D) noise | 0.105
| Siltronic Singapore | 2024-05-31 | Image offset after moving from -Y to +Y (Y scan) | 54.9316
| Siltronic Singapore | 2024-05-31 | Image offset after moving from -Y to +Y (X scan) | 7.3242
| Siltronic Singapore | 2024-05-31 | Image offset after moving from -X to +X (Y scan) | 48.3398
| Siltronic Singapore | 2024-05-31 | Image offset after moving from -X to +X (X scan) | 41.748
| Siltronic Singapore | 2024-05-31 | Image offset after moving from +Y to -Y (Y scan) | 63.7207
| Siltronic Singapore | 2024-05-31 | Image offset after moving from +Y to -Y (X scan) | 16.1133
| Siltronic Singapore | 2024-05-31 | Image offset after moving from +X to -X (Y scan) | 26.3672
| Siltronic Singapore | 2024-05-31 | Image offset after moving from +X to -X (X scan) | 32.959
| Siltronic Singapore | 2024-05-31 | Full range of Z Height during Z Scan sweep | 19.01
| Siltronic Singapore | 2024-05-31 | Full range of Z Detector during Z Scan sweep | 14.11
| Siltronic Singapore | 2024-05-31 | Full range of Y2 Detector Raw during Y Scan sweep | 14.11
| Siltronic Singapore | 2024-05-31 | Full range of Y2 Detector during Y Scan sweep | 213.4
| Siltronic Singapore | 2024-05-31 | Full range of Y1 Detector Raw during Y Scan sweep | 14.32
| Siltronic Singapore | 2024-05-31 | Full range of Y1 Detector during Y Scan sweep | 199.8
| Siltronic Singapore | 2024-05-31 | Full range of Y Detector during Y Scan sweep | 206.6
| Siltronic Singapore | 2024-05-31 | Full range of X2 Detector Raw during X Scan sweep | 14.15
| Siltronic Singapore | 2024-05-31 | Full range of X2 Detector during X Scan sweep | 218.8
| Siltronic Singapore | 2024-05-31 | Full range of X1 Detector Raw during X Scan sweep | 14.28
| Siltronic Singapore | 2024-05-31 | Full range of X1 Detector during X Scan sweep | 221.2
| Siltronic Singapore | 2024-05-31 | Full range of X Detector during X Scan sweep | 220
| Siltronic Singapore | 2024-05-31 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Siltronic Singapore | 2024-05-31 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Siltronic Singapore | 2024-05-31 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.32
| Siltronic Singapore | 2024-05-31 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| Siltronic Singapore | 2024-05-31 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Siltronic Singapore | 2024-05-31 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.4
| Siltronic Singapore | 2024-05-31 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 7
| Siltronic Singapore | 2024-05-31 | %GRR w/ roughness of non-patterened wafer | 8.9
| 한국표준과학연구원 | 2024-05-30 | Z Scanner position at -3μm | 0.0054
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Z Scanner position at -3μm | 0.0071
| 한국표준과학연구원 | 2024-05-30 | Z Scanner position at +3μm | 0.0132
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Z Scanner position at +3μm | 0.0145
| 한국표준과학연구원 | 2024-05-30 | Z linearity by sample height | 0.995
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Z linearity by sample height | 1.018
| 한국표준과학연구원 | 2024-05-30 | Z Height noise | 0.028
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Z Height noise | 0.03
| 한국표준과학연구원 | 2024-05-30 | Y Scan Stroke | 55
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Y Scan Stroke | 114.9
| 한국표준과학연구원 | 2024-05-30 | Y Detector noise (XY Servo on) | 0.26
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Y Detector noise (XY Servo on) | 0.28
| 한국표준과학연구원 | 2024-05-30 | Y Detector noise (XY Servo off) | 0.25
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Y Detector noise (XY Servo off) | 0.27
| 한국표준과학연구원 | 2024-05-30 | Y axis OPM | 0.3
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Y axis OPM | 1.57
| 한국표준과학연구원 | 2024-05-30 | X Scan Stroke | 55
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | X Scan Stroke | 114.6
| 한국표준과학연구원 | 2024-05-30 | X Detector noise (XY Servo on) | 0.26
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | X Detector noise (XY Servo on) | 0.28
| 한국표준과학연구원 | 2024-05-30 | X Detector noise (XY Servo off) | 0.25
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | X Detector noise (XY Servo off) | 0.28
| 한국표준과학연구원 | 2024-05-30 | X axis OPM | 0.68
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | X axis OPM | 0.72
| 한국표준과학연구원 | 2024-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 한국표준과학연구원 | 2024-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 한국표준과학연구원 | 2024-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 한국표준과학연구원 | 2024-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 한국표준과학연구원 | 2024-05-30 | NCM Amplitude noise | 15.02
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | NCM Amplitude noise | 14.06
| 한국표준과학연구원 | 2024-05-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.198
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.769
| 한국표준과학연구원 | 2024-05-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.32
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.185
| 한국표준과학연구원 | 2024-05-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 110
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 162.8
| 한국표준과학연구원 | 2024-05-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 108.3
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 170
| 한국표준과학연구원 | 2024-05-30 | Lateral(C-D) sensitivity | 15.7
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Lateral(C-D) sensitivity | 15.4
| 한국표준과학연구원 | 2024-05-30 | Lateral(C-D) noise | 0.101
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Lateral(C-D) noise | 0.081
| 한국표준과학연구원 | 2024-05-30 | Full range of Z Height during Z Scan sweep | 17.8
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Full range of Z Height during Z Scan sweep | 17.2
| 한국표준과학연구원 | 2024-05-30 | Full range of Z Detector during Z Scan sweep | 12.6
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Full range of Z Detector during Z Scan sweep | 14.2
| 한국표준과학연구원 | 2024-05-30 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| 한국표준과학연구원 | 2024-05-30 | Full range of Y Detector during Y Scan sweep | 74.8
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Full range of Y Detector during Y Scan sweep | 114.9
| 한국표준과학연구원 | 2024-05-30 | Full range of X Detector Raw during X Scan sweep | 13.7
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Full range of X Detector Raw during X Scan sweep | 13.8
| 한국표준과학연구원 | 2024-05-30 | Full range of X Detector during X Scan sweep | 74
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | Full range of X Detector during X Scan sweep | 114.6
| 한국표준과학연구원 | 2024-05-30 | A-B noise | 7.48
| National Research Council of Italy, Bari (CNR-IFN, Di Franco) | 2024-05-30 | A-B noise | 7.62
| CXT #5 | 2024-05-24 | Z Linearity by Z scanner position | 1.049
| CXT #5 | 2024-05-24 | Z Linearity by 4 different sample height | 0.99993
| CXT #5 | 2024-05-24 | Z Height noise | 0.029
| CXT #5 | 2024-05-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.1117
| CXT #5 | 2024-05-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.2361
| CXT #5 | 2024-05-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5816
| CXT #5 | 2024-05-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.7108
| CXT #5 | 2024-05-24 | Z Drift | 6.7748
| CXT #5 | 2024-05-24 | Z Detector offset | 23.5
| CXT #5 | 2024-05-24 | Y2 Detector noise (XY Servo on) | 0.558
| CXT #5 | 2024-05-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 251.452
| CXT #5 | 2024-05-24 | Y1 Detector noise (XY Servo on) | 0.529
| CXT #5 | 2024-05-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 249.9653
| CXT #5 | 2024-05-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 186.852
| CXT #5 | 2024-05-24 | Y Ringing(Rpv) | 0.215
| CXT #5 | 2024-05-24 | Y Non-linearity error | 2.2197
| CXT #5 | 2024-05-24 | Y Detector noise (XY Servo on) | 0.516
| CXT #5 | 2024-05-24 | Y axis, Repeatability @1st position | 0.03
| CXT #5 | 2024-05-24 | Y axis, Repeatability @ 4th position | 0.02
| CXT #5 | 2024-05-24 | Y axis, Repeatability @ 3rd position | 0.03
| CXT #5 | 2024-05-24 | Y axis, Repeatability @ 2nd position | 0.02
| CXT #5 | 2024-05-24 | Y axis, Accuracy @1st position | 0.29
| CXT #5 | 2024-05-24 | Y axis, Accuracy @ 4th position | 0.31
| CXT #5 | 2024-05-24 | Y axis, Accuracy @ 3rd position | 0.26
| CXT #5 | 2024-05-24 | Y axis, Accuracy @ 2nd position | 0.19
| CXT #5 | 2024-05-24 | Y axis OPM | 0.836
| CXT #5 | 2024-05-24 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.76
| CXT #5 | 2024-05-24 | XY Stage positioning Y offset (LLC Translation) | 2.83
| CXT #5 | 2024-05-24 | XY Stage positioning Y offset (Global Align) | 1.75
| CXT #5 | 2024-05-24 | XY Stage positioning X offset (Vision Pattern Matching) | 3.03
| CXT #5 | 2024-05-24 | XY Stage positioning X offset (LLC Translation) | 2.15
| CXT #5 | 2024-05-24 | XY Stage positioning X offset (Global Align) | 1.99
| CXT #5 | 2024-05-24 | XY Orthogonality | 90.1
| CXT #5 | 2024-05-24 | X2 Detector noise (XY Servo on) | 0.347
| CXT #5 | 2024-05-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 262.5702
| CXT #5 | 2024-05-24 | X1 Detector noise (XY Servo on) | 0.345
| CXT #5 | 2024-05-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 264.4684
| CXT #5 | 2024-05-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.071
| CXT #5 | 2024-05-24 | X Ringing(Rpv) | 0.174
| CXT #5 | 2024-05-24 | X Non-linearity error | 2.8717
| CXT #5 | 2024-05-24 | X Detector noise (XY Servo on) | 0.292
| CXT #5 | 2024-05-24 | X axis, Repeatability @1st position | 0.03
| CXT #5 | 2024-05-24 | X axis, Repeatability @ 4th position | 0.02
| CXT #5 | 2024-05-24 | X axis, Repeatability @ 3rd position | 0.02
| CXT #5 | 2024-05-24 | X axis, Repeatability @ 2nd position | 0.02
| CXT #5 | 2024-05-24 | X axis, Accuracy @1st position | 0.23
| CXT #5 | 2024-05-24 | X axis, Accuracy @ 4th position | 0.13
| CXT #5 | 2024-05-24 | X axis, Accuracy @ 3rd position | 0.11
| CXT #5 | 2024-05-24 | X axis, Accuracy @ 2nd position | 0.14
| CXT #5 | 2024-05-24 | X axis OPM | 0.752
| CXT #5 | 2024-05-24 | Wafer placement(Loading) error on the load port2 (Y offset) | 17.66
| CXT #5 | 2024-05-24 | Wafer placement(Loading) error on the load port2 (X offset) | 35.95
| CXT #5 | 2024-05-24 | Wafer placement(Loading) error on the load port1 (Y offset) | 27.16
| CXT #5 | 2024-05-24 | Wafer placement(Loading) error on the load port1 (X offset) | 32.86
| CXT #5 | 2024-05-24 | Torque Profile (50mm, 20um/s) | 2.4272
| CXT #5 | 2024-05-24 | Torque Profile (50mm, 200um/s) | 1.8055
| CXT #5 | 2024-05-24 | StrokeYMin | -153264
| CXT #5 | 2024-05-24 | StrokeYMax | 150199
| CXT #5 | 2024-05-24 | StrokeY | 303463
| CXT #5 | 2024-05-24 | StrokeXMin | -204137
| CXT #5 | 2024-05-24 | StrokeXMax | 203397
| CXT #5 | 2024-05-24 | StrokeX | 407534
| CXT #5 | 2024-05-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.74364
| CXT #5 | 2024-05-24 | Roll motion to X+ full scan | 16.7
| CXT #5 | 2024-05-24 | RMS of ZeroScan | 0.043
| CXT #5 | 2024-05-24 | Position repeatability(10repeats, Position Based) | 0.24
| CXT #5 | 2024-05-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4.71e-05
| CXT #5 | 2024-05-24 | OpticalAxisShiftPerFocusY | -0.0011
| CXT #5 | 2024-05-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 4.83e-05
| CXT #5 | 2024-05-24 | OpticalAxisShiftPerFocusX | 0.0062
| CXT #5 | 2024-05-24 | OPM Repeatability @5mm | 0.812
| CXT #5 | 2024-05-24 | OPM Repeatability @25mm | 1.87
| CXT #5 | 2024-05-24 | OPM Repeatability @1mm | 0.37
| CXT #5 | 2024-05-24 | OPM Repeatability @10mm | 1.727
| CXT #5 | 2024-05-24 | NCM driving% with AC160TS cantilever | 28.95
| CXT #5 | 2024-05-24 | Min. Vibration Transmissbility (~100Hz) | 0.036775
| CXT #5 | 2024-05-24 | Max OPM @5mm | 8.969
| CXT #5 | 2024-05-24 | Max OPM @25mm | 75.022
| CXT #5 | 2024-05-24 | Max OPM @1mm | 3.914
| CXT #5 | 2024-05-24 | Max OPM @10mm | 42.602
| CXT #5 | 2024-05-24 | LowerLimit | 2349
| CXT #5 | 2024-05-24 | LowerLimit | 4148
| CXT #5 | 2024-05-24 | Lateral(C-D) sensitivity | 14.2
| CXT #5 | 2024-05-24 | Lateral(C-D) noise | 0.2283
| CXT #5 | 2024-05-24 | Image offset after moving from -Y to +Y (Y scan) | 96.6797
| CXT #5 | 2024-05-24 | Image offset after moving from -Y to +Y (X scan) | 19.7754
| CXT #5 | 2024-05-24 | Image offset after moving from -X to +X (Y scan) | 52.002
| CXT #5 | 2024-05-24 | Image offset after moving from -X to +X (X scan) | 79.1016
| CXT #5 | 2024-05-24 | Image offset after moving from +Y to -Y (Y scan) | 43.9453
| CXT #5 | 2024-05-24 | Image offset after moving from +Y to -Y (X scan) | 15.3809
| CXT #5 | 2024-05-24 | Image offset after moving from +X to -X (Y scan) | 15.3809
| CXT #5 | 2024-05-24 | Image offset after moving from +X to -X (X scan) | 46.875
| CXT #5 | 2024-05-24 | Full range of Z Height during Z Scan sweep | 16.8041
| CXT #5 | 2024-05-24 | Full range of Z Detector during Z Scan sweep | 12.34
| CXT #5 | 2024-05-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.28
| CXT #5 | 2024-05-24 | Full range of Y2 Detector during Y Scan sweep | 178.3634
| CXT #5 | 2024-05-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.21
| CXT #5 | 2024-05-24 | Full range of Y1 Detector during Y Scan sweep | 179.1802
| CXT #5 | 2024-05-24 | Full range of Y Detector during Y Scan sweep | 179.1802
| CXT #5 | 2024-05-24 | Full range of X2 Detector Raw during X Scan sweep | 14.14
| CXT #5 | 2024-05-24 | Full range of X2 Detector during X Scan sweep | 183.6346
| CXT #5 | 2024-05-24 | Full range of X1 Detector Raw during X Scan sweep | 14.26
| CXT #5 | 2024-05-24 | Full range of X1 Detector during X Scan sweep | 186.0372
| CXT #5 | 2024-05-24 | Full range of X Detector during X Scan sweep | 186.0372
| CXT #5 | 2024-05-24 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| CXT #5 | 2024-05-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12075
| CXT #5 | 2024-05-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.66825
| CXT #5 | 2024-05-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.65739
| CXT #5 | 2024-05-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11894
| CXT #5 | 2024-05-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.050005
| CXT #5 | 2024-05-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27584
| CXT #5 | 2024-05-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31036
| CXT #5 | 2024-05-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.056258
| CXT #5 | 2024-05-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 0.88823
| CXT #5 | 2024-05-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.3466
| CXT #5 | 2024-05-24 | %GRR w/ roughness of non-patterened wafer | 3.4625
| Linton Kayex Technology | 2024-05-22 | Z Scanner position at -3μm | 0.0092
| Jiangsu Leadmicro | 2024-05-22 | Z Scanner position at -3μm | 0.0025
| Linton Kayex Technology | 2024-05-22 | Z Scanner position at +3μm | 0.0075
| Jiangsu Leadmicro | 2024-05-22 | Z Scanner position at +3μm | 0.0062
| Linton Kayex Technology | 2024-05-22 | Z linearity by sample height | 0.994
| Jiangsu Leadmicro | 2024-05-22 | Z linearity by sample height | 1.004
| Linton Kayex Technology | 2024-05-22 | Z Height noise | 0.028
| Jiangsu Leadmicro | 2024-05-22 | Z Height noise | 0.027
| Linton Kayex Technology | 2024-05-22 | Y2 Detector noise (XY Servo on) | 0.22
| Jiangsu Leadmicro | 2024-05-22 | Y2 Detector noise (XY Servo on) | 0.23
| Linton Kayex Technology | 2024-05-22 | Y2 Detector noise (XY Servo off) | 0.22
| Jiangsu Leadmicro | 2024-05-22 | Y2 Detector noise (XY Servo off) | 0.23
| Linton Kayex Technology | 2024-05-22 | Y1 Detector noise (XY Servo on) | 0.23
| Jiangsu Leadmicro | 2024-05-22 | Y1 Detector noise (XY Servo on) | 0.23
| Linton Kayex Technology | 2024-05-22 | Y1 Detector noise (XY Servo off) | 0.23
| Jiangsu Leadmicro | 2024-05-22 | Y1 Detector noise (XY Servo off) | 0.23
| Linton Kayex Technology | 2024-05-22 | Y Scan Stroke | 181.5
| Jiangsu Leadmicro | 2024-05-22 | Y Scan Stroke | 173.2
| Linton Kayex Technology | 2024-05-22 | Y Detector noise (XY Servo on) | 0.16
| Jiangsu Leadmicro | 2024-05-22 | Y Detector noise (XY Servo on) | 0.16
| Linton Kayex Technology | 2024-05-22 | Y Detector noise (XY Servo off) | 0.16
| Jiangsu Leadmicro | 2024-05-22 | Y Detector noise (XY Servo off) | 0.16
| Linton Kayex Technology | 2024-05-22 | Y axis OPM | 0.65
| Jiangsu Leadmicro | 2024-05-22 | Y axis OPM | 1.68
| Linton Kayex Technology | 2024-05-22 | X2 Detector noise (XY Servo on) | 0.26
| Jiangsu Leadmicro | 2024-05-22 | X2 Detector noise (XY Servo on) | 0.25
| Linton Kayex Technology | 2024-05-22 | X2 Detector noise (XY Servo off) | 0.25
| Jiangsu Leadmicro | 2024-05-22 | X2 Detector noise (XY Servo off) | 0.24
| Linton Kayex Technology | 2024-05-22 | X1 Detector noise (XY Servo on) | 0.24
| Jiangsu Leadmicro | 2024-05-22 | X1 Detector noise (XY Servo on) | 0.26
| Linton Kayex Technology | 2024-05-22 | X1 Detector noise (XY Servo off) | 0.25
| Jiangsu Leadmicro | 2024-05-22 | X1 Detector noise (XY Servo off) | 0.25
| Linton Kayex Technology | 2024-05-22 | X Scan Stroke | 182.7
| Jiangsu Leadmicro | 2024-05-22 | X Scan Stroke | 179.6
| Linton Kayex Technology | 2024-05-22 | X Detector noise (XY Servo on) | 0.19
| Jiangsu Leadmicro | 2024-05-22 | X Detector noise (XY Servo on) | 0.19
| Linton Kayex Technology | 2024-05-22 | X Detector noise (XY Servo off) | 0.19
| Jiangsu Leadmicro | 2024-05-22 | X Detector noise (XY Servo off) | 0.19
| Linton Kayex Technology | 2024-05-22 | X axis OPM | 0.92
| Jiangsu Leadmicro | 2024-05-22 | X axis OPM | 0.95
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.39
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.28
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.28
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.24
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.34
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.23
| Linton Kayex Technology | 2024-05-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.23
| Linton Kayex Technology | 2024-05-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Linton Kayex Technology | 2024-05-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Jiangsu Leadmicro | 2024-05-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Linton Kayex Technology | 2024-05-22 | NCM Amplitude noise | 11.08
| Jiangsu Leadmicro | 2024-05-22 | NCM Amplitude noise | 12.72
| Linton Kayex Technology | 2024-05-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.216
| Jiangsu Leadmicro | 2024-05-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.333
| Linton Kayex Technology | 2024-05-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.713
| Jiangsu Leadmicro | 2024-05-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.082
| Linton Kayex Technology | 2024-05-22 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 251.9
| Jiangsu Leadmicro | 2024-05-22 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 257.7
| Linton Kayex Technology | 2024-05-22 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 253.1
| Jiangsu Leadmicro | 2024-05-22 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 251.7
| Linton Kayex Technology | 2024-05-22 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 258.4
| Jiangsu Leadmicro | 2024-05-22 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 258.4
| Linton Kayex Technology | 2024-05-22 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 250.1
| Jiangsu Leadmicro | 2024-05-22 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 262.2
| Linton Kayex Technology | 2024-05-22 | Maximum Y Scan range (Full Stroke) | 180.3
| Jiangsu Leadmicro | 2024-05-22 | Maximum Y Scan range (Full Stroke) | 172.9
| Linton Kayex Technology | 2024-05-22 | Maximum X Scan range (Full Stroke) | 180.5
| Jiangsu Leadmicro | 2024-05-22 | Maximum X Scan range (Full Stroke) | 179.9
| Linton Kayex Technology | 2024-05-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Jiangsu Leadmicro | 2024-05-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Linton Kayex Technology | 2024-05-22 | Lateral(C-D) sensitivity | 11.6
| Jiangsu Leadmicro | 2024-05-22 | Lateral(C-D) sensitivity | 18.5
| Linton Kayex Technology | 2024-05-22 | Lateral(C-D) noise | 0.116
| Jiangsu Leadmicro | 2024-05-22 | Lateral(C-D) noise | 0.134
| Linton Kayex Technology | 2024-05-22 | Full range of Z Height during Z Scan sweep | 18.2
| Jiangsu Leadmicro | 2024-05-22 | Full range of Z Height during Z Scan sweep | 18
| Linton Kayex Technology | 2024-05-22 | Full range of Z Detector during Z Scan sweep | 14.2
| Jiangsu Leadmicro | 2024-05-22 | Full range of Z Detector during Z Scan sweep | 14.4
| Linton Kayex Technology | 2024-05-22 | Full range of Y2 Detector raw | 14.3
| Jiangsu Leadmicro | 2024-05-22 | Full range of Y2 Detector raw | 13.5
| Linton Kayex Technology | 2024-05-22 | Full range of Y1 Detector raw | 14.3
| Jiangsu Leadmicro | 2024-05-22 | Full range of Y1 Detector raw | 13.7
| Linton Kayex Technology | 2024-05-22 | Full range of X2 Detector raw | 14.1
| Jiangsu Leadmicro | 2024-05-22 | Full range of X2 Detector raw | 13.9
| Linton Kayex Technology | 2024-05-22 | Full range of X1 Detector raw | 14.3
| Jiangsu Leadmicro | 2024-05-22 | Full range of X1 Detector raw | 13.7
| Linton Kayex Technology | 2024-05-22 | A-B noise | 7.29
| Jiangsu Leadmicro | 2024-05-22 | A-B noise | 7.01
| NAURA, Beijing | 2024-05-20 | Z Linearity by Z scanner position | 0.8
| NAURA, Beijing | 2024-05-20 | Z Linearity by 4 different sample height | 0.999
| NAURA, Beijing | 2024-05-20 | Z Height noise | 0.032
| NAURA, Beijing | 2024-05-20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.1
| NAURA, Beijing | 2024-05-20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.1
| NAURA, Beijing | 2024-05-20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2
| NAURA, Beijing | 2024-05-20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2
| NAURA, Beijing | 2024-05-20 | Z Drift | 10.4
| NAURA, Beijing | 2024-05-20 | Z Detector offset | 1.2
| NAURA, Beijing | 2024-05-20 | Y2 Detector noise (XY Servo on) | 0.28
| NAURA, Beijing | 2024-05-20 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 274.7
| NAURA, Beijing | 2024-05-20 | Y1 Detector noise (XY Servo on) | 0.29
| NAURA, Beijing | 2024-05-20 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 272.6
| NAURA, Beijing | 2024-05-20 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 177.2
| NAURA, Beijing | 2024-05-20 | Y Ringing(Rpv) | 0.15
| NAURA, Beijing | 2024-05-20 | Y Detector noise (XY Servo on) | 0.21
| NAURA, Beijing | 2024-05-20 | Y axis, Repeatability @1st position | 0.02
| NAURA, Beijing | 2024-05-20 | Y axis, Repeatability @ 4th position | 0.02
| NAURA, Beijing | 2024-05-20 | Y axis, Repeatability @ 3rd position | 0.03
| NAURA, Beijing | 2024-05-20 | Y axis, Repeatability @ 2nd position | 0.02
| NAURA, Beijing | 2024-05-20 | Y axis, Accuracy @1st position | 0.32
| NAURA, Beijing | 2024-05-20 | Y axis, Accuracy @ 4th position | 0.39
| NAURA, Beijing | 2024-05-20 | Y axis, Accuracy @ 3rd position | 0.34
| NAURA, Beijing | 2024-05-20 | Y axis, Accuracy @ 2nd position | 0.26
| NAURA, Beijing | 2024-05-20 | Y axis OPM | 1.55
| NAURA, Beijing | 2024-05-20 | XY Stage positioning Y offset (Vision Pattern Matching) | 8.14
| NAURA, Beijing | 2024-05-20 | XY Stage positioning Y offset (LLC Translation) | 3.39
| NAURA, Beijing | 2024-05-20 | XY Stage positioning Y offset (Global Align) | 7.97
| NAURA, Beijing | 2024-05-20 | XY Stage positioning X offset (Vision Pattern Matching) | 4.8
| NAURA, Beijing | 2024-05-20 | XY Stage positioning X offset (LLC Translation) | 2.05
| NAURA, Beijing | 2024-05-20 | XY Stage positioning X offset (Global Align) | 8.95
| NAURA, Beijing | 2024-05-20 | XY Orthogonality | 89.9
| NAURA, Beijing | 2024-05-20 | X2 Detector noise (XY Servo on) | 0.29
| NAURA, Beijing | 2024-05-20 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 266.1
| NAURA, Beijing | 2024-05-20 | X1 Detector noise (XY Servo on) | 0.28
| NAURA, Beijing | 2024-05-20 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 260.2
| NAURA, Beijing | 2024-05-20 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 173.5
| NAURA, Beijing | 2024-05-20 | X Ringing(Rpv) | 0.06
| NAURA, Beijing | 2024-05-20 | X Detector noise (XY Servo on) | 0.21
| NAURA, Beijing | 2024-05-20 | X axis, Repeatability @1st position | 0.04
| NAURA, Beijing | 2024-05-20 | X axis, Repeatability @ 4th position | 0.02
| NAURA, Beijing | 2024-05-20 | X axis, Repeatability @ 3rd position | 0.03
| NAURA, Beijing | 2024-05-20 | X axis, Repeatability @ 2nd position | 0.02
| NAURA, Beijing | 2024-05-20 | X axis, Accuracy @1st position | 0.19
| NAURA, Beijing | 2024-05-20 | X axis, Accuracy @ 4th position | 0.09
| NAURA, Beijing | 2024-05-20 | X axis, Accuracy @ 3rd position | 0.29
| NAURA, Beijing | 2024-05-20 | X axis, Accuracy @ 2nd position | 0.17
| NAURA, Beijing | 2024-05-20 | X axis OPM | 0.6
| NAURA, Beijing | 2024-05-20 | Wafer placement(Loading) error on the load port2 (Y offset) | 11.3
| NAURA, Beijing | 2024-05-20 | Wafer placement(Loading) error on the load port2 (X offset) | 6.6
| NAURA, Beijing | 2024-05-20 | Wafer placement(Loading) error on the load port1 (Y offset) | 13.6
| NAURA, Beijing | 2024-05-20 | Wafer placement(Loading) error on the load port1 (X offset) | 8.1
| NAURA, Beijing | 2024-05-20 | Torque Profile (50mm, 20um/s) | 3.805
| NAURA, Beijing | 2024-05-20 | Torque Profile (50mm, 200um/s) | 2.299
| NAURA, Beijing | 2024-05-20 | StrokeYMin | -152213
| NAURA, Beijing | 2024-05-20 | StrokeYMax | 150093
| NAURA, Beijing | 2024-05-20 | StrokeY | 302306
| NAURA, Beijing | 2024-05-20 | StrokeXMin | -203570
| NAURA, Beijing | 2024-05-20 | StrokeXMax | 203680
| NAURA, Beijing | 2024-05-20 | StrokeX | 407250
| NAURA, Beijing | 2024-05-20 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7
| NAURA, Beijing | 2024-05-20 | Roll motion to X+ full scan | 16.1
| NAURA, Beijing | 2024-05-20 | RMS of ZeroScan | 0.052
| NAURA, Beijing | 2024-05-20 | Position repeatability(10repeats, Position Based) | 0.48
| NAURA, Beijing | 2024-05-20 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| NAURA, Beijing | 2024-05-20 | OpticalAxisShiftPerFocusY | 0.0012
| NAURA, Beijing | 2024-05-20 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| NAURA, Beijing | 2024-05-20 | OpticalAxisShiftPerFocusX | 0.002
| NAURA, Beijing | 2024-05-20 | OPM Repeatability @5mm | 0.49
| NAURA, Beijing | 2024-05-20 | OPM Repeatability @25mm | 7.74
| NAURA, Beijing | 2024-05-20 | OPM Repeatability @1mm | 0.42
| NAURA, Beijing | 2024-05-20 | OPM Repeatability @10mm | 1.15
| NAURA, Beijing | 2024-05-20 | NCM driving% with AC160TS cantilever | 10.4
| NAURA, Beijing | 2024-05-20 | Min. Vibration Transmissbility (~100Hz) | 0.043
| NAURA, Beijing | 2024-05-20 | Max OPM @5mm | 7.77
| NAURA, Beijing | 2024-05-20 | Max OPM @25mm | 61.18
| NAURA, Beijing | 2024-05-20 | Max OPM @1mm | 4.28
| NAURA, Beijing | 2024-05-20 | Max OPM @10mm | 14.28
| NAURA, Beijing | 2024-05-20 | LowerLimit | 1955.12
| NAURA, Beijing | 2024-05-20 | LowerLimit | 4345
| NAURA, Beijing | 2024-05-20 | Lateral(C-D) sensitivity | 16.6
| NAURA, Beijing | 2024-05-20 | Lateral(C-D) noise | 0.2
| NAURA, Beijing | 2024-05-20 | Image offset after moving from -Y to +Y (Y scan) | 90.09
| NAURA, Beijing | 2024-05-20 | Image offset after moving from -Y to +Y (X scan) | 9.52
| NAURA, Beijing | 2024-05-20 | Image offset after moving from -X to +X (Y scan) | 42.48
| NAURA, Beijing | 2024-05-20 | Image offset after moving from -X to +X (X scan) | 34.42
| NAURA, Beijing | 2024-05-20 | Image offset after moving from +Y to -Y (Y scan) | 17.58
| NAURA, Beijing | 2024-05-20 | Image offset after moving from +Y to -Y (X scan) | 10.25
| NAURA, Beijing | 2024-05-20 | Image offset after moving from +X to -X (Y scan) | 10.99
| NAURA, Beijing | 2024-05-20 | Image offset after moving from +X to -X (X scan) | 34.42
| NAURA, Beijing | 2024-05-20 | Full range of Z Height during Z Scan sweep | 16.2
| NAURA, Beijing | 2024-05-20 | Full range of Z Detector during Z Scan sweep | 12.9
| NAURA, Beijing | 2024-05-20 | Full range of Y2 Detector Raw during Y Scan sweep | 14.56
| NAURA, Beijing | 2024-05-20 | Full range of Y2 Detector during Y Scan sweep | 208.4
| NAURA, Beijing | 2024-05-20 | Full range of Y1 Detector Raw during Y Scan sweep | 14.35
| NAURA, Beijing | 2024-05-20 | Full range of Y1 Detector during Y Scan sweep | 203.3
| NAURA, Beijing | 2024-05-20 | Full range of Y Detector during Y Scan sweep | 203.4
| NAURA, Beijing | 2024-05-20 | Full range of X2 Detector Raw during X Scan sweep | 14.68
| NAURA, Beijing | 2024-05-20 | Full range of X2 Detector during X Scan sweep | 222.4
| NAURA, Beijing | 2024-05-20 | Full range of X1 Detector Raw during X Scan sweep | 14.77
| NAURA, Beijing | 2024-05-20 | Full range of X1 Detector during X Scan sweep | 223.7
| NAURA, Beijing | 2024-05-20 | Full range of X Detector during X Scan sweep | 223.7
| NAURA, Beijing | 2024-05-20 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| NAURA, Beijing | 2024-05-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| NAURA, Beijing | 2024-05-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.49
| NAURA, Beijing | 2024-05-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.49
| NAURA, Beijing | 2024-05-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| NAURA, Beijing | 2024-05-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| NAURA, Beijing | 2024-05-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| NAURA, Beijing | 2024-05-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.64
| NAURA, Beijing | 2024-05-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| NAURA, Beijing | 2024-05-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.3
| NAURA, Beijing | 2024-05-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.4
| NAURA, Beijing | 2024-05-20 | %GRR w/ roughness of non-patterened wafer | 3.3
| Getinge Atrium Medical(Dao) | 2024-05-17 | Z Scanner position at -3μm | 0.0084
| Getinge Atrium Medical(Dao) | 2024-05-17 | Z Scanner position at +3μm | 0.014
| Getinge Atrium Medical(Dao) | 2024-05-17 | Z linearity by sample height | 1.015
| Getinge Atrium Medical(Dao) | 2024-05-17 | Z Height noise | 0.016
| Getinge Atrium Medical(Dao) | 2024-05-17 | Y Scan Stroke | 116.2
| Getinge Atrium Medical(Dao) | 2024-05-17 | Y Detector noise (XY Servo on) | 0.28
| Getinge Atrium Medical(Dao) | 2024-05-17 | Y Detector noise (XY Servo off) | 0.26
| Getinge Atrium Medical(Dao) | 2024-05-17 | Y axis OPM | 0.95
| Getinge Atrium Medical(Dao) | 2024-05-17 | XY orthogonality | 89.9
| Getinge Atrium Medical(Dao) | 2024-05-17 | X Scan Stroke | 118.8
| Getinge Atrium Medical(Dao) | 2024-05-17 | X Detector noise (XY Servo on) | 0.27
| Getinge Atrium Medical(Dao) | 2024-05-17 | X Detector noise (XY Servo off) | 0.26
| Getinge Atrium Medical(Dao) | 2024-05-17 | X axis OPM | 0.9
| Getinge Atrium Medical(Dao) | 2024-05-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Getinge Atrium Medical(Dao) | 2024-05-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Getinge Atrium Medical(Dao) | 2024-05-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Getinge Atrium Medical(Dao) | 2024-05-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Getinge Atrium Medical(Dao) | 2024-05-17 | NCM Amplitude noise | 8.06
| Getinge Atrium Medical(Dao) | 2024-05-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.22
| Getinge Atrium Medical(Dao) | 2024-05-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.491
| Getinge Atrium Medical(Dao) | 2024-05-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 161.6
| Getinge Atrium Medical(Dao) | 2024-05-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.4
| Getinge Atrium Medical(Dao) | 2024-05-17 | Lateral(C-D) sensitivity | 9.3
| Getinge Atrium Medical(Dao) | 2024-05-17 | Lateral(C-D) noise | 0.124
| Getinge Atrium Medical(Dao) | 2024-05-17 | Full range of Z Height during Z Scan sweep | 16.1
| Getinge Atrium Medical(Dao) | 2024-05-17 | Full range of Z Detector during Z Scan sweep | 15
| Getinge Atrium Medical(Dao) | 2024-05-17 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| Getinge Atrium Medical(Dao) | 2024-05-17 | Full range of Y Detector during Y Scan sweep | 118.4
| Getinge Atrium Medical(Dao) | 2024-05-17 | Full range of X Detector Raw during X Scan sweep | 14.7
| Getinge Atrium Medical(Dao) | 2024-05-17 | Full range of X Detector during X Scan sweep | 119.9
| Getinge Atrium Medical(Dao) | 2024-05-17 | A-B noise | 13.23
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Z Scanner position at -3μm | 0.011
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Z Scanner position at +3μm | 0.0119
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Z linearity by sample height | 1.003
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Z Height noise | 0.017
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Y Scan Stroke | 115
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Y Detector noise (XY Servo on) | 0.25
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Y Detector noise (XY Servo off) | 0.26
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Y axis OPM | 1.45
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | XY orthogonality | 90
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | X Scan Stroke | 113.8
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | X Detector noise (XY Servo on) | 0.25
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | X Detector noise (XY Servo off) | 0.25
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | X axis OPM | 1.03
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | NCM Amplitude noise | 11.53
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.707
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.065
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 162.9
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.3
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Lateral(C-D) sensitivity | 6.9
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Lateral(C-D) noise | 0.095
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Full range of Z Height during Z Scan sweep | 16.6
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Full range of Z Detector during Z Scan sweep | 14.4
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Full range of Y Detector during Y Scan sweep | 115
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Full range of X Detector Raw during X Scan sweep | 14.3
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | Full range of X Detector during X Scan sweep | 113.8
| 삼성전자 파운드리 분석기술팀 | 2024-05-16 | A-B noise | 16.47
| Park Systems Inc. (PSI) | 2024-05-09 | Z Scanner position at -3μm | 0.0057
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Z Scanner position at -3μm | 0.0027
| Park Systems Inc. (PSI) | 2024-05-09 | Z Scanner position at +3μm | 0.011
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Z Scanner position at +3μm | 0.0105
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Z linearity by sample height | 1.007
| Park Systems Inc. (PSI) | 2024-05-09 | Z linearity by sample height | 1.016
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Z Height noise | 0.016
| Park Systems Inc. (PSI) | 2024-05-09 | Z Height noise | 0.027
| Park Systems Inc. (PSI) | 2024-05-09 | Y2 Detector noise (XY Servo on) | 0.33
| Park Systems Inc. (PSI) | 2024-05-09 | Y2 Detector noise (XY Servo off) | 0.22
| Park Systems Inc. (PSI) | 2024-05-09 | Y1 Detector noise (XY Servo on) | 0.32
| Park Systems Inc. (PSI) | 2024-05-09 | Y1 Detector noise (XY Servo off) | 0.21
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Y Scan Stroke | 116.7
| Park Systems Inc. (PSI) | 2024-05-09 | Y Scan Stroke | 171.2
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Y Detector noise (XY Servo on) | 0.26
| Park Systems Inc. (PSI) | 2024-05-09 | Y Detector noise (XY Servo on) | 0.25
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Y Detector noise (XY Servo off) | 0.25
| Park Systems Inc. (PSI) | 2024-05-09 | Y Detector noise (XY Servo off) | 0.15
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Y axis OPM | 0.59
| Park Systems Inc. (PSI) | 2024-05-09 | Y axis OPM | 1.81
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | XY orthogonality | 90
| Park Systems Inc. (PSI) | 2024-05-09 | X2 Detector noise (XY Servo on) | 0.33
| Park Systems Inc. (PSI) | 2024-05-09 | X2 Detector noise (XY Servo off) | 0.22
| Park Systems Inc. (PSI) | 2024-05-09 | X1 Detector noise (XY Servo on) | 0.38
| Park Systems Inc. (PSI) | 2024-05-09 | X1 Detector noise (XY Servo off) | 0.28
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | X Scan Stroke | 113.2
| Park Systems Inc. (PSI) | 2024-05-09 | X Scan Stroke | 171
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | X Detector noise (XY Servo on) | 0.26
| Park Systems Inc. (PSI) | 2024-05-09 | X Detector noise (XY Servo on) | 0.28
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | X Detector noise (XY Servo off) | 0.26
| Park Systems Inc. (PSI) | 2024-05-09 | X Detector noise (XY Servo off) | 0.18
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | X axis OPM | 0.98
| Park Systems Inc. (PSI) | 2024-05-09 | X axis OPM | 1.24
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.32
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Park Systems Inc. (PSI) | 2024-05-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | NCM Amplitude noise | 5.02
| Park Systems Inc. (PSI) | 2024-05-09 | NCM Amplitude noise | 14.65
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.897
| Park Systems Inc. (PSI) | 2024-05-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.242
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.005
| Park Systems Inc. (PSI) | 2024-05-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.79
| Park Systems Inc. (PSI) | 2024-05-09 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 267.3
| Park Systems Inc. (PSI) | 2024-05-09 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.7
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.6
| Park Systems Inc. (PSI) | 2024-05-09 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 253.2
| Park Systems Inc. (PSI) | 2024-05-09 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 253.2
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.9
| Park Systems Inc. (PSI) | 2024-05-09 | Maximum Y Scan range (Full Stroke) | 171.3
| Park Systems Inc. (PSI) | 2024-05-09 | Maximum X Scan range (Full Stroke) | 170.7
| Park Systems Inc. (PSI) | 2024-05-09 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Lateral(C-D) sensitivity | 11.7
| Park Systems Inc. (PSI) | 2024-05-09 | Lateral(C-D) sensitivity | 20
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Lateral(C-D) noise | 0.112
| Park Systems Inc. (PSI) | 2024-05-09 | Lateral(C-D) noise | 0.072
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Full range of Z Height during Z Scan sweep | 17.2
| Park Systems Inc. (PSI) | 2024-05-09 | Full range of Z Height during Z Scan sweep | 18.2
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Full range of Z Detector during Z Scan sweep | 16.4
| Park Systems Inc. (PSI) | 2024-05-09 | Full range of Z Detector during Z Scan sweep | 13.6
| Park Systems Inc. (PSI) | 2024-05-09 | Full range of Y2 Detector raw | 13.2
| Park Systems Inc. (PSI) | 2024-05-09 | Full range of Y1 Detector raw | 13.6
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Full range of Y Detector during Y Scan sweep | 116.7
| Park Systems Inc. (PSI) | 2024-05-09 | Full range of X2 Detector raw | 13.5
| Park Systems Inc. (PSI) | 2024-05-09 | Full range of X1 Detector raw | 13.5
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Full range of X Detector Raw during X Scan sweep | 14
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | Full range of X Detector during X Scan sweep | 113.2
| Park Systems China_Beijing(Demo 장비) | 2024-05-09 | A-B noise | 8.83
| Park Systems Inc. (PSI) | 2024-05-09 | A-B noise | 9.48
| University of Texas, Arlington | 2024-05-08 | Z Scanner position at -3μm | 0.0007
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Z Scanner position at -3μm | 0.0022
| Cenimat (Tomas) | 2024-05-08 | Z Scanner position at -3μm | 0.011
| University of Texas, Arlington | 2024-05-08 | Z Scanner position at +3μm | 0.0057
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Z Scanner position at +3μm | 0.0085
| Cenimat (Tomas) | 2024-05-08 | Z Scanner position at +3μm | 0.0119
| Cenimat (Tomas) | 2024-05-08 | Z linearity by sample height | 1.005
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Z linearity by sample height | 0.998
| University of Texas, Arlington | 2024-05-08 | Z linearity by sample height | 1.018
| Cenimat (Tomas) | 2024-05-08 | Z Height noise | 0.011
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Z Height noise | 0.037
| University of Texas, Arlington | 2024-05-08 | Z Height noise | 0.028
| University of Texas, Arlington | 2024-05-08 | Y2 Detector noise (XY Servo on) | 0.22
| University of Texas, Arlington | 2024-05-08 | Y2 Detector noise (XY Servo off) | 0.22
| University of Texas, Arlington | 2024-05-08 | Y1 Detector noise (XY Servo on) | 0.21
| University of Texas, Arlington | 2024-05-08 | Y1 Detector noise (XY Servo off) | 0.21
| Cenimat (Tomas) | 2024-05-08 | Y Scan Stroke | 119
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Y Scan Stroke | 55
| University of Texas, Arlington | 2024-05-08 | Y Scan Stroke | 166.4
| Cenimat (Tomas) | 2024-05-08 | Y Detector noise (XY Servo on) | 0.14
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Y Detector noise (XY Servo on) | 0.26
| University of Texas, Arlington | 2024-05-08 | Y Detector noise (XY Servo on) | 0.15
| Cenimat (Tomas) | 2024-05-08 | Y Detector noise (XY Servo off) | 0.14
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Y Detector noise (XY Servo off) | 0.25
| University of Texas, Arlington | 2024-05-08 | Y Detector noise (XY Servo off) | 0.15
| Cenimat (Tomas) | 2024-05-08 | Y axis OPM | 1.1
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Y axis OPM | 1.59
| University of Texas, Arlington | 2024-05-08 | Y axis OPM | 1.76
| Cenimat (Tomas) | 2024-05-08 | XY orthogonality | 90.1
| University of Texas, Arlington | 2024-05-08 | X2 Detector noise (XY Servo on) | 0.23
| University of Texas, Arlington | 2024-05-08 | X2 Detector noise (XY Servo off) | 0.23
| University of Texas, Arlington | 2024-05-08 | X1 Detector noise (XY Servo on) | 0.24
| University of Texas, Arlington | 2024-05-08 | X1 Detector noise (XY Servo off) | 0.23
| Cenimat (Tomas) | 2024-05-08 | X Scan Stroke | 114.2
| 충남테크노파크-디스플레이센터 | 2024-05-08 | X Scan Stroke | 55
| University of Texas, Arlington | 2024-05-08 | X Scan Stroke | 180.3
| Cenimat (Tomas) | 2024-05-08 | X Detector noise (XY Servo on) | 0.14
| 충남테크노파크-디스플레이센터 | 2024-05-08 | X Detector noise (XY Servo on) | 0.27
| University of Texas, Arlington | 2024-05-08 | X Detector noise (XY Servo on) | 0.17
| Cenimat (Tomas) | 2024-05-08 | X Detector noise (XY Servo off) | 0.14
| 충남테크노파크-디스플레이센터 | 2024-05-08 | X Detector noise (XY Servo off) | 0.26
| University of Texas, Arlington | 2024-05-08 | X Detector noise (XY Servo off) | 0.17
| Cenimat (Tomas) | 2024-05-08 | X axis OPM | 0.58
| 충남테크노파크-디스플레이센터 | 2024-05-08 | X axis OPM | 0.76
| University of Texas, Arlington | 2024-05-08 | X axis OPM | 1.2
| Cenimat (Tomas) | 2024-05-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 충남테크노파크-디스플레이센터 | 2024-05-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| Cenimat (Tomas) | 2024-05-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 충남테크노파크-디스플레이센터 | 2024-05-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.4
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.39
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.38
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.37
| University of Texas, Arlington | 2024-05-08 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.4
| Cenimat (Tomas) | 2024-05-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 충남테크노파크-디스플레이센터 | 2024-05-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| University of Texas, Arlington | 2024-05-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Cenimat (Tomas) | 2024-05-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 충남테크노파크-디스플레이센터 | 2024-05-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| University of Texas, Arlington | 2024-05-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Cenimat (Tomas) | 2024-05-08 | NCM Amplitude noise | 7.12
| 충남테크노파크-디스플레이센터 | 2024-05-08 | NCM Amplitude noise | 13.76
| University of Texas, Arlington | 2024-05-08 | NCM Amplitude noise | 14.83
| Cenimat (Tomas) | 2024-05-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.223
| 충남테크노파크-디스플레이센터 | 2024-05-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.47
| University of Texas, Arlington | 2024-05-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.154
| Cenimat (Tomas) | 2024-05-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 19.773
| 충남테크노파크-디스플레이센터 | 2024-05-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.64
| University of Texas, Arlington | 2024-05-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.835
| University of Texas, Arlington | 2024-05-08 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 244.9
| University of Texas, Arlington | 2024-05-08 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 235.7
| Cenimat (Tomas) | 2024-05-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.1
| 충남테크노파크-디스플레이센터 | 2024-05-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 109.4
| University of Texas, Arlington | 2024-05-08 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 260.5
| University of Texas, Arlington | 2024-05-08 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 260.3
| Cenimat (Tomas) | 2024-05-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.6
| 충남테크노파크-디스플레이센터 | 2024-05-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.4
| University of Texas, Arlington | 2024-05-08 | Maximum Y Scan range (Full Stroke) | 163.9
| University of Texas, Arlington | 2024-05-08 | Maximum X Scan range (Full Stroke) | 180.3
| University of Texas, Arlington | 2024-05-08 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Cenimat (Tomas) | 2024-05-08 | Lateral(C-D) sensitivity | 9
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Lateral(C-D) sensitivity | 6.9
| University of Texas, Arlington | 2024-05-08 | Lateral(C-D) sensitivity | 13.8
| Cenimat (Tomas) | 2024-05-08 | Lateral(C-D) noise | 0.139
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Lateral(C-D) noise | 0.1
| University of Texas, Arlington | 2024-05-08 | Lateral(C-D) noise | 0.121
| Cenimat (Tomas) | 2024-05-08 | Full range of Z Height during Z Scan sweep | 16.8
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Full range of Z Height during Z Scan sweep | 18.3
| University of Texas, Arlington | 2024-05-08 | Full range of Z Height during Z Scan sweep | 18
| Cenimat (Tomas) | 2024-05-08 | Full range of Z Detector during Z Scan sweep | 17.2
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Full range of Z Detector during Z Scan sweep | 14.2
| University of Texas, Arlington | 2024-05-08 | Full range of Z Detector during Z Scan sweep | 13.4
| University of Texas, Arlington | 2024-05-08 | Full range of Y2 Detector raw | 13.8
| University of Texas, Arlington | 2024-05-08 | Full range of Y1 Detector raw | 13.9
| Cenimat (Tomas) | 2024-05-08 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Cenimat (Tomas) | 2024-05-08 | Full range of Y Detector during Y Scan sweep | 119.2
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Full range of Y Detector during Y Scan sweep | 77.3
| University of Texas, Arlington | 2024-05-08 | Full range of X2 Detector raw | 13.9
| University of Texas, Arlington | 2024-05-08 | Full range of X1 Detector raw | 13.8
| Cenimat (Tomas) | 2024-05-08 | Full range of X Detector Raw during X Scan sweep | 14.1
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Full range of X Detector Raw during X Scan sweep | 14.4
| Cenimat (Tomas) | 2024-05-08 | Full range of X Detector during X Scan sweep | 114.3
| 충남테크노파크-디스플레이센터 | 2024-05-08 | Full range of X Detector during X Scan sweep | 75
| Cenimat (Tomas) | 2024-05-08 | A-B noise | 11.49
| 충남테크노파크-디스플레이센터 | 2024-05-08 | A-B noise | 8.68
| University of Texas, Arlington | 2024-05-08 | A-B noise | 6.55
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Z Scanner position at -3μm | 0.0012
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Z Scanner position at -3μm | 0.0005
| Technolutions | 2024-05-07 | Z Scanner position at -3μm | 0.0014
| Technolutions | 2024-05-07 | Z Scanner position at -3μm | 0.0053
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Z Scanner position at +3μm | 0.005
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Z Scanner position at +3μm | 0.0005
| Technolutions | 2024-05-07 | Z Scanner position at +3μm | 0.0053
| Technolutions | 2024-05-07 | Z Scanner position at +3μm | 0.0018
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Z Linearity by Z Scanner position
| Technolutions | 2024-05-07 | Z linearity by sample height | 0.999
| Technolutions | 2024-05-07 | Z linearity by sample height | 1.003
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Z linearity by sample height | 1.008
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Z Linearity by 3 different sample height | 0.999
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Z Height noise | 0.029
| Technolutions | 2024-05-07 | Z Height noise | 0.029
| Technolutions | 2024-05-07 | Z Height noise | 0.026
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Z Height noise | 0.028
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y2 Detector noise (XY Servo on) | 0.23
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y2 Detector noise (XY Servo off) | 0.23
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y1 Detector noise (XY Servo on) | 0.28
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y1 Detector noise (XY Servo off) | 0.28
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Y Scan Stroke | 137.7
| Technolutions | 2024-05-07 | Y Scan Stroke | 55
| Technolutions | 2024-05-07 | Y Scan Stroke | 55
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y Scan Stroke | 177.3
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Y Detector noise (XY Servo on) | 0.42
| Technolutions | 2024-05-07 | Y Detector noise (XY Servo on) | 0.26
| Technolutions | 2024-05-07 | Y Detector noise (XY Servo on) | 0.34
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y Detector noise (XY Servo on) | 0.18
| Technolutions | 2024-05-07 | Y Detector noise (XY Servo off) | 0.25
| Technolutions | 2024-05-07 | Y Detector noise (XY Servo off) | 0.39
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y Detector noise (XY Servo off) | 0.18
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Y Detector noise (XY Servo off) | 0.47
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Y axis OPM | 1.35
| Technolutions | 2024-05-07 | Y axis OPM | 0.35
| Technolutions | 2024-05-07 | Y axis OPM | 0.96
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Y axis OPM | 0.53
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X2 Detector noise (XY Servo on) | 0.26
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X2 Detector noise (XY Servo off) | 0.24
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X1 Detector noise (XY Servo on) | 0.26
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X1 Detector noise (XY Servo off) | 0.24
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | X Scan Stroke | 139.3
| Technolutions | 2024-05-07 | X Scan Stroke | 55
| Technolutions | 2024-05-07 | X Scan Stroke | 55
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X Scan Stroke | 180.4
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | X Detector noise (XY Servo on) | 0.44
| Technolutions | 2024-05-07 | X Detector noise (XY Servo on) | 0.26
| Technolutions | 2024-05-07 | X Detector noise (XY Servo on) | 0.35
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X Detector noise (XY Servo on) | 0.2
| Technolutions | 2024-05-07 | X Detector noise (XY Servo off) | 0.25
| Technolutions | 2024-05-07 | X Detector noise (XY Servo off) | 0.39
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X Detector noise (XY Servo off) | 0.18
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | X Detector noise (XY Servo off) | 0.47
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | X axis OPM | 1
| Technolutions | 2024-05-07 | X axis OPM | 0.74
| Technolutions | 2024-05-07 | X axis OPM | 0.65
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | X axis OPM | 0.58
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Technolutions | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Technolutions | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.54
| Technolutions | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Technolutions | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.27
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.25
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.24
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.25
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Technolutions | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Technolutions | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| Technolutions | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Technolutions | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | NCM Amplitude noise | 15.13
| Technolutions | 2024-05-07 | NCM Amplitude noise | 14.31
| Technolutions | 2024-05-07 | NCM Amplitude noise | 16.22
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | NCM Amplitude noise | 12.27
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.03
| Technolutions | 2024-05-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.584
| Technolutions | 2024-05-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.238
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.099
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.3
| Technolutions | 2024-05-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.742
| Technolutions | 2024-05-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.605
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.038
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 259.4
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 260.1
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 201.4
| Technolutions | 2024-05-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105.4
| Technolutions | 2024-05-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 259.7
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 265.2
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 207.2
| Technolutions | 2024-05-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.9
| Technolutions | 2024-05-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Maximum Y Scan range (Full Stroke) | 180
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Maximum X Scan range (Full Stroke) | 180.2
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Lateral(C-D) sensitivity | 14
| Technolutions | 2024-05-07 | Lateral(C-D) sensitivity | 18.7
| Technolutions | 2024-05-07 | Lateral(C-D) sensitivity | 3.1
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Lateral(C-D) sensitivity | 55.535
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Lateral(C-D) noise | 0.07
| Technolutions | 2024-05-07 | Lateral(C-D) noise | 0.055
| Technolutions | 2024-05-07 | Lateral(C-D) noise | 0.068
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Lateral(C-D) noise | 0.102
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Full range of Z Height during Z Scan sweep | 17.8
| Technolutions | 2024-05-07 | Full range of Z Height during Z Scan sweep | 18.5
| Technolutions | 2024-05-07 | Full range of Z Height during Z Scan sweep | 17.9
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Full range of Z Height during Z Scan sweep | 18.1
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Full range of Z Detector during Z Scan sweep | 13.5
| Technolutions | 2024-05-07 | Full range of Z Detector during Z Scan sweep | 13.3
| Technolutions | 2024-05-07 | Full range of Z Detector during Z Scan sweep | 14.6
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Full range of Z Detector during Z Scan sweep | 12.9
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Full range of Y2 Detector raw | 13.5
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Full range of Y1 Detector raw | 13.8
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| Technolutions | 2024-05-07 | Full range of Y Detector Raw during Y Scan sweep | 14
| Technolutions | 2024-05-07 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Full range of Y Detector during Y Scan sweep | 136.8
| Technolutions | 2024-05-07 | Full range of Y Detector during Y Scan sweep | 74
| Technolutions | 2024-05-07 | Full range of Y Detector during Y Scan sweep | 67.2
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Full range of X2 Detector raw | 13.9
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | Full range of X1 Detector raw | 13.6
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Full range of X Detector Raw during X Scan sweep | 13.4
| Technolutions | 2024-05-07 | Full range of X Detector Raw during X Scan sweep | 14.3
| Technolutions | 2024-05-07 | Full range of X Detector Raw during X Scan sweep | 13.4
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | Full range of X Detector during X Scan sweep | 138.5
| Technolutions | 2024-05-07 | Full range of X Detector during X Scan sweep | 74
| Technolutions | 2024-05-07 | Full range of X Detector during X Scan sweep | 67.2
| Xinxing Jihua (Shanghai) Engneering Science and Technology Research Institute | 2024-05-07 | A-B noise | 9.07
| Technolutions | 2024-05-07 | A-B noise | 11.61
| Technolutions | 2024-05-07 | A-B noise | 7.49
| Park Systems Europe GmbH (PSE)_NX20 | 2024-05-07 | A-B noise | 9.42
| SMSC TD | 2024-04-30 | Z Linearity by Z scanner position | 0.28
| SMSC TD | 2024-04-30 | Z Linearity by 3 different sample height | 1.0017
| SMSC TD | 2024-04-30 | Z Height noise | 0.024
| SMSC TD | 2024-04-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 23.9441
| SMSC TD | 2024-04-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.9257
| SMSC TD | 2024-04-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2665
| SMSC TD | 2024-04-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2791
| SMSC TD | 2024-04-30 | Z Drift | 16.8
| SMSC TD | 2024-04-30 | Z Detector offset | 65.6
| SMSC TD | 2024-04-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 264.575
| SMSC TD | 2024-04-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 244.2206
| SMSC TD | 2024-04-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.439
| SMSC TD | 2024-04-30 | Y Ringing(Rpv) | 0.194
| SMSC TD | 2024-04-30 | Y Non-linearity error | 3.54
| SMSC TD | 2024-04-30 | Y axis, Repeatability @1st position | 0.05
| SMSC TD | 2024-04-30 | Y axis, Repeatability @ 4th position | 0.04
| SMSC TD | 2024-04-30 | Y axis, Repeatability @ 3rd position | 0.03
| SMSC TD | 2024-04-30 | Y axis, Repeatability @ 2nd position | 0.03
| SMSC TD | 2024-04-30 | Y axis, Accuracy @1st position | 0.3
| SMSC TD | 2024-04-30 | Y axis, Accuracy @ 4th position | 0.19
| SMSC TD | 2024-04-30 | Y axis, Accuracy @ 3rd position | 0.16
| SMSC TD | 2024-04-30 | Y axis, Accuracy @ 2nd position | 0.15
| SMSC TD | 2024-04-30 | Y axis OPM | 1.933
| SMSC TD | 2024-04-30 | XY Stage positioning Y offset (Vision Pattern Matching) | 2.29
| SMSC TD | 2024-04-30 | XY Stage positioning Y offset (LLC Translation) | 12.34
| SMSC TD | 2024-04-30 | XY Stage positioning Y offset (Global Align) | 8.68
| SMSC TD | 2024-04-30 | XY Stage positioning X offset (Vision Pattern Matching) | 2.14
| SMSC TD | 2024-04-30 | XY Stage positioning X offset (LLC Translation) | 2.95
| SMSC TD | 2024-04-30 | XY Stage positioning X offset (Global Align) | 6.91
| SMSC TD | 2024-04-30 | XY Orthogonality | 90.2
| SMSC TD | 2024-04-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 262.018
| SMSC TD | 2024-04-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 264.2196
| SMSC TD | 2024-04-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 175.402
| SMSC TD | 2024-04-30 | X Ringing(Rpv) | 0.291
| SMSC TD | 2024-04-30 | X Non-linearity error | 3.71
| SMSC TD | 2024-04-30 | X axis, Repeatability @1st position | 0.04
| SMSC TD | 2024-04-30 | X axis, Repeatability @ 4th position | 0.04
| SMSC TD | 2024-04-30 | X axis, Repeatability @ 3rd position | 0.04
| SMSC TD | 2024-04-30 | X axis, Repeatability @ 2nd position | 0.04
| SMSC TD | 2024-04-30 | X axis, Accuracy @1st position | 0.21
| SMSC TD | 2024-04-30 | X axis, Accuracy @ 4th position | 0.17
| SMSC TD | 2024-04-30 | X axis, Accuracy @ 3rd position | 0.18
| SMSC TD | 2024-04-30 | X axis, Accuracy @ 2nd position | 0.2
| SMSC TD | 2024-04-30 | X axis OPM | 0.608
| SMSC TD | 2024-04-30 | Wafer placement(Loading) error on the load port2 (Y offset) | 22.86
| SMSC TD | 2024-04-30 | Wafer placement(Loading) error on the load port2 (X offset) | 26.95
| SMSC TD | 2024-04-30 | Wafer placement(Loading) error on the load port1 (Y offset) | 13.2
| SMSC TD | 2024-04-30 | Wafer placement(Loading) error on the load port1 (X offset) | 18.8
| SMSC TD | 2024-04-30 | Torque Profile (50mm, 20um/s) | 2.226
| SMSC TD | 2024-04-30 | Torque Profile (50mm, 200um/s) | 1.649
| SMSC TD | 2024-04-30 | StrokeYMin | -153683
| SMSC TD | 2024-04-30 | StrokeYMax | 150430
| SMSC TD | 2024-04-30 | StrokeY | 304113
| SMSC TD | 2024-04-30 | StrokeXMin | -203771
| SMSC TD | 2024-04-30 | StrokeXMax | 203673
| SMSC TD | 2024-04-30 | StrokeX | 407444
| SMSC TD | 2024-04-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.42
| SMSC TD | 2024-04-30 | Roll motion to X+ full scan | 21.5
| SMSC TD | 2024-04-30 | Position repeatability(10repeats, Position Based) | 0.22
| SMSC TD | 2024-04-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| SMSC TD | 2024-04-30 | OpticalAxisShiftPerFocusY | 0.0011
| SMSC TD | 2024-04-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 3e-05
| SMSC TD | 2024-04-30 | OpticalAxisShiftPerFocusX | -0.0027
| SMSC TD | 2024-04-30 | OPM Repeatability @5mm | 0.983
| SMSC TD | 2024-04-30 | OPM Repeatability @25mm | 3.835
| SMSC TD | 2024-04-30 | OPM Repeatability @1mm | 0.813
| SMSC TD | 2024-04-30 | OPM Repeatability @10mm | 1.339
| SMSC TD | 2024-04-30 | NCM driving% with AC160TS cantilever | 15.68
| SMSC TD | 2024-04-30 | Min. Vibration Transmissbility (~100Hz) | 0.02
| SMSC TD | 2024-04-30 | Max OPM @5mm | 24.233
| SMSC TD | 2024-04-30 | Max OPM @25mm | 136.642
| SMSC TD | 2024-04-30 | Max OPM @1mm | 7.859
| SMSC TD | 2024-04-30 | Max OPM @10mm | 40.132
| SMSC TD | 2024-04-30 | LowerLimit | 2393
| SMSC TD | 2024-04-30 | LowerLimit | 5119
| SMSC TD | 2024-04-30 | Lateral(C-D) sensitivity | 13.7
| SMSC TD | 2024-04-30 | Lateral(C-D) noise | 0.163
| SMSC TD | 2024-04-30 | Image offset after moving from -Y to +Y (Y scan) | 38.0859
| SMSC TD | 2024-04-30 | Image offset after moving from -Y to +Y (X scan) | 15.3809
| SMSC TD | 2024-04-30 | Image offset after moving from -X to +X (Y scan) | 34.4238
| SMSC TD | 2024-04-30 | Image offset after moving from -X to +X (X scan) | 10.9863
| SMSC TD | 2024-04-30 | Image offset after moving from +Y to -Y (Y scan) | 11.7188
| SMSC TD | 2024-04-30 | Image offset after moving from +Y to -Y (X scan) | 11.7188
| SMSC TD | 2024-04-30 | Image offset after moving from +X to -X (Y scan) | 36.6211
| SMSC TD | 2024-04-30 | Image offset after moving from +X to -X (X scan) | 8.7891
| SMSC TD | 2024-04-30 | Full range of Z Height during Z Scan sweep | 16.3
| SMSC TD | 2024-04-30 | Full range of Z Detector during Z Scan sweep | 13.62
| SMSC TD | 2024-04-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.03
| SMSC TD | 2024-04-30 | Full range of Y2 Detector during Y Scan sweep | 196.31
| SMSC TD | 2024-04-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.46
| SMSC TD | 2024-04-30 | Full range of Y1 Detector during Y Scan sweep | 189.33
| SMSC TD | 2024-04-30 | Full range of Y Detector during Y Scan sweep | 192.82
| SMSC TD | 2024-04-30 | Full range of X2 Detector Raw during X Scan sweep | 14.65
| SMSC TD | 2024-04-30 | Full range of X2 Detector during X Scan sweep | 203.12
| SMSC TD | 2024-04-30 | Full range of X1 Detector Raw during X Scan sweep | 14.67
| SMSC TD | 2024-04-30 | Full range of X1 Detector during X Scan sweep | 204.24
| SMSC TD | 2024-04-30 | Full range of X Detector during X Scan sweep | 203.68
| SMSC TD | 2024-04-30 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| SMSC TD | 2024-04-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.068646
| SMSC TD | 2024-04-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.38147
| SMSC TD | 2024-04-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.39483
| SMSC TD | 2024-04-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.071032
| SMSC TD | 2024-04-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| SMSC TD | 2024-04-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| SMSC TD | 2024-04-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1645
| SMSC TD | 2024-04-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.029689
| SMSC TD | 2024-04-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.2
| SMSC TD | 2024-04-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 1.4
| SMSC TD | 2024-04-30 | %GRR w/ roughness of non-patterened wafer | 2.6
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Z Scanner position at -3μm | 0.0017
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Z Scanner position at +3μm | 0.0162
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Z linearity by sample height | 1.003
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Z Height noise | 0.016
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Y Scan Stroke | 112.8
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Y Detector noise (XY Servo on) | 0.26
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Y Detector noise (XY Servo off) | 0.25
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Y axis OPM | 1.52
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | XY orthogonality | 89.8
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | X Scan Stroke | 111.7
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | X Detector noise (XY Servo on) | 0.25
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | X Detector noise (XY Servo off) | 0.24
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | X axis OPM | 1.19
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | NCM Amplitude noise | 9.43
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.871
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.16
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Lateral(C-D) sensitivity | 9.1
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Lateral(C-D) noise | 0.112
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Full range of Z Height during Z Scan sweep | 17.3
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Full range of Z Detector during Z Scan sweep | 15.6
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Full range of Y Detector during Y Scan sweep | 112.8
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Full range of X Detector Raw during X Scan sweep | 14
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | Full range of X Detector during X Scan sweep | 111.7
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #2 | 2024-04-29 | A-B noise | 12.16
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Linearity by Z scanner position | 1.21
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Linearity by 4 different sample height | 0.998
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Height noise | 0.036
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.2744
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.3146
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1743
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.1188
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Drift | 6.073
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Z Detector offset | 42
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y2 Detector noise (XY Servo on) | 0.312
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 280.6584
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y1 Detector noise (XY Servo on) | 0.307
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 269.8468
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 184.024
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y Ringing(Rpv) | 0.253
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y Detector noise (XY Servo on) | 0.244
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Repeatability @1st position | 0.04
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Repeatability @ 4th position | 0.03
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Repeatability @ 3rd position | 0.05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Repeatability @ 2nd position | 0.03
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Accuracy @1st position | 0.18
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Accuracy @ 4th position | 0.14
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Accuracy @ 3rd position | 0.23
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis, Accuracy @ 2nd position | 0.15
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Y axis OPM | 0.549
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Stage positioning Y offset (Vision Pattern Matching) | 3.6279
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Stage positioning Y offset (LLC Translation) | 2.17
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Stage positioning Y offset (Global Align) | 3.86
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Stage positioning X offset (Vision Pattern Matching) | 1.5186
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Stage positioning X offset (LLC Translation) | 3.1
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Stage positioning X offset (Global Align) | 4.3
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | XY Orthogonality | 90.1
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X2 Detector noise (XY Servo on) | 0.334
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 282.1026
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X1 Detector noise (XY Servo on) | 0.324
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 269.9462
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 179.705
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X Ringing(Rpv) | 0.252
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X Detector noise (XY Servo on) | 0.269
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Repeatability @1st position | 0.04
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Repeatability @ 4th position | 0.05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Repeatability @ 3rd position | 0.05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Repeatability @ 2nd position | 0.08
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Accuracy @1st position | 0.22
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Accuracy @ 4th position | 0.28
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Accuracy @ 3rd position | 0.21
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis, Accuracy @ 2nd position | 0.36
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | X axis OPM | 1.389
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Wafer placement(Loading) error on the load port2 (Y offset) | 33.31
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Wafer placement(Loading) error on the load port2 (X offset) | 48.1
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Wafer placement(Loading) error on the load port1 (Y offset) | 18.6
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Wafer placement(Loading) error on the load port1 (X offset) | 47.5
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Torque Profile (50mm, 20um/s) | 3.446
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Torque Profile (50mm, 200um/s) | 1.655
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | StrokeYMin | -152250
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | StrokeYMax | 155190
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | StrokeY | 307440
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | StrokeXMin | -205900
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | StrokeXMax | 207000
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | StrokeX | 412900
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.553
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Roll motion to X+ full scan | 2.8
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | RMS of ZeroScan | 0.047
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Position repeatability(10repeats, Position Based) | 0.67
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OpticalAxisShiftPerFocusY | -0.0007
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 4e-05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OpticalAxisShiftPerFocusX | 0.0003
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OPM Repeatability @5mm | 0.869
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OPM Repeatability @25mm | 1.784
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OPM Repeatability @1mm | 0.736
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | OPM Repeatability @10mm | 1.105
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | NCM driving% with AC160TS cantilever | 27.91
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Min. Vibration Transmissbility (~100Hz) | 0.003
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Max OPM @5mm | 8.923
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Max OPM @25mm | 97.947
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Max OPM @1mm | 5.714
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Max OPM @10mm | 34.52
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | LowerLimit | 2168
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | LowerLimit | 5399
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Lateral(C-D) sensitivity | 10.4
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Lateral(C-D) noise | 0.238
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from -Y to +Y (Y scan) | 88.6231
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from -Y to +Y (X scan) | 12.4512
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from -X to +X (Y scan) | 40.2832
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from -X to +X (X scan) | 54.9316
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from +Y to -Y (Y scan) | 64.4531
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from +Y to -Y (X scan) | 30.0293
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from +X to -X (Y scan) | 52.002
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Image offset after moving from +X to -X (X scan) | 79.834
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Z Height during Z Scan sweep | 16.1
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Z Detector during Z Scan sweep | 11.3
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.33
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Y2 Detector during Y Scan sweep | 204.86
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Y1 Detector during Y Scan sweep | 205.65
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of Y Detector during Y Scan sweep | 205.65
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of X2 Detector Raw during X Scan sweep | 14.34
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of X2 Detector during X Scan sweep | 216.05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of X1 Detector Raw during X Scan sweep | 14.38
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of X1 Detector during X Scan sweep | 215.81
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Full range of X Detector during X Scan sweep | 216.05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.35
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.25
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 9.6
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 1.1
| Runpeng Semiconductor (Shenzhen) Co., LTD. | 2024-04-26 | %GRR w/ roughness of non-patterened wafer | 9.7
| Aixtron SE | 2024-04-25 | Z Linearity by Z scanner position | 0.63
| Aixtron SE | 2024-04-25 | Z Linearity by 4 different sample height | 1.001
| Aixtron SE | 2024-04-25 | Z Height noise | 0.038
| Aixtron SE | 2024-04-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.8
| Aixtron SE | 2024-04-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.8
| Aixtron SE | 2024-04-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.3
| Aixtron SE | 2024-04-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3
| Aixtron SE | 2024-04-25 | Z Drift (Short Traveling) | 0.03
| Aixtron SE | 2024-04-25 | Y2 Detector noise (XY Servo on) | 0.63
| Aixtron SE | 2024-04-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.8
| Aixtron SE | 2024-04-25 | Y1 Detector noise (XY Servo on) | 0.62
| Aixtron SE | 2024-04-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 268.4
| Aixtron SE | 2024-04-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 191.5
| Aixtron SE | 2024-04-25 | Y Ringing(Rpv) | 0.22
| Aixtron SE | 2024-04-25 | Y Non-linearity error | 2.62
| Aixtron SE | 2024-04-25 | Y Detector noise (XY Servo on) | 0.57
| Aixtron SE | 2024-04-25 | Y axis, Repeatability @1st position | 0.05
| Aixtron SE | 2024-04-25 | Y axis, Repeatability @ 4th position | 0.04
| Aixtron SE | 2024-04-25 | Y axis, Repeatability @ 3rd position | 0.03
| Aixtron SE | 2024-04-25 | Y axis, Repeatability @ 2nd position | 0.06
| Aixtron SE | 2024-04-25 | Y axis, Accuracy @1st position | 0.53
| Aixtron SE | 2024-04-25 | Y axis, Accuracy @ 4th position | 0.42
| Aixtron SE | 2024-04-25 | Y axis, Accuracy @ 3rd position | 0.33
| Aixtron SE | 2024-04-25 | Y axis, Accuracy @ 2nd position | 0.42
| Aixtron SE | 2024-04-25 | Y axis OPM | 0.65
| Aixtron SE | 2024-04-25 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.8848
| Aixtron SE | 2024-04-25 | XY Stage positioning Y offset (LLC Translation) | 1.94
| Aixtron SE | 2024-04-25 | XY Stage positioning Y offset (Global Align) | 4.21
| Aixtron SE | 2024-04-25 | XY Stage positioning X offset (Vision Pattern Matching) | 0.75195
| Aixtron SE | 2024-04-25 | XY Stage positioning X offset (LLC Translation) | 2.09
| Aixtron SE | 2024-04-25 | XY Stage positioning X offset (Global Align) | 2.42
| Aixtron SE | 2024-04-25 | XY Orthogonality | 89.7
| Aixtron SE | 2024-04-25 | X2 Detector noise (XY Servo on) | 0.42
| Aixtron SE | 2024-04-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 267
| Aixtron SE | 2024-04-25 | X1 Detector noise (XY Servo on) | 0.41
| Aixtron SE | 2024-04-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 263
| Aixtron SE | 2024-04-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 181.2
| Aixtron SE | 2024-04-25 | X Ringing(Rpv) | 0.46
| Aixtron SE | 2024-04-25 | X Non-linearity error | 2.18
| Aixtron SE | 2024-04-25 | X Detector noise (XY Servo on) | 0.32
| Aixtron SE | 2024-04-25 | X axis, Repeatability @1st position | 0.03
| Aixtron SE | 2024-04-25 | X axis, Repeatability @ 4th position | 0.04
| Aixtron SE | 2024-04-25 | X axis, Repeatability @ 3rd position | 0.03
| Aixtron SE | 2024-04-25 | X axis, Repeatability @ 2nd position | 0.05
| Aixtron SE | 2024-04-25 | X axis, Accuracy @1st position | 0.17
| Aixtron SE | 2024-04-25 | X axis, Accuracy @ 4th position | 0.24
| Aixtron SE | 2024-04-25 | X axis, Accuracy @ 3rd position | 0.18
| Aixtron SE | 2024-04-25 | X axis, Accuracy @ 2nd position | 0.26
| Aixtron SE | 2024-04-25 | X axis OPM | 1.44
| Aixtron SE | 2024-04-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 25.77
| Aixtron SE | 2024-04-25 | Wafer placement(Loading) error on the load port2 (X offset) | 20.3
| Aixtron SE | 2024-04-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 34.57
| Aixtron SE | 2024-04-25 | Wafer placement(Loading) error on the load port1 (X offset) | 59
| Aixtron SE | 2024-04-25 | Under 50nm Step Height Sample (10 Repeats) | 18.06
| Aixtron SE | 2024-04-25 | Torque Profile (50mm, 20um/s)
| Aixtron SE | 2024-04-25 | Torque Profile (50mm, 200um/s)
| Aixtron SE | 2024-04-25 | StrokeYMin | -153069
| Aixtron SE | 2024-04-25 | StrokeYMax | 150269.7
| Aixtron SE | 2024-04-25 | StrokeY | 303338.6
| Aixtron SE | 2024-04-25 | StrokeXMin | -204032
| Aixtron SE | 2024-04-25 | StrokeXMax | 202797.3
| Aixtron SE | 2024-04-25 | StrokeX | 406829.5
| Aixtron SE | 2024-04-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.931
| Aixtron SE | 2024-04-25 | Roll motion to X+ full scan
| Aixtron SE | 2024-04-25 | RMS of ZeroScan | 0.045
| Aixtron SE | 2024-04-25 | Profiler noise @200um
| Aixtron SE | 2024-04-25 | Position repeatability(10repeats, Position Based)
| Aixtron SE | 2024-04-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| Aixtron SE | 2024-04-25 | OpticalAxisShiftPerFocusY | 0.0001
| Aixtron SE | 2024-04-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| Aixtron SE | 2024-04-25 | OpticalAxisShiftPerFocusX | -0.0028
| Aixtron SE | 2024-04-25 | OPM Repeatability @5mm
| Aixtron SE | 2024-04-25 | OPM Repeatability @25mm
| Aixtron SE | 2024-04-25 | OPM Repeatability @1mm
| Aixtron SE | 2024-04-25 | OPM Repeatability @10mm
| Aixtron SE | 2024-04-25 | NCM driving% with AC160TS cantilever | 8.1
| Aixtron SE | 2024-04-25 | Min. Vibration Transmissbility (~100Hz) | 0.021
| Aixtron SE | 2024-04-25 | Max OPM @5mm
| Aixtron SE | 2024-04-25 | Max OPM @25mm
| Aixtron SE | 2024-04-25 | Max OPM @1mm
| Aixtron SE | 2024-04-25 | Max OPM @10mm
| Aixtron SE | 2024-04-25 | LowerLimit | 3273
| Aixtron SE | 2024-04-25 | Lift pins synchronization offset check (Y offset)
| Aixtron SE | 2024-04-25 | Lift pins synchronization offset check (X offset)
| Aixtron SE | 2024-04-25 | Lateral(C-D) sensitivity | 14.4
| Aixtron SE | 2024-04-25 | Lateral(C-D) noise | 0.233
| Aixtron SE | 2024-04-25 | Image offset after moving from -Y to +Y (Y scan) | 38.09
| Aixtron SE | 2024-04-25 | Image offset after moving from -Y to +Y (X scan) | 9.52
| Aixtron SE | 2024-04-25 | Image offset after moving from -X to +X (Y scan) | 39.55
| Aixtron SE | 2024-04-25 | Image offset after moving from -X to +X (X scan) | 64.45
| Aixtron SE | 2024-04-25 | Image offset after moving from +Y to -Y (Y scan) | 34.42
| Aixtron SE | 2024-04-25 | Image offset after moving from +Y to -Y (X scan) | 7.32
| Aixtron SE | 2024-04-25 | Image offset after moving from +X to -X (Y scan) | 20.51
| Aixtron SE | 2024-04-25 | Image offset after moving from +X to -X (X scan) | 35.16
| Aixtron SE | 2024-04-25 | Full range of Z Height during Z Scan sweep | 16.6
| Aixtron SE | 2024-04-25 | Full range of Z Detector during Z Scan sweep | 12.5
| Aixtron SE | 2024-04-25 | Full range of Y2 Detector Raw during Y Scan sweep | 13.62
| Aixtron SE | 2024-04-25 | Full range of Y2 Detector during Y Scan sweep | 203.6
| Aixtron SE | 2024-04-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.34
| Aixtron SE | 2024-04-25 | Full range of Y1 Detector during Y Scan sweep | 209.8
| Aixtron SE | 2024-04-25 | Full range of X2 Detector Raw during X Scan sweep | 13.73
| Aixtron SE | 2024-04-25 | Full range of X2 Detector during X Scan sweep | 199.5
| Aixtron SE | 2024-04-25 | Full range of X1 Detector Raw during X Scan sweep | 13.85
| Aixtron SE | 2024-04-25 | Full range of X1 Detector during X Scan sweep | 195
| Aixtron SE | 2024-04-25 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Aixtron SE | 2024-04-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.16
| Aixtron SE | 2024-04-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.9
| Aixtron SE | 2024-04-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.54
| Aixtron SE | 2024-04-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| Aixtron SE | 2024-04-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Aixtron SE | 2024-04-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.51
| Aixtron SE | 2024-04-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.32
| Aixtron SE | 2024-04-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| Aixtron SE | 2024-04-25 | 940nm Step Height Sample (10 Repeats) | 948.92
| Aixtron SE | 2024-04-25 | 50~200nm Step Height Sample (10 Repeats) | 182.94
| Aixtron SE | 2024-04-25 | 200~500nm Step Height Sample (10 Repeats) | 462.17
| Aixtron SE | 2024-04-25 | * 1800nm Step Height Sample (10 Repeats) | 1795.61
| Aixtron SE | 2024-04-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.6
| Aixtron SE | 2024-04-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.1
| Aixtron SE | 2024-04-25 | %GRR w/ roughness of non-patterened wafer | 10
| Shenzhen Pengjin High-Tech | 2024-04-25 | Z Scanner position at -3μm | 0.0033
| Shenzhen Pengjin High-Tech | 2024-04-25 | Z Scanner position at +3μm | 0.0066
| Shenzhen Pengjin High-Tech | 2024-04-25 | Z linearity by sample height | 1.001
| Shenzhen Pengjin High-Tech | 2024-04-25 | Z Height noise | 0.027
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y2 Detector noise (XY Servo on) | 0.28
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y2 Detector noise (XY Servo off) | 0.27
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y1 Detector noise (XY Servo on) | 0.28
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y1 Detector noise (XY Servo off) | 0.28
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y Scan Stroke | 173.9
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y Detector noise (XY Servo on) | 0.2
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y Detector noise (XY Servo off) | 0.19
| Shenzhen Pengjin High-Tech | 2024-04-25 | Y axis OPM | 1.18
| Shenzhen Pengjin High-Tech | 2024-04-25 | X2 Detector noise (XY Servo on) | 0.22
| Shenzhen Pengjin High-Tech | 2024-04-25 | X2 Detector noise (XY Servo off) | 0.23
| Shenzhen Pengjin High-Tech | 2024-04-25 | X1 Detector noise (XY Servo on) | 0.22
| Shenzhen Pengjin High-Tech | 2024-04-25 | X1 Detector noise (XY Servo off) | 0.22
| Shenzhen Pengjin High-Tech | 2024-04-25 | X Scan Stroke | 177.1
| Shenzhen Pengjin High-Tech | 2024-04-25 | X Detector noise (XY Servo on) | 0.15
| Shenzhen Pengjin High-Tech | 2024-04-25 | X Detector noise (XY Servo off) | 0.16
| Shenzhen Pengjin High-Tech | 2024-04-25 | X axis OPM | 0.96
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.39
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.41
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.37
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.33
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.34
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Shenzhen Pengjin High-Tech | 2024-04-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Shenzhen Pengjin High-Tech | 2024-04-25 | NCM Amplitude noise | 15.579
| Shenzhen Pengjin High-Tech | 2024-04-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.437
| Shenzhen Pengjin High-Tech | 2024-04-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.971
| Shenzhen Pengjin High-Tech | 2024-04-25 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 236.4
| Shenzhen Pengjin High-Tech | 2024-04-25 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.2
| Shenzhen Pengjin High-Tech | 2024-04-25 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 246.7
| Shenzhen Pengjin High-Tech | 2024-04-25 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 240.3
| Shenzhen Pengjin High-Tech | 2024-04-25 | Maximum Y Scan range (Full Stroke) | 173.7
| Shenzhen Pengjin High-Tech | 2024-04-25 | Maximum X Scan range (Full Stroke) | 177.1
| Shenzhen Pengjin High-Tech | 2024-04-25 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Shenzhen Pengjin High-Tech | 2024-04-25 | Lateral(C-D) sensitivity | 14.2
| Shenzhen Pengjin High-Tech | 2024-04-25 | Lateral(C-D) noise | 0.126
| Shenzhen Pengjin High-Tech | 2024-04-25 | Full range of Z Height during Z Scan sweep | 16.5
| Shenzhen Pengjin High-Tech | 2024-04-25 | Full range of Z Detector during Z Scan sweep | 12.7
| Shenzhen Pengjin High-Tech | 2024-04-25 | Full range of Y2 Detector raw | 14.5
| Shenzhen Pengjin High-Tech | 2024-04-25 | Full range of Y1 Detector raw | 14.6
| Shenzhen Pengjin High-Tech | 2024-04-25 | Full range of X2 Detector raw | 14.5
| Shenzhen Pengjin High-Tech | 2024-04-25 | Full range of X1 Detector raw | 14.6
| Shenzhen Pengjin High-Tech | 2024-04-25 | A-B noise | 15.02
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Scanner position at -3μm | 0.0004
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Scanner position at -3μm | 0.0022
| IQS Nanoptiqs | 2024-04-24 | Z Scanner position at -3μm | 0.0012
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Scanner position at +3μm | 0.0079
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Scanner position at +3μm | 0.0139
| IQS Nanoptiqs | 2024-04-24 | Z Scanner position at +3μm | 0.0055
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Linearity by Z Scanner position
| University of Illinois at Urbana Champaign | 2024-04-24 | Z linearity by sample height | 1.0113
| IQS Nanoptiqs | 2024-04-24 | Z linearity by sample height | 1.006
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Linearity by 3 different sample height | 0.995
| IQS Nanoptiqs | 2024-04-24 | Z Linearity by 3 different sample height | 1.003
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Height noise | 0.027
| University of Illinois at Urbana Champaign | 2024-04-24 | Z Height noise | 0.032
| IQS Nanoptiqs | 2024-04-24 | Z Height noise | 0.021
| IQS Nanoptiqs | 2024-04-24 | Y2 Detector noise (XY Servo on) | 0.28
| IQS Nanoptiqs | 2024-04-24 | Y2 Detector noise (XY Servo off) | 0.27
| IQS Nanoptiqs | 2024-04-24 | Y1 Detector noise (XY Servo on) | 0.32
| IQS Nanoptiqs | 2024-04-24 | Y1 Detector noise (XY Servo off) | 0.32
| University of Illinois at Urbana Champaign | 2024-04-24 | Y Scan Stroke | 141.7
| University of Illinois at Urbana Champaign | 2024-04-24 | Y Scan Stroke | 132.6
| IQS Nanoptiqs | 2024-04-24 | Y Scan Stroke | 179.7
| University of Illinois at Urbana Champaign | 2024-04-24 | -Y Direction software Limit | -75687.5
| University of Illinois at Urbana Champaign | 2024-04-24 | Y Detector noise (XY Servo on) | 0.26
| University of Illinois at Urbana Champaign | 2024-04-24 | Y Detector noise (XY Servo on) | 0.39
| IQS Nanoptiqs | 2024-04-24 | Y Detector noise (XY Servo on) | 0.22
| University of Illinois at Urbana Champaign | 2024-04-24 | Y Detector noise (XY Servo off) | 0.25
| University of Illinois at Urbana Champaign | 2024-04-24 | Y Detector noise (XY Servo off) | 0.4
| IQS Nanoptiqs | 2024-04-24 | Y Detector noise (XY Servo off) | 0.22
| University of Illinois at Urbana Champaign | 2024-04-24 | Y axis OPM | 0.47
| University of Illinois at Urbana Champaign | 2024-04-24 | Y axis OPM | 1.43
| IQS Nanoptiqs | 2024-04-24 | Y axis OPM | 1.05
| IQS Nanoptiqs | 2024-04-24 | X2 Detector noise (XY Servo on) | 0.36
| IQS Nanoptiqs | 2024-04-24 | X2 Detector noise (XY Servo off) | 0.35
| IQS Nanoptiqs | 2024-04-24 | X1 Detector noise (XY Servo on) | 0.23
| IQS Nanoptiqs | 2024-04-24 | X1 Detector noise (XY Servo off) | 0.23
| University of Illinois at Urbana Champaign | 2024-04-24 | X Scan Stroke | 134.7
| University of Illinois at Urbana Champaign | 2024-04-24 | X Scan Stroke | 133.6
| IQS Nanoptiqs | 2024-04-24 | X Scan Stroke | 177.1
| University of Illinois at Urbana Champaign | 2024-04-24 | -X Direction software Limit | -76094.4
| University of Illinois at Urbana Champaign | 2024-04-24 | X Detector noise (XY Servo on) | 0.25
| University of Illinois at Urbana Champaign | 2024-04-24 | X Detector noise (XY Servo on) | 0.4
| IQS Nanoptiqs | 2024-04-24 | X Detector noise (XY Servo on) | 0.22
| University of Illinois at Urbana Champaign | 2024-04-24 | X Detector noise (XY Servo off) | 0.25
| University of Illinois at Urbana Champaign | 2024-04-24 | X Detector noise (XY Servo off) | 0.42
| IQS Nanoptiqs | 2024-04-24 | X Detector noise (XY Servo off) | 0.22
| University of Illinois at Urbana Champaign | 2024-04-24 | X axis OPM | 0.46
| University of Illinois at Urbana Champaign | 2024-04-24 | X axis OPM | 0.86
| IQS Nanoptiqs | 2024-04-24 | X axis OPM | 1.8
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.57
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.31
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.38
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.3
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.37
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.32
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.37
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.3
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.36
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.37
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| University of Illinois at Urbana Champaign | 2024-04-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| IQS Nanoptiqs | 2024-04-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| University of Illinois at Urbana Champaign | 2024-04-24 | NCM Amplitude noise | 18.305
| University of Illinois at Urbana Champaign | 2024-04-24 | NCM Amplitude noise | 12.572
| IQS Nanoptiqs | 2024-04-24 | NCM Amplitude noise | 8.343
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.712
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.182
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.999
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 24.957
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.758
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.707
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.138
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 59.863
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 258.7
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 249.2
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 172.4
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 195.3
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 250.7
| IQS Nanoptiqs | 2024-04-24 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 244.7
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 166.7
| University of Illinois at Urbana Champaign | 2024-04-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 202.1
| IQS Nanoptiqs | 2024-04-24 | Maximum Y Scan range (Full Stroke) | 179.6
| IQS Nanoptiqs | 2024-04-24 | Maximum X Scan range (Full Stroke) | 176.8
| University of Illinois at Urbana Champaign | 2024-04-24 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| IQS Nanoptiqs | 2024-04-24 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| University of Illinois at Urbana Champaign | 2024-04-24 | Lateral(C-D) sensitivity | 4.2
| University of Illinois at Urbana Champaign | 2024-04-24 | Lateral(C-D) sensitivity | 14.1
| IQS Nanoptiqs | 2024-04-24 | Lateral(C-D) sensitivity | 15.8
| University of Illinois at Urbana Champaign | 2024-04-24 | Lateral(C-D) noise | 0.13
| University of Illinois at Urbana Champaign | 2024-04-24 | Lateral(C-D) noise | 0.114
| IQS Nanoptiqs | 2024-04-24 | Lateral(C-D) noise | 0.116
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Z Height during Z Scan sweep | 16.8
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Z Height during Z Scan sweep | 17.9
| IQS Nanoptiqs | 2024-04-24 | Full range of Z Height during Z Scan sweep | 17.4
| IQS Nanoptiqs | 2024-04-24 | Full range of Z Height during Z Scan sweep | 40.6
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Z Detector during Z Scan sweep | 13.1
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Z Detector during Z Scan sweep | 13.9
| IQS Nanoptiqs | 2024-04-24 | Full range of Z Detector during Z Scan sweep | 13.3
| IQS Nanoptiqs | 2024-04-24 | Full range of Z Detector during Z Scan sweep | 13.6
| IQS Nanoptiqs | 2024-04-24 | Full range of Y2 Detector raw | 13.9
| IQS Nanoptiqs | 2024-04-24 | Full range of Y1 Detector raw | 14.4
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Y Detector Raw during Y Scan sweep | 15.8
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Y Detector during Y Scan sweep | 135.8
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of Y Detector during Y Scan sweep | 134.4
| IQS Nanoptiqs | 2024-04-24 | Full range of X2 Detector raw | 14.2
| IQS Nanoptiqs | 2024-04-24 | Full range of X1 Detector raw | 14.3
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of X Detector Raw during X Scan sweep | 15.5
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of X Detector Raw during X Scan sweep | 13.4
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of X Detector during X Scan sweep | 129
| University of Illinois at Urbana Champaign | 2024-04-24 | Full range of X Detector during X Scan sweep | 135.4
| IQS Nanoptiqs | 2024-04-24 | Approach position at -5μm | 0.0002
| IQS Nanoptiqs | 2024-04-24 | Approach position at +5μm | 0.0065
| University of Illinois at Urbana Champaign | 2024-04-24 | A-B noise | 9.074
| University of Illinois at Urbana Champaign | 2024-04-24 | A-B noise | 10.186
| IQS Nanoptiqs | 2024-04-24 | A-B noise | 13.594
| HK SAE co.,ltd | 2024-04-19 | Z linearity by sample height | 0.997
| HK SAE co.,ltd | 2024-04-19 | Z Height noise | 0.027
| HK SAE co.,ltd | 2024-04-19 | Y2 Detector noise (XY Servo on) | 0.24
| HK SAE co.,ltd | 2024-04-19 | Y2 Detector noise (XY Servo off) | 0.28
| HK SAE co.,ltd | 2024-04-19 | Y1 Detector noise (XY Servo on) | 0.23
| HK SAE co.,ltd | 2024-04-19 | Y1 Detector noise (XY Servo off) | 0.24
| HK SAE co.,ltd | 2024-04-19 | Y Scan Stroke | 178.2
| HK SAE co.,ltd | 2024-04-19 | Y Detector noise (XY Servo on) | 0.18
| HK SAE co.,ltd | 2024-04-19 | Y Detector noise (XY Servo off) | 0.2
| HK SAE co.,ltd | 2024-04-19 | Y axis OPM | 0.99
| HK SAE co.,ltd | 2024-04-19 | X2 Detector noise (XY Servo on) | 0.23
| HK SAE co.,ltd | 2024-04-19 | X2 Detector noise (XY Servo off) | 0.23
| HK SAE co.,ltd | 2024-04-19 | X1 Detector noise (XY Servo on) | 0.23
| HK SAE co.,ltd | 2024-04-19 | X1 Detector noise (XY Servo off) | 0.24
| HK SAE co.,ltd | 2024-04-19 | X Scan Stroke | 178.1
| HK SAE co.,ltd | 2024-04-19 | X Detector noise (XY Servo on) | 0.17
| HK SAE co.,ltd | 2024-04-19 | X Detector noise (XY Servo off) | 0.18
| HK SAE co.,ltd | 2024-04-19 | X axis OPM | 1.31
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.28
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.34
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.29
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.27
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| HK SAE co.,ltd | 2024-04-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| HK SAE co.,ltd | 2024-04-19 | NCM Amplitude noise | 12.31
| HK SAE co.,ltd | 2024-04-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.14
| HK SAE co.,ltd | 2024-04-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.644
| HK SAE co.,ltd | 2024-04-19 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 250.7
| HK SAE co.,ltd | 2024-04-19 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 244.3
| HK SAE co.,ltd | 2024-04-19 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 247.4
| HK SAE co.,ltd | 2024-04-19 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 247.2
| HK SAE co.,ltd | 2024-04-19 | Maximum Y Scan range (Full Stroke) | 177.4
| HK SAE co.,ltd | 2024-04-19 | Maximum X Scan range (Full Stroke) | 179.2
| HK SAE co.,ltd | 2024-04-19 | Full range of Z Height during Z Scan sweep | 18.2
| HK SAE co.,ltd | 2024-04-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| HK SAE co.,ltd | 2024-04-19 | Lateral(C-D) sensitivity | 10
| HK SAE co.,ltd | 2024-04-19 | Lateral(C-D) noise | 0.117
| HK SAE co.,ltd | 2024-04-19 | Full range of Z Detector raw | 14.8
| HK SAE co.,ltd | 2024-04-19 | Full range of Y2 Detector raw | 14.4
| HK SAE co.,ltd | 2024-04-19 | Full range of Y1 Detector raw | 14.5
| HK SAE co.,ltd | 2024-04-19 | Full range of X2 Detector raw | 14.4
| HK SAE co.,ltd | 2024-04-19 | Full range of X1 Detector raw | 14.4
| HK SAE co.,ltd | 2024-04-19 | Approach position at -3μm | 0.0015
| HK SAE co.,ltd | 2024-04-19 | Approach position at +3μm | 0.0024
| HK SAE co.,ltd | 2024-04-19 | A-B noise | 7.2
| Jiangsu Jixin Advance Materials | 2024-04-11 | Z linearity by sample height | 1.005
| Jiangsu Jixin Advance Materials | 2024-04-11 | Z Height noise | 0.029
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y2 Detector noise (XY Servo on) | 0.24
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y2 Detector noise (XY Servo off) | 0.23
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y1 Detector noise (XY Servo on) | 0.25
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y1 Detector noise (XY Servo off) | 0.25
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y Scan Stroke | 181.5
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y Detector noise (XY Servo on) | 0.17
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y Detector noise (XY Servo off) | 0.17
| Jiangsu Jixin Advance Materials | 2024-04-11 | Y axis OPM | 1.24
| Jiangsu Jixin Advance Materials | 2024-04-11 | X2 Detector noise (XY Servo on) | 0.25
| Jiangsu Jixin Advance Materials | 2024-04-11 | X2 Detector noise (XY Servo off) | 0.25
| Jiangsu Jixin Advance Materials | 2024-04-11 | X1 Detector noise (XY Servo on) | 0.26
| Jiangsu Jixin Advance Materials | 2024-04-11 | X1 Detector noise (XY Servo off) | 0.25
| Jiangsu Jixin Advance Materials | 2024-04-11 | X Scan Stroke | 182.7
| Jiangsu Jixin Advance Materials | 2024-04-11 | X Detector noise (XY Servo on) | 0.19
| Jiangsu Jixin Advance Materials | 2024-04-11 | X Detector noise (XY Servo off) | 0.19
| Jiangsu Jixin Advance Materials | 2024-04-11 | X axis OPM | 0.73
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.36
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.3
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.32
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Jiangsu Jixin Advance Materials | 2024-04-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Jiangsu Jixin Advance Materials | 2024-04-11 | NCM Amplitude noise | 14.14
| Jiangsu Jixin Advance Materials | 2024-04-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.355
| Jiangsu Jixin Advance Materials | 2024-04-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.713
| Jiangsu Jixin Advance Materials | 2024-04-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 251.9
| Jiangsu Jixin Advance Materials | 2024-04-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 250.6
| Jiangsu Jixin Advance Materials | 2024-04-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 252.6
| Jiangsu Jixin Advance Materials | 2024-04-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 251.3
| Jiangsu Jixin Advance Materials | 2024-04-11 | Maximum Y Scan range (Full Stroke) | 181.6
| Jiangsu Jixin Advance Materials | 2024-04-11 | Maximum X Scan range (Full Stroke) | 182.1
| Jiangsu Jixin Advance Materials | 2024-04-11 | Full range of Z Height during Z Scan sweep | 18.3
| Jiangsu Jixin Advance Materials | 2024-04-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Jiangsu Jixin Advance Materials | 2024-04-11 | Lateral(C-D) sensitivity | 13
| Jiangsu Jixin Advance Materials | 2024-04-11 | Lateral(C-D) noise | 0.139
| Jiangsu Jixin Advance Materials | 2024-04-11 | Full range of Z Detector raw | 14.3
| Jiangsu Jixin Advance Materials | 2024-04-11 | Full range of Y2 Detector raw | 14.4
| Jiangsu Jixin Advance Materials | 2024-04-11 | Full range of Y1 Detector raw | 14.5
| Jiangsu Jixin Advance Materials | 2024-04-11 | Full range of X2 Detector raw | 14.5
| Jiangsu Jixin Advance Materials | 2024-04-11 | Full range of X1 Detector raw | 14.5
| Jiangsu Jixin Advance Materials | 2024-04-11 | Approach position at -3μm | 0.0055
| Jiangsu Jixin Advance Materials | 2024-04-11 | Approach position at +3μm | 0.0033
| Jiangsu Jixin Advance Materials | 2024-04-11 | A-B noise | 9.49
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Z Scanner position at -3μm | 0.0014
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Z Scanner position at +3μm | 0.0003
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Z linearity by sample height | 1.001
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Z Height noise | 0.027
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y2 Detector noise (XY Servo on) | 0.28
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y2 Detector noise (XY Servo off) | 0.28
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y1 Detector noise (XY Servo on) | 0.27
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y1 Detector noise (XY Servo off) | 0.27
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y Scan Stroke | 182.2
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y Detector noise (XY Servo on) | 0.22
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y Detector noise (XY Servo off) | 0.22
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Y axis OPM | 0.59
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X2 Detector noise (XY Servo on) | 0.27
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X2 Detector noise (XY Servo off) | 0.26
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X1 Detector noise (XY Servo on) | 0.26
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X1 Detector noise (XY Servo off) | 0.25
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X Scan Stroke | 186
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X Detector noise (XY Servo on) | 0.2
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X Detector noise (XY Servo off) | 0.19
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | X axis OPM | 3.49
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of Wafer Zeroscan(-120000,0) | 0.32
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of Wafer Zeroscan(120000,0) | 0.31
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of Wafer Zeroscan(0,-120000) | 0.29
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of Wafer Zeroscan(0,120000) | 0.27
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | NCM Amplitude noise | 12.55
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.396
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.72
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 261.1
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 259.5
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 265.1
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 264.5
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Lateral(C-D) sensitivity | 11.1
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Lateral(C-D) noise | 0.137
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Z Height during Z Scan sweep | 16.8
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Z Detector during Z Scan sweep | 12.6
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Y2 Detector Raw during Y Scan sweep | 13.8
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Y2 Detector during Y Scan sweep | 179.9
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Y1 Detector during Y Scan sweep | 184.3
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of Y Detector during Y Scan sweep | 182.1
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of X2 Detector Raw during X Scan sweep | 14
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of X2 Detector during X Scan sweep | 185.6
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of X1 Detector during X Scan sweep | 187
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | Full range of X Detector during X Scan sweep | 186.3
| SUSS MicroTec Solutions GmbH & Co.KG | 2024-04-09 | A-B noise | 8.41
| Southern Illinois University(Shamsi) | 2024-04-08 | Z Scanner position at -3μm | 0.0042
| Southern Illinois University(Shamsi) | 2024-04-08 | Z Scanner position at +3μm | 0.0097
| Southern Illinois University(Shamsi) | 2024-04-08 | Z linearity by sample height | 0.998
| Southern Illinois University(Shamsi) | 2024-04-08 | Z Height noise | 0.026
| Southern Illinois University(Shamsi) | 2024-04-08 | Y Scan Stroke | 55
| Southern Illinois University(Shamsi) | 2024-04-08 | Y Detector noise (XY Servo on) | 0.35
| Southern Illinois University(Shamsi) | 2024-04-08 | Y Detector noise (XY Servo off) | 0.37
| Southern Illinois University(Shamsi) | 2024-04-08 | Y axis OPM | 0.43
| Southern Illinois University(Shamsi) | 2024-04-08 | XY orthogonality | 89.9
| Southern Illinois University(Shamsi) | 2024-04-08 | X Scan Stroke | 55
| Southern Illinois University(Shamsi) | 2024-04-08 | X Detector noise (XY Servo on) | 0.35
| Southern Illinois University(Shamsi) | 2024-04-08 | X Detector noise (XY Servo off) | 0.37
| Southern Illinois University(Shamsi) | 2024-04-08 | X axis OPM | 0.48
| Southern Illinois University(Shamsi) | 2024-04-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Southern Illinois University(Shamsi) | 2024-04-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Southern Illinois University(Shamsi) | 2024-04-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Southern Illinois University(Shamsi) | 2024-04-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Southern Illinois University(Shamsi) | 2024-04-08 | NCM Amplitude noise | 12.77
| Southern Illinois University(Shamsi) | 2024-04-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.311
| Southern Illinois University(Shamsi) | 2024-04-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.566
| Southern Illinois University(Shamsi) | 2024-04-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.6
| Southern Illinois University(Shamsi) | 2024-04-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.4
| Southern Illinois University(Shamsi) | 2024-04-08 | Lateral(C-D) sensitivity | 12.2
| Southern Illinois University(Shamsi) | 2024-04-08 | Lateral(C-D) noise | 0.061
| Southern Illinois University(Shamsi) | 2024-04-08 | Full range of Z Height during Z Scan sweep | 18.3
| Southern Illinois University(Shamsi) | 2024-04-08 | Full range of Z Detector during Z Scan sweep | 13.8
| Southern Illinois University(Shamsi) | 2024-04-08 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Southern Illinois University(Shamsi) | 2024-04-08 | Full range of Y Detector during Y Scan sweep | 66.4
| Southern Illinois University(Shamsi) | 2024-04-08 | Full range of X Detector Raw during X Scan sweep | 13.6
| Southern Illinois University(Shamsi) | 2024-04-08 | Full range of X Detector during X Scan sweep | 67.1
| Southern Illinois University(Shamsi) | 2024-04-08 | A-B noise | 11.23
| University of Delaware(Rachel) | 2024-04-05 | Z Scanner position at -3μm | 0.0039
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Z Scanner position at -3μm | 0.0059
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Z Scanner position at +3μm | 0.0149
| University of Delaware(Rachel) | 2024-04-05 | Z Scanner position at +3μm | 0.0012
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Z linearity by sample height | 1.006
| University of Delaware(Rachel) | 2024-04-05 | Z linearity by sample height | 1.01
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Z Height noise | 0.016
| University of Delaware(Rachel) | 2024-04-05 | Z Height noise | 0.024
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Y Scan Stroke | 110
| University of Delaware(Rachel) | 2024-04-05 | Y Scan Stroke | 186.406
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Y Detector noise (XY Servo on) | 0.27
| University of Delaware(Rachel) | 2024-04-05 | Y Detector noise (XY Servo on) | 0.35
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Y Detector noise (XY Servo off) | 0.27
| University of Delaware(Rachel) | 2024-04-05 | Y Detector noise (XY Servo off) | 0.35
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Y axis OPM | 1.42
| University of Delaware(Rachel) | 2024-04-05 | Y axis OPM | 1.55
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | XY orthogonality | 90.1
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | X Scan Stroke | 109.2
| University of Delaware(Rachel) | 2024-04-05 | X Scan Stroke | 175.867
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | X Detector noise (XY Servo on) | 0.27
| University of Delaware(Rachel) | 2024-04-05 | X Detector noise (XY Servo on) | 0.35
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | X Detector noise (XY Servo off) | 0.26
| University of Delaware(Rachel) | 2024-04-05 | X Detector noise (XY Servo off) | 0.35
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | X axis OPM | 0.99
| University of Delaware(Rachel) | 2024-04-05 | X axis OPM | 1.71
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| University of Delaware(Rachel) | 2024-04-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.49
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| University of Delaware(Rachel) | 2024-04-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.63
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| University of Delaware(Rachel) | 2024-04-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.48
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| University of Delaware(Rachel) | 2024-04-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.62
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | NCM Amplitude noise | 11.31
| University of Delaware(Rachel) | 2024-04-05 | NCM Amplitude noise | 13.36
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.248
| University of Delaware(Rachel) | 2024-04-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.948
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 20.609
| University of Delaware(Rachel) | 2024-04-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.938
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 156
| University of Delaware(Rachel) | 2024-04-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 251.9
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.4
| University of Delaware(Rachel) | 2024-04-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 247.7
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Lateral(C-D) sensitivity | 16.8
| University of Delaware(Rachel) | 2024-04-05 | Lateral(C-D) sensitivity | 1.5
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Lateral(C-D) noise | 0.104
| University of Delaware(Rachel) | 2024-04-05 | Lateral(C-D) noise | 0.168
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Full range of Z Height during Z Scan sweep | 16
| University of Delaware(Rachel) | 2024-04-05 | Full range of Z Height during Z Scan sweep | 17.7
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Full range of Z Detector during Z Scan sweep | 15.5
| University of Delaware(Rachel) | 2024-04-05 | Full range of Z Detector during Z Scan sweep | 14.2
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| University of Delaware(Rachel) | 2024-04-05 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Full range of Y Detector during Y Scan sweep | 110
| University of Delaware(Rachel) | 2024-04-05 | Full range of Y Detector during Y Scan sweep | 186
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Full range of X Detector Raw during X Scan sweep | 14.1
| University of Delaware(Rachel) | 2024-04-05 | Full range of X Detector Raw during X Scan sweep | 14.2
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | Full range of X Detector during X Scan sweep | 109.2
| University of Delaware(Rachel) | 2024-04-05 | Full range of X Detector during X Scan sweep | 175.9
| International Iberian Nanotechnology Laboratory (INL), in Braga (PSE) #1 | 2024-04-05 | A-B noise | 15.99
| University of Delaware(Rachel) | 2024-04-05 | A-B noise | 7.09
| Samsung Display A1 | 2024-03-29 | Z Linearity by Z scanner position | 1.3
| Samsung Display A1 | 2024-03-29 | Z Linearity by 4 different sample height | 0.999
| Samsung Display A1 | 2024-03-29 | Z Height noise | 0.035
| Samsung Display A1 | 2024-03-29 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.4
| Samsung Display A1 | 2024-03-29 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7
| Samsung Display A1 | 2024-03-29 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| Samsung Display A1 | 2024-03-29 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11
| Samsung Display A1 | 2024-03-29 | Z Drift | 16.56
| Samsung Display A1 | 2024-03-29 | Y2 Detector noise (XY Servo on) | 0.33
| Samsung Display A1 | 2024-03-29 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 262
| Samsung Display A1 | 2024-03-29 | Y1 Detector noise (XY Servo on) | 0.35
| Samsung Display A1 | 2024-03-29 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 254.7
| Samsung Display A1 | 2024-03-29 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.3
| Samsung Display A1 | 2024-03-29 | Y Ringing(Rpv) | 0.17
| Samsung Display A1 | 2024-03-29 | Y Non-linearity error | 3.65
| Samsung Display A1 | 2024-03-29 | Y Detector noise (XY Servo on) | 0.29
| Samsung Display A1 | 2024-03-29 | Y axis, Repeatability @1st position | 0.03
| Samsung Display A1 | 2024-03-29 | Y axis, Repeatability @ 4th position | 0.04
| Samsung Display A1 | 2024-03-29 | Y axis, Repeatability @ 3rd position | 0.04
| Samsung Display A1 | 2024-03-29 | Y axis, Repeatability @ 2nd position | 0.04
| Samsung Display A1 | 2024-03-29 | Y axis, Accuracy @1st position | 0.84
| Samsung Display A1 | 2024-03-29 | Y axis, Accuracy @ 4th position | 0.97
| Samsung Display A1 | 2024-03-29 | Y axis, Accuracy @ 3rd position | 0.89
| Samsung Display A1 | 2024-03-29 | Y axis, Accuracy @ 2nd position | 0.79
| Samsung Display A1 | 2024-03-29 | Y axis OPM | 1.54
| Samsung Display A1 | 2024-03-29 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.35
| Samsung Display A1 | 2024-03-29 | XY Stage positioning Y offset (LLC Translation) | 3.54
| Samsung Display A1 | 2024-03-29 | XY Stage positioning Y offset (Global Align) | 4.49
| Samsung Display A1 | 2024-03-29 | XY Stage positioning X offset (Vision Pattern Matching) | 4.21
| Samsung Display A1 | 2024-03-29 | XY Stage positioning X offset (LLC Translation) | 11.67
| Samsung Display A1 | 2024-03-29 | XY Stage positioning X offset (Global Align) | 5.56
| Samsung Display A1 | 2024-03-29 | X2 Detector noise (XY Servo on) | 0.41
| Samsung Display A1 | 2024-03-29 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 273.2
| Samsung Display A1 | 2024-03-29 | X1 Detector noise (XY Servo on) | 0.34
| Samsung Display A1 | 2024-03-29 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265
| Samsung Display A1 | 2024-03-29 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 168.9
| Samsung Display A1 | 2024-03-29 | X Ringing(Rpv) | 0.14
| Samsung Display A1 | 2024-03-29 | X Non-linearity error | 3.47
| Samsung Display A1 | 2024-03-29 | X Detector noise (XY Servo on) | 0.3
| Samsung Display A1 | 2024-03-29 | X axis, Repeatability @1st position | 0.04
| Samsung Display A1 | 2024-03-29 | X axis, Repeatability @ 4th position | 0.04
| Samsung Display A1 | 2024-03-29 | X axis, Repeatability @ 3rd position | 0.05
| Samsung Display A1 | 2024-03-29 | X axis, Repeatability @ 2nd position | 0.04
| Samsung Display A1 | 2024-03-29 | X axis, Accuracy @1st position | 0.36
| Samsung Display A1 | 2024-03-29 | X axis, Accuracy @ 4th position | 0.41
| Samsung Display A1 | 2024-03-29 | X axis, Accuracy @ 3rd position | 0.46
| Samsung Display A1 | 2024-03-29 | X axis, Accuracy @ 2nd position | 0.43
| Samsung Display A1 | 2024-03-29 | X axis OPM | 1.28
| Samsung Display A1 | 2024-03-29 | Wafer placement(Loading) error on the load port2 (Y offset) | 26.21
| Samsung Display A1 | 2024-03-29 | Wafer placement(Loading) error on the load port2 (X offset) | 25.5
| Samsung Display A1 | 2024-03-29 | Wafer placement(Loading) error on the load port1 (Y offset) | 30.92
| Samsung Display A1 | 2024-03-29 | Wafer placement(Loading) error on the load port1 (X offset) | 19.55
| Samsung Display A1 | 2024-03-29 | Under 50nm Step Height Sample (10 Repeats) | 18.16
| Samsung Display A1 | 2024-03-29 | StrokeYMin | -153440
| Samsung Display A1 | 2024-03-29 | StrokeYMax | 150278
| Samsung Display A1 | 2024-03-29 | StrokeY | 303718
| Samsung Display A1 | 2024-03-29 | StrokeXMin | -2034943
| Samsung Display A1 | 2024-03-29 | StrokeXMax | 204627
| Samsung Display A1 | 2024-03-29 | StrokeX | 2239570
| Samsung Display A1 | 2024-03-29 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.46
| Samsung Display A1 | 2024-03-29 | RMS of ZeroScan | 0.047
| Samsung Display A1 | 2024-03-29 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| Samsung Display A1 | 2024-03-29 | OpticalAxisShiftPerFocusY | -0.0007
| Samsung Display A1 | 2024-03-29 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00046
| Samsung Display A1 | 2024-03-29 | OpticalAxisShiftPerFocusX | -0.0046
| Samsung Display A1 | 2024-03-29 | NCM driving% with AC160TS cantilever | 26.33
| Samsung Display A1 | 2024-03-29 | Min. Vibration Transmissbility (~100Hz) | 0.024
| Samsung Display A1 | 2024-03-29 | LowerLimit | 4915
| Samsung Display A1 | 2024-03-29 | Lift pins synchronization offset check (Y offset) | 58.53
| Samsung Display A1 | 2024-03-29 | Lift pins synchronization offset check (X offset) | 13.29
| Samsung Display A1 | 2024-03-29 | Lateral(C-D) sensitivity | 11.7
| Samsung Display A1 | 2024-03-29 | Lateral(C-D) noise | 0.118
| Samsung Display A1 | 2024-03-29 | Image offset after moving from -Y to +Y (Y scan) | 9.52
| Samsung Display A1 | 2024-03-29 | Image offset after moving from -Y to +Y (X scan) | 66.65
| Samsung Display A1 | 2024-03-29 | Image offset after moving from -X to +X (Y scan) | 19.77
| Samsung Display A1 | 2024-03-29 | Image offset after moving from -X to +X (X scan) | 46.14
| Samsung Display A1 | 2024-03-29 | Image offset after moving from +Y to -Y (Y scan) | 23.43
| Samsung Display A1 | 2024-03-29 | Image offset after moving from +Y to -Y (X scan) | 67.38
| Samsung Display A1 | 2024-03-29 | Image offset after moving from +X to -X (Y scan) | 50.53
| Samsung Display A1 | 2024-03-29 | Image offset after moving from +X to -X (X scan) | 10.25
| Samsung Display A1 | 2024-03-29 | Full range of Z Height during Z Scan sweep | 16
| Samsung Display A1 | 2024-03-29 | Full range of Z Detector during Z Scan sweep | 12
| Samsung Display A1 | 2024-03-29 | Full range of Y2 Detector Raw during Y Scan sweep | 13.64
| Samsung Display A1 | 2024-03-29 | Full range of Y2 Detector during Y Scan sweep | 195.88
| Samsung Display A1 | 2024-03-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.01
| Samsung Display A1 | 2024-03-29 | Full range of Y1 Detector during Y Scan sweep | 185.23
| Samsung Display A1 | 2024-03-29 | Full range of Y Detector during Y Scan sweep | 185.23
| Samsung Display A1 | 2024-03-29 | Full range of X2 Detector Raw during X Scan sweep | 13.92
| Samsung Display A1 | 2024-03-29 | Full range of X2 Detector during X Scan sweep | 200.76
| Samsung Display A1 | 2024-03-29 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Samsung Display A1 | 2024-03-29 | Full range of X1 Detector during X Scan sweep | 202.58
| Samsung Display A1 | 2024-03-29 | Full range of X Detector during X Scan sweep | 200.76
| Samsung Display A1 | 2024-03-29 | Cycle time for 15-point measurement | 1677.799
| Samsung Display A1 | 2024-03-29 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Samsung Display A1 | 2024-03-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Samsung Display A1 | 2024-03-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.28
| Samsung Display A1 | 2024-03-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Samsung Display A1 | 2024-03-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Samsung Display A1 | 2024-03-29 | 940nm Step Height Sample (10 Repeats) | 946.45
| Samsung Display A1 | 2024-03-29 | 50~200nm Step Height Sample (10 Repeats) | 184.46
| Samsung Display A1 | 2024-03-29 | 200~500nm Step Height Sample (10 Repeats) | 462.13
| Samsung Display A1 | 2024-03-29 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 0.6
| Samsung Display A1 | 2024-03-29 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.5
| Samsung Display A1 | 2024-03-29 | %GRR w/ roughness of non-patterened wafer | 3.5
| University of Washington | 2024-03-29 | Z Scanner position at -3μm | 0.0025
| University of Washington | 2024-03-29 | Z Scanner position at +3μm | 0.0053
| University of Washington | 2024-03-29 | Z linearity by sample height | 1.001
| University of Washington | 2024-03-29 | Z Height noise | 0.032
| University of Washington | 2024-03-29 | Y Scan Stroke | 13.7
| University of Washington | 2024-03-29 | Y Detector noise (XY Servo on) | 0.25
| University of Washington | 2024-03-29 | Y Detector noise (XY Servo off) | 0.25
| University of Washington | 2024-03-29 | Y axis OPM | 0.77
| University of Washington | 2024-03-29 | X Scan Stroke | 14.2
| University of Washington | 2024-03-29 | X Detector noise (XY Servo on) | 0.26
| University of Washington | 2024-03-29 | X Detector noise (XY Servo off) | 0.25
| University of Washington | 2024-03-29 | X axis OPM | 1.08
| University of Washington | 2024-03-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| University of Washington | 2024-03-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| University of Washington | 2024-03-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| University of Washington | 2024-03-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| University of Washington | 2024-03-29 | NCM Amplitude noise | 17.24
| University of Washington | 2024-03-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.294
| University of Washington | 2024-03-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.698
| University of Washington | 2024-03-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 177.7
| University of Washington | 2024-03-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 173.6
| University of Washington | 2024-03-29 | Lateral(C-D) sensitivity | 14.8
| University of Washington | 2024-03-29 | Lateral(C-D) noise | 0.098
| University of Washington | 2024-03-29 | Full range of Z Height during Z Scan sweep | 18.4
| University of Washington | 2024-03-29 | Full range of Z Detector during Z Scan sweep | 11.7
| University of Washington | 2024-03-29 | Full range of Y Detector Raw during Y Scan sweep | 121.9
| University of Washington | 2024-03-29 | Full range of Y Detector during Y Scan sweep | 121.9
| University of Washington | 2024-03-29 | Full range of X Detector Raw during X Scan sweep | 123.2
| University of Washington | 2024-03-29 | Full range of X Detector during X Scan sweep | 123.2
| University of Washington | 2024-03-29 | A-B noise | 13.19
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Repeatability @1st position | 0.06
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Repeatability @ 4th position | 0.07
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Repeatability @ 3rd position | 0.07
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Repeatability @ 2nd position | 0.06
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Accuracy @1st position | 0.42
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Accuracy @ 4th position | 0.31
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Accuracy @ 3rd position | 0.34
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | Y axis, Accuracy @ 2nd position | 0.32
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Repeatability @1st position | 0.06
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Repeatability @ 4th position | 0.08
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Repeatability @ 3rd position | 0.06
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Repeatability @ 2nd position | 0.06
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Accuracy @1st position | 0.26
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Accuracy @ 4th position | 0.29
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Accuracy @ 3rd position | 0.28
| SMSC (SMIC Shanghai Mask shop) | 2024-03-28 | X axis, Accuracy @ 2nd position | 0.24
| Qorvo (Thomas Russel) | 2024-03-25 | Z Linearity by Z scanner position | 0.43
| Qorvo (Thomas Russel) | 2024-03-25 | Z Linearity by 3 different sample height | 0.999
| Qorvo (Thomas Russel) | 2024-03-25 | Z Height noise | 0.062
| Qorvo (Thomas Russel) | 2024-03-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 65.4783
| Qorvo (Thomas Russel) | 2024-03-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 65.4797
| Qorvo (Thomas Russel) | 2024-03-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.8754
| Qorvo (Thomas Russel) | 2024-03-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.8336
| Qorvo (Thomas Russel) | 2024-03-25 | Z Drift | 7.32
| Qorvo (Thomas Russel) | 2024-03-25 | Y2 Detector noise (XY Servo on) | 0.24
| Qorvo (Thomas Russel) | 2024-03-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 247.4
| Qorvo (Thomas Russel) | 2024-03-25 | Y1 Detector noise (XY Servo on) | 0.24
| Qorvo (Thomas Russel) | 2024-03-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 240.7
| Qorvo (Thomas Russel) | 2024-03-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 169.9
| Qorvo (Thomas Russel) | 2024-03-25 | Y Ringing(Rpv) | 0.08
| Qorvo (Thomas Russel) | 2024-03-25 | Y Non-linearity error | 3.34
| Qorvo (Thomas Russel) | 2024-03-25 | Y Detector noise (XY Servo on) | 0.18
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Repeatability @1st position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Repeatability @ 4th position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Repeatability @ 3rd position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Repeatability @ 2nd position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Accuracy @1st position | 1.35
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Accuracy @ 4th position | 1.58
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Accuracy @ 3rd position | 1.41
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis, Accuracy @ 2nd position | 1.17
| Qorvo (Thomas Russel) | 2024-03-25 | Y axis OPM | 0.36
| Qorvo (Thomas Russel) | 2024-03-25 | X2 Detector noise (XY Servo on) | 0.24
| Qorvo (Thomas Russel) | 2024-03-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 253.8
| Qorvo (Thomas Russel) | 2024-03-25 | X1 Detector noise (XY Servo on) | 0.23
| Qorvo (Thomas Russel) | 2024-03-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 245
| Qorvo (Thomas Russel) | 2024-03-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 170.8
| Qorvo (Thomas Russel) | 2024-03-25 | X Ringing(Rpv) | 0.15
| Qorvo (Thomas Russel) | 2024-03-25 | X Non-linearity error | 3.55
| Qorvo (Thomas Russel) | 2024-03-25 | X Detector noise (XY Servo on) | 0.17
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Repeatability @1st position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Repeatability @ 4th position | 0.02
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Repeatability @ 3rd position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Repeatability @ 2nd position | 0.03
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Accuracy @1st position | 0.22
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Accuracy @ 4th position | 0.09
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Accuracy @ 3rd position | 0.36
| Qorvo (Thomas Russel) | 2024-03-25 | X axis, Accuracy @ 2nd position | 0.17
| Qorvo (Thomas Russel) | 2024-03-25 | X axis OPM | 0.39
| Qorvo (Thomas Russel) | 2024-03-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 25.16
| Qorvo (Thomas Russel) | 2024-03-25 | Wafer placement(Loading) error on the load port2 (X offset) | 26.55
| Qorvo (Thomas Russel) | 2024-03-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 40.12
| Qorvo (Thomas Russel) | 2024-03-25 | Wafer placement(Loading) error on the load port1 (X offset) | 28.9
| Qorvo (Thomas Russel) | 2024-03-25 | Torque Profile (50mm, 20um/s) | 2.14
| Qorvo (Thomas Russel) | 2024-03-25 | Torque Profile (50mm, 200um/s) | 2.08
| Qorvo (Thomas Russel) | 2024-03-25 | StrokeYMin | -102137
| Qorvo (Thomas Russel) | 2024-03-25 | StrokeYMax | 106859.7
| Qorvo (Thomas Russel) | 2024-03-25 | StrokeY | 208996.8
| Qorvo (Thomas Russel) | 2024-03-25 | StrokeXMin | -153207
| Qorvo (Thomas Russel) | 2024-03-25 | StrokeXMax | 154018.7
| Qorvo (Thomas Russel) | 2024-03-25 | StrokeX | 307225.6
| Qorvo (Thomas Russel) | 2024-03-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.4
| Qorvo (Thomas Russel) | 2024-03-25 | Roll motion to X+ full scan | 13.4
| Qorvo (Thomas Russel) | 2024-03-25 | RMS of ZeroScan | 0.021
| Qorvo (Thomas Russel) | 2024-03-25 | Profiler noise @200um | 0.192
| Qorvo (Thomas Russel) | 2024-03-25 | PosMin | -25.9
| Qorvo (Thomas Russel) | 2024-03-25 | PosMax | 25.9
| Qorvo (Thomas Russel) | 2024-03-25 | Position repeatability(10repeats, Position Based) | 0.38
| Qorvo (Thomas Russel) | 2024-03-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Qorvo (Thomas Russel) | 2024-03-25 | OpticalAxisShiftPerFocusY | -0.0026
| Qorvo (Thomas Russel) | 2024-03-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| Qorvo (Thomas Russel) | 2024-03-25 | OpticalAxisShiftPerFocusX | -0.004
| Qorvo (Thomas Russel) | 2024-03-25 | OPM Repeatability @5mm | 1.21
| Qorvo (Thomas Russel) | 2024-03-25 | OPM Repeatability @25mm | 0.72
| Qorvo (Thomas Russel) | 2024-03-25 | OPM Repeatability @1mm | 0.71
| Qorvo (Thomas Russel) | 2024-03-25 | OPM Repeatability @10mm | 1.23
| Qorvo (Thomas Russel) | 2024-03-25 | NCM driving% with AC160TS cantilever | 15.69
| Qorvo (Thomas Russel) | 2024-03-25 | Min. Vibration Transmissbility (~100Hz) | 0.019
| Qorvo (Thomas Russel) | 2024-03-25 | MaxLength | 51
| Qorvo (Thomas Russel) | 2024-03-25 | Max OPM @5mm | 13.78
| Qorvo (Thomas Russel) | 2024-03-25 | Max OPM @25mm | 27.9
| Qorvo (Thomas Russel) | 2024-03-25 | Max OPM @1mm | 6.61
| Qorvo (Thomas Russel) | 2024-03-25 | Max OPM @10mm | 14.04
| Qorvo (Thomas Russel) | 2024-03-25 | LowerLimit | 4050
| Qorvo (Thomas Russel) | 2024-03-25 | Lateral(C-D) sensitivity | 11.7
| Qorvo (Thomas Russel) | 2024-03-25 | Lateral(C-D) noise | 0.118
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from -Y to +Y (Y scan) | 62.26
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from -Y to +Y (X scan) | 22.71
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from -X to +X (Y scan) | 8.06
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from -X to +X (X scan) | 18.31
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from +Y to -Y (Y scan) | 13.18
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from +Y to -Y (X scan) | 3.66
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from +X to -X (Y scan) | 68.85
| Qorvo (Thomas Russel) | 2024-03-25 | Image offset after moving from +X to -X (X scan) | 22.71
| Qorvo (Thomas Russel) | 2024-03-25 | FullRecoveryProfileLength | 50
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Z Height during Z Scan sweep | 43.6073
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Z Detector during Z Scan sweep | 13.3
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Y2 Detector during Y Scan sweep | 180.9
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.17
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Y1 Detector during Y Scan sweep | 183.7
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of Y Detector during Y Scan sweep | 183.7
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of X2 Detector Raw during X Scan sweep | 14.04
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of X2 Detector during X Scan sweep | 182
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of X1 Detector Raw during X Scan sweep | 14.14
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of X1 Detector during X Scan sweep | 193.1
| Qorvo (Thomas Russel) | 2024-03-25 | Full range of X Detector during X Scan sweep | 182
| Qorvo (Thomas Russel) | 2024-03-25 | Cycle time for 15-point measurement | 1771.458
| Qorvo (Thomas Russel) | 2024-03-25 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Qorvo (Thomas Russel) | 2024-03-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Qorvo (Thomas Russel) | 2024-03-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| Qorvo (Thomas Russel) | 2024-03-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| Qorvo (Thomas Russel) | 2024-03-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Qorvo (Thomas Russel) | 2024-03-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Qorvo (Thomas Russel) | 2024-03-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Qorvo (Thomas Russel) | 2024-03-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.2
| Qorvo (Thomas Russel) | 2024-03-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| Qorvo (Thomas Russel) | 2024-03-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.6
| Qorvo (Thomas Russel) | 2024-03-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.1
| Qorvo (Thomas Russel) | 2024-03-25 | %GRR w/ roughness of non-patterened wafer | 9.2
| 성균관대학교(기계공학부) | 2024-03-25 | Z Scanner position at -3μm | 0.002
| 성균관대학교(기계공학부) | 2024-03-25 | Z Scanner position at +3μm | 0.0036
| 성균관대학교(기계공학부) | 2024-03-25 | Z linearity by sample height | 1.01
| 성균관대학교(기계공학부) | 2024-03-25 | Z Height noise | 0.026
| 성균관대학교(기계공학부) | 2024-03-25 | Y Scan Stroke | 55
| 성균관대학교(기계공학부) | 2024-03-25 | Y Detector noise (XY Servo on) | 0.25
| 성균관대학교(기계공학부) | 2024-03-25 | Y Detector noise (XY Servo off) | 0.24
| 성균관대학교(기계공학부) | 2024-03-25 | Y axis OPM | 0.42
| 성균관대학교(기계공학부) | 2024-03-25 | X Scan Stroke | 55
| 성균관대학교(기계공학부) | 2024-03-25 | X Detector noise (XY Servo on) | 0.24
| 성균관대학교(기계공학부) | 2024-03-25 | X Detector noise (XY Servo off) | 0.24
| 성균관대학교(기계공학부) | 2024-03-25 | X axis OPM | 0.42
| 성균관대학교(기계공학부) | 2024-03-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| 성균관대학교(기계공학부) | 2024-03-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 성균관대학교(기계공학부) | 2024-03-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| 성균관대학교(기계공학부) | 2024-03-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| 성균관대학교(기계공학부) | 2024-03-25 | NCM Amplitude noise | 13.87
| 성균관대학교(기계공학부) | 2024-03-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.71
| 성균관대학교(기계공학부) | 2024-03-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.785
| 성균관대학교(기계공학부) | 2024-03-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.4
| 성균관대학교(기계공학부) | 2024-03-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.8
| 성균관대학교(기계공학부) | 2024-03-25 | Lateral(C-D) sensitivity | 14.9
| 성균관대학교(기계공학부) | 2024-03-25 | Lateral(C-D) noise | 0.083
| 성균관대학교(기계공학부) | 2024-03-25 | Full range of Z Height during Z Scan sweep | 18.9
| 성균관대학교(기계공학부) | 2024-03-25 | Full range of Z Detector during Z Scan sweep | 16.6
| 성균관대학교(기계공학부) | 2024-03-25 | Full range of Y Detector Raw during Y Scan sweep | 14
| 성균관대학교(기계공학부) | 2024-03-25 | Full range of Y Detector during Y Scan sweep | 70.4
| 성균관대학교(기계공학부) | 2024-03-25 | Full range of X Detector Raw during X Scan sweep | 14.3
| 성균관대학교(기계공학부) | 2024-03-25 | Full range of X Detector during X Scan sweep | 70
| 성균관대학교(기계공학부) | 2024-03-25 | A-B noise | 8.92
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Z Scanner position at -3μm | 0.0064
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Z Scanner position at -3μm | 0.0073
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Z Scanner position at +3μm | 0.0107
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Z Scanner position at +3μm | 0.0006
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Z linearity by sample height | 1.005
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Z linearity by sample height | 0.995
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Z Height noise | 0.016
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Z Height noise | 0.02
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Y Scan Stroke | 115.9
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Y Scan Stroke | 110.2
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Y Detector noise (XY Servo on) | 0.26
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Y Detector noise (XY Servo on) | 0.27
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Y Detector noise (XY Servo off) | 0.25
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Y Detector noise (XY Servo off) | 0.26
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Y axis OPM | 0.64
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Y axis OPM | 1.05
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | X Scan Stroke | 114.4
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | X Scan Stroke | 110.2
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | X Detector noise (XY Servo on) | 0.26
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | X Detector noise (XY Servo on) | 0.25
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | X Detector noise (XY Servo off) | 0.26
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | X Detector noise (XY Servo off) | 0.24
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | X axis OPM | 0.99
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | X axis OPM | 1.03
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | SLD beam power at the end of the lens | 0.25
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | SLD beam power at the end of the lens | 0.2
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | NCM driving% with AC160TS cantilever | 20
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | NCM driving% with AC160TS cantilever | 2.56
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | NCM Amplitude noise | 18.35
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | NCM Amplitude noise | 10.06
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.731
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.182
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.208
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.438
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 156.4
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.3
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Lateral(C-D) sensitivity | 8.2
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Lateral(C-D) sensitivity | 5.7
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Lateral(C-D) noise | 0.12
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Lateral(C-D) noise | 0.1
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Full range of Z Height during Z Scan sweep | 15.3
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Full range of Z Height during Z Scan sweep | 15.5
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Full range of Z Detector during Z Scan sweep | 13.8
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Full range of Z Detector during Z Scan sweep | 12.6
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Full range of Y Detector during Y Scan sweep | 115.9
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Full range of Y Detector during Y Scan sweep | 110.2
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Full range of X Detector Raw during X Scan sweep | 14.3
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Full range of X Detector Raw during X Scan sweep | 14
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | Full range of X Detector during X Scan sweep | 114.4
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | Full range of X Detector during X Scan sweep | 110.2
| Institute of Materials Science of Barcelona (ICMAB) # 1 | 2024-03-22 | A-B noise | 11.79
| Institute of Materials Science of Barcelona (ICMAB) # 2 | 2024-03-22 | A-B noise | 16.82
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Z Scanner position at -3μm | 0.0059
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Z Scanner position at +3μm | 0.0092
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Z linearity by sample height | 1.009
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Z Height noise | 0.027
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Y Scan Stroke | 116.4
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Y Detector noise (XY Servo on) | 0.27
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Y Detector noise (XY Servo off) | 0.27
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Y axis OPM | 1.36
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | X Scan Stroke | 117.4
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | X Detector noise (XY Servo on) | 0.26
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | X Detector noise (XY Servo off) | 0.26
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | X axis OPM | 1.07
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | NCM driving% with AC160TS cantilever | 2.89
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | NCM Amplitude noise | 11.46
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.382
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.469
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.4
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 168.3
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Lateral(C-D) sensitivity | 44.9
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Lateral(C-D) noise | 0.035
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Full range of Z Height during Z Scan sweep | 18.4
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Full range of Z Detector during Z Scan sweep | 13.9
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Full range of Y Detector during Y Scan sweep | 115.9
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Full range of X Detector Raw during X Scan sweep | 13.9
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | Full range of X Detector during X Scan sweep | 116.6
| Merseburg University of Applied Sciences (HS Merseburg Krause) | 2024-03-18 | A-B noise | 7.06
| 한국나노마이스터고등학교 | 2024-03-12 | Z linearity by sample height | 0.998
| 한국나노마이스터고등학교 | 2024-03-12 | Z Height noise | 0.024
| 한국나노마이스터고등학교 | 2024-03-12 | Y Scan Stroke | 55
| 한국나노마이스터고등학교 | 2024-03-12 | Y Detector noise (XY Servo on) | 0.39
| 한국나노마이스터고등학교 | 2024-03-12 | Y Detector noise (XY Servo off) | 0.41
| 한국나노마이스터고등학교 | 2024-03-12 | Y axis OPM | 0.69
| 한국나노마이스터고등학교 | 2024-03-12 | X Scan Stroke | 55
| 한국나노마이스터고등학교 | 2024-03-12 | X Detector noise (XY Servo on) | 0.4
| 한국나노마이스터고등학교 | 2024-03-12 | X Detector noise (XY Servo off) | 0.43
| 한국나노마이스터고등학교 | 2024-03-12 | X axis OPM | 0.44
| 한국나노마이스터고등학교 | 2024-03-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| 한국나노마이스터고등학교 | 2024-03-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| 한국나노마이스터고등학교 | 2024-03-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한국나노마이스터고등학교 | 2024-03-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| 한국나노마이스터고등학교 | 2024-03-12 | NCM driving% with AC160TS cantilever | 3.19
| 한국나노마이스터고등학교 | 2024-03-12 | NCM Amplitude noise | 16.86
| 한국나노마이스터고등학교 | 2024-03-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.815
| 한국나노마이스터고등학교 | 2024-03-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.264
| 한국나노마이스터고등학교 | 2024-03-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98
| 한국나노마이스터고등학교 | 2024-03-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102
| 한국나노마이스터고등학교 | 2024-03-12 | Maximum Y Scan range (Full Stroke) | 67.6
| 한국나노마이스터고등학교 | 2024-03-12 | Maximum X Scan range (Full Stroke) | 67.5
| 한국나노마이스터고등학교 | 2024-03-12 | Full range of Z Height during Z Scan sweep | 16.9
| 한국나노마이스터고등학교 | 2024-03-12 | Lateral(C-D) sensitivity | 13.9
| 한국나노마이스터고등학교 | 2024-03-12 | Lateral(C-D) noise | 0.055
| 한국나노마이스터고등학교 | 2024-03-12 | Full range of Z Detector raw | 13.9
| 한국나노마이스터고등학교 | 2024-03-12 | Full range of Y Detector raw | 13.8
| 한국나노마이스터고등학교 | 2024-03-12 | Full range of X Detector raw | 13.2
| 한국나노마이스터고등학교 | 2024-03-12 | Approach position at -3μm | 0.0063
| 한국나노마이스터고등학교 | 2024-03-12 | Approach position at +3μm | 0.0123
| 한국나노마이스터고등학교 | 2024-03-12 | A-B noise | 7.51
| 단국대학교(전기전자공학부) | 2024-03-08 | Z linearity by sample height | 0.999
| NIPER Raebareli | 2024-03-08 | Z linearity by sample height | 0.995
| 단국대학교(전기전자공학부) | 2024-03-08 | Z Height noise | 0.027
| NIPER Raebareli | 2024-03-08 | Z Height noise | 0.029
| 단국대학교(전기전자공학부) | 2024-03-08 | Y Scan Stroke | 55
| NIPER Raebareli | 2024-03-08 | Y Scan Stroke | 188.748
| 단국대학교(전기전자공학부) | 2024-03-08 | Y Detector noise (XY Servo on) | 0.26
| NIPER Raebareli | 2024-03-08 | Y Detector noise (XY Servo on) | 0.3
| 단국대학교(전기전자공학부) | 2024-03-08 | Y Detector noise (XY Servo off) | 0.25
| NIPER Raebareli | 2024-03-08 | Y Detector noise (XY Servo off) | 0.29
| 단국대학교(전기전자공학부) | 2024-03-08 | Y axis OPM | 0.63
| NIPER Raebareli | 2024-03-08 | Y axis OPM | 1.02
| 단국대학교(전기전자공학부) | 2024-03-08 | X Scan Stroke | 55
| NIPER Raebareli | 2024-03-08 | X Scan Stroke | 189.225
| 단국대학교(전기전자공학부) | 2024-03-08 | X Detector noise (XY Servo on) | 0.26
| NIPER Raebareli | 2024-03-08 | X Detector noise (XY Servo on) | 0.29
| 단국대학교(전기전자공학부) | 2024-03-08 | X Detector noise (XY Servo off) | 0.26
| NIPER Raebareli | 2024-03-08 | X Detector noise (XY Servo off) | 0.29
| 단국대학교(전기전자공학부) | 2024-03-08 | X axis OPM | 0.91
| NIPER Raebareli | 2024-03-08 | X axis OPM | 1.01
| 단국대학교(전기전자공학부) | 2024-03-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| NIPER Raebareli | 2024-03-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.53
| 단국대학교(전기전자공학부) | 2024-03-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| NIPER Raebareli | 2024-03-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.63
| 단국대학교(전기전자공학부) | 2024-03-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| NIPER Raebareli | 2024-03-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.51
| 단국대학교(전기전자공학부) | 2024-03-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| NIPER Raebareli | 2024-03-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.61
| 단국대학교(전기전자공학부) | 2024-03-08 | NCM driving% with AC160TS cantilever | 0.7
| NIPER Raebareli | 2024-03-08 | NCM driving% with AC160TS cantilever | 1.07
| 단국대학교(전기전자공학부) | 2024-03-08 | NCM Amplitude noise | 12.93
| NIPER Raebareli | 2024-03-08 | NCM Amplitude noise | 12.25
| 단국대학교(전기전자공학부) | 2024-03-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.086
| NIPER Raebareli | 2024-03-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.924
| 단국대학교(전기전자공학부) | 2024-03-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.655
| NIPER Raebareli | 2024-03-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.913
| 단국대학교(전기전자공학부) | 2024-03-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 107.3
| NIPER Raebareli | 2024-03-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 256.8
| 단국대학교(전기전자공학부) | 2024-03-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.2
| NIPER Raebareli | 2024-03-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 252.3
| 단국대학교(전기전자공학부) | 2024-03-08 | Maximum Y Scan range (Full Stroke) | 75.6
| NIPER Raebareli | 2024-03-08 | Maximum Y Scan range (Full Stroke) | 188.25
| 단국대학교(전기전자공학부) | 2024-03-08 | Maximum X Scan range (Full Stroke) | 76.3
| NIPER Raebareli | 2024-03-08 | Maximum X Scan range (Full Stroke) | 188.6
| 단국대학교(전기전자공학부) | 2024-03-08 | Full range of Z Height during Z Scan sweep | 18
| NIPER Raebareli | 2024-03-08 | Full range of Z Height during Z Scan sweep | 17.8
| 단국대학교(전기전자공학부) | 2024-03-08 | Lateral(C-D) sensitivity | 14.6
| NIPER Raebareli | 2024-03-08 | Lateral(C-D) sensitivity | 14.7
| 단국대학교(전기전자공학부) | 2024-03-08 | Lateral(C-D) noise | 0.086
| NIPER Raebareli | 2024-03-08 | Lateral(C-D) noise | 0.099
| 단국대학교(전기전자공학부) | 2024-03-08 | Full range of Z Detector raw | 14.6
| NIPER Raebareli | 2024-03-08 | Full range of Z Detector raw | 12.3
| 단국대학교(전기전자공학부) | 2024-03-08 | Full range of Y Detector raw | 14.1
| NIPER Raebareli | 2024-03-08 | Full range of Y Detector raw | 14.7
| 단국대학교(전기전자공학부) | 2024-03-08 | Full range of X Detector raw | 14.5
| NIPER Raebareli | 2024-03-08 | Full range of X Detector raw | 15
| 단국대학교(전기전자공학부) | 2024-03-08 | Approach position at -3μm | 0.0014
| NIPER Raebareli | 2024-03-08 | Approach position at -3μm | 0.0001
| 단국대학교(전기전자공학부) | 2024-03-08 | Approach position at +3μm | 0.0064
| NIPER Raebareli | 2024-03-08 | Approach position at +3μm | 0.0046
| 단국대학교(전기전자공학부) | 2024-03-08 | A-B noise | 8.88
| NIPER Raebareli | 2024-03-08 | A-B noise | 10.05
| China Academy of Engineering Physics | 2024-03-07 | Z Scanner position at -3μm | 0.0082
| China Academy of Engineering Physics | 2024-03-07 | Z Scanner position at +3μm | 0.0193
| China Academy of Engineering Physics | 2024-03-07 | Z linearity by sample height | 0.998
| China Academy of Engineering Physics | 2024-03-07 | Z Height noise | 0.035
| China Academy of Engineering Physics | 2024-03-07 | Y Scan Stroke | 55
| China Academy of Engineering Physics | 2024-03-07 | Y Detector noise (XY Servo on) | 0.38
| China Academy of Engineering Physics | 2024-03-07 | Y Detector noise (XY Servo off) | 0.4
| China Academy of Engineering Physics | 2024-03-07 | Y axis OPM | 0.35
| China Academy of Engineering Physics | 2024-03-07 | X Scan Stroke | 55
| China Academy of Engineering Physics | 2024-03-07 | X Detector noise (XY Servo on) | 0.4
| China Academy of Engineering Physics | 2024-03-07 | X Detector noise (XY Servo off) | 0.41
| China Academy of Engineering Physics | 2024-03-07 | X axis OPM | 0.57
| China Academy of Engineering Physics | 2024-03-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| China Academy of Engineering Physics | 2024-03-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| China Academy of Engineering Physics | 2024-03-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| China Academy of Engineering Physics | 2024-03-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| China Academy of Engineering Physics | 2024-03-07 | NCM driving% with AC160TS cantilever | 0.38
| China Academy of Engineering Physics | 2024-03-07 | NCM Amplitude noise | 13.63
| China Academy of Engineering Physics | 2024-03-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.232
| China Academy of Engineering Physics | 2024-03-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.173
| China Academy of Engineering Physics | 2024-03-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 93.9
| China Academy of Engineering Physics | 2024-03-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.1
| China Academy of Engineering Physics | 2024-03-07 | Lateral(C-D) sensitivity | 15.9
| China Academy of Engineering Physics | 2024-03-07 | Lateral(C-D) noise | 0.034
| China Academy of Engineering Physics | 2024-03-07 | Full range of Z Height during Z Scan sweep | 16.2
| China Academy of Engineering Physics | 2024-03-07 | Full range of Z Detector during Z Scan sweep | 11.5
| China Academy of Engineering Physics | 2024-03-07 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| China Academy of Engineering Physics | 2024-03-07 | Full range of Y Detector during Y Scan sweep | 68
| China Academy of Engineering Physics | 2024-03-07 | Full range of X Detector Raw during X Scan sweep | 14.5
| China Academy of Engineering Physics | 2024-03-07 | Full range of X Detector during X Scan sweep | 70.3
| China Academy of Engineering Physics | 2024-03-07 | A-B noise | 6.76
| State Grid Tianjin | 2024-03-04 | Z Scanner position at -3μm | 0.0009
| State Grid Tianjin | 2024-03-04 | Z Scanner position at +3μm | 0.0059
| State Grid Tianjin | 2024-03-04 | Z linearity by sample height | 1.005
| State Grid Tianjin | 2024-03-04 | Z Height noise | 0.032
| State Grid Tianjin | 2024-03-04 | Y Scan Stroke | 17.2
| State Grid Tianjin | 2024-03-04 | Y Detector noise (XY Servo on) | 0.033
| State Grid Tianjin | 2024-03-04 | Y Detector noise (XY Servo off) | 0.036
| State Grid Tianjin | 2024-03-04 | Y axis OPM | 0.47
| State Grid Tianjin | 2024-03-04 | X Scan Stroke | 16.3
| State Grid Tianjin | 2024-03-04 | X Detector noise (XY Servo on) | 0.026
| State Grid Tianjin | 2024-03-04 | X Detector noise (XY Servo off) | 0.027
| State Grid Tianjin | 2024-03-04 | X axis OPM | 0.91
| State Grid Tianjin | 2024-03-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| State Grid Tianjin | 2024-03-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| State Grid Tianjin | 2024-03-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| State Grid Tianjin | 2024-03-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| State Grid Tianjin | 2024-03-04 | NCM Amplitude noise | 12.74
| State Grid Tianjin | 2024-03-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.553
| State Grid Tianjin | 2024-03-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.684
| State Grid Tianjin | 2024-03-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 27.2
| State Grid Tianjin | 2024-03-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 25.4
| State Grid Tianjin | 2024-03-04 | Lateral(C-D) sensitivity | 12
| State Grid Tianjin | 2024-03-04 | Lateral(C-D) noise | 0.077
| State Grid Tianjin | 2024-03-04 | Full range of Z Height during Z Scan sweep | 18.8
| State Grid Tianjin | 2024-03-04 | Full range of Z Detector during Z Scan sweep | 15.9
| State Grid Tianjin | 2024-03-04 | Full range of Y Detector Raw during Y Scan sweep | 12.6
| State Grid Tianjin | 2024-03-04 | Full range of Y Detector during Y Scan sweep | 17.2
| State Grid Tianjin | 2024-03-04 | Full range of X Detector Raw during X Scan sweep | 12.8
| State Grid Tianjin | 2024-03-04 | Full range of X Detector during X Scan sweep | 16.3
| State Grid Tianjin | 2024-03-04 | A-B noise | 7.77
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Z Scanner position at -3μm | 0.0071
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Z Scanner position at +3μm | 0.0156
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Z Linearity by Z Scanner position
| 한국전자기술연구원 | 2024-02-29 | Z linearity by sample height | 0.998
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Z Linearity by 3 different sample height | 1.01
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Z Height noise | 0.027
| 한국전자기술연구원 | 2024-02-29 | Z Height noise | 0.028
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Y Scan Stroke | 139.3
| 한국전자기술연구원 | 2024-02-29 | Y Scan Stroke | 55
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | -Y Direction software Limit | -76117.5
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Y Detector noise (XY Servo on) | 0.26
| 한국전자기술연구원 | 2024-02-29 | Y Detector noise (XY Servo on) | 0.27
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Y Detector noise (XY Servo off) | 0.27
| 한국전자기술연구원 | 2024-02-29 | Y Detector noise (XY Servo off) | 0.27
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Y axis OPM | 1.13
| 한국전자기술연구원 | 2024-02-29 | Y axis OPM | 0.31
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | X Scan Stroke | 138.2
| 한국전자기술연구원 | 2024-02-29 | X Scan Stroke | 55
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | -X Direction software Limit | -76077.5
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | X Detector noise (XY Servo on) | 0.27
| 한국전자기술연구원 | 2024-02-29 | X Detector noise (XY Servo on) | 0.26
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | X Detector noise (XY Servo off) | 0.26
| 한국전자기술연구원 | 2024-02-29 | X Detector noise (XY Servo off) | 0.25
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | X axis OPM | 1.18
| 한국전자기술연구원 | 2024-02-29 | X axis OPM | 0.61
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 한국전자기술연구원 | 2024-02-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 한국전자기술연구원 | 2024-02-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.24
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.25
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.24
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.24
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.24
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 한국전자기술연구원 | 2024-02-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| 한국전자기술연구원 | 2024-02-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | NCM driving% with AC160TS cantilever | 1.06
| 한국전자기술연구원 | 2024-02-29 | NCM driving% with AC160TS cantilever | 0.76
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | NCM Amplitude noise | 13.33
| 한국전자기술연구원 | 2024-02-29 | NCM Amplitude noise | 6.95
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.995
| 한국전자기술연구원 | 2024-02-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.675
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.286
| 한국전자기술연구원 | 2024-02-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.025
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 190.8
| 한국전자기술연구원 | 2024-02-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.8
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 191.9
| 한국전자기술연구원 | 2024-02-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 96.9
| 한국전자기술연구원 | 2024-02-29 | Maximum Y Scan range (Full Stroke) | 73.4
| 한국전자기술연구원 | 2024-02-29 | Maximum X Scan range (Full Stroke) | 69.2
| 한국전자기술연구원 | 2024-02-29 | Full range of Z Height during Z Scan sweep | 19
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Lateral(C-D) sensitivity | 15.3
| 한국전자기술연구원 | 2024-02-29 | Lateral(C-D) sensitivity | 15.7
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Lateral(C-D) noise | 0.113
| 한국전자기술연구원 | 2024-02-29 | Lateral(C-D) noise | 0.083
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Full range of Z Height during Z Scan sweep | 17.6
| 한국전자기술연구원 | 2024-02-29 | Full range of Z Detector raw | 15.2
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Full range of Z Detector during Z Scan sweep | 13.9
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| 한국전자기술연구원 | 2024-02-29 | Full range of Y Detector raw | 14.1
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Full range of Y Detector during Y Scan sweep | 138.9
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Full range of X Detector Raw during X Scan sweep | 14.4
| 한국전자기술연구원 | 2024-02-29 | Full range of X Detector raw | 14.3
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | Full range of X Detector during X Scan sweep | 137.9
| 한국전자기술연구원 | 2024-02-29 | Approach position at -3μm | 0.0031
| 한국전자기술연구원 | 2024-02-29 | Approach position at +3μm | 0.0145
| Xiamen Tianma Optoelectronics Co.,Ltd | 2024-02-29 | A-B noise | 11.01
| 한국전자기술연구원 | 2024-02-29 | A-B noise | 17.45
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Z Linearity by Z scanner position | 1.57
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Z Linearity by Z scanner position | 1.64
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Z Linearity by 4 different sample height | 1.009
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Z Linearity by 4 different sample height | 1.004
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Z Height noise | 0.034
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Z Height noise | 0.027
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.4
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.4
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.8
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.3
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Z Drift (Short Traveling) | 0.26
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Z Drift | 27.05
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 63.1
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 61.2
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y Ringing(Rpv) | 0.27
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Y Ringing(Rpv) | 0.66
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y Detector noise (XY Servo on) | 0.17
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Y detector noise (XY Servo on) | 0.148
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 93.4
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 84.84
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Repeatability @1st position | 0.03
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Repeatability @ 4th position | 0.03
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Repeatability @ 3rd position | 0.03
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Repeatability @ 2nd position | 0.02
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Accuracy @1st position | 0.16
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Accuracy @ 4th position | 0.16
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Accuracy @ 3rd position | 0.16
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis, Accuracy @ 2nd position | 0.14
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Y axis OPM | 1.22
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Y axis OPM | 0.61
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | XY Orthogonality | 90.2
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | XY Orthogonality | 90.1
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 62.2
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 61.7
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X Ringing(Rpv) | 0.31
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | X Ringing(Rpv) | 0.7
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X Detector noise (XY Servo on) | 0.17
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | X detector noise (XY Servo on) | 0.147
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 88.3
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 79.41
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Repeatability @1st position | 0.03
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Repeatability @ 4th position | 0.02
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Repeatability @ 3rd position | 0.02
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Repeatability @ 2nd position | 0.02
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Accuracy @1st position | 0.17
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Accuracy @ 4th position | 0.1
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Accuracy @ 3rd position | 0.09
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis, Accuracy @ 2nd position | 0.13
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | X axis OPM | 1.25
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | X axis OPM | 1.71
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Under 50nm Step Height Sample (10 Repeats) | 20.19
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | StrokeYMin | -153633
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | StrokeYMax | 150623.3
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | StrokeY | 304256.6
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | StrokeXMin | -203915
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | StrokeXMax | 203689.1
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | StrokeX | 407603.6
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Scan offset accuracy (Precision) | 0.4
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Scan offset accuracy (Precision) | 0.516
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | RMS of ZeroScan | 0.043
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | RMS of ZeroScan | 0.063
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | RMS error after Stage Calibration | 0.8
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | OpticalAxisShiftPerFocusY | -0.0001
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | OpticalAxisShiftPerFocusX | -0.0004
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | NCM driving% with AC160TS cantilever | 4.6
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Min. Vibration Transmissbility (~100Hz) | 0.032
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Min. Vibration Transmissbility (~100Hz) | 0.015
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Mask Roughness Repeatability (10Repeats, σ ) | 0.004
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Mask Roughness Repeatability (10Repeats, σ ) | 0.02
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Mask Roughness Repeatability (10Repeats, RSD%) | 2.68
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Mask Roughness Repeatability (10Repeats, RSD%) | 0.0025
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Mask exchange time | 4"44
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Mask exchange time | 4.91
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | LowerLimit | 8222
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from -Y to +Y (Y scan) | 43.95
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from -Y to +Y (Y scan) | 14.64
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from -Y to +Y (X scan) | 58.59
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from -Y to +Y (X scan) | 39.06
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from -X to +X (Y scan) | 87.89
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from -X to +X (Y scan) | 34.17
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from -X to +X (X scan) | 48.83
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from -X to +X (X scan) | 9.76
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from +Y to -Y (Y scan) | 43.95
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from +Y to -Y (Y scan) | 29.29
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from +Y to -Y (X scan) | 29.3
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from +Y to -Y (X scan) | 14.64
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from +X to -X (Y scan) | 34.18
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from +X to -X (Y scan) | 87.89
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Image offset after moving from +X to -X (X scan) | 43.95
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Image offset after moving from +X to -X (X scan) | 19.53
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Imag offset after Lock stage (Y scan) | 34.18
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Imag offset after Lock stage (Y scan) | 73.2
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Imag offset after Lock stage (X scan) | 185.55
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Imag offset after Lock stage (X scan) | 19.5
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Imag offset after EFEM loading (Y scan) | 63.48
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Imag offset after EFEM loading (Y scan) | 136.716
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Imag offset after EFEM loading (X scan) | 190.43
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Imag offset after EFEM loading (X scan) | 107.422
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Full range of Z Height during Z Scan sweep | 17.8
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Full range of Z Height during Z Scan sweep | 18.2
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Full range of Z Detector during Z Scan sweep | 15
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Full range of Z Detector during Z Scan sweep | 14.3
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Full range of Y Detector during Y Scan sweep | 63.6
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Full range of Y Detector during Y Scan sweep | 73.2
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Full range of X Detector Raw during X Scan sweep | 14.1
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Full range of X Detector during X Scan sweep | 62.4
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Full range of X Detector during X Scan sweep | 65.7
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.05
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB) | 50.05
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.05
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB) | 50.05
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (Y offset-Point Align3) | 14.75
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (Y offset-Point Align3) | 1.74
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (Y offset-Point Align2) | 0.83
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (Y offset-Point Align2) | 1.07
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (Y offset-Point Align1) | 0.74
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (Y offset-Point Align1) | 0.79
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (Y offset) | -0.73
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (Y offset) | 0.64
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (X offset-Point Align3) | 0.97
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (X offset-Point Align3) | 0.72
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (X offset-Point Align2) | 1.96
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (X offset-Point Align2) | 0.6
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (X offset-Point Align1) | 13.6
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (X offset-Point Align1) | 1.59
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | Aligned stage accuracy (X offset) | 0.88
| SMSC (SMIC Shanghai Mask shop) | 2024-02-28 | Aligned stage accuracy (X offset) | 0.95
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | 940nm Step Height Sample (10 Repeats) | 948.47
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | 50~200nm Step Height Sample (10 Repeats) | 183.34
| QuanYi Mask Optoelectronics Technology #2 | 2024-02-28 | 200~500nm Step Height Sample (10 Repeats) | 461
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Z Scanner position at -3μm | 0.0012
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Z Scanner position at +3μm | 0.0112
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Z Linearity by Z Scanner position
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Z Linearity by Z Scanner position
| 김해의생명산업진흥원 | 2024-02-16 | Z Linearity by Z Scanner position
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Z Linearity by 3 different sample height | 1.007
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Z Linearity by 3 different sample height | 1.006
| 김해의생명산업진흥원 | 2024-02-16 | Z Linearity by 3 different sample height | 0.998
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Z Height noise | 0.028
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Z Height noise | 0.019
| 김해의생명산업진흥원 | 2024-02-16 | Z Height noise | 0.02
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y2 Detector noise (XY Servo on) | 0.22
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y2 Detector noise (XY Servo off) | 0.22
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y1 Detector noise (XY Servo on) | 0.24
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y1 Detector noise (XY Servo off) | 0.24
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y Scan Stroke | 177.2
| Nagoya Univ(Prof. Kato) | 2024-02-16 | -Y Direction software Limit | 101669
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y Detector noise (XY Servo on) | 0.17
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Y Detector noise (XY Servo on) | 0.26
| 김해의생명산업진흥원 | 2024-02-16 | Y Detector noise (XY Servo on) | 0.3
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y Detector noise (XY Servo off) | 0.16
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Y Detector noise (XY Servo off) | 0.26
| 김해의생명산업진흥원 | 2024-02-16 | Y Detector noise (XY Servo off) | 0.28
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Y axis OPM | 0.82
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Y axis OPM | 0.78
| 김해의생명산업진흥원 | 2024-02-16 | Y axis OPM | 0.96
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X2 Detector noise (XY Servo on) | 0.23
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X2 Detector noise (XY Servo off) | 0.23
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X1 Detector noise (XY Servo on) | 0.24
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X1 Detector noise (XY Servo off) | 0.25
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X Scan Stroke | 177.9
| Nagoya Univ(Prof. Kato) | 2024-02-16 | -X Direction software Limit | 103230
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X Detector noise (XY Servo on) | 0.18
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | X Detector noise (XY Servo on) | 0.28
| 김해의생명산업진흥원 | 2024-02-16 | X Detector noise (XY Servo on) | 0.28
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X Detector noise (XY Servo off) | 0.18
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | X Detector noise (XY Servo off) | 0.28
| 김해의생명산업진흥원 | 2024-02-16 | X Detector noise (XY Servo off) | 0.27
| Nagoya Univ(Prof. Kato) | 2024-02-16 | X axis OPM | 0.79
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | X axis OPM | 0.78
| 김해의생명산업진흥원 | 2024-02-16 | X axis OPM | 1.47
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | SLD beam power at the end of the lens | 0.24
| 김해의생명산업진흥원 | 2024-02-16 | SLD beam power at the end of the lens | 0.21
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.38
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| 김해의생명산업진흥원 | 2024-02-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 김해의생명산업진흥원 | 2024-02-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.44
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.44
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.41
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.39
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.38
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| 김해의생명산업진흥원 | 2024-02-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Nagoya Univ(Prof. Kato) | 2024-02-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.53
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 김해의생명산업진흥원 | 2024-02-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Nagoya Univ(Prof. Kato) | 2024-02-16 | NCM driving% with AC160TS cantilever | 0.44
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | NCM driving% with AC160TS cantilever | 4.05
| 김해의생명산업진흥원 | 2024-02-16 | NCM driving% with AC160TS cantilever | 2.08
| Nagoya Univ(Prof. Kato) | 2024-02-16 | NCM Amplitude noise | 12.47
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | NCM Amplitude noise | 10.41
| 김해의생명산업진흥원 | 2024-02-16 | NCM Amplitude noise | 14.97
| Nagoya Univ(Prof. Kato) | 2024-02-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.23
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.343
| 김해의생명산업진흥원 | 2024-02-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.148
| Nagoya Univ(Prof. Kato) | 2024-02-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.08
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.24
| 김해의생명산업진흥원 | 2024-02-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.875
| Nagoya Univ(Prof. Kato) | 2024-02-16 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 237.3
| Nagoya Univ(Prof. Kato) | 2024-02-16 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 248.2
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 159.2
| 김해의생명산업진흥원 | 2024-02-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 169.6
| Nagoya Univ(Prof. Kato) | 2024-02-16 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 236.4
| Nagoya Univ(Prof. Kato) | 2024-02-16 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 242.7
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156
| 김해의생명산업진흥원 | 2024-02-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 169.9
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Full range of Z Height during Z Scan sweep | 16.4
| 김해의생명산업진흥원 | 2024-02-16 | Full range of Z Height during Z Scan sweep | 16
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | LD beam power at the end of the lens | 2.5
| 김해의생명산업진흥원 | 2024-02-16 | LD beam power at the end of the lens | 1.62
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Lateral(C-D) sensitivity | 15.7
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Lateral(C-D) sensitivity | 30.9
| 김해의생명산업진흥원 | 2024-02-16 | Lateral(C-D) sensitivity | 4.2
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Lateral(C-D) noise | 0.131
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Lateral(C-D) noise | 0.153
| 김해의생명산업진흥원 | 2024-02-16 | Lateral(C-D) noise | 0.136
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of Z Height during Z Scan sweep | 18.1
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of Z Detector during Z Scan sweep | 13
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Full range of Z Detector during Z Scan sweep | 15.4
| 김해의생명산업진흥원 | 2024-02-16 | Full range of Z Detector during Z Scan sweep | 14
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of Y1 Detector Raw during Y Scan sweep | 14.8
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| 김해의생명산업진흥원 | 2024-02-16 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of Y Detector during Y Scan sweep | 177.7
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Full range of Y Detector during Y Scan sweep | 112.1
| 김해의생명산업진흥원 | 2024-02-16 | Full range of Y Detector during Y Scan sweep | 122.2
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of X2 Detector Raw during X Scan sweep | 14.8
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of X1 Detector Raw during X Scan sweep | 14.9
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Full range of X Detector Raw during X Scan sweep | 14.1
| 김해의생명산업진흥원 | 2024-02-16 | Full range of X Detector Raw during X Scan sweep | 14.3
| Nagoya Univ(Prof. Kato) | 2024-02-16 | Full range of X Detector during X Scan sweep | 177.8
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Full range of X Detector during X Scan sweep | 110
| 김해의생명산업진흥원 | 2024-02-16 | Full range of X Detector during X Scan sweep | 121.4
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | EFM Phase : For the positive bias | 90.9
| 김해의생명산업진흥원 | 2024-02-16 | EFM Phase : For the positive bias | 90.2
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | EFM Phase : For the negative bias | -90.1
| 김해의생명산업진흥원 | 2024-02-16 | EFM Phase : For the negative bias | -90.2
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Approach position at -3μm | 0.0011
| 김해의생명산업진흥원 | 2024-02-16 | Approach position at -3μm | 0.0101
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | Approach position at +3μm | 0.0003
| 김해의생명산업진흥원 | 2024-02-16 | Approach position at +3μm | 0.0001
| Nagoya Univ(Prof. Kato) | 2024-02-16 | A-B noise | 8.49
| Universidad Autonoma de Madrid (PSE) (UA Madrid) | 2024-02-16 | A-B noise | 19.33
| 김해의생명산업진흥원 | 2024-02-16 | A-B noise | 16.06
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Z Scanner position at -3μm | 0.0005
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Z Scanner position at +3μm | 0.0098
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Z Linearity by Z Scanner position
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Z Linearity by 3 different sample height | 1.002
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Z Height noise | 0.031
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Y Detector noise (XY Servo on) | 0.26
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Y Detector noise (XY Servo off) | 0.25
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Y axis OPM | 0.89
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | X Detector noise (XY Servo on) | 0.26
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | X Detector noise (XY Servo off) | 0.26
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | X axis OPM | 1.18
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | NCM driving% with AC160TS cantilever | 1.63
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | NCM Amplitude noise | 14.55
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.577
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.81
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.9
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 171.3
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Lateral(C-D) sensitivity | 16
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Lateral(C-D) noise | 0.099
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Full range of Z Height during Z Scan sweep | 18.8
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Full range of Z Detector during Z Scan sweep | 15.2
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Full range of Y Detector during Y Scan sweep | 122
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Full range of X Detector Raw during X Scan sweep | 14.1
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | Full range of X Detector during X Scan sweep | 120.8
| National Research Council, Institute for photonics and nanotechnologies - (CNR, IFN) (Gembetti) Zupella | 2024-02-08 | A-B noise | 8.52
| University of Luxembourg (ST Instruments) | 2024-02-02 | Z Scanner position at -3μm | 0.0006
| University of Luxembourg (ST Instruments) | 2024-02-02 | Z Scanner position at +3μm | 0.0061
| University of Luxembourg (ST Instruments) | 2024-02-02 | Z Linearity by Z Scanner position
| University of Luxembourg (ST Instruments) | 2024-02-02 | Z Linearity by 3 different sample height | 0.991
| University of Luxembourg (ST Instruments) | 2024-02-02 | Z Height noise | 0.031
| University of Luxembourg (ST Instruments) | 2024-02-02 | Y Scan Stroke | 55
| University of Luxembourg (ST Instruments) | 2024-02-02 | Y Detector noise (XY Servo on) | 0.29
| University of Luxembourg (ST Instruments) | 2024-02-02 | Y Detector noise (XY Servo off) | 0.28
| University of Luxembourg (ST Instruments) | 2024-02-02 | Y axis OPM | 1.12
| University of Luxembourg (ST Instruments) | 2024-02-02 | X Scan Stroke | 55
| University of Luxembourg (ST Instruments) | 2024-02-02 | X Detector noise (XY Servo on) | 0.29
| University of Luxembourg (ST Instruments) | 2024-02-02 | X Detector noise (XY Servo off) | 0.28
| University of Luxembourg (ST Instruments) | 2024-02-02 | X axis OPM | 1.51
| University of Luxembourg (ST Instruments) | 2024-02-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| University of Luxembourg (ST Instruments) | 2024-02-02 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| University of Luxembourg (ST Instruments) | 2024-02-02 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| University of Luxembourg (ST Instruments) | 2024-02-02 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| University of Luxembourg (ST Instruments) | 2024-02-02 | NCM driving% with AC160TS cantilever | 3.67
| University of Luxembourg (ST Instruments) | 2024-02-02 | NCM Amplitude noise | 10.23
| University of Luxembourg (ST Instruments) | 2024-02-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.97
| University of Luxembourg (ST Instruments) | 2024-02-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.628
| University of Luxembourg (ST Instruments) | 2024-02-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102
| University of Luxembourg (ST Instruments) | 2024-02-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101
| University of Luxembourg (ST Instruments) | 2024-02-02 | Lateral(C-D) sensitivity | 42.4
| University of Luxembourg (ST Instruments) | 2024-02-02 | Lateral(C-D) noise | 0.059
| University of Luxembourg (ST Instruments) | 2024-02-02 | Full range of Z Height during Z Scan sweep | 15.6
| University of Luxembourg (ST Instruments) | 2024-02-02 | Full range of Z Detector during Z Scan sweep | 11.3
| University of Luxembourg (ST Instruments) | 2024-02-02 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| University of Luxembourg (ST Instruments) | 2024-02-02 | Full range of Y Detector during Y Scan sweep | 73.6
| University of Luxembourg (ST Instruments) | 2024-02-02 | Full range of X Detector Raw during X Scan sweep | 14.5
| University of Luxembourg (ST Instruments) | 2024-02-02 | Full range of X Detector during X Scan sweep | 73.3
| University of Luxembourg (ST Instruments) | 2024-02-02 | A-B noise | 7.4
| 경남테크노파크 | 2024-01-31 | Z Scanner position at -3μm | 0.0035
| 경남테크노파크 | 2024-01-31 | Z Scanner position at +3μm | 0.0121
| 경남테크노파크 | 2024-01-31 | Z Linearity by Z Scanner position
| 베스트그래핀 | 2024-01-31 | Z Linearity by Z Scanner position
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Z linearity by sample height | 0.994
| 경남테크노파크 | 2024-01-31 | Z Linearity by 3 different sample height | 0.995
| 베스트그래핀 | 2024-01-31 | Z Linearity by 3 different sample height | 1.006
| 경남테크노파크 | 2024-01-31 | Z Height noise | 0.03
| 베스트그래핀 | 2024-01-31 | Z Height noise | 0.022
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Z Height noise | 0.033
| 경남테크노파크 | 2024-01-31 | Y2 Detector noise (XY Servo on) | 0.24
| 경남테크노파크 | 2024-01-31 | Y2 Detector noise (XY Servo off) | 0.24
| 경남테크노파크 | 2024-01-31 | Y1 Detector noise (XY Servo on) | 0.24
| 경남테크노파크 | 2024-01-31 | Y1 Detector noise (XY Servo off) | 0.25
| 경남테크노파크 | 2024-01-31 | Y Scan Stroke | 178.8
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Y Scan Stroke | 55
| 경남테크노파크 | 2024-01-31 | -Y Direction software Limit | -102856
| 경남테크노파크 | 2024-01-31 | Y Detector noise (XY Servo on) | 0.18
| 베스트그래핀 | 2024-01-31 | Y Detector noise (XY Servo on) | 0.26
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Y Detector noise (XY Servo on) | 0.25
| 경남테크노파크 | 2024-01-31 | Y Detector noise (XY Servo off) | 0.18
| 베스트그래핀 | 2024-01-31 | Y Detector noise (XY Servo off) | 0.26
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Y Detector noise (XY Servo off) | 0.25
| 경남테크노파크 | 2024-01-31 | Y axis OPM | 1.11
| 베스트그래핀 | 2024-01-31 | Y axis OPM | 0.55
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Y axis OPM | 0.95
| 경남테크노파크 | 2024-01-31 | X2 Detector noise (XY Servo on) | 0.24
| 경남테크노파크 | 2024-01-31 | X2 Detector noise (XY Servo off) | 0.24
| 경남테크노파크 | 2024-01-31 | X1 Detector noise (XY Servo on) | 0.24
| 경남테크노파크 | 2024-01-31 | X1 Detector noise (XY Servo off) | 0.24
| 경남테크노파크 | 2024-01-31 | X Scan Stroke | 178.6
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | X Scan Stroke | 55
| 경남테크노파크 | 2024-01-31 | -X Direction software Limit | -102176
| 경남테크노파크 | 2024-01-31 | X Detector noise (XY Servo on) | 0.18
| 베스트그래핀 | 2024-01-31 | X Detector noise (XY Servo on) | 0.27
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | X Detector noise (XY Servo on) | 0.26
| 경남테크노파크 | 2024-01-31 | X Detector noise (XY Servo off) | 0.17
| 베스트그래핀 | 2024-01-31 | X Detector noise (XY Servo off) | 0.26
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | X Detector noise (XY Servo off) | 0.25
| 경남테크노파크 | 2024-01-31 | X axis OPM | 1.73
| 베스트그래핀 | 2024-01-31 | X axis OPM | 1.34
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | X axis OPM | 0.79
| 베스트그래핀 | 2024-01-31 | SLD beam power at the end of the lens | 0.21
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| 베스트그래핀 | 2024-01-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 베스트그래핀 | 2024-01-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.37
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.33
| 경남테크노파크 | 2024-01-31 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.34
| 경남테크노파크 | 2024-01-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| 베스트그래핀 | 2024-01-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 경남테크노파크 | 2024-01-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| 베스트그래핀 | 2024-01-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 경남테크노파크 | 2024-01-31 | NCM driving% with AC160TS cantilever | 1.29
| 베스트그래핀 | 2024-01-31 | NCM driving% with AC160TS cantilever | 3.16
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | NCM driving% with AC160TS cantilever | 0.81
| 경남테크노파크 | 2024-01-31 | NCM Amplitude noise | 15.98
| 베스트그래핀 | 2024-01-31 | NCM Amplitude noise | 14.09
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | NCM Amplitude noise | 13.48
| 경남테크노파크 | 2024-01-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.964
| 베스트그래핀 | 2024-01-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.315
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.939
| 경남테크노파크 | 2024-01-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.079
| 베스트그래핀 | 2024-01-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.608
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.083
| 경남테크노파크 | 2024-01-31 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 257.1
| 경남테크노파크 | 2024-01-31 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 253.8
| 베스트그래핀 | 2024-01-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 151.8
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102.4
| 경남테크노파크 | 2024-01-31 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 259.1
| 경남테크노파크 | 2024-01-31 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 253.8
| 베스트그래핀 | 2024-01-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 158.1
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Maximum Y Scan range (Full Stroke) | 73.4
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Maximum X Scan range (Full Stroke) | 73.8
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Full range of Z Height during Z Scan sweep | 17.5
| 베스트그래핀 | 2024-01-31 | Full range of Z Height during Z Scan sweep | 16.4
| 경남테크노파크 | 2024-01-31 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| 베스트그래핀 | 2024-01-31 | LD beam power at the end of the lens | 1.7
| 경남테크노파크 | 2024-01-31 | Lateral(C-D) sensitivity | 14.3
| 베스트그래핀 | 2024-01-31 | Lateral(C-D) sensitivity | 24.6
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Lateral(C-D) sensitivity | 13.1
| 경남테크노파크 | 2024-01-31 | Lateral(C-D) noise | 0.15
| 베스트그래핀 | 2024-01-31 | Lateral(C-D) noise | 0.23
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Lateral(C-D) noise | 0.096
| 경남테크노파크 | 2024-01-31 | Full range of Z Height during Z Scan sweep | 17.8
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Full range of Z Detector raw | 14
| 경남테크노파크 | 2024-01-31 | Full range of Z Detector during Z Scan sweep | 14.2
| 베스트그래핀 | 2024-01-31 | Full range of Z Detector during Z Scan sweep | 14.5
| 경남테크노파크 | 2024-01-31 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| 경남테크노파크 | 2024-01-31 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| 베스트그래핀 | 2024-01-31 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Full range of Y Detector raw | 14.3
| 경남테크노파크 | 2024-01-31 | Full range of Y Detector during Y Scan sweep | 178.9
| 베스트그래핀 | 2024-01-31 | Full range of Y Detector during Y Scan sweep | 111.8
| 경남테크노파크 | 2024-01-31 | Full range of X2 Detector Raw during X Scan sweep | 13.8
| 경남테크노파크 | 2024-01-31 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| 베스트그래핀 | 2024-01-31 | Full range of X Detector Raw during X Scan sweep | 14
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Full range of X Detector raw | 14.5
| 경남테크노파크 | 2024-01-31 | Full range of X Detector during X Scan sweep | 178.6
| 베스트그래핀 | 2024-01-31 | Full range of X Detector during X Scan sweep | 110.3
| 베스트그래핀 | 2024-01-31 | EFM Phase : For the positive bias | 90.1
| 베스트그래핀 | 2024-01-31 | EFM Phase : For the negative bias | -90.8
| 베스트그래핀 | 2024-01-31 | Approach position at -3μm | 0.0014
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Approach position at -3μm | 0.0002
| 베스트그래핀 | 2024-01-31 | Approach position at +3μm | 0.0094
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | Approach position at +3μm | 0.0086
| 경남테크노파크 | 2024-01-31 | A-B noise | 8.09
| 베스트그래핀 | 2024-01-31 | A-B noise | 18.12
| 한국해양대학교 나노반도체공학과 | 2024-01-31 | A-B noise | 9.63
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Z Scanner position at -3μm | 0.0038
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Z Scanner position at +3μm | 0.0048
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Z Linearity by Z Scanner position
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Z Linearity by Z Scanner position
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Z Linearity by 3 different sample height | 1
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Z Linearity by 3 different sample height | 1
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Z Height noise | 0.031
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Z Height noise | 0.022
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y2 Detector noise (XY Servo on) | 0.23
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y2 Detector noise (XY Servo off) | 0.22
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y1 Detector noise (XY Servo on) | 0.23
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y1 Detector noise (XY Servo off) | 0.23
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y Scan Stroke | 178.7
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y Detector noise (XY Servo on) | 0.16
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Y Detector noise (XY Servo on) | 0.28
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y Detector noise (XY Servo off) | 0.16
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Y Detector noise (XY Servo off) | 0.27
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Y axis OPM | 0.88
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Y axis OPM | 1.44
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X2 Detector noise (XY Servo on) | 0.25
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X2 Detector noise (XY Servo off) | 0.25
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X1 Detector noise (XY Servo on) | 0.23
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X1 Detector noise (XY Servo off) | 0.22
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X Scan Stroke | 174.3
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X Detector noise (XY Servo on) | 0.18
| Tianjin De Quan Xing Ye technology | 2024-01-29 | X Detector noise (XY Servo on) | 0.27
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X Detector noise (XY Servo off) | 0.18
| Tianjin De Quan Xing Ye technology | 2024-01-29 | X Detector noise (XY Servo off) | 0.27
| Park Systems Europe GmbH (PSE) | 2024-01-29 | X axis OPM | 1.04
| Tianjin De Quan Xing Ye technology | 2024-01-29 | X axis OPM | 1.07
| Tianjin De Quan Xing Ye technology | 2024-01-29 | SLD beam power at the end of the lens | 0.26
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.34
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.28
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Park Systems Europe GmbH (PSE) | 2024-01-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Park Systems Europe GmbH (PSE) | 2024-01-29 | NCM driving% with AC160TS cantilever | 0.32
| Tianjin De Quan Xing Ye technology | 2024-01-29 | NCM driving% with AC160TS cantilever | 2.1
| Park Systems Europe GmbH (PSE) | 2024-01-29 | NCM Amplitude noise | 12.8
| Park Systems Europe GmbH (PSE) | 2024-01-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.11
| Tianjin De Quan Xing Ye technology | 2024-01-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.331
| Park Systems Europe GmbH (PSE) | 2024-01-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.331
| Tianjin De Quan Xing Ye technology | 2024-01-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.759
| Park Systems Europe GmbH (PSE) | 2024-01-29 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 245.8
| Park Systems Europe GmbH (PSE) | 2024-01-29 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 242.1
| Tianjin De Quan Xing Ye technology | 2024-01-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.3
| Park Systems Europe GmbH (PSE) | 2024-01-29 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 249.4
| Park Systems Europe GmbH (PSE) | 2024-01-29 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238.4
| Tianjin De Quan Xing Ye technology | 2024-01-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 147.8
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Full range of Z Height during Z Scan sweep | 16
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Tianjin De Quan Xing Ye technology | 2024-01-29 | LD beam power at the end of the lens | 2.51
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Lateral(C-D) sensitivity | 12.2
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Lateral(C-D) noise | 0.128
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of Z Height during Z Scan sweep | 18
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of Z Detector during Z Scan sweep | 14.2
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Full range of Z Detector during Z Scan sweep | 13.4
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of Y Detector during Y Scan sweep | 179.1
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Full range of Y Detector during Y Scan sweep | 111.9
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of X2 Detector Raw during X Scan sweep | 14
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of X1 Detector Raw during X Scan sweep | 14
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Full range of X Detector Raw during X Scan sweep | 15.2
| Park Systems Europe GmbH (PSE) | 2024-01-29 | Full range of X Detector during X Scan sweep | 174
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Full range of X Detector during X Scan sweep | 112.3
| Tianjin De Quan Xing Ye technology | 2024-01-29 | EFM Phase : For the positive bias | 90.4
| Tianjin De Quan Xing Ye technology | 2024-01-29 | EFM Phase : For the negative bias | -90.2
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Approach position at -3μm | 0.0018
| Tianjin De Quan Xing Ye technology | 2024-01-29 | Approach position at +3μm | 0.013
| Park Systems Europe GmbH (PSE) | 2024-01-29 | A-B noise | 9.23
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Linearity by Z scanner position | 0.7
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Linearity by Z scanner position | 0.52
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Linearity by 4 different sample height | 0.999
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Linearity by 4 different sample height | 1.006
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Height noise | 0.028
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Height noise | 0.035
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.17
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.18
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.07
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.68
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Z Drift (Short Traveling) | 0.12
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Z Drift (Short Traveling) | 0.03
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y2 Detector noise (XY Servo on) | 0.4
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y2 Detector noise (XY Servo on) | 0.63
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 265.6
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.8
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y1 Detector noise (XY Servo on) | 0.4
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y1 Detector noise (XY Servo on) | 0.62
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 255.8
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 268.4
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 190.6
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 191.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y Ringing(Rpv) | 0.23
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y Ringing(Rpv) | 0.22
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y Non-linearity error | 2.73
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y Non-linearity error | 2.62
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y Detector noise (XY Servo on) | 0.32
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y Detector noise (XY Servo on) | 0.57
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Repeatability @1st position | 0.03
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Repeatability @1st position | 0.05
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Repeatability @ 4th position | 0.06
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Repeatability @ 4th position | 0.04
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Repeatability @ 3rd position | 0.04
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Repeatability @ 3rd position | 0.03
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Repeatability @ 2nd position | 0.05
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Repeatability @ 2nd position | 0.06
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Accuracy @1st position | 0.33
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Accuracy @1st position | 0.53
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Accuracy @ 4th position | 0.65
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Accuracy @ 4th position | 0.42
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Accuracy @ 3rd position | 0.39
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Accuracy @ 3rd position | 0.33
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis, Accuracy @ 2nd position | 0.4
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis, Accuracy @ 2nd position | 0.42
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Y axis OPM | 1.21
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Y axis OPM | 0.65
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Stage positioning Y offset (Vision Pattern Matching) | 0.62
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Stage positioning Y offset (Vision Pattern Matching) | 1.67
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Stage positioning Y offset (LLC Translation) | 1.94
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Stage positioning Y offset (LLC Translation) | 1.79
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Stage positioning Y offset (Global Align) | 4.21
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Stage positioning Y offset (Global Align) | 2.92
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Stage positioning X offset (Vision Pattern Matching) | 1.89
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Stage positioning X offset (Vision Pattern Matching) | 1.82
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Stage positioning X offset (LLC Translation) | 2.09
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Stage positioning X offset (LLC Translation) | 1.64
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Stage positioning X offset (Global Align) | 2.42
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Stage positioning X offset (Global Align) | 1.75
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | XY Orthogonality | 90
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | XY Orthogonality | 89.7
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X2 Detector noise (XY Servo on) | 0.46
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X2 Detector noise (XY Servo on) | 0.42
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 266.4
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 267
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X1 Detector noise (XY Servo on) | 0.41
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X1 Detector noise (XY Servo on) | 0.41
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 262.8
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 263
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 190
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 181.2
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X Ringing(Rpv) | 0.21
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X Ringing(Rpv) | 0.46
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X Non-linearity error | 2.16
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X Non-linearity error | 2.18
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X Detector noise (XY Servo on) | 0.37
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X Detector noise (XY Servo on) | 0.32
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Repeatability @1st position | 0.03
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Repeatability @1st position | 0.03
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Repeatability @ 4th position | 0.03
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Repeatability @ 4th position | 0.04
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Repeatability @ 3rd position | 0.03
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Repeatability @ 3rd position | 0.03
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Repeatability @ 2nd position | 0.03
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Repeatability @ 2nd position | 0.05
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Accuracy @1st position | 0.17
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Accuracy @1st position | 0.17
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Accuracy @ 4th position | 0.25
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Accuracy @ 4th position | 0.24
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Accuracy @ 3rd position | 0.14
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Accuracy @ 3rd position | 0.18
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis, Accuracy @ 2nd position | 0.19
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis, Accuracy @ 2nd position | 0.26
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | X axis OPM | 0.45
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | X axis OPM | 1.44
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 26.8
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Wafer placement(Loading) error on the load port2 (Y offset) | 9.95
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Wafer placement(Loading) error on the load port2 (X offset) | 13.45
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Wafer placement(Loading) error on the load port2 (X offset) | 8.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 63.77
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Wafer placement(Loading) error on the load port1 (Y offset) | 9.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Wafer placement(Loading) error on the load port1 (X offset) | 13.05
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Wafer placement(Loading) error on the load port1 (X offset) | 19.8
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Under 50nm Step Height Sample (10 Repeats) | 18.15
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Under 50nm Step Height Sample (10 Repeats) | 18.21
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Torque Profile (50mm, 20um/s) | 2.482
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Torque Profile (50mm, 20um/s) | 3.078
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Torque Profile (50mm, 200um/s) | 1.692
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Torque Profile (50mm, 200um/s) | 1.846
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | StrokeYMin | -153536
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | StrokeYMin | -153069
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | StrokeYMax | 152127
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | StrokeYMax | 150269.7
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | StrokeY | 305663
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | StrokeY | 303338.6
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | StrokeXMin | -205345
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | StrokeXMin | -204032
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | StrokeXMax | 204220
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | StrokeXMax | 202797.3
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | StrokeX | 409565
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | StrokeX | 406829.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.763
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.931
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Roll motion to X+ full scan | 2.4
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Roll motion to X+ full scan | 5.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | RMS of ZeroScan | 0.059
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | RMS of ZeroScan | 0.051
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Profiler noise @200um | 0.133
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Profiler noise @200um | 0.115
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Position repeatability(10repeats, Position Based) | 0.28
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Position repeatability(10repeats, Position Based) | 0.2
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00015
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OpticalAxisShiftPerFocusY | -0.0015
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OpticalAxisShiftPerFocusY | -0.0024
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.0002
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.0001
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OpticalAxisShiftPerFocusX | -0.0011
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OpticalAxisShiftPerFocusX | -0.0023
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OPM Repeatability @5mm | 0.67
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OPM Repeatability @5mm | 0.77
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OPM Repeatability @25mm | 1.09
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OPM Repeatability @25mm | 6.43
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OPM Repeatability @1mm | 0.67
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OPM Repeatability @1mm | 1.18
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | OPM Repeatability @10mm | 0.94
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | OPM Repeatability @10mm | 1.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | NCM driving% with AC160TS cantilever | 27.2
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | NCM driving% with AC160TS cantilever | 4.2
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Min. Vibration Transmissbility (~100Hz) | 0.02
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Min. Vibration Transmissbility (~100Hz) | 0.022
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Max OPM @5mm | 9.72
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Max OPM @5mm | 19.85
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Max OPM @25mm | 39.27
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Max OPM @25mm | 139.51
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Max OPM @1mm | 4.54
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Max OPM @1mm | 8.32
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Max OPM @10mm | 15.79
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Max OPM @10mm | 50.82
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | LowerLimit | 4348
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | LowerLimit | 3273
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Lift pins synchronization offset check (Y offset) | 14.7
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Lift pins synchronization offset check (Y offset) | 2.74
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Lift pins synchronization offset check (X offset) | 2.2
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Lift pins synchronization offset check (X offset) | 30.26
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Lateral(C-D) sensitivity | 29.5
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Lateral(C-D) sensitivity | 7.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Lateral(C-D) noise | 0.238
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Lateral(C-D) noise | 0.19
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from -Y to +Y (Y scan) | 90.09
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from -Y to +Y (Y scan) | 38.09
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from -Y to +Y (X scan) | 12.45
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from -Y to +Y (X scan) | 9.52
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from -X to +X (Y scan) | 17.58
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from -X to +X (Y scan) | 39.55
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from -X to +X (X scan) | 60.79
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from -X to +X (X scan) | 64.45
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from +Y to -Y (Y scan) | 21.97
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from +Y to -Y (Y scan) | 34.42
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from +Y to -Y (X scan) | 13.92
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from +Y to -Y (X scan) | 7.32
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from +X to -X (Y scan) | 24.17
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from +X to -X (Y scan) | 20.51
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Image offset after moving from +X to -X (X scan) | 18.31
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Image offset after moving from +X to -X (X scan) | 35.16
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of Z Height during Z Scan sweep | 15.7
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of Z Height during Z Scan sweep | 16.9
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of Z Detector during Z Scan sweep | 11.2
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of Z Detector during Z Scan sweep | 12.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.8
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of Y2 Detector Raw during Y Scan sweep | 13.62
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of Y2 Detector during Y Scan sweep | 214.6
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of Y2 Detector during Y Scan sweep | 203.6
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.49
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.34
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of Y1 Detector during Y Scan sweep | 202.8
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of Y1 Detector during Y Scan sweep | 209.8
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of X2 Detector Raw during X Scan sweep | 14.23
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of X2 Detector Raw during X Scan sweep | 13.73
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of X2 Detector during X Scan sweep | 203.7
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of X2 Detector during X Scan sweep | 199.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of X1 Detector Raw during X Scan sweep | 14.49
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of X1 Detector Raw during X Scan sweep | 13.85
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Full range of X1 Detector during X Scan sweep | 206.3
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Full range of X1 Detector during X Scan sweep | 195
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.17
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.16
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.92
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.9
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.56
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.54
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.36
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.51
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.45
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.32
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | 940nm Step Height Sample (10 Repeats) | 947.34
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | 940nm Step Height Sample (10 Repeats) | 950.53
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | 50~200nm Step Height Sample (10 Repeats) | 184.71
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | 50~200nm Step Height Sample (10 Repeats) | 182.62
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | 200~500nm Step Height Sample (10 Repeats) | 462.67
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | 200~500nm Step Height Sample (10 Repeats) | 457.5
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | * 1800nm Step Height Sample (10 Repeats) | 1805.82
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | * 1800nm Step Height Sample (10 Repeats) | 1792.33
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 1.9
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.6
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.9
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.1
| Shenzhen Pensun Technology Co., Ltd #1 | 2024-01-25 | %GRR w/ roughness of non-patterened wafer | 5.2
| Shenzhen Pensun Technology Co., Ltd #2 | 2024-01-25 | %GRR w/ roughness of non-patterened wafer | 10
| Xidian University | 2024-01-24 | Z Scanner position at -3μm | 0.0076
| 서강대학교 물리학과 | 2024-01-24 | Z Scanner position at -3μm | 0.0132
| Xidian University | 2024-01-24 | Z Scanner position at +3μm | 0.0084
| 서강대학교 물리학과 | 2024-01-24 | Z Scanner position at +3μm | 0.0118
| Xidian University | 2024-01-24 | Z Linearity by Z Scanner position
| 서강대학교 물리학과 | 2024-01-24 | Z Linearity by Z Scanner position
| Xidian University | 2024-01-24 | Z Linearity by 3 different sample height | 1.006
| 서강대학교 물리학과 | 2024-01-24 | Z Linearity by 3 different sample height | 1.003
| Xidian University | 2024-01-24 | Z Height noise | 0.028
| 서강대학교 물리학과 | 2024-01-24 | Z Height noise | 0.026
| Xidian University | 2024-01-24 | Y2 Detector noise (XY Servo on) | 0.23
| Xidian University | 2024-01-24 | Y2 Detector noise (XY Servo off) | 0.23
| Xidian University | 2024-01-24 | Y1 Detector noise (XY Servo on) | 0.22
| Xidian University | 2024-01-24 | Y1 Detector noise (XY Servo off) | 0.21
| Xidian University | 2024-01-24 | Y Scan Stroke | 161.2
| 서강대학교 물리학과 | 2024-01-24 | Y Scan Stroke | 55
| Xidian University | 2024-01-24 | Y Detector noise (XY Servo on) | 0.16
| 서강대학교 물리학과 | 2024-01-24 | Y Detector noise (XY Servo on) | 0.27
| Xidian University | 2024-01-24 | Y Detector noise (XY Servo off) | 0.15
| 서강대학교 물리학과 | 2024-01-24 | Y Detector noise (XY Servo off) | 0.26
| Xidian University | 2024-01-24 | Y axis OPM | 1.18
| 서강대학교 물리학과 | 2024-01-24 | Y axis OPM | 1.14
| Xidian University | 2024-01-24 | X2 Detector noise (XY Servo on) | 0.26
| Xidian University | 2024-01-24 | X2 Detector noise (XY Servo off) | 0.26
| Xidian University | 2024-01-24 | X1 Detector noise (XY Servo on) | 0.24
| Xidian University | 2024-01-24 | X1 Detector noise (XY Servo off) | 0.24
| Xidian University | 2024-01-24 | X Scan Stroke | 157.5
| 서강대학교 물리학과 | 2024-01-24 | X Scan Stroke | 55
| Xidian University | 2024-01-24 | X Detector noise (XY Servo on) | 0.2
| 서강대학교 물리학과 | 2024-01-24 | X Detector noise (XY Servo on) | 0.26
| Xidian University | 2024-01-24 | X Detector noise (XY Servo off) | 0.19
| 서강대학교 물리학과 | 2024-01-24 | X Detector noise (XY Servo off) | 0.25
| Xidian University | 2024-01-24 | X axis OPM | 1.52
| 서강대학교 물리학과 | 2024-01-24 | X axis OPM | 1.22
| Xidian University | 2024-01-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| 서강대학교 물리학과 | 2024-01-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Xidian University | 2024-01-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| 서강대학교 물리학과 | 2024-01-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Xidian University | 2024-01-24 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.35
| Xidian University | 2024-01-24 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.35
| Xidian University | 2024-01-24 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.3
| Xidian University | 2024-01-24 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| Xidian University | 2024-01-24 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Xidian University | 2024-01-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 서강대학교 물리학과 | 2024-01-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.09
| Xidian University | 2024-01-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 서강대학교 물리학과 | 2024-01-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Xidian University | 2024-01-24 | NCM driving% with AC160TS cantilever | 1.53
| 서강대학교 물리학과 | 2024-01-24 | NCM driving% with AC160TS cantilever | 0.53
| Xidian University | 2024-01-24 | NCM Amplitude noise | 17.6
| 서강대학교 물리학과 | 2024-01-24 | NCM Amplitude noise | 12.53
| Xidian University | 2024-01-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.725
| 서강대학교 물리학과 | 2024-01-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.289
| Xidian University | 2024-01-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.247
| 서강대학교 물리학과 | 2024-01-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.356
| Xidian University | 2024-01-24 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 222.3
| Xidian University | 2024-01-24 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 222.4
| 서강대학교 물리학과 | 2024-01-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105.2
| Xidian University | 2024-01-24 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 220.1
| Xidian University | 2024-01-24 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 217.3
| 서강대학교 물리학과 | 2024-01-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.8
| Xidian University | 2024-01-24 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Xidian University | 2024-01-24 | Lateral(C-D) sensitivity | 45.5
| 서강대학교 물리학과 | 2024-01-24 | Lateral(C-D) sensitivity | 57.2
| Xidian University | 2024-01-24 | Lateral(C-D) noise | 0.139
| 서강대학교 물리학과 | 2024-01-24 | Lateral(C-D) noise | 0.043
| Xidian University | 2024-01-24 | Full range of Z Height during Z Scan sweep | 16.9
| 서강대학교 물리학과 | 2024-01-24 | Full range of Z Height during Z Scan sweep | 15.9
| Xidian University | 2024-01-24 | Full range of Z Detector during Z Scan sweep | 12.8
| 서강대학교 물리학과 | 2024-01-24 | Full range of Z Detector during Z Scan sweep | 12
| Xidian University | 2024-01-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| Xidian University | 2024-01-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| 서강대학교 물리학과 | 2024-01-24 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Xidian University | 2024-01-24 | Full range of Y Detector during Y Scan sweep | 161.4
| 서강대학교 물리학과 | 2024-01-24 | Full range of Y Detector during Y Scan sweep | 73.3
| Xidian University | 2024-01-24 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Xidian University | 2024-01-24 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| 서강대학교 물리학과 | 2024-01-24 | Full range of X Detector Raw during X Scan sweep | 14.4
| Xidian University | 2024-01-24 | Full range of X Detector during X Scan sweep | 157.3
| 서강대학교 물리학과 | 2024-01-24 | Full range of X Detector during X Scan sweep | 73.9
| Xidian University | 2024-01-24 | A-B noise | 7.45
| 서강대학교 물리학과 | 2024-01-24 | A-B noise | 8.5
| 중앙대학교 전기전자공학부 | 2024-01-18 | Z Scanner position at -3μm | 0.0035
| 중앙대학교 전기전자공학부 | 2024-01-18 | Z Scanner position at +3μm | 0.0049
| 중앙대학교 전기전자공학부 | 2024-01-18 | Z Linearity by Z Scanner position
| 중앙대학교 전기전자공학부 | 2024-01-18 | Z Linearity by 3 different sample height | 0.993
| 중앙대학교 전기전자공학부 | 2024-01-18 | Z Height noise | 0.028
| 중앙대학교 전기전자공학부 | 2024-01-18 | Y Scan Stroke | 55
| 중앙대학교 전기전자공학부 | 2024-01-18 | Y Detector noise (XY Servo on) | 0.39
| 중앙대학교 전기전자공학부 | 2024-01-18 | Y Detector noise (XY Servo off) | 0.41
| 중앙대학교 전기전자공학부 | 2024-01-18 | Y axis OPM | 0.66
| 중앙대학교 전기전자공학부 | 2024-01-18 | X Scan Stroke | 55
| 중앙대학교 전기전자공학부 | 2024-01-18 | X Detector noise (XY Servo on) | 0.39
| 중앙대학교 전기전자공학부 | 2024-01-18 | X Detector noise (XY Servo off) | 0.43
| 중앙대학교 전기전자공학부 | 2024-01-18 | X axis OPM | 1.04
| 중앙대학교 전기전자공학부 | 2024-01-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 중앙대학교 전기전자공학부 | 2024-01-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.57
| 중앙대학교 전기전자공학부 | 2024-01-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 중앙대학교 전기전자공학부 | 2024-01-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| 중앙대학교 전기전자공학부 | 2024-01-18 | NCM driving% with AC160TS cantilever | 0.36
| 중앙대학교 전기전자공학부 | 2024-01-18 | NCM Amplitude noise | 12.01
| 중앙대학교 전기전자공학부 | 2024-01-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.631
| 중앙대학교 전기전자공학부 | 2024-01-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.204
| 중앙대학교 전기전자공학부 | 2024-01-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 88
| 중앙대학교 전기전자공학부 | 2024-01-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 90.9
| 중앙대학교 전기전자공학부 | 2024-01-18 | Lateral(C-D) sensitivity | 64.6
| 중앙대학교 전기전자공학부 | 2024-01-18 | Lateral(C-D) noise | 0.01
| 중앙대학교 전기전자공학부 | 2024-01-18 | Full range of Z Height during Z Scan sweep | 18.5
| 중앙대학교 전기전자공학부 | 2024-01-18 | Full range of Z Detector during Z Scan sweep | 15.3
| 중앙대학교 전기전자공학부 | 2024-01-18 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| 중앙대학교 전기전자공학부 | 2024-01-18 | Full range of Y Detector during Y Scan sweep | 66.9
| 중앙대학교 전기전자공학부 | 2024-01-18 | Full range of X Detector Raw during X Scan sweep | 14.4
| 중앙대학교 전기전자공학부 | 2024-01-18 | Full range of X Detector during X Scan sweep | 65.4
| 중앙대학교 전기전자공학부 | 2024-01-18 | A-B noise | 9.62
| PSI(Mexico) Demo | 2024-01-17 | Z Scanner position at -3μm | 0.0025
| PSI(Mexico) Demo | 2024-01-17 | Z Scanner position at +3μm | 0.0043
| PSI(Mexico) Demo | 2024-01-17 | Z Linearity by Z Scanner position
| PSI(Mexico) Demo | 2024-01-17 | Z Linearity by 3 different sample height | 1.006
| PSI(Mexico) Demo | 2024-01-17 | Z Height noise | 0.029
| PSI(Mexico) Demo | 2024-01-17 | Y Detector noise (XY Servo on) | 0.26
| PSI(Mexico) Demo | 2024-01-17 | Y Detector noise (XY Servo off) | 0.26
| PSI(Mexico) Demo | 2024-01-17 | Y axis OPM | 0.26
| PSI(Mexico) Demo | 2024-01-17 | X Detector noise (XY Servo on) | 0.27
| PSI(Mexico) Demo | 2024-01-17 | X Detector noise (XY Servo off) | 0.27
| PSI(Mexico) Demo | 2024-01-17 | X axis OPM | 0.88
| PSI(Mexico) Demo | 2024-01-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| PSI(Mexico) Demo | 2024-01-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| PSI(Mexico) Demo | 2024-01-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| PSI(Mexico) Demo | 2024-01-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| PSI(Mexico) Demo | 2024-01-17 | NCM driving% with AC160TS cantilever | 3.52
| PSI(Mexico) Demo | 2024-01-17 | NCM Amplitude noise | 14.03
| PSI(Mexico) Demo | 2024-01-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.065
| PSI(Mexico) Demo | 2024-01-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.668
| PSI(Mexico) Demo | 2024-01-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.5
| PSI(Mexico) Demo | 2024-01-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167
| PSI(Mexico) Demo | 2024-01-17 | Lateral(C-D) sensitivity | 6.9
| PSI(Mexico) Demo | 2024-01-17 | Lateral(C-D) noise | 0.1
| PSI(Mexico) Demo | 2024-01-17 | Full range of Z Height during Z Scan sweep | 17.4
| PSI(Mexico) Demo | 2024-01-17 | Full range of Z Detector during Z Scan sweep | 13.6
| PSI(Mexico) Demo | 2024-01-17 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| PSI(Mexico) Demo | 2024-01-17 | Full range of Y Detector during Y Scan sweep | 124.3
| PSI(Mexico) Demo | 2024-01-17 | Full range of X Detector Raw during X Scan sweep | 14.4
| PSI(Mexico) Demo | 2024-01-17 | Full range of X Detector during X Scan sweep | 122.6
| PSI(Mexico) Demo | 2024-01-17 | A-B noise | 8.22
| Intel Hillsboro #1 | 2024-01-15 | Z Linearity by Z scanner position | 0.49
| Intel Hillsboro #1 | 2024-01-15 | Z Linearity by 3 different sample height | 1.002
| Intel Hillsboro #1 | 2024-01-15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 59.9
| Intel Hillsboro #1 | 2024-01-15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 60.2
| Intel Hillsboro #1 | 2024-01-15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 26.4
| Intel Hillsboro #1 | 2024-01-15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 26.6
| Intel Hillsboro #1 | 2024-01-15 | Z Drift (Short Traveling) | 0.21
| Intel Hillsboro #1 | 2024-01-15 | Z Detector Min | -0.28
| Intel Hillsboro #1 | 2024-01-15 | Z Detector Max | 0.38
| Intel Hillsboro #1 | 2024-01-15 | Y2 Detector noise (XY Servo on) | 0.29
| Intel Hillsboro #1 | 2024-01-15 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 242.7
| Intel Hillsboro #1 | 2024-01-15 | Y1 Detector noise (XY Servo on) | 0.29
| Intel Hillsboro #1 | 2024-01-15 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 241.6
| Intel Hillsboro #1 | 2024-01-15 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 177.2
| Intel Hillsboro #1 | 2024-01-15 | Y Ringing(Rpv) | 0.27
| Intel Hillsboro #1 | 2024-01-15 | Y Non-linearity error | 2.66
| Intel Hillsboro #1 | 2024-01-15 | Y Detector noise (XY Servo on) | 0.24
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Repeatability @1st position | 0.07
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Repeatability @ 4th position | 0.08
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Repeatability @ 3rd position | 0.07
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Repeatability @ 2nd position | 0.05
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Accuracy @1st position | 0.35
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Accuracy @ 4th position | 0.43
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Accuracy @ 3rd position | 0.43
| Intel Hillsboro #1 | 2024-01-15 | Y axis, Accuracy @ 2nd position | 0.27
| Intel Hillsboro #1 | 2024-01-15 | Y axis OPM | 2.61
| Intel Hillsboro #1 | 2024-01-15 | XY Orthogonality | 90.1
| Intel Hillsboro #1 | 2024-01-15 | X2 Detector noise (XY Servo on) | 0.31
| Intel Hillsboro #1 | 2024-01-15 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 254.5
| Intel Hillsboro #1 | 2024-01-15 | X1 Detector noise (XY Servo on) | 0.39
| Intel Hillsboro #1 | 2024-01-15 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 252.4
| Intel Hillsboro #1 | 2024-01-15 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.7
| Intel Hillsboro #1 | 2024-01-15 | X Ringing(Rpv) | 0.12
| Intel Hillsboro #1 | 2024-01-15 | X Non-linearity error | 3.22
| Intel Hillsboro #1 | 2024-01-15 | X Detector noise (XY Servo on) | 0.27
| Intel Hillsboro #1 | 2024-01-15 | X axis, Repeatability @1st position | 0.03
| Intel Hillsboro #1 | 2024-01-15 | X axis, Repeatability @ 4th position | 0.03
| Intel Hillsboro #1 | 2024-01-15 | X axis, Repeatability @ 3rd position | 0.03
| Intel Hillsboro #1 | 2024-01-15 | X axis, Repeatability @ 2nd position | 0.03
| Intel Hillsboro #1 | 2024-01-15 | X axis, Accuracy @1st position | 0.19
| Intel Hillsboro #1 | 2024-01-15 | X axis, Accuracy @ 4th position | 0.18
| Intel Hillsboro #1 | 2024-01-15 | X axis, Accuracy @ 3rd position | 0.18
| Intel Hillsboro #1 | 2024-01-15 | X axis, Accuracy @ 2nd position | 0.15
| Intel Hillsboro #1 | 2024-01-15 | X axis OPM | 1.77
| Intel Hillsboro #1 | 2024-01-15 | WLI_Scale(Z Calibration) | 33.8
| Intel Hillsboro #1 | 2024-01-15 | Wafer placement(Loading) error on the load port2 (Y offset) | 18.4
| Intel Hillsboro #1 | 2024-01-15 | Wafer placement(Loading) error on the load port2 (X offset) | 23
| Intel Hillsboro #1 | 2024-01-15 | Wafer placement(Loading) error on the load port1 (Y offset) | 10.5
| Intel Hillsboro #1 | 2024-01-15 | Wafer placement(Loading) error on the load port1 (X offset) | 4.6
| Intel Hillsboro #1 | 2024-01-15 | Under 50nm Step Height Sample (10 Repeats) | 18.21
| Intel Hillsboro #1 | 2024-01-15 | Torque Profile (50mm, 20um/s) | 5.26
| Intel Hillsboro #1 | 2024-01-15 | Torque Profile (50mm, 200um/s) | 2.55
| Intel Hillsboro #1 | 2024-01-15 | StrokeYMin | -153158
| Intel Hillsboro #1 | 2024-01-15 | StrokeYMax | 150888
| Intel Hillsboro #1 | 2024-01-15 | StrokeY | 304046
| Intel Hillsboro #1 | 2024-01-15 | StrokeXMin | -203751
| Intel Hillsboro #1 | 2024-01-15 | StrokeXMax | 203691
| Intel Hillsboro #1 | 2024-01-15 | StrokeX | 407442
| Intel Hillsboro #1 | 2024-01-15 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer) | 0.03
| Intel Hillsboro #1 | 2024-01-15 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer) | 0.18
| Intel Hillsboro #1 | 2024-01-15 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.6
| Intel Hillsboro #1 | 2024-01-15 | Roll motion to X+ full scan | 48.2
| Intel Hillsboro #1 | 2024-01-15 | RMS of ZeroScan | 0.056
| Intel Hillsboro #1 | 2024-01-15 | Profiler noise @200um | 0.156
| Intel Hillsboro #1 | 2024-01-15 | Position repeatability(10repeats, Position Based) | 0.38
| Intel Hillsboro #1 | 2024-01-15 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8e-05
| Intel Hillsboro #1 | 2024-01-15 | OpticalAxisShiftPerFocusY | 0.0069
| Intel Hillsboro #1 | 2024-01-15 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| Intel Hillsboro #1 | 2024-01-15 | OpticalAxisShiftPerFocusX | -0.0032
| Intel Hillsboro #1 | 2024-01-15 | OPM Repeatability @5mm | 0.97
| Intel Hillsboro #1 | 2024-01-15 | OPM Repeatability @25mm | 8.87
| Intel Hillsboro #1 | 2024-01-15 | OPM Repeatability @1mm | 0.73
| Intel Hillsboro #1 | 2024-01-15 | OPM Repeatability @10mm | 3.15
| Intel Hillsboro #1 | 2024-01-15 | NCM driving% with AC160TS cantilever | 16.54
| Intel Hillsboro #1 | 2024-01-15 | Min. Vibration Transmissbility (~100Hz) | 0.061
| Intel Hillsboro #1 | 2024-01-15 | Max OPM @5mm | 32.8
| Intel Hillsboro #1 | 2024-01-15 | Max OPM @25mm | 158.2
| Intel Hillsboro #1 | 2024-01-15 | Max OPM @1mm | 7.3
| Intel Hillsboro #1 | 2024-01-15 | Max OPM @10mm | 61.6
| Intel Hillsboro #1 | 2024-01-15 | LowerLimit | 5045
| Intel Hillsboro #1 | 2024-01-15 | Lift pins synchronization offset check (Y offset) | 34.2
| Intel Hillsboro #1 | 2024-01-15 | Lift pins synchronization offset check (X offset) | 49.2
| Intel Hillsboro #1 | 2024-01-15 | Lateral(C-D) sensitivity | 10.6
| Intel Hillsboro #1 | 2024-01-15 | Lateral(C-D) noise | 0.214
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from -Y to +Y (Y scan) | 108.39
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from -Y to +Y (X scan) | 43.94
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from -X to +X (Y scan) | 24.16
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from -X to +X (X scan) | 24.9
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from +Y to -Y (Y scan) | 109.86
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from +Y to -Y (X scan) | 25.63
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from +X to -X (Y scan) | 46.14
| Intel Hillsboro #1 | 2024-01-15 | Image offset after moving from +X to -X (X scan) | 24.9
| Intel Hillsboro #1 | 2024-01-15 | Full range of Z Height during Z Scan sweep | 41.8
| Intel Hillsboro #1 | 2024-01-15 | Full range of Z Detector during Z Scan sweep | 14.2
| Intel Hillsboro #1 | 2024-01-15 | Full range of Y2 Detector Raw during Y Scan sweep | 14.9
| Intel Hillsboro #1 | 2024-01-15 | Full range of Y2 Detector during Y Scan sweep | 187.6
| Intel Hillsboro #1 | 2024-01-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| Intel Hillsboro #1 | 2024-01-15 | Full range of Y1 Detector during Y Scan sweep | 184.2
| Intel Hillsboro #1 | 2024-01-15 | Full range of X2 Detector Raw during X Scan sweep | 14.8
| Intel Hillsboro #1 | 2024-01-15 | Full range of X2 Detector during X Scan sweep | 199.2
| Intel Hillsboro #1 | 2024-01-15 | Full range of X1 Detector Raw during X Scan sweep | 14.8
| Intel Hillsboro #1 | 2024-01-15 | Full range of X1 Detector during X Scan sweep | 200.6
| Intel Hillsboro #1 | 2024-01-15 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Intel Hillsboro #1 | 2024-01-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Intel Hillsboro #1 | 2024-01-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.41
| Intel Hillsboro #1 | 2024-01-15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.28
| Intel Hillsboro #1 | 2024-01-15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Intel Hillsboro #1 | 2024-01-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Intel Hillsboro #1 | 2024-01-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.21
| Intel Hillsboro #1 | 2024-01-15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| Intel Hillsboro #1 | 2024-01-15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Intel Hillsboro #1 | 2024-01-15 | 940nm Step Height Sample (10 Repeats) | 949.98
| Intel Hillsboro #1 | 2024-01-15 | 50~200nm Step Height Sample (10 Repeats) | 178.37
| Intel Hillsboro #1 | 2024-01-15 | 200~500nm Step Height Sample (10 Repeats) | 458.68
| Intel Hillsboro #1 | 2024-01-15 | * 1800nm Step Height Sample (10 Repeats) | 1792.81
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Z Scanner position at -3μm | 0.0024
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Z Scanner position at +3μm | 0.0151
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Z Linearity by Z Scanner position
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Z Linearity by 3 different sample height | 1
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Z Height noise | 0.03
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Y Scan Stroke | 55
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Y Detector noise (XY Servo on) | 0.37
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Y Detector noise (XY Servo off) | 0.37
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Y axis OPM | 0.7
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | X Scan Stroke | 55
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | X Detector noise (XY Servo on) | 0.37
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | X Detector noise (XY Servo off) | 0.36
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | X axis OPM | 0.47
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | NCM driving% with AC160TS cantilever | 0.15
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | NCM Amplitude noise | 9.17
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.27
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.953
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.2
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Lateral(C-D) sensitivity | 17.6
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Lateral(C-D) noise | 0.052
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Full range of Z Height during Z Scan sweep | 18.1
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Full range of Z Detector during Z Scan sweep | 13
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Full range of Y Detector during Y Scan sweep | 73.4
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Full range of X Detector Raw during X Scan sweep | 14.1
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | Full range of X Detector during X Scan sweep | 69.8
| 한밭대학교 반도체소재부품장비센터 | 2023-12-29 | A-B noise | 7.26
| 전북대학교 물리학과 | 2023-12-28 | Z Scanner position at -3μm | 0
| 전북대학교 물리학과 | 2023-12-28 | Z Scanner position at +3μm | 0.0168
| 전북대학교 물리학과 | 2023-12-28 | Z Linearity by Z Scanner position
| QST | 2023-12-28 | Z linearity by sample height | 1.001
| 전북대학교 물리학과 | 2023-12-28 | Z Linearity by 3 different sample height | 0.991
| 전북대학교 물리학과 | 2023-12-28 | Z Height noise | 0.026
| QST | 2023-12-28 | Z Height noise | 0.029
| 전북대학교 물리학과 | 2023-12-28 | Y Scan Stroke | 55
| QST | 2023-12-28 | Y Scan Stroke | 55
| QST | 2023-12-28 | Y Detector noise (XY Servo on) | 0.25
| 전북대학교 물리학과 | 2023-12-28 | Y Detector noise (XY Servo on) | 0.27
| QST | 2023-12-28 | Y Detector noise (XY Servo off) | 0.25
| 전북대학교 물리학과 | 2023-12-28 | Y Detector noise (XY Servo off) | 0.26
| 전북대학교 물리학과 | 2023-12-28 | Y axis OPM | 0.25
| QST | 2023-12-28 | Y axis OPM | 1.18
| 전북대학교 물리학과 | 2023-12-28 | X Scan Stroke | 55
| QST | 2023-12-28 | X Scan Stroke | 55
| QST | 2023-12-28 | X Detector noise (XY Servo on) | 0.25
| 전북대학교 물리학과 | 2023-12-28 | X Detector noise (XY Servo on) | 0.27
| QST | 2023-12-28 | X Detector noise (XY Servo off) | 0.25
| 전북대학교 물리학과 | 2023-12-28 | X Detector noise (XY Servo off) | 0.26
| 전북대학교 물리학과 | 2023-12-28 | X axis OPM | 0.89
| QST | 2023-12-28 | X axis OPM | 0.7
| 전북대학교 물리학과 | 2023-12-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| QST | 2023-12-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 전북대학교 물리학과 | 2023-12-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| QST | 2023-12-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| 전북대학교 물리학과 | 2023-12-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| QST | 2023-12-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| 전북대학교 물리학과 | 2023-12-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| QST | 2023-12-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.12
| 전북대학교 물리학과 | 2023-12-28 | NCM driving% with AC160TS cantilever | 0.78
| QST | 2023-12-28 | NCM driving% with AC160TS cantilever | 0.63
| 전북대학교 물리학과 | 2023-12-28 | NCM Amplitude noise | 13.97
| QST | 2023-12-28 | NCM Amplitude noise | 15.61
| 전북대학교 물리학과 | 2023-12-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.65
| QST | 2023-12-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.86
| 전북대학교 물리학과 | 2023-12-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.358
| QST | 2023-12-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.028
| 전북대학교 물리학과 | 2023-12-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105
| QST | 2023-12-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.3
| 전북대학교 물리학과 | 2023-12-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.2
| QST | 2023-12-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.6
| QST | 2023-12-28 | Maximum Y Scan range (Full Stroke) | 73.4
| QST | 2023-12-28 | Maximum X Scan range (Full Stroke) | 70
| QST | 2023-12-28 | Full range of Z Height during Z Scan sweep | 17.1
| 전북대학교 물리학과 | 2023-12-28 | Lateral(C-D) sensitivity | 7.2
| QST | 2023-12-28 | Lateral(C-D) sensitivity | 14.7
| 전북대학교 물리학과 | 2023-12-28 | Lateral(C-D) noise | 0.08
| QST | 2023-12-28 | Lateral(C-D) noise | 0.067
| 전북대학교 물리학과 | 2023-12-28 | Full range of Z Height during Z Scan sweep | 16.2
| QST | 2023-12-28 | Full range of Z Detector raw | 13.7
| 전북대학교 물리학과 | 2023-12-28 | Full range of Z Detector during Z Scan sweep | 13.9
| 전북대학교 물리학과 | 2023-12-28 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| QST | 2023-12-28 | Full range of Y Detector raw | 14.6
| 전북대학교 물리학과 | 2023-12-28 | Full range of Y Detector during Y Scan sweep | 72.6
| 전북대학교 물리학과 | 2023-12-28 | Full range of X Detector Raw during X Scan sweep | 13.9
| QST | 2023-12-28 | Full range of X Detector raw | 14.2
| 전북대학교 물리학과 | 2023-12-28 | Full range of X Detector during X Scan sweep | 71
| QST | 2023-12-28 | Approach position at -3μm | 0.0033
| QST | 2023-12-28 | Approach position at +3μm | 0.0135
| 전북대학교 물리학과 | 2023-12-28 | A-B noise | 7.03
| QST | 2023-12-28 | A-B noise | 8.34
| 한국폴리텍대학교_인천 | 2023-12-27 | Z Scanner position at -3μm | 0.0035
| Missouri State University(Rishi Patel) | 2023-12-27 | Z Scanner position at -3μm | 0.0028
| 한국폴리텍대학교_인천 | 2023-12-27 | Z Scanner position at +3μm | 0.0096
| Missouri State University(Rishi Patel) | 2023-12-27 | Z Scanner position at +3μm | 0.0118
| 한국폴리텍대학교_인천 | 2023-12-27 | Z Linearity by Z Scanner position
| Missouri State University(Rishi Patel) | 2023-12-27 | Z Linearity by Z Scanner position
| 한국폴리텍대학교_인천 | 2023-12-27 | Z Linearity by 3 different sample height | 0.998
| Missouri State University(Rishi Patel) | 2023-12-27 | Z Linearity by 3 different sample height | 1.011
| 한국폴리텍대학교_인천 | 2023-12-27 | Z Height noise | 0.033
| Missouri State University(Rishi Patel) | 2023-12-27 | Z Height noise | 0.028
| Missouri State University(Rishi Patel) | 2023-12-27 | Y2 Detector noise (XY Servo on) | 0.37
| Missouri State University(Rishi Patel) | 2023-12-27 | Y2 Detector noise (XY Servo off) | 0.21
| Missouri State University(Rishi Patel) | 2023-12-27 | Y1 Detector noise (XY Servo on) | 0.36
| Missouri State University(Rishi Patel) | 2023-12-27 | Y1 Detector noise (XY Servo off) | 0.21
| 한국폴리텍대학교_인천 | 2023-12-27 | Y Scan Stroke | 55
| Missouri State University(Rishi Patel) | 2023-12-27 | Y Scan Stroke | 171
| Missouri State University(Rishi Patel) | 2023-12-27 | -Y Direction software Limit | -103098
| 한국폴리텍대학교_인천 | 2023-12-27 | Y Detector noise (XY Servo on) | 0.24
| Missouri State University(Rishi Patel) | 2023-12-27 | Y Detector noise (XY Servo on) | 0.29
| 한국폴리텍대학교_인천 | 2023-12-27 | Y Detector noise (XY Servo off) | 0.23
| Missouri State University(Rishi Patel) | 2023-12-27 | Y Detector noise (XY Servo off) | 0.15
| 한국폴리텍대학교_인천 | 2023-12-27 | Y axis OPM | 1.18
| Missouri State University(Rishi Patel) | 2023-12-27 | Y axis OPM | 1.34
| Missouri State University(Rishi Patel) | 2023-12-27 | X2 Detector noise (XY Servo on) | 0.36
| Missouri State University(Rishi Patel) | 2023-12-27 | X2 Detector noise (XY Servo off) | 0.22
| Missouri State University(Rishi Patel) | 2023-12-27 | X1 Detector noise (XY Servo on) | 0.39
| Missouri State University(Rishi Patel) | 2023-12-27 | X1 Detector noise (XY Servo off) | 0.25
| 한국폴리텍대학교_인천 | 2023-12-27 | X Scan Stroke | 55
| Missouri State University(Rishi Patel) | 2023-12-27 | X Scan Stroke | 171.2
| Missouri State University(Rishi Patel) | 2023-12-27 | -X Direction software Limit | -102147
| 한국폴리텍대학교_인천 | 2023-12-27 | X Detector noise (XY Servo on) | 0.24
| Missouri State University(Rishi Patel) | 2023-12-27 | X Detector noise (XY Servo on) | 0.3
| 한국폴리텍대학교_인천 | 2023-12-27 | X Detector noise (XY Servo off) | 0.24
| Missouri State University(Rishi Patel) | 2023-12-27 | X Detector noise (XY Servo off) | 0.17
| 한국폴리텍대학교_인천 | 2023-12-27 | X axis OPM | 1.31
| Missouri State University(Rishi Patel) | 2023-12-27 | X axis OPM | 1.33
| 한국폴리텍대학교_인천 | 2023-12-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| 한국폴리텍대학교_인천 | 2023-12-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.32
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.3
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| 한국폴리텍대학교_인천 | 2023-12-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| 한국폴리텍대학교_인천 | 2023-12-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| Missouri State University(Rishi Patel) | 2023-12-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 한국폴리텍대학교_인천 | 2023-12-27 | NCM driving% with AC160TS cantilever | 2.18
| Missouri State University(Rishi Patel) | 2023-12-27 | NCM driving% with AC160TS cantilever | 0.88
| 한국폴리텍대학교_인천 | 2023-12-27 | NCM Amplitude noise | 13.25
| Missouri State University(Rishi Patel) | 2023-12-27 | NCM Amplitude noise | 14.76
| 한국폴리텍대학교_인천 | 2023-12-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.993
| Missouri State University(Rishi Patel) | 2023-12-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.951
| 한국폴리텍대학교_인천 | 2023-12-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.348
| Missouri State University(Rishi Patel) | 2023-12-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.297
| Missouri State University(Rishi Patel) | 2023-12-27 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 238.6
| Missouri State University(Rishi Patel) | 2023-12-27 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 242.9
| 한국폴리텍대학교_인천 | 2023-12-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.1
| Missouri State University(Rishi Patel) | 2023-12-27 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 240.8
| Missouri State University(Rishi Patel) | 2023-12-27 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 248.3
| 한국폴리텍대학교_인천 | 2023-12-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100.9
| Missouri State University(Rishi Patel) | 2023-12-27 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| 한국폴리텍대학교_인천 | 2023-12-27 | Lateral(C-D) sensitivity | 13.667
| Missouri State University(Rishi Patel) | 2023-12-27 | Lateral(C-D) sensitivity | 12.5
| 한국폴리텍대학교_인천 | 2023-12-27 | Lateral(C-D) noise | 0.054
| Missouri State University(Rishi Patel) | 2023-12-27 | Lateral(C-D) noise | 0.112
| 한국폴리텍대학교_인천 | 2023-12-27 | Full range of Z Height during Z Scan sweep | 17.1
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of Z Height during Z Scan sweep | 17.6
| 한국폴리텍대학교_인천 | 2023-12-27 | Full range of Z Detector during Z Scan sweep | 11.6
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of Z Detector during Z Scan sweep | 13.4
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| 한국폴리텍대학교_인천 | 2023-12-27 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| 한국폴리텍대학교_인천 | 2023-12-27 | Full range of Y Detector during Y Scan sweep | 72.1
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of Y Detector during Y Scan sweep | 174.3
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| 한국폴리텍대학교_인천 | 2023-12-27 | Full range of X Detector Raw during X Scan sweep | 14.1
| 한국폴리텍대학교_인천 | 2023-12-27 | Full range of X Detector during X Scan sweep | 71
| Missouri State University(Rishi Patel) | 2023-12-27 | Full range of X Detector during X Scan sweep | 173.2
| 한국폴리텍대학교_인천 | 2023-12-27 | A-B noise | 7.15
| Missouri State University(Rishi Patel) | 2023-12-27 | A-B noise | 8.4
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Z Scanner position at -3μm | 0.0006
| IIPE Visakhapatnam | 2023-12-22 | Z Scanner position at -3μm | 0.0034
| PowerEpi Semiconductor | 2023-12-22 | Z Scanner position at -3μm | 0.007
| Hon Young Semiconductor | 2023-12-22 | Z Scanner position at -3μm | 0.0059
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Z Scanner position at +3μm | 0.0042
| IIPE Visakhapatnam | 2023-12-22 | Z Scanner position at +3μm | 0.0116
| PowerEpi Semiconductor | 2023-12-22 | Z Scanner position at +3μm | 0
| Hon Young Semiconductor | 2023-12-22 | Z Scanner position at +3μm | 0.0011
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Z Linearity by Z Scanner position
| IIPE Visakhapatnam | 2023-12-22 | Z Linearity by Z Scanner position
| PowerEpi Semiconductor | 2023-12-22 | Z Linearity by Z Scanner position
| Hon Young Semiconductor | 2023-12-22 | Z Linearity by Z Scanner position
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Z Linearity by 3 different sample height | 1.004
| IIPE Visakhapatnam | 2023-12-22 | Z Linearity by 3 different sample height | 1.012
| PowerEpi Semiconductor | 2023-12-22 | Z Linearity by 3 different sample height | 0.985
| Hon Young Semiconductor | 2023-12-22 | Z Linearity by 3 different sample height | 0.994
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Z Height noise | 0.024
| IIPE Visakhapatnam | 2023-12-22 | Z Height noise | 0.027
| PowerEpi Semiconductor | 2023-12-22 | Z Height noise | 0.029
| Hon Young Semiconductor | 2023-12-22 | Z Height noise | 0.029
| IIPE Visakhapatnam | 2023-12-22 | Y2 Detector noise (XY Servo on) | 0.33
| PowerEpi Semiconductor | 2023-12-22 | Y2 Detector noise (XY Servo on) | 0.23
| IIPE Visakhapatnam | 2023-12-22 | Y2 Detector noise (XY Servo off) | 0.32
| PowerEpi Semiconductor | 2023-12-22 | Y2 Detector noise (XY Servo off) | 0.24
| IIPE Visakhapatnam | 2023-12-22 | Y1 Detector noise (XY Servo on) | 0.23
| PowerEpi Semiconductor | 2023-12-22 | Y1 Detector noise (XY Servo on) | 0.22
| IIPE Visakhapatnam | 2023-12-22 | Y1 Detector noise (XY Servo off) | 0.23
| PowerEpi Semiconductor | 2023-12-22 | Y1 Detector noise (XY Servo off) | 0.21
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Y Scan Stroke | 55
| IIPE Visakhapatnam | 2023-12-22 | Y Scan Stroke | 178.6
| PowerEpi Semiconductor | 2023-12-22 | Y Scan Stroke | 177.509
| Hon Young Semiconductor | 2023-12-22 | Y Scan Stroke | 134.3
| IIPE Visakhapatnam | 2023-12-22 | -Y Direction software Limit | -75130
| PowerEpi Semiconductor | 2023-12-22 | -Y Direction software Limit | -103022
| Hon Young Semiconductor | 2023-12-22 | -Y Direction software Limit | -100363
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Y Detector noise (XY Servo on) | 0.28
| IIPE Visakhapatnam | 2023-12-22 | Y Detector noise (XY Servo on) | 0.22
| PowerEpi Semiconductor | 2023-12-22 | Y Detector noise (XY Servo on) | 0.16
| Hon Young Semiconductor | 2023-12-22 | Y Detector noise (XY Servo on) | 0.27
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Y Detector noise (XY Servo off) | 0.26
| IIPE Visakhapatnam | 2023-12-22 | Y Detector noise (XY Servo off) | 0.2
| PowerEpi Semiconductor | 2023-12-22 | Y Detector noise (XY Servo off) | 0.16
| Hon Young Semiconductor | 2023-12-22 | Y Detector noise (XY Servo off) | 0.26
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Y axis OPM | 0.83
| IIPE Visakhapatnam | 2023-12-22 | Y axis OPM | 1.36
| PowerEpi Semiconductor | 2023-12-22 | Y axis OPM | 1.35
| Hon Young Semiconductor | 2023-12-22 | Y axis OPM | 0.98
| IIPE Visakhapatnam | 2023-12-22 | X2 Detector noise (XY Servo on) | 0.23
| PowerEpi Semiconductor | 2023-12-22 | X2 Detector noise (XY Servo on) | 0.22
| IIPE Visakhapatnam | 2023-12-22 | X2 Detector noise (XY Servo off) | 0.22
| PowerEpi Semiconductor | 2023-12-22 | X2 Detector noise (XY Servo off) | 0.22
| IIPE Visakhapatnam | 2023-12-22 | X1 Detector noise (XY Servo on) | 0.23
| PowerEpi Semiconductor | 2023-12-22 | X1 Detector noise (XY Servo on) | 0.22
| IIPE Visakhapatnam | 2023-12-22 | X1 Detector noise (XY Servo off) | 0.22
| PowerEpi Semiconductor | 2023-12-22 | X1 Detector noise (XY Servo off) | 0.22
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | X Scan Stroke | 55
| IIPE Visakhapatnam | 2023-12-22 | X Scan Stroke | 171.6
| PowerEpi Semiconductor | 2023-12-22 | X Scan Stroke | 169.1
| Hon Young Semiconductor | 2023-12-22 | X Scan Stroke | 135.5
| IIPE Visakhapatnam | 2023-12-22 | -X Direction software Limit | -75355
| PowerEpi Semiconductor | 2023-12-22 | -X Direction software Limit | -101862
| Hon Young Semiconductor | 2023-12-22 | -X Direction software Limit | -100449
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | X Detector noise (XY Servo on) | 0.28
| IIPE Visakhapatnam | 2023-12-22 | X Detector noise (XY Servo on) | 0.17
| PowerEpi Semiconductor | 2023-12-22 | X Detector noise (XY Servo on) | 0.16
| Hon Young Semiconductor | 2023-12-22 | X Detector noise (XY Servo on) | 0.28
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | X Detector noise (XY Servo off) | 0.26
| IIPE Visakhapatnam | 2023-12-22 | X Detector noise (XY Servo off) | 0.16
| PowerEpi Semiconductor | 2023-12-22 | X Detector noise (XY Servo off) | 0.16
| Hon Young Semiconductor | 2023-12-22 | X Detector noise (XY Servo off) | 0.27
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | X axis OPM | 1.68
| IIPE Visakhapatnam | 2023-12-22 | X axis OPM | 0.45
| PowerEpi Semiconductor | 2023-12-22 | X axis OPM | 0.81
| Hon Young Semiconductor | 2023-12-22 | X axis OPM | 1.12
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.27
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.32
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.29
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.32
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.3
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.27
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.27
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.34
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.29
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.32
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| IIPE Visakhapatnam | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| PowerEpi Semiconductor | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Hon Young Semiconductor | 2023-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | NCM driving% with AC160TS cantilever | 0.61
| IIPE Visakhapatnam | 2023-12-22 | NCM driving% with AC160TS cantilever | 0.49
| PowerEpi Semiconductor | 2023-12-22 | NCM driving% with AC160TS cantilever | 1.36
| Hon Young Semiconductor | 2023-12-22 | NCM driving% with AC160TS cantilever | 3.26
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | NCM Amplitude noise | 12.8
| IIPE Visakhapatnam | 2023-12-22 | NCM Amplitude noise | 14.35
| PowerEpi Semiconductor | 2023-12-22 | NCM Amplitude noise | 12.39
| Hon Young Semiconductor | 2023-12-22 | NCM Amplitude noise | 12.66
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.442
| IIPE Visakhapatnam | 2023-12-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.033
| PowerEpi Semiconductor | 2023-12-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.49
| Hon Young Semiconductor | 2023-12-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.234
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.151
| IIPE Visakhapatnam | 2023-12-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.711
| PowerEpi Semiconductor | 2023-12-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.073
| Hon Young Semiconductor | 2023-12-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.874
| IIPE Visakhapatnam | 2023-12-22 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 258.5
| PowerEpi Semiconductor | 2023-12-22 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 253.3
| IIPE Visakhapatnam | 2023-12-22 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 255.1
| PowerEpi Semiconductor | 2023-12-22 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.3
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.4
| Hon Young Semiconductor | 2023-12-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 185.2
| IIPE Visakhapatnam | 2023-12-22 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 248.8
| PowerEpi Semiconductor | 2023-12-22 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 244.3
| IIPE Visakhapatnam | 2023-12-22 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 243
| PowerEpi Semiconductor | 2023-12-22 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 237.2
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99.5
| Hon Young Semiconductor | 2023-12-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 181.9
| IIPE Visakhapatnam | 2023-12-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| PowerEpi Semiconductor | 2023-12-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Hon Young Semiconductor | 2023-12-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Lateral(C-D) sensitivity | 11.2
| IIPE Visakhapatnam | 2023-12-22 | Lateral(C-D) sensitivity | 7.4
| PowerEpi Semiconductor | 2023-12-22 | Lateral(C-D) sensitivity | 7.8
| Hon Young Semiconductor | 2023-12-22 | Lateral(C-D) sensitivity | 11.2
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Lateral(C-D) noise | 0.111
| IIPE Visakhapatnam | 2023-12-22 | Lateral(C-D) noise | 0.128
| PowerEpi Semiconductor | 2023-12-22 | Lateral(C-D) noise | 0.088
| Hon Young Semiconductor | 2023-12-22 | Lateral(C-D) noise | 0.131
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Full range of Z Height during Z Scan sweep | 16.5
| IIPE Visakhapatnam | 2023-12-22 | Full range of Z Height during Z Scan sweep | 17.5
| PowerEpi Semiconductor | 2023-12-22 | Full range of Z Height during Z Scan sweep | 18.5
| Hon Young Semiconductor | 2023-12-22 | Full range of Z Height during Z Scan sweep | 18.1
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Full range of Z Detector during Z Scan sweep | 13.7
| IIPE Visakhapatnam | 2023-12-22 | Full range of Z Detector during Z Scan sweep | 14.2
| PowerEpi Semiconductor | 2023-12-22 | Full range of Z Detector during Z Scan sweep | 13.6
| Hon Young Semiconductor | 2023-12-22 | Full range of Z Detector during Z Scan sweep | 12.5
| IIPE Visakhapatnam | 2023-12-22 | Full range of Y2 Detector Raw during Y Scan sweep | 13.8
| PowerEpi Semiconductor | 2023-12-22 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| IIPE Visakhapatnam | 2023-12-22 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| PowerEpi Semiconductor | 2023-12-22 | Full range of Y1 Detector Raw during Y Scan sweep | 14.9
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Full range of Y Detector Raw during Y Scan sweep | 15
| Hon Young Semiconductor | 2023-12-22 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Full range of Y Detector during Y Scan sweep | 74.4
| IIPE Visakhapatnam | 2023-12-22 | Full range of Y Detector during Y Scan sweep | 178.6
| PowerEpi Semiconductor | 2023-12-22 | Full range of Y Detector during Y Scan sweep | 182.3
| Hon Young Semiconductor | 2023-12-22 | Full range of Y Detector during Y Scan sweep | 134.6
| IIPE Visakhapatnam | 2023-12-22 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| PowerEpi Semiconductor | 2023-12-22 | Full range of X2 Detector Raw during X Scan sweep | 15
| IIPE Visakhapatnam | 2023-12-22 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| PowerEpi Semiconductor | 2023-12-22 | Full range of X1 Detector Raw during X Scan sweep | 14.9
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Full range of X Detector Raw during X Scan sweep | 14.1
| Hon Young Semiconductor | 2023-12-22 | Full range of X Detector Raw during X Scan sweep | 14.9
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | Full range of X Detector during X Scan sweep | 70.3
| IIPE Visakhapatnam | 2023-12-22 | Full range of X Detector during X Scan sweep | 171.6
| PowerEpi Semiconductor | 2023-12-22 | Full range of X Detector during X Scan sweep | 179.8
| Hon Young Semiconductor | 2023-12-22 | Full range of X Detector during X Scan sweep | 135.4
| Institute of Industrial Science, the Univ. of Tokyo | 2023-12-22 | A-B noise | 7.96
| IIPE Visakhapatnam | 2023-12-22 | A-B noise | 14.42
| PowerEpi Semiconductor | 2023-12-22 | A-B noise | 7.37
| Hon Young Semiconductor | 2023-12-22 | A-B noise | 7.45
| Google(Melissa Archer) | 2023-12-20 | Z Linearity by Z scanner position | 0.54
| Google(Melissa Archer) | 2023-12-20 | Z Linearity by 3 different sample height | 1.006
| Google(Melissa Archer) | 2023-12-20 | Z Height noise | 0.034
| Google(Melissa Archer) | 2023-12-20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.6
| Google(Melissa Archer) | 2023-12-20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.6
| Google(Melissa Archer) | 2023-12-20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4
| Google(Melissa Archer) | 2023-12-20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3
| Google(Melissa Archer) | 2023-12-20 | Z Drift (Short Traveling) | 0.25
| Google(Melissa Archer) | 2023-12-20 | Y2 Detector noise (XY Servo on) | 0.35
| Google(Melissa Archer) | 2023-12-20 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 261.6
| Google(Melissa Archer) | 2023-12-20 | Y1 Detector noise (XY Servo on) | 0.33
| Google(Melissa Archer) | 2023-12-20 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 250.5
| Google(Melissa Archer) | 2023-12-20 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.7
| Google(Melissa Archer) | 2023-12-20 | Y Ringing(Rpv) | 0.12
| Google(Melissa Archer) | 2023-12-20 | Y Non-linearity error | -1.85
| Google(Melissa Archer) | 2023-12-20 | Y Detector noise (XY Servo on) | 0.25
| Google(Melissa Archer) | 2023-12-20 | Y axis, Repeatability @1st position | 0.04
| Google(Melissa Archer) | 2023-12-20 | Y axis, Repeatability @ 4th position | 0.04
| Google(Melissa Archer) | 2023-12-20 | Y axis, Repeatability @ 3rd position | 0.02
| Google(Melissa Archer) | 2023-12-20 | Y axis, Repeatability @ 2nd position | 0.03
| Google(Melissa Archer) | 2023-12-20 | Y axis, Accuracy @1st position | 0.94
| Google(Melissa Archer) | 2023-12-20 | Y axis, Accuracy @ 4th position | 0.87
| Google(Melissa Archer) | 2023-12-20 | Y axis, Accuracy @ 3rd position | 0.73
| Google(Melissa Archer) | 2023-12-20 | Y axis, Accuracy @ 2nd position | 0.74
| Google(Melissa Archer) | 2023-12-20 | Y axis OPM | 0.75
| Google(Melissa Archer) | 2023-12-20 | XY Orthogonality | 90
| Google(Melissa Archer) | 2023-12-20 | X2 Detector noise (XY Servo on) | 0.27
| Google(Melissa Archer) | 2023-12-20 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 260.1
| Google(Melissa Archer) | 2023-12-20 | X1 Detector noise (XY Servo on) | 0.27
| Google(Melissa Archer) | 2023-12-20 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 256
| Google(Melissa Archer) | 2023-12-20 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.7
| Google(Melissa Archer) | 2023-12-20 | X Ringing(Rpv) | 0.16
| Google(Melissa Archer) | 2023-12-20 | X Non-linearity error | -2.17
| Google(Melissa Archer) | 2023-12-20 | X Detector noise (XY Servo on) | 0.2
| Google(Melissa Archer) | 2023-12-20 | X axis, Repeatability @1st position | 0.03
| Google(Melissa Archer) | 2023-12-20 | X axis, Repeatability @ 4th position | 0.04
| Google(Melissa Archer) | 2023-12-20 | X axis, Repeatability @ 3rd position | 0.03
| Google(Melissa Archer) | 2023-12-20 | X axis, Repeatability @ 2nd position | 0.04
| Google(Melissa Archer) | 2023-12-20 | X axis, Accuracy @1st position | 0.21
| Google(Melissa Archer) | 2023-12-20 | X axis, Accuracy @ 4th position | 0.19
| Google(Melissa Archer) | 2023-12-20 | X axis, Accuracy @ 3rd position | 0.18
| Google(Melissa Archer) | 2023-12-20 | X axis, Accuracy @ 2nd position | 0.34
| Google(Melissa Archer) | 2023-12-20 | X axis OPM | 0.76
| Google(Melissa Archer) | 2023-12-20 | Wafer placement(Loading) error on the load port1 (Y offset) | -26.4
| Google(Melissa Archer) | 2023-12-20 | Wafer placement(Loading) error on the load port1 (X offset) | 24.8
| Google(Melissa Archer) | 2023-12-20 | Torque Profile (10mm, 20um/s) | 2.382
| Google(Melissa Archer) | 2023-12-20 | Torque Profile (10mm, 200um/s) | 2.005
| Google(Melissa Archer) | 2023-12-20 | StrokeYMin | -109541
| Google(Melissa Archer) | 2023-12-20 | StrokeYMax | 103162.4
| Google(Melissa Archer) | 2023-12-20 | StrokeY | 212703.5
| Google(Melissa Archer) | 2023-12-20 | StrokeXMin | -140827
| Google(Melissa Archer) | 2023-12-20 | StrokeXMax | 137662.5
| Google(Melissa Archer) | 2023-12-20 | StrokeX | 278489.7
| Google(Melissa Archer) | 2023-12-20 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.92
| Google(Melissa Archer) | 2023-12-20 | Roll motion to X+ full scan | 8
| Google(Melissa Archer) | 2023-12-20 | RMS of ZeroScan | 0.031
| Google(Melissa Archer) | 2023-12-20 | Profiler noise @200um | 0.18
| Google(Melissa Archer) | 2023-12-20 | Position repeatability(10repeats, Position Based) | 0.36
| Google(Melissa Archer) | 2023-12-20 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8e-05
| Google(Melissa Archer) | 2023-12-20 | OpticalAxisShiftPerFocusY | -0.0011
| Google(Melissa Archer) | 2023-12-20 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00019
| Google(Melissa Archer) | 2023-12-20 | OpticalAxisShiftPerFocusX | -0.0038
| Google(Melissa Archer) | 2023-12-20 | OPM Repeatability @5mm | 0.36
| Google(Melissa Archer) | 2023-12-20 | OPM Repeatability @1mm | 0.47
| Google(Melissa Archer) | 2023-12-20 | OPM Repeatability @10mm | 0.67
| Google(Melissa Archer) | 2023-12-20 | NCM driving% with AC160TS cantilever | 7.7
| Google(Melissa Archer) | 2023-12-20 | Min. Vibration Transmissbility (~100Hz) | 0.039
| Google(Melissa Archer) | 2023-12-20 | Max OPM @5mm | 31.11
| Google(Melissa Archer) | 2023-12-20 | Max OPM @1mm | 4.76
| Google(Melissa Archer) | 2023-12-20 | Max OPM @10mm | 76.42
| Google(Melissa Archer) | 2023-12-20 | LowerLimit | 4002
| Google(Melissa Archer) | 2023-12-20 | Lateral(C-D) sensitivity | 9
| Google(Melissa Archer) | 2023-12-20 | Lateral(C-D) noise | 0.15
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from -Y to +Y (Y scan) | 59.33
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from -Y to +Y (X scan) | 33.69
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from -X to +X (Y scan) | 32.23
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from -X to +X (X scan) | 43.95
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from +Y to -Y (Y scan) | 101.81
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from +Y to -Y (X scan) | 65.92
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from +X to -X (Y scan) | 126.71
| Google(Melissa Archer) | 2023-12-20 | Image offset after moving from +X to -X (X scan) | 68.85
| Google(Melissa Archer) | 2023-12-20 | Full range of Z Height during Z Scan sweep | 16.5
| Google(Melissa Archer) | 2023-12-20 | Full range of Z Detector during Z Scan sweep | 11.6
| Google(Melissa Archer) | 2023-12-20 | Full range of Y2 Detector Raw during Y Scan sweep | 13.86
| Google(Melissa Archer) | 2023-12-20 | Full range of Y2 Detector during Y Scan sweep | 190
| Google(Melissa Archer) | 2023-12-20 | Full range of Y1 Detector Raw during Y Scan sweep | 14.13
| Google(Melissa Archer) | 2023-12-20 | Full range of Y1 Detector during Y Scan sweep | 190
| Google(Melissa Archer) | 2023-12-20 | Full range of X2 Detector Raw during X Scan sweep | 14.01
| Google(Melissa Archer) | 2023-12-20 | Full range of X2 Detector during X Scan sweep | 188.5
| Google(Melissa Archer) | 2023-12-20 | Full range of X1 Detector Raw during X Scan sweep | 14.07
| Google(Melissa Archer) | 2023-12-20 | Full range of X1 Detector during X Scan sweep | 198.9
| Google(Melissa Archer) | 2023-12-20 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Google(Melissa Archer) | 2023-12-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Google(Melissa Archer) | 2023-12-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.46
| Google(Melissa Archer) | 2023-12-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.48
| Google(Melissa Archer) | 2023-12-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Google(Melissa Archer) | 2023-12-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Google(Melissa Archer) | 2023-12-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| Google(Melissa Archer) | 2023-12-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.28
| Google(Melissa Archer) | 2023-12-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Google(Melissa Archer) | 2023-12-20 | 940nm Step Height Sample (10 Repeats) | 948.18
| Google(Melissa Archer) | 2023-12-20 | 50~200nm Step Height Sample (10 Repeats) | 176.29
| Google(Melissa Archer) | 2023-12-20 | 200~500nm Step Height Sample (10 Repeats) | 452.75
| Google(Melissa Archer) | 2023-12-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.9
| Google(Melissa Archer) | 2023-12-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 1.9
| Google(Melissa Archer) | 2023-12-20 | %GRR w/ roughness of non-patterened wafer | 6.2
| National Taiwan Normal University | 2023-12-20 | Z Scanner position at -3μm | 0.0012
| China University of Petroleum | 2023-12-20 | Z Scanner position at -3μm | 0.0065
| National Taiwan Normal University | 2023-12-20 | Z Scanner position at +3μm | 0.0086
| China University of Petroleum | 2023-12-20 | Z Scanner position at +3μm | 0.0121
| China University of Petroleum | 2023-12-20 | Z Linearity by Z Scanner position
| National Taiwan Normal University | 2023-12-20 | Z Linearity by 3 different sample height | 1.009
| China University of Petroleum | 2023-12-20 | Z Linearity by 3 different sample height | 0.996
| National Taiwan Normal University | 2023-12-20 | Z Height noise | 0.027
| China University of Petroleum | 2023-12-20 | Z Height noise | 0.029
| National Taiwan Normal University | 2023-12-20 | Y Detector noise (XY Servo on) | 0.26
| China University of Petroleum | 2023-12-20 | Y Detector noise (XY Servo on) | 0.27
| National Taiwan Normal University | 2023-12-20 | Y Detector noise (XY Servo off) | 0.26
| China University of Petroleum | 2023-12-20 | Y Detector noise (XY Servo off) | 0.26
| National Taiwan Normal University | 2023-12-20 | Y axis OPM | 0.96
| China University of Petroleum | 2023-12-20 | Y axis OPM | 0.46
| National Taiwan Normal University | 2023-12-20 | X Detector noise (XY Servo on) | 0.26
| China University of Petroleum | 2023-12-20 | X Detector noise (XY Servo on) | 0.27
| National Taiwan Normal University | 2023-12-20 | X Detector noise (XY Servo off) | 0.26
| China University of Petroleum | 2023-12-20 | X Detector noise (XY Servo off) | 0.26
| National Taiwan Normal University | 2023-12-20 | X axis OPM | 0.75
| China University of Petroleum | 2023-12-20 | X axis OPM | 0.83
| National Taiwan Normal University | 2023-12-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| China University of Petroleum | 2023-12-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| National Taiwan Normal University | 2023-12-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| China University of Petroleum | 2023-12-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| National Taiwan Normal University | 2023-12-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| China University of Petroleum | 2023-12-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| National Taiwan Normal University | 2023-12-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| China University of Petroleum | 2023-12-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| National Taiwan Normal University | 2023-12-20 | NCM driving% with AC160TS cantilever | 2.04
| China University of Petroleum | 2023-12-20 | NCM driving% with AC160TS cantilever | 1.76
| National Taiwan Normal University | 2023-12-20 | NCM Amplitude noise | 8.6
| China University of Petroleum | 2023-12-20 | NCM Amplitude noise | 13.33
| National Taiwan Normal University | 2023-12-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.535
| China University of Petroleum | 2023-12-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.648
| National Taiwan Normal University | 2023-12-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.455
| China University of Petroleum | 2023-12-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.857
| National Taiwan Normal University | 2023-12-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.4
| China University of Petroleum | 2023-12-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 169.7
| National Taiwan Normal University | 2023-12-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 166.4
| China University of Petroleum | 2023-12-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 164.3
| National Taiwan Normal University | 2023-12-20 | Lateral(C-D) sensitivity | 14.5
| China University of Petroleum | 2023-12-20 | Lateral(C-D) sensitivity | 14.4
| National Taiwan Normal University | 2023-12-20 | Lateral(C-D) noise | 0.111
| China University of Petroleum | 2023-12-20 | Lateral(C-D) noise | 0.113
| National Taiwan Normal University | 2023-12-20 | Full range of Z Height during Z Scan sweep | 16.3
| China University of Petroleum | 2023-12-20 | Full range of Z Height during Z Scan sweep | 16.7
| National Taiwan Normal University | 2023-12-20 | Full range of Z Detector during Z Scan sweep | 12.3
| China University of Petroleum | 2023-12-20 | Full range of Z Detector during Z Scan sweep | 12
| National Taiwan Normal University | 2023-12-20 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| China University of Petroleum | 2023-12-20 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| National Taiwan Normal University | 2023-12-20 | Full range of Y Detector during Y Scan sweep | 114.9
| China University of Petroleum | 2023-12-20 | Full range of Y Detector during Y Scan sweep | 116.9
| National Taiwan Normal University | 2023-12-20 | Full range of X Detector Raw during X Scan sweep | 13.8
| China University of Petroleum | 2023-12-20 | Full range of X Detector Raw during X Scan sweep | 14.5
| National Taiwan Normal University | 2023-12-20 | Full range of X Detector during X Scan sweep | 114.7
| China University of Petroleum | 2023-12-20 | Full range of X Detector during X Scan sweep | 119
| National Taiwan Normal University | 2023-12-20 | A-B noise | 13.99
| China University of Petroleum | 2023-12-20 | A-B noise | 7.64
| Chulalongkorn University | 2023-12-19 | Z Scanner position at -3μm | 0.0026
| Chulalongkorn University | 2023-12-19 | Z Scanner position at +3μm | 0.0097
| Chulalongkorn University | 2023-12-19 | Z Linearity by Z Scanner position
| Chulalongkorn University | 2023-12-19 | Z Linearity by 3 different sample height | 1.008
| Chulalongkorn University | 2023-12-19 | Z Height noise | 0.028
| Chulalongkorn University | 2023-12-19 | Y Scan Stroke | 55
| Chulalongkorn University | 2023-12-19 | Y Detector noise (XY Servo on) | 0.29
| Chulalongkorn University | 2023-12-19 | Y Detector noise (XY Servo off) | 0.26
| Chulalongkorn University | 2023-12-19 | Y axis OPM | 0.4
| Chulalongkorn University | 2023-12-19 | X Scan Stroke | 55
| Chulalongkorn University | 2023-12-19 | X Detector noise (XY Servo on) | 0.28
| Chulalongkorn University | 2023-12-19 | X Detector noise (XY Servo off) | 0.26
| Chulalongkorn University | 2023-12-19 | X axis OPM | 0.72
| Chulalongkorn University | 2023-12-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Chulalongkorn University | 2023-12-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Chulalongkorn University | 2023-12-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| Chulalongkorn University | 2023-12-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| Chulalongkorn University | 2023-12-19 | NCM driving% with AC160TS cantilever | 0.45
| Chulalongkorn University | 2023-12-19 | NCM Amplitude noise | 13.46
| Chulalongkorn University | 2023-12-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.644
| Chulalongkorn University | 2023-12-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 27
| Chulalongkorn University | 2023-12-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 104.6
| Chulalongkorn University | 2023-12-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.1
| Chulalongkorn University | 2023-12-19 | Lateral(C-D) sensitivity | 7.8
| Chulalongkorn University | 2023-12-19 | Lateral(C-D) noise | 0.059
| Chulalongkorn University | 2023-12-19 | Full range of Z Height during Z Scan sweep | 17.5
| Chulalongkorn University | 2023-12-19 | Full range of Z Detector during Z Scan sweep | 12.9
| Chulalongkorn University | 2023-12-19 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Chulalongkorn University | 2023-12-19 | Full range of Y Detector during Y Scan sweep | 71.9
| Chulalongkorn University | 2023-12-19 | Full range of X Detector Raw during X Scan sweep | 13.9
| Chulalongkorn University | 2023-12-19 | Full range of X Detector during X Scan sweep | 72.5
| Chulalongkorn University | 2023-12-19 | A-B noise | 8.22
| XFAB France #2 | 2023-12-15 | Z Linearity by Z scanner position | 0.53
| XFAB France #2 | 2023-12-15 | Z Linearity by Z scanner position | 1.07
| XFAB France #2 | 2023-12-15 | Z Linearity by 3 different sample height | 1
| XFAB France #2 | 2023-12-15 | Z Height noise | 0.025
| XFAB France #2 | 2023-12-15 | Z Height noise | 0.036
| XFAB France #2 | 2023-12-15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.6
| XFAB France #2 | 2023-12-15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.1
| XFAB France #2 | 2023-12-15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.6
| XFAB France #2 | 2023-12-15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.1
| XFAB France #2 | 2023-12-15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11
| XFAB France #2 | 2023-12-15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2
| XFAB France #2 | 2023-12-15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11
| XFAB France #2 | 2023-12-15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2
| XFAB France #2 | 2023-12-15 | Z Drift (Short Traveling) | 0.13
| XFAB France #2 | 2023-12-15 | Y2 Detector noise (XY Servo on) | 0.36
| XFAB France #2 | 2023-12-15 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 254.7
| XFAB France #2 | 2023-12-15 | Y1 Detector noise (XY Servo on) | 0.29
| XFAB France #2 | 2023-12-15 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 247.3
| XFAB France #2 | 2023-12-15 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 151.4
| XFAB France #2 | 2023-12-15 | Y Ringing(Rpv) | 0.14
| XFAB France #2 | 2023-12-15 | Y Non-linearity error | 2.36
| XFAB France #2 | 2023-12-15 | Y Detector noise (XY Servo on) | 0.23
| XFAB France #2 | 2023-12-15 | Y axis, Repeatability @1st position | 0.03
| XFAB France #2 | 2023-12-15 | Y axis, Repeatability @ 4th position | 0.08
| XFAB France #2 | 2023-12-15 | Y axis, Repeatability @ 3rd position | 0.03
| XFAB France #2 | 2023-12-15 | Y axis, Repeatability @ 2nd position | 0.08
| XFAB France #2 | 2023-12-15 | Y axis, Accuracy @1st position | 0.34
| XFAB France #2 | 2023-12-15 | Y axis, Accuracy @ 4th position | 0.43
| XFAB France #2 | 2023-12-15 | Y axis, Accuracy @ 3rd position | 0.32
| XFAB France #2 | 2023-12-15 | Y axis, Accuracy @ 2nd position | 0.51
| XFAB France #2 | 2023-12-15 | Y axis OPM | 1.24
| XFAB France #2 | 2023-12-15 | XY Orthogonality | 89.9
| XFAB France #2 | 2023-12-15 | X2 Detector noise (XY Servo on) | 0.27
| XFAB France #2 | 2023-12-15 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 246.4
| XFAB France #2 | 2023-12-15 | X1 Detector noise (XY Servo on) | 0.24
| XFAB France #2 | 2023-12-15 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 234.5
| XFAB France #2 | 2023-12-15 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 156.1
| XFAB France #2 | 2023-12-15 | X Ringing(Rpv) | 0.14
| XFAB France #2 | 2023-12-15 | X Non-linearity error | 2.37
| XFAB France #2 | 2023-12-15 | X Detector noise (XY Servo on) | 0.19
| XFAB France #2 | 2023-12-15 | X axis, Repeatability @1st position | 0.05
| XFAB France #2 | 2023-12-15 | X axis, Repeatability @ 4th position | 0.03
| XFAB France #2 | 2023-12-15 | X axis, Repeatability @ 3rd position | 0.03
| XFAB France #2 | 2023-12-15 | X axis, Repeatability @ 2nd position | 0.04
| XFAB France #2 | 2023-12-15 | X axis, Accuracy @1st position | 0.24
| XFAB France #2 | 2023-12-15 | X axis, Accuracy @ 4th position | 0.16
| XFAB France #2 | 2023-12-15 | X axis, Accuracy @ 3rd position | 0.18
| XFAB France #2 | 2023-12-15 | X axis, Accuracy @ 2nd position | 0.23
| XFAB France #2 | 2023-12-15 | X axis OPM | 0.84
| XFAB France #2 | 2023-12-15 | Wafer placement(Loading) error on the load port2 (Y offset) | 33.71
| XFAB France #2 | 2023-12-15 | Wafer placement(Loading) error on the load port2 (X offset) | 68.15
| XFAB France #2 | 2023-12-15 | Wafer placement(Loading) error on the load port1 (Y offset) | 27.81
| XFAB France #2 | 2023-12-15 | Wafer placement(Loading) error on the load port1 (X offset) | 49.24
| XFAB France #2 | 2023-12-15 | Under 50nm Step Height Sample (10 Repeats) | 14.818
| XFAB France #2 | 2023-12-15 | Torque Profile (50mm, 20um/s) | 2.623
| XFAB France #2 | 2023-12-15 | Torque Profile (50mm, 200um/s) | 2.048
| XFAB France #2 | 2023-12-15 | StrokeYMin | -102312
| XFAB France #2 | 2023-12-15 | StrokeYMax | 104424
| XFAB France #2 | 2023-12-15 | StrokeY | 206736
| XFAB France #2 | 2023-12-15 | StrokeXMin | -152909
| XFAB France #2 | 2023-12-15 | StrokeXMax | 152341
| XFAB France #2 | 2023-12-15 | StrokeX | 305250
| XFAB France #2 | 2023-12-15 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.4
| XFAB France #2 | 2023-12-15 | Roll motion to X+ full scan | 19.6
| XFAB France #2 | 2023-12-15 | RMS of ZeroScan | 0.033
| XFAB France #2 | 2023-12-15 | Profiler noise @200um | 0.135
| XFAB France #2 | 2023-12-15 | Position repeatability(10repeats, Position Based) | 0.14
| XFAB France #2 | 2023-12-15 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| XFAB France #2 | 2023-12-15 | OpticalAxisShiftPerFocusY | -0.0007
| XFAB France #2 | 2023-12-15 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00023
| XFAB France #2 | 2023-12-15 | OpticalAxisShiftPerFocusX | -0.0002
| XFAB France #2 | 2023-12-15 | OPM Repeatability @5mm | 0.51
| XFAB France #2 | 2023-12-15 | OPM Repeatability @25mm | 2.18
| XFAB France #2 | 2023-12-15 | OPM Repeatability @1mm | 0.25
| XFAB France #2 | 2023-12-15 | OPM Repeatability @10mm | 0.36
| XFAB France #2 | 2023-12-15 | NCM driving% with AC160TS cantilever | 10.83
| XFAB France #2 | 2023-12-15 | NCM driving% with AC160TS cantilever | 31.04
| XFAB France #2 | 2023-12-15 | Min. Vibration Transmissbility (~100Hz) | 0.024
| XFAB France #2 | 2023-12-15 | Max OPM @5mm | 9.92
| XFAB France #2 | 2023-12-15 | Max OPM @25mm | 36.21
| XFAB France #2 | 2023-12-15 | Max OPM @1mm | 4.94
| XFAB France #2 | 2023-12-15 | Max OPM @10mm | 8.64
| XFAB France #2 | 2023-12-15 | LowerLimit | 4300
| XFAB France #2 | 2023-12-15 | Lift pins synchronization offset check (Y offset) | -7.2
| XFAB France #2 | 2023-12-15 | Lift pins synchronization offset check (X offset) | 34.24
| XFAB France #2 | 2023-12-15 | Lateral(C-D) sensitivity | 4.9
| XFAB France #2 | 2023-12-15 | Lateral(C-D) noise | 0.117
| XFAB France #2 | 2023-12-15 | Image offset after moving from -Y to +Y (Y scan) | 133.3
| XFAB France #2 | 2023-12-15 | Image offset after moving from -Y to +Y (X scan) | 33.69
| XFAB France #2 | 2023-12-15 | Image offset after moving from -X to +X (Y scan) | 20.51
| XFAB France #2 | 2023-12-15 | Image offset after moving from -X to +X (X scan) | 13.18
| XFAB France #2 | 2023-12-15 | Image offset after moving from +Y to -Y (Y scan) | 154.54
| XFAB France #2 | 2023-12-15 | Image offset after moving from +Y to -Y (X scan) | 8.79
| XFAB France #2 | 2023-12-15 | Image offset after moving from +X to -X (Y scan) | 121.58
| XFAB France #2 | 2023-12-15 | Image offset after moving from +X to -X (X scan) | 52.73
| XFAB France #2 | 2023-12-15 | Full range of Z Height during Z Scan sweep | 15.9
| XFAB France #2 | 2023-12-15 | Full range of Z Height during Z Scan sweep | 17.2
| XFAB France #2 | 2023-12-15 | Full range of Z Detector during Z Scan sweep | 12.3
| XFAB France #2 | 2023-12-15 | Full range of Z Detector during Z Scan sweep | 13.4
| XFAB France #2 | 2023-12-15 | Full range of Y2 Detector Raw during Y Scan sweep | 13.87
| XFAB France #2 | 2023-12-15 | Full range of Y2 Detector during Y Scan sweep | 186.4
| XFAB France #2 | 2023-12-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14.02
| XFAB France #2 | 2023-12-15 | Full range of Y1 Detector during Y Scan sweep | 186.6
| XFAB France #2 | 2023-12-15 | Full range of X2 Detector Raw during X Scan sweep | 13.99
| XFAB France #2 | 2023-12-15 | Full range of X2 Detector during X Scan sweep | 177.6
| XFAB France #2 | 2023-12-15 | Full range of X1 Detector Raw during X Scan sweep | 14.02
| XFAB France #2 | 2023-12-15 | Full range of X1 Detector during X Scan sweep | 187.4
| XFAB France #2 | 2023-12-15 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| XFAB France #2 | 2023-12-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| XFAB France #2 | 2023-12-15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.6
| XFAB France #2 | 2023-12-15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.46
| XFAB France #2 | 2023-12-15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| XFAB France #2 | 2023-12-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| XFAB France #2 | 2023-12-15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.35
| XFAB France #2 | 2023-12-15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.33
| XFAB France #2 | 2023-12-15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| XFAB France #2 | 2023-12-15 | 940nm Step Height Sample (10 Repeats) | 945.78
| XFAB France #2 | 2023-12-15 | 940nm Step Height Sample (10 Repeats) | 945.18
| XFAB France #2 | 2023-12-15 | 50~200nm Step Height Sample (10 Repeats) | 184.58
| XFAB France #2 | 2023-12-15 | 50~200nm Step Height Sample (10 Repeats) | 184.25
| XFAB France #2 | 2023-12-15 | 200~500nm Step Height Sample (10 Repeats) | 451.75
| XFAB France #2 | 2023-12-15 | 200~500nm Step Height Sample (10 Repeats) | 451.58
| XFAB France #2 | 2023-12-15 | * 1800nm Step Height Sample (10 Repeats) | 1779.54
| XFAB France #2 | 2023-12-15 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.5
| XFAB France #2 | 2023-12-15 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.1
| XFAB France #2 | 2023-12-15 | %GRR w/ roughness of non-patterened wafer | 6.1
| Shandong University | 2023-12-15 | Z Scanner position at -3μm | 0.0074
| 한양대학교 나노광전자학과 | 2023-12-15 | Z Scanner position at -3μm | 0.0027
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Z Scanner position at -3μm | 0.001
| Shandong University | 2023-12-15 | Z Scanner position at +3μm | 0.0019
| 한양대학교 나노광전자학과 | 2023-12-15 | Z Scanner position at +3μm | 0.0004
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Z Scanner position at +3μm | 0.0087
| Shandong University | 2023-12-15 | Z Linearity by Z Scanner position
| 한양대학교 나노광전자학과 | 2023-12-15 | Z Linearity by Z Scanner position
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Z Linearity by Z Scanner position
| Shandong University | 2023-12-15 | Z Linearity by 3 different sample height | 0.999
| 한양대학교 나노광전자학과 | 2023-12-15 | Z Linearity by 3 different sample height | 0.999
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Z Linearity by 3 different sample height | 1
| Shandong University | 2023-12-15 | Z Height noise | 0.027
| 한양대학교 나노광전자학과 | 2023-12-15 | Z Height noise | 0.029
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Z Height noise | 0.027
| Shandong University | 2023-12-15 | Y2 Detector noise (XY Servo on) | 0.22
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y2 Detector noise (XY Servo on) | 0.23
| Shandong University | 2023-12-15 | Y2 Detector noise (XY Servo off) | 0.22
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y2 Detector noise (XY Servo off) | 0.24
| Shandong University | 2023-12-15 | Y1 Detector noise (XY Servo on) | 0.31
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y1 Detector noise (XY Servo on) | 0.24
| Shandong University | 2023-12-15 | Y1 Detector noise (XY Servo off) | 0.3
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y1 Detector noise (XY Servo off) | 0.25
| Shandong University | 2023-12-15 | Y Scan Stroke | 174
| 한양대학교 나노광전자학과 | 2023-12-15 | Y Scan Stroke | 55
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y Scan Stroke | 178
| Shandong University | 2023-12-15 | -Y Direction software Limit | -102588
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | -Y Direction software Limit | -75310
| Shandong University | 2023-12-15 | Y Detector noise (XY Servo on) | 0.19
| 한양대학교 나노광전자학과 | 2023-12-15 | Y Detector noise (XY Servo on) | 0.25
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y Detector noise (XY Servo on) | 0.17
| Shandong University | 2023-12-15 | Y Detector noise (XY Servo off) | 0.19
| 한양대학교 나노광전자학과 | 2023-12-15 | Y Detector noise (XY Servo off) | 0.24
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y Detector noise (XY Servo off) | 0.18
| Shandong University | 2023-12-15 | Y axis OPM | 0.95
| 한양대학교 나노광전자학과 | 2023-12-15 | Y axis OPM | 0.77
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Y axis OPM | 1.19
| Shandong University | 2023-12-15 | X2 Detector noise (XY Servo on) | 0.23
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X2 Detector noise (XY Servo on) | 0.26
| Shandong University | 2023-12-15 | X2 Detector noise (XY Servo off) | 0.22
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X2 Detector noise (XY Servo off) | 0.27
| Shandong University | 2023-12-15 | X1 Detector noise (XY Servo on) | 0.23
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X1 Detector noise (XY Servo on) | 0.23
| Shandong University | 2023-12-15 | X1 Detector noise (XY Servo off) | 0.23
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X1 Detector noise (XY Servo off) | 0.24
| Shandong University | 2023-12-15 | X Scan Stroke | 178.7
| 한양대학교 나노광전자학과 | 2023-12-15 | X Scan Stroke | 55
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X Scan Stroke | 172.3
| Shandong University | 2023-12-15 | -X Direction software Limit | -101879
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | -X Direction software Limit | -75365.6
| Shandong University | 2023-12-15 | X Detector noise (XY Servo on) | 0.17
| 한양대학교 나노광전자학과 | 2023-12-15 | X Detector noise (XY Servo on) | 0.26
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X Detector noise (XY Servo on) | 0.18
| Shandong University | 2023-12-15 | X Detector noise (XY Servo off) | 0.17
| 한양대학교 나노광전자학과 | 2023-12-15 | X Detector noise (XY Servo off) | 0.25
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X Detector noise (XY Servo off) | 0.19
| Shandong University | 2023-12-15 | X axis OPM | 1.1
| 한양대학교 나노광전자학과 | 2023-12-15 | X axis OPM | 0.54
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | X axis OPM | 1.29
| Shandong University | 2023-12-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| 한양대학교 나노광전자학과 | 2023-12-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Shandong University | 2023-12-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 한양대학교 나노광전자학과 | 2023-12-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Shandong University | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Shandong University | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.36
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.34
| Shandong University | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| Shandong University | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.34
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.34
| Shandong University | 2023-12-15 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.34
| Shandong University | 2023-12-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한양대학교 나노광전자학과 | 2023-12-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Shandong University | 2023-12-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 한양대학교 나노광전자학과 | 2023-12-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Shandong University | 2023-12-15 | NCM driving% with AC160TS cantilever | 1.62
| 한양대학교 나노광전자학과 | 2023-12-15 | NCM driving% with AC160TS cantilever | 1.87
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | NCM driving% with AC160TS cantilever | 1.75
| Shandong University | 2023-12-15 | NCM Amplitude noise | 11.92
| 한양대학교 나노광전자학과 | 2023-12-15 | NCM Amplitude noise | 12.38
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | NCM Amplitude noise | 14.69
| Shandong University | 2023-12-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.09
| 한양대학교 나노광전자학과 | 2023-12-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.195
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.852
| Shandong University | 2023-12-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.488
| 한양대학교 나노광전자학과 | 2023-12-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.825
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.837
| Shandong University | 2023-12-15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 234.8
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 261.8
| Shandong University | 2023-12-15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 241.9
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 258
| 한양대학교 나노광전자학과 | 2023-12-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.2
| Shandong University | 2023-12-15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 244.4
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 258.9
| Shandong University | 2023-12-15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 248.7
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 244.2
| 한양대학교 나노광전자학과 | 2023-12-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 108.5
| Shandong University | 2023-12-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Shandong University | 2023-12-15 | Lateral(C-D) sensitivity | 8.5
| 한양대학교 나노광전자학과 | 2023-12-15 | Lateral(C-D) sensitivity | 4.6
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Lateral(C-D) sensitivity | 6.2
| Shandong University | 2023-12-15 | Lateral(C-D) noise | 0.106
| 한양대학교 나노광전자학과 | 2023-12-15 | Lateral(C-D) noise | 0.067
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Lateral(C-D) noise | 0.137
| Shandong University | 2023-12-15 | Full range of Z Height during Z Scan sweep | 15.6
| 한양대학교 나노광전자학과 | 2023-12-15 | Full range of Z Height during Z Scan sweep | 17.9
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of Z Height during Z Scan sweep | 17.1
| Shandong University | 2023-12-15 | Full range of Z Detector during Z Scan sweep | 12.2
| 한양대학교 나노광전자학과 | 2023-12-15 | Full range of Z Detector during Z Scan sweep | 14.4
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of Z Detector during Z Scan sweep | 13.3
| Shandong University | 2023-12-15 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of Y2 Detector Raw during Y Scan sweep | 13.7
| Shandong University | 2023-12-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of Y1 Detector Raw during Y Scan sweep | 13.6
| 한양대학교 나노광전자학과 | 2023-12-15 | Full range of Y Detector Raw during Y Scan sweep | 15
| Shandong University | 2023-12-15 | Full range of Y Detector during Y Scan sweep | 173.7
| 한양대학교 나노광전자학과 | 2023-12-15 | Full range of Y Detector during Y Scan sweep | 75.9
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of Y Detector during Y Scan sweep | 177.2
| Shandong University | 2023-12-15 | Full range of X2 Detector Raw during X Scan sweep | 14.6
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of X2 Detector Raw during X Scan sweep | 13.6
| Shandong University | 2023-12-15 | Full range of X1 Detector Raw during X Scan sweep | 14.6
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of X1 Detector Raw during X Scan sweep | 13.8
| 한양대학교 나노광전자학과 | 2023-12-15 | Full range of X Detector Raw during X Scan sweep | 14
| Shandong University | 2023-12-15 | Full range of X Detector during X Scan sweep | 179.5
| 한양대학교 나노광전자학과 | 2023-12-15 | Full range of X Detector during X Scan sweep | 75.7
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | Full range of X Detector during X Scan sweep | 172.3
| Shandong University | 2023-12-15 | A-B noise | 16.77
| 한양대학교 나노광전자학과 | 2023-12-15 | A-B noise | 8.32
| Suzhou Loongspeed Semiconductor Technology | 2023-12-15 | A-B noise | 8.23
| Tokyo Electron Kyushu | 2023-12-14 | Z Linearity by Z scanner position | 0.45
| Tokyo Electron Kyushu | 2023-12-14 | Z Linearity by 4 different sample height | 1.005
| Tokyo Electron Kyushu | 2023-12-14 | Z Height noise | 0.025
| Tokyo Electron Kyushu | 2023-12-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.2
| Tokyo Electron Kyushu | 2023-12-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.3
| Tokyo Electron Kyushu | 2023-12-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1
| Tokyo Electron Kyushu | 2023-12-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.1
| Tokyo Electron Kyushu | 2023-12-14 | Z Drift (Short Traveling) | 0.23
| Tokyo Electron Kyushu | 2023-12-14 | Y2 Detector noise (XY Servo on) | 0.34
| Tokyo Electron Kyushu | 2023-12-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 256.4
| Tokyo Electron Kyushu | 2023-12-14 | Y1 Detector noise (XY Servo on) | 0.33
| Tokyo Electron Kyushu | 2023-12-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 255.9
| Tokyo Electron Kyushu | 2023-12-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.4
| Tokyo Electron Kyushu | 2023-12-14 | Y Ringing(Rpv) | 0.17
| Tokyo Electron Kyushu | 2023-12-14 | Y Non-linearity error | 3.62
| Tokyo Electron Kyushu | 2023-12-14 | Y Detector noise (XY Servo on) | 0.28
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Repeatability @1st position | 0.03
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Repeatability @ 4th position | 0.03
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Repeatability @ 3rd position | 0.02
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Repeatability @ 2nd position | 0.04
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Accuracy @1st position | 0.31
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Accuracy @ 4th position | 0.29
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Accuracy @ 3rd position | 0.27
| Tokyo Electron Kyushu | 2023-12-14 | Y axis, Accuracy @ 2nd position | 0.29
| Tokyo Electron Kyushu | 2023-12-14 | Y axis OPM | 0.59
| Tokyo Electron Kyushu | 2023-12-14 | XY Orthogonality | 90.1
| Tokyo Electron Kyushu | 2023-12-14 | X2 Detector noise (XY Servo on) | 0.36
| Tokyo Electron Kyushu | 2023-12-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 259.9
| Tokyo Electron Kyushu | 2023-12-14 | X1 Detector noise (XY Servo on) | 0.39
| Tokyo Electron Kyushu | 2023-12-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.5
| Tokyo Electron Kyushu | 2023-12-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 163.1
| Tokyo Electron Kyushu | 2023-12-14 | X Ringing(Rpv) | 0.08
| Tokyo Electron Kyushu | 2023-12-14 | X Non-linearity error | 3.01
| Tokyo Electron Kyushu | 2023-12-14 | X Detector noise (XY Servo on) | 0.32
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Repeatability @1st position | 0.05
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Repeatability @ 4th position | 0.03
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Repeatability @ 3rd position | 0.05
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Repeatability @ 2nd position | 0.03
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Accuracy @1st position | 0.26
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Accuracy @ 4th position | 0.17
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Accuracy @ 3rd position | 0.26
| Tokyo Electron Kyushu | 2023-12-14 | X axis, Accuracy @ 2nd position | 0.19
| Tokyo Electron Kyushu | 2023-12-14 | X axis OPM | 0.41
| Tokyo Electron Kyushu | 2023-12-14 | Wafer placement(Loading) error on the load port2 (Y offset) | 21.36
| Tokyo Electron Kyushu | 2023-12-14 | Wafer placement(Loading) error on the load port2 (X offset) | 20.85
| Tokyo Electron Kyushu | 2023-12-14 | Wafer placement(Loading) error on the load port1 (Y offset) | 22.06
| Tokyo Electron Kyushu | 2023-12-14 | Wafer placement(Loading) error on the load port1 (X offset) | 8.85
| Tokyo Electron Kyushu | 2023-12-14 | Under 50nm Step Height Sample (10 Repeats) | 17.92
| Tokyo Electron Kyushu | 2023-12-14 | StrokeYMin | -153453
| Tokyo Electron Kyushu | 2023-12-14 | StrokeYMax | 150740.2
| Tokyo Electron Kyushu | 2023-12-14 | StrokeY | 304193.3
| Tokyo Electron Kyushu | 2023-12-14 | StrokeXMin | -204051
| Tokyo Electron Kyushu | 2023-12-14 | StrokeXMax | 203739.6
| Tokyo Electron Kyushu | 2023-12-14 | StrokeX | 407790.3
| Tokyo Electron Kyushu | 2023-12-14 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7
| Tokyo Electron Kyushu | 2023-12-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00014
| Tokyo Electron Kyushu | 2023-12-14 | OpticalAxisShiftPerFocusY | -0.0029
| Tokyo Electron Kyushu | 2023-12-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00039
| Tokyo Electron Kyushu | 2023-12-14 | OpticalAxisShiftPerFocusX | -0.0042
| Tokyo Electron Kyushu | 2023-12-14 | NCM driving% with AC160TS cantilever | 42.74
| Tokyo Electron Kyushu | 2023-12-14 | LowerLimit | 4774
| Tokyo Electron Kyushu | 2023-12-14 | Lift pins synchronization offset check (Y offset) | 1.34
| Tokyo Electron Kyushu | 2023-12-14 | Lift pins synchronization offset check (X offset) | 5.32
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from -Y to +Y (Y scan) | 60.06
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from -Y to +Y (X scan) | 27.1
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from -X to +X (Y scan) | 177.98
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from -X to +X (X scan) | 6.59
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from +Y to -Y (Y scan) | 86.43
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from +Y to -Y (X scan) | 13.92
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from +X to -X (Y scan) | 62.99
| Tokyo Electron Kyushu | 2023-12-14 | Image offset after moving from +X to -X (X scan) | 17.58
| Tokyo Electron Kyushu | 2023-12-14 | Full range of Z Height during Z Scan sweep | 16.1
| Tokyo Electron Kyushu | 2023-12-14 | Full range of Z Detector during Z Scan sweep | 12.3
| Tokyo Electron Kyushu | 2023-12-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.21
| Tokyo Electron Kyushu | 2023-12-14 | Full range of Y2 Detector during Y Scan sweep | 204.1
| Tokyo Electron Kyushu | 2023-12-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.45
| Tokyo Electron Kyushu | 2023-12-14 | Full range of Y1 Detector during Y Scan sweep | 199.5
| Tokyo Electron Kyushu | 2023-12-14 | Full range of X2 Detector Raw during X Scan sweep | 14.16
| Tokyo Electron Kyushu | 2023-12-14 | Full range of X2 Detector during X Scan sweep | 203.7
| Tokyo Electron Kyushu | 2023-12-14 | Full range of X1 Detector Raw during X Scan sweep | 14.25
| Tokyo Electron Kyushu | 2023-12-14 | Full range of X1 Detector during X Scan sweep | 199.5
| Tokyo Electron Kyushu | 2023-12-14 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Tokyo Electron Kyushu | 2023-12-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| Tokyo Electron Kyushu | 2023-12-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.15
| Tokyo Electron Kyushu | 2023-12-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| Tokyo Electron Kyushu | 2023-12-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| Tokyo Electron Kyushu | 2023-12-14 | 940nm Step Height Sample (10 Repeats) | 951.76
| Tokyo Electron Kyushu | 2023-12-14 | 50~200nm Step Height Sample (10 Repeats) | 182.48
| Tokyo Electron Kyushu | 2023-12-14 | 200~500nm Step Height Sample (10 Repeats) | 462.37
| Tokyo Electron Kyushu | 2023-12-14 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 3.9
| Tokyo Electron Kyushu | 2023-12-14 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.5
| Tokyo Electron Kyushu | 2023-12-14 | %GRR w/ roughness of non-patterened wafer | 5.9
| 서울대학교 전기정보공학과 | 2023-12-14 | Z Scanner position at -3μm | 0.0034
| 서울대학교 전기정보공학과 | 2023-12-14 | Z Scanner position at +3μm | 0.0166
| 서울대학교 전기정보공학과 | 2023-12-14 | Z Linearity by Z Scanner position
| 서울대학교 전기정보공학과 | 2023-12-14 | Z Linearity by 3 different sample height | 0.996
| 서울대학교 전기정보공학과 | 2023-12-14 | Z Height noise | 0.025
| 서울대학교 전기정보공학과 | 2023-12-14 | Y Scan Stroke | 55
| 서울대학교 전기정보공학과 | 2023-12-14 | Y Detector noise (XY Servo on) | 0.26
| 서울대학교 전기정보공학과 | 2023-12-14 | Y Detector noise (XY Servo off) | 0.25
| 서울대학교 전기정보공학과 | 2023-12-14 | Y axis OPM | 0.68
| 서울대학교 전기정보공학과 | 2023-12-14 | X Scan Stroke | 55
| 서울대학교 전기정보공학과 | 2023-12-14 | X Detector noise (XY Servo on) | 0.26
| 서울대학교 전기정보공학과 | 2023-12-14 | X Detector noise (XY Servo off) | 0.25
| 서울대학교 전기정보공학과 | 2023-12-14 | X axis OPM | 0.75
| 서울대학교 전기정보공학과 | 2023-12-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 서울대학교 전기정보공학과 | 2023-12-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 서울대학교 전기정보공학과 | 2023-12-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| 서울대학교 전기정보공학과 | 2023-12-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 서울대학교 전기정보공학과 | 2023-12-14 | NCM driving% with AC160TS cantilever | 1.58
| 서울대학교 전기정보공학과 | 2023-12-14 | NCM Amplitude noise | 12.54
| 서울대학교 전기정보공학과 | 2023-12-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.699
| 서울대학교 전기정보공학과 | 2023-12-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.002
| 서울대학교 전기정보공학과 | 2023-12-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 104.2
| 서울대학교 전기정보공학과 | 2023-12-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.7
| 서울대학교 전기정보공학과 | 2023-12-14 | Lateral(C-D) sensitivity | 7.3
| 서울대학교 전기정보공학과 | 2023-12-14 | Lateral(C-D) noise | 0.074
| 서울대학교 전기정보공학과 | 2023-12-14 | Full range of Z Height during Z Scan sweep | 17
| 서울대학교 전기정보공학과 | 2023-12-14 | Full range of Z Detector during Z Scan sweep | 13.6
| 서울대학교 전기정보공학과 | 2023-12-14 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 서울대학교 전기정보공학과 | 2023-12-14 | Full range of Y Detector during Y Scan sweep | 72.5
| 서울대학교 전기정보공학과 | 2023-12-14 | Full range of X Detector Raw during X Scan sweep | 13.8
| 서울대학교 전기정보공학과 | 2023-12-14 | Full range of X Detector during X Scan sweep | 72.8
| 서울대학교 전기정보공학과 | 2023-12-14 | A-B noise | 7.88
| Harland Medical Systems | 2023-12-12 | Z Scanner position at -3μm | 0.0009
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Z Scanner position at -3μm | 0.0028
| Harland Medical Systems | 2023-12-12 | Z Scanner position at +3μm | 0.0053
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Z Scanner position at +3μm | 0.0108
| Harland Medical Systems | 2023-12-12 | Z Linearity by Z Scanner position
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Z Linearity by Z Scanner position
| Harland Medical Systems | 2023-12-12 | Z Linearity by 3 different sample height | 0.996
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Z Linearity by 3 different sample height | 0.996
| Harland Medical Systems | 2023-12-12 | Z Height noise | 0.031
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Z Height noise | 0.033
| Harland Medical Systems | 2023-12-12 | Y Scan Stroke | 55
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Y Scan Stroke | 55
| Harland Medical Systems | 2023-12-12 | Y Detector noise (XY Servo on) | 0.28
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Y Detector noise (XY Servo on) | 0.25
| Harland Medical Systems | 2023-12-12 | Y Detector noise (XY Servo off) | 0.27
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Y Detector noise (XY Servo off) | 0.23
| Harland Medical Systems | 2023-12-12 | Y axis OPM | 0.83
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Y axis OPM | 1.03
| Harland Medical Systems | 2023-12-12 | X Scan Stroke | 55
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | X Scan Stroke | 55
| Harland Medical Systems | 2023-12-12 | X Detector noise (XY Servo on) | 0.29
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | X Detector noise (XY Servo on) | 0.25
| Harland Medical Systems | 2023-12-12 | X Detector noise (XY Servo off) | 0.27
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | X Detector noise (XY Servo off) | 0.24
| Harland Medical Systems | 2023-12-12 | X axis OPM | 0.91
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | X axis OPM | 0.7
| Harland Medical Systems | 2023-12-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Harland Medical Systems | 2023-12-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Harland Medical Systems | 2023-12-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Harland Medical Systems | 2023-12-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Harland Medical Systems | 2023-12-12 | NCM driving% with AC160TS cantilever | 1.15
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | NCM driving% with AC160TS cantilever | 0.7
| Harland Medical Systems | 2023-12-12 | NCM Amplitude noise | 12.36
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | NCM Amplitude noise | 11.85
| Harland Medical Systems | 2023-12-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.398
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.858
| Harland Medical Systems | 2023-12-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.258
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.426
| Harland Medical Systems | 2023-12-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.3
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 104.5
| Harland Medical Systems | 2023-12-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.4
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 107.1
| Harland Medical Systems | 2023-12-12 | Lateral(C-D) sensitivity | 6.3
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Lateral(C-D) sensitivity | 5.9
| Harland Medical Systems | 2023-12-12 | Lateral(C-D) noise | 0.097
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Lateral(C-D) noise | 0.105
| Harland Medical Systems | 2023-12-12 | Full range of Z Height during Z Scan sweep | 16.3
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Full range of Z Height during Z Scan sweep | 17.1
| Harland Medical Systems | 2023-12-12 | Full range of Z Detector during Z Scan sweep | 11.6
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Full range of Z Detector during Z Scan sweep | 12.1
| Harland Medical Systems | 2023-12-12 | Full range of Y Detector Raw during Y Scan sweep | 14
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Harland Medical Systems | 2023-12-12 | Full range of Y Detector during Y Scan sweep | 72.5
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Full range of Y Detector during Y Scan sweep | 77
| Harland Medical Systems | 2023-12-12 | Full range of X Detector Raw during X Scan sweep | 14.1
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Full range of X Detector Raw during X Scan sweep | 14.5
| Harland Medical Systems | 2023-12-12 | Full range of X Detector during X Scan sweep | 72.7
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | Full range of X Detector during X Scan sweep | 77.7
| Harland Medical Systems | 2023-12-12 | A-B noise | 12.09
| Federal University of Juiz de Fora(Benjamin) | 2023-12-12 | A-B noise | 10.04
| Brock University(Kaur) | 2023-12-11 | Z Scanner position at -3μm | 0.0022
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Z Scanner position at -3μm | 0.0059
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Z Scanner position at -3μm | 0.0035
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Z Scanner position at -3μm | 0.0039
| Brock University(Kaur) | 2023-12-11 | Z Scanner position at +3μm | 0.0037
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Z Scanner position at +3μm | 0.0148
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Z Scanner position at +3μm | 0.0075
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Z Scanner position at +3μm | 0.0076
| Brock University(Kaur) | 2023-12-11 | Z Linearity by Z Scanner position
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Z Linearity by Z Scanner position
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Z Linearity by Z Scanner position
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Z Linearity by Z Scanner position
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Z Linearity by Z Scanner position
| Brock University(Kaur) | 2023-12-11 | Z Linearity by 3 different sample height | 0.987
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Z Linearity by 3 different sample height | 1.003
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Z Linearity by 3 different sample height | 0.996
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Z Linearity by 3 different sample height | 1.008
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Z Linearity by 3 different sample height | 1.002
| Brock University(Kaur) | 2023-12-11 | Z Height noise | 0.032
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Z Height noise | 0.021
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Z Height noise | 0.031
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Z Height noise | 0.032
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Z Height noise | 0.027
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y2 Detector noise (XY Servo on) | 0.26
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y2 Detector noise (XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y2 Detector noise (XY Servo on) | 0.25
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y2 Detector noise (XY Servo off) | 0.26
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y2 Detector noise (XY Servo off) | 0.24
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y2 Detector noise (XY Servo off) | 0.26
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y1 Detector noise (XY Servo on) | 0.29
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y1 Detector noise (XY Servo on) | 0.22
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y1 Detector noise (XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y1 Detector noise (XY Servo off) | 0.28
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y1 Detector noise (XY Servo off) | 0.22
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y1 Detector noise (XY Servo off) | 0.23
| Brock University(Kaur) | 2023-12-11 | Y Scan Stroke | 55
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y Scan Stroke | 176.1
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y Scan Stroke | 176.5
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y Scan Stroke | 178.6
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | -Y Direction software Limit | -102796
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | -Y Direction software Limit | -102516
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | -Y Direction software Limit | -102741
| Brock University(Kaur) | 2023-12-11 | Y Detector noise (XY Servo on) | 0.28
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Y Detector noise (XY Servo on) | 0.26
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y Detector noise (XY Servo on) | 0.2
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y Detector noise (XY Servo on) | 0.17
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y Detector noise (XY Servo on) | 0.18
| Brock University(Kaur) | 2023-12-11 | Y Detector noise (XY Servo off) | 0.28
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Y Detector noise (XY Servo off) | 0.26
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y Detector noise (XY Servo off) | 0.19
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y Detector noise (XY Servo off) | 0.17
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y Detector noise (XY Servo off) | 0.18
| Brock University(Kaur) | 2023-12-11 | Y axis OPM | 0.82
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Y axis OPM | 0.72
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Y axis OPM | 1.02
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Y axis OPM | 1.63
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Y axis OPM | 1.06
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X2 Detector noise (XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X2 Detector noise (XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X2 Detector noise (XY Servo on) | 0.25
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X2 Detector noise (XY Servo off) | 0.24
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X2 Detector noise (XY Servo off) | 0.23
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X2 Detector noise (XY Servo off) | 0.25
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X1 Detector noise (XY Servo on) | 0.25
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X1 Detector noise (XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X1 Detector noise (XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X1 Detector noise (XY Servo off) | 0.25
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X1 Detector noise (XY Servo off) | 0.23
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X1 Detector noise (XY Servo off) | 0.24
| Brock University(Kaur) | 2023-12-11 | X Scan Stroke | 55
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X Scan Stroke | 177.8
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X Scan Stroke | 178.3
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X Scan Stroke | 182
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | -X Direction software Limit | -101791
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | -X Direction software Limit | -101968
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | -X Direction software Limit | -102034
| Brock University(Kaur) | 2023-12-11 | X Detector noise (XY Servo on) | 0.29
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | X Detector noise (XY Servo on) | 0.25
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X Detector noise (XY Servo on) | 0.18
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X Detector noise (XY Servo on) | 0.17
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X Detector noise (XY Servo on) | 0.18
| Brock University(Kaur) | 2023-12-11 | X Detector noise (XY Servo off) | 0.28
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | X Detector noise (XY Servo off) | 0.25
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X Detector noise (XY Servo off) | 0.19
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X Detector noise (XY Servo off) | 0.17
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X Detector noise (XY Servo off) | 0.18
| Brock University(Kaur) | 2023-12-11 | X axis OPM | 0.89
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | X axis OPM | 1.32
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | X axis OPM | 0.89
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | X axis OPM | 1.55
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | X axis OPM | 0.9
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | SLD beam power at the end of the lens | 0.22
| Brock University(Kaur) | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Brock University(Kaur) | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.28
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.31
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.35
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.27
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.3
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Brock University(Kaur) | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Brock University(Kaur) | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.18
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Brock University(Kaur) | 2023-12-11 | NCM driving% with AC160TS cantilever | 0.36
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | NCM driving% with AC160TS cantilever | 5.95
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | NCM driving% with AC160TS cantilever | 1.5
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | NCM driving% with AC160TS cantilever | 1.69
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | NCM driving% with AC160TS cantilever | 0.92
| Brock University(Kaur) | 2023-12-11 | NCM Amplitude noise | 14.47
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | NCM Amplitude noise | 13.6
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | NCM Amplitude noise | 12.44
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | NCM Amplitude noise | 11.99
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | NCM Amplitude noise | 11.81
| Brock University(Kaur) | 2023-12-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.753
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.499
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.002
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.079
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.863
| Brock University(Kaur) | 2023-12-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.228
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.595
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.737
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.353
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.35
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 248.2
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 261.3
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 268.7
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 254.9
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 250.8
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 258.6
| Brock University(Kaur) | 2023-12-11 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.8
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.6
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 246.7
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 251.9
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 269.7
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 261.1
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 257.9
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 267.4
| Brock University(Kaur) | 2023-12-11 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100.5
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.6
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Full range of Z Height during Z Scan sweep | 16.3
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | LD beam power at the end of the lens | 2.97
| Brock University(Kaur) | 2023-12-11 | Lateral(C-D) sensitivity | 5.9
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Lateral(C-D) sensitivity | 3.6
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Lateral(C-D) sensitivity | 4.1
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Lateral(C-D) sensitivity | 14.8
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Lateral(C-D) sensitivity | 9.6
| Brock University(Kaur) | 2023-12-11 | Lateral(C-D) noise | 0.072
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Lateral(C-D) noise | 0.124
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Lateral(C-D) noise | 0.151
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Lateral(C-D) noise | 0.095
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Lateral(C-D) noise | 0.117
| Brock University(Kaur) | 2023-12-11 | Full range of Z Height during Z Scan sweep | 17
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of Z Height during Z Scan sweep | 17.3
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of Z Height during Z Scan sweep | 15.9
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of Z Height during Z Scan sweep | 15.8
| Brock University(Kaur) | 2023-12-11 | Full range of Z Detector during Z Scan sweep | 12
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Full range of Z Detector during Z Scan sweep | 15.1
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of Z Detector during Z Scan sweep | 14
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of Z Detector during Z Scan sweep | 12.6
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of Z Detector during Z Scan sweep | 11.2
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of Y2 Detector Raw during Y Scan sweep | 13.7
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of Y2 Detector Raw during Y Scan sweep | 13.4
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of Y1 Detector Raw during Y Scan sweep | 13.9
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of Y1 Detector Raw during Y Scan sweep | 13.7
| Brock University(Kaur) | 2023-12-11 | Full range of Y Detector Raw during Y Scan sweep | 14
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Brock University(Kaur) | 2023-12-11 | Full range of Y Detector during Y Scan sweep | 70
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Full range of Y Detector during Y Scan sweep | 116.3
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of Y Detector during Y Scan sweep | 175.6
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of Y Detector during Y Scan sweep | 178
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of Y Detector during Y Scan sweep | 178.6
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of X2 Detector Raw during X Scan sweep | 14
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of X2 Detector Raw during X Scan sweep | 13.7
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of X2 Detector Raw during X Scan sweep | 13.5
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of X1 Detector Raw during X Scan sweep | 14
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of X1 Detector Raw during X Scan sweep | 13.5
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of X1 Detector Raw during X Scan sweep | 13.6
| Brock University(Kaur) | 2023-12-11 | Full range of X Detector Raw during X Scan sweep | 14.3
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Full range of X Detector Raw during X Scan sweep | 14.7
| Brock University(Kaur) | 2023-12-11 | Full range of X Detector during X Scan sweep | 72.1
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Full range of X Detector during X Scan sweep | 115.9
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | Full range of X Detector during X Scan sweep | 178.1
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | Full range of X Detector during X Scan sweep | 173.2
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | Full range of X Detector during X Scan sweep | 182.1
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | EFM Phase : For the positive bias | 89.6
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | EFM Phase : For the negative bias | -89.6
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Approach position at -3μm | 0.005
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | Approach position at +3μm | 0.0074
| Brock University(Kaur) | 2023-12-11 | A-B noise | 8.35
| Heriot-Watt University (Prof.Brian Gerardot) | 2023-12-11 | A-B noise | 14.41
| EpiWorld International Co.,Ltd. #3 | 2023-12-11 | A-B noise | 7.7
| EpiWorld International Co.,Ltd. #4 | 2023-12-11 | A-B noise | 7.42
| EpiWorld International Co.,Ltd. #5 | 2023-12-11 | A-B noise | 6.67
| Zhongke Dexin Sensing Technology | 2023-12-08 | Z Scanner position at -3μm | 0.0087
| Zhongke Dexin Sensing Technology | 2023-12-08 | Z Scanner position at +3μm | 0.0035
| Zhongke Dexin Sensing Technology | 2023-12-08 | Z Linearity by Z Scanner position
| Zhongke Dexin Sensing Technology | 2023-12-08 | Z Linearity by 3 different sample height | 1.003
| Zhongke Dexin Sensing Technology | 2023-12-08 | Z Height noise | 0.029
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y2 Detector noise (XY Servo on) | 0.22
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y2 Detector noise (XY Servo off) | 0.23
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y1 Detector noise (XY Servo on) | 0.22
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y1 Detector noise (XY Servo off) | 0.22
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y Scan Stroke | 178.29
| Zhongke Dexin Sensing Technology | 2023-12-08 | -Y Direction software Limit | -75138
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y Detector noise (XY Servo on) | 0.17
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y Detector noise (XY Servo off) | 0.17
| Zhongke Dexin Sensing Technology | 2023-12-08 | Y axis OPM | 1.44
| Zhongke Dexin Sensing Technology | 2023-12-08 | X2 Detector noise (XY Servo on) | 0.23
| Zhongke Dexin Sensing Technology | 2023-12-08 | X2 Detector noise (XY Servo off) | 0.22
| Zhongke Dexin Sensing Technology | 2023-12-08 | X1 Detector noise (XY Servo on) | 0.23
| Zhongke Dexin Sensing Technology | 2023-12-08 | X1 Detector noise (XY Servo off) | 0.23
| Zhongke Dexin Sensing Technology | 2023-12-08 | X Scan Stroke | 182
| Zhongke Dexin Sensing Technology | 2023-12-08 | -X Direction software Limit | -75493
| Zhongke Dexin Sensing Technology | 2023-12-08 | X Detector noise (XY Servo on) | 0.17
| Zhongke Dexin Sensing Technology | 2023-12-08 | X Detector noise (XY Servo off) | 0.17
| Zhongke Dexin Sensing Technology | 2023-12-08 | X axis OPM | 1.61
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.32
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.32
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.32
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.32
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Zhongke Dexin Sensing Technology | 2023-12-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| Zhongke Dexin Sensing Technology | 2023-12-08 | NCM driving% with AC160TS cantilever | 1.18
| Zhongke Dexin Sensing Technology | 2023-12-08 | NCM Amplitude noise | 10.62
| Zhongke Dexin Sensing Technology | 2023-12-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.568
| Zhongke Dexin Sensing Technology | 2023-12-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.733
| Zhongke Dexin Sensing Technology | 2023-12-08 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 263.9
| Zhongke Dexin Sensing Technology | 2023-12-08 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 250.1
| Zhongke Dexin Sensing Technology | 2023-12-08 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 258
| Zhongke Dexin Sensing Technology | 2023-12-08 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 257.4
| Zhongke Dexin Sensing Technology | 2023-12-08 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Zhongke Dexin Sensing Technology | 2023-12-08 | Lateral(C-D) sensitivity | 14.9
| Zhongke Dexin Sensing Technology | 2023-12-08 | Lateral(C-D) noise | 0.145
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of Z Height during Z Scan sweep | 18.4
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of Z Detector during Z Scan sweep | 12.8
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of Y2 Detector Raw during Y Scan sweep | 13.8
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of Y Detector during Y Scan sweep | 181.4
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of X2 Detector Raw during X Scan sweep | 14
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Zhongke Dexin Sensing Technology | 2023-12-08 | Full range of X Detector during X Scan sweep | 180.5
| Zhongke Dexin Sensing Technology | 2023-12-08 | A-B noise | 6.64
| 충북대학교 물리학과 | 2023-12-05 | Z Scanner position at -3μm | 0.0029
| Sabaragamuwa University | 2023-12-05 | Z Scanner position at -3μm | 0.0027
| 충북대학교 물리학과 | 2023-12-05 | Z Scanner position at +3μm | 0.0104
| Sabaragamuwa University | 2023-12-05 | Z Scanner position at +3μm | 0.0077
| 충북대학교 물리학과 | 2023-12-05 | Z Linearity by Z Scanner position
| Sabaragamuwa University | 2023-12-05 | Z Linearity by Z Scanner position
| 충북대학교 물리학과 | 2023-12-05 | Z Linearity by 3 different sample height | 0.995
| Sabaragamuwa University | 2023-12-05 | Z Linearity by 3 different sample height | 1
| 충북대학교 물리학과 | 2023-12-05 | Z Height noise | 0.027
| Sabaragamuwa University | 2023-12-05 | Z Height noise | 0.033
| 충북대학교 물리학과 | 2023-12-05 | Y Scan Stroke | 55
| Sabaragamuwa University | 2023-12-05 | Y Scan Stroke | 55
| 충북대학교 물리학과 | 2023-12-05 | Y Detector noise (XY Servo on) | 0.27
| Sabaragamuwa University | 2023-12-05 | Y Detector noise (XY Servo on) | 0.25
| 충북대학교 물리학과 | 2023-12-05 | Y Detector noise (XY Servo off) | 0.26
| Sabaragamuwa University | 2023-12-05 | Y Detector noise (XY Servo off) | 0.24
| 충북대학교 물리학과 | 2023-12-05 | Y axis OPM | 0.47
| Sabaragamuwa University | 2023-12-05 | Y axis OPM | 0.6
| 충북대학교 물리학과 | 2023-12-05 | X Scan Stroke | 55
| Sabaragamuwa University | 2023-12-05 | X Scan Stroke | 55
| 충북대학교 물리학과 | 2023-12-05 | X Detector noise (XY Servo on) | 0.28
| Sabaragamuwa University | 2023-12-05 | X Detector noise (XY Servo on) | 0.25
| 충북대학교 물리학과 | 2023-12-05 | X Detector noise (XY Servo off) | 0.28
| Sabaragamuwa University | 2023-12-05 | X Detector noise (XY Servo off) | 0.25
| 충북대학교 물리학과 | 2023-12-05 | X axis OPM | 0.48
| Sabaragamuwa University | 2023-12-05 | X axis OPM | 0.75
| 충북대학교 물리학과 | 2023-12-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Sabaragamuwa University | 2023-12-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 충북대학교 물리학과 | 2023-12-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Sabaragamuwa University | 2023-12-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| 충북대학교 물리학과 | 2023-12-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| Sabaragamuwa University | 2023-12-05 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| 충북대학교 물리학과 | 2023-12-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Sabaragamuwa University | 2023-12-05 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| 충북대학교 물리학과 | 2023-12-05 | NCM driving% with AC160TS cantilever | 0.43
| Sabaragamuwa University | 2023-12-05 | NCM driving% with AC160TS cantilever | 0.47
| 충북대학교 물리학과 | 2023-12-05 | NCM Amplitude noise | 13.5
| Sabaragamuwa University | 2023-12-05 | NCM Amplitude noise | 16.87
| 충북대학교 물리학과 | 2023-12-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.467
| Sabaragamuwa University | 2023-12-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.326
| 충북대학교 물리학과 | 2023-12-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.527
| Sabaragamuwa University | 2023-12-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.121
| 충북대학교 물리학과 | 2023-12-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.4
| Sabaragamuwa University | 2023-12-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.8
| 충북대학교 물리학과 | 2023-12-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 104.1
| Sabaragamuwa University | 2023-12-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.4
| 충북대학교 물리학과 | 2023-12-05 | Lateral(C-D) sensitivity | 6.7
| Sabaragamuwa University | 2023-12-05 | Lateral(C-D) sensitivity | 13.1
| 충북대학교 물리학과 | 2023-12-05 | Lateral(C-D) noise | 0.077
| Sabaragamuwa University | 2023-12-05 | Lateral(C-D) noise | 0.06
| 충북대학교 물리학과 | 2023-12-05 | Full range of Z Height during Z Scan sweep | 16
| Sabaragamuwa University | 2023-12-05 | Full range of Z Height during Z Scan sweep | 18.3
| 충북대학교 물리학과 | 2023-12-05 | Full range of Z Detector during Z Scan sweep | 12.7
| Sabaragamuwa University | 2023-12-05 | Full range of Z Detector during Z Scan sweep | 13.5
| 충북대학교 물리학과 | 2023-12-05 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Sabaragamuwa University | 2023-12-05 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| 충북대학교 물리학과 | 2023-12-05 | Full range of Y Detector during Y Scan sweep | 73.6
| Sabaragamuwa University | 2023-12-05 | Full range of Y Detector during Y Scan sweep | 74.7
| 충북대학교 물리학과 | 2023-12-05 | Full range of X Detector Raw during X Scan sweep | 14
| Sabaragamuwa University | 2023-12-05 | Full range of X Detector Raw during X Scan sweep | 14.1
| 충북대학교 물리학과 | 2023-12-05 | Full range of X Detector during X Scan sweep | 73
| Sabaragamuwa University | 2023-12-05 | Full range of X Detector during X Scan sweep | 71.4
| 충북대학교 물리학과 | 2023-12-05 | A-B noise | 10.5
| Sabaragamuwa University | 2023-12-05 | A-B noise | 8.52
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Z Scanner position at -3μm | 0.0041
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Z Scanner position at +3μm | 0.0097
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Z Linearity by Z Scanner position
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Z Linearity by 3 different sample height | 1.003
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Z Height noise | 0.027
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y2 Detector noise (XY Servo on) | 0.23
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y2 Detector noise (XY Servo off) | 0.22
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y1 Detector noise (XY Servo on) | 0.25
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y1 Detector noise (XY Servo off) | 0.25
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y Scan Stroke | 179.4
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | -Y Direction software Limit | -102701
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y Detector noise (XY Servo on) | 0.17
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y Detector noise (XY Servo off) | 0.17
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Y axis OPM | 1
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X2 Detector noise (XY Servo on) | 0.24
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X2 Detector noise (XY Servo off) | 0.24
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X1 Detector noise (XY Servo on) | 0.24
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X1 Detector noise (XY Servo off) | 0.23
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X Scan Stroke | 179.1
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | -X Direction software Limit | -102350
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X Detector noise (XY Servo on) | 0.18
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X Detector noise (XY Servo off) | 0.17
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | X axis OPM | 1.4
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.32
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | NCM driving% with AC160TS cantilever | 0.2
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | NCM Amplitude noise | 17.49
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.688
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.506
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 258.7
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 266.2
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 267.5
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 263.2
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Lateral(C-D) sensitivity | 6.4
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Lateral(C-D) noise | 0.088
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of Z Height during Z Scan sweep | 17.1
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of Z Detector during Z Scan sweep | 12.4
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of Y2 Detector Raw during Y Scan sweep | 13.7
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of Y1 Detector Raw during Y Scan sweep | 13.6
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of Y Detector during Y Scan sweep | 179.4
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of X2 Detector Raw during X Scan sweep | 13.5
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of X1 Detector Raw during X Scan sweep | 13.5
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | Full range of X Detector during X Scan sweep | 179.1
| Jinan Guangwei Semiconductor Technology | 2023-12-04 | A-B noise | 8.42
| Samsung P3 Foundry S5 | 2023-12-01 | Z Linearity by Z scanner position | 0.18
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Linearity by Z scanner position | 1.14
| Samsung P3 Foundry S5 | 2023-12-01 | Z Linearity by 5 different sample height | 1.002
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Linearity by 4 different sample height | 1
| Samsung P3 Foundry S5 | 2023-12-01 | Z Height noise | 0.03
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Height noise | 0.039
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.7
| Samsung P3 Foundry S5 | 2023-12-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.7265
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.8
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.9
| Samsung P3 Foundry S5 | 2023-12-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4008
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.9
| Samsung P3 Foundry S5 | 2023-12-01 | Z Drift (Short Traveling) | 0.24
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Z Drift (Short Traveling) | 0.19
| Samsung P3 Foundry S5 | 2023-12-01 | Z Detector offset | 8.7
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y2 Detector noise (XY Servo on) | 0.32
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 270.1
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y1 Detector noise (XY Servo on) | 0.29
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 262.1
| Samsung P3 Foundry S5 | 2023-12-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 62.2318
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 182.8
| Samsung P3 Foundry S5 | 2023-12-01 | Y Ringing(Rpv) | 0.242
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y Ringing(Rpv) | 0.63
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y Non-linearity error | 3.33
| Samsung P3 Foundry S5 | 2023-12-01 | Y Detector noise (XY Servo on) | 0.157
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y Detector noise (XY Servo on) | 0.24
| Samsung P3 Foundry S5 | 2023-12-01 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 89.6503
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Repeatability @1st position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Repeatability @1st position | 0.03
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Repeatability @ 4th position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Repeatability @ 4th position | 0.03
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Repeatability @ 3rd position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Repeatability @ 3rd position | 0.03
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Repeatability @ 2nd position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Repeatability @ 2nd position | 0.03
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Accuracy @1st position | 0.41
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Accuracy @1st position | 0.21
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Accuracy @ 4th position | 0.41
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Accuracy @ 4th position | 0.19
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Accuracy @ 3rd position | 0.49
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Accuracy @ 3rd position | 0.24
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis, Accuracy @ 2nd position | 0.43
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis, Accuracy @ 2nd position | 0.24
| Samsung P3 Foundry S5 | 2023-12-01 | Y axis OPM | 0.78
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Y axis OPM | 0.72
| Samsung P3 Foundry S5 | 2023-12-01 | XY Orthogonality | 89.8
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | XY Orthogonality | 89.8
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X2 Detector noise (XY Servo on) | 0.37
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 261.4
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X1 Detector noise (XY Servo on) | 0.35
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 269.4
| Samsung P3 Foundry S5 | 2023-12-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 62.8689
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 184.2
| Samsung P3 Foundry S5 | 2023-12-01 | X Ringing(Rpv) | 0.239
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X Ringing(Rpv) | 0.77
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X Non-linearity error | -2.48
| Samsung P3 Foundry S5 | 2023-12-01 | X Detector noise (XY Servo on) | 0.154
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X Detector noise (XY Servo on) | 0.3
| Samsung P3 Foundry S5 | 2023-12-01 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 87.3722
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Repeatability @1st position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Repeatability @1st position | 0.02
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Repeatability @ 4th position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Repeatability @ 4th position | 0.03
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Repeatability @ 3rd position | 0.02
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Repeatability @ 3rd position | 0.02
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Repeatability @ 2nd position | 0.01
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Repeatability @ 2nd position | 0.02
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Accuracy @1st position | 0.15
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Accuracy @1st position | 0.11
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Accuracy @ 4th position | 0.12
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Accuracy @ 4th position | 0.13
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Accuracy @ 3rd position | 0.14
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Accuracy @ 3rd position | 0.14
| Samsung P3 Foundry S5 | 2023-12-01 | X axis, Accuracy @ 2nd position | 0.23
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis, Accuracy @ 2nd position | 0.12
| Samsung P3 Foundry S5 | 2023-12-01 | X axis OPM | 0.84
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | X axis OPM | 1.51
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Wafer placement(Loading) error on the load port1 (Y offset) | 8.05
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Wafer placement(Loading) error on the load port1 (X offset) | 6.95
| Samsung P3 Foundry S5 | 2023-12-01 | Under 50nm Step Height Sample (10 Repeats) | 19.783
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Under 50nm Step Height Sample (10 Repeats) | 18.24
| Samsung P3 Foundry S5 | 2023-12-01 | StrokeYMin | -153639
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | StrokeYMin | -153900
| Samsung P3 Foundry S5 | 2023-12-01 | StrokeYMax | 150250.6
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | StrokeYMax | 150350
| Samsung P3 Foundry S5 | 2023-12-01 | StrokeY | 303889.6
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | StrokeY | 304250
| Samsung P3 Foundry S5 | 2023-12-01 | StrokeXMin | -203257
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | StrokeXMin | -203330
| Samsung P3 Foundry S5 | 2023-12-01 | StrokeXMax | 203771
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | StrokeXMax | 203019
| Samsung P3 Foundry S5 | 2023-12-01 | StrokeX | 407028
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | StrokeX | 406349
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.85
| Samsung P3 Foundry S5 | 2023-12-01 | RMS of ZeroScan | 0.055
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | RMS of ZeroScan | 0.048
| Samsung P3 Foundry S5 | 2023-12-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00017
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| Samsung P3 Foundry S5 | 2023-12-01 | OpticalAxisShiftPerFocusY | 0.0008
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | OpticalAxisShiftPerFocusY | 0.0012
| Samsung P3 Foundry S5 | 2023-12-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00024
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00021
| Samsung P3 Foundry S5 | 2023-12-01 | OpticalAxisShiftPerFocusX | 0.0028
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | OpticalAxisShiftPerFocusX | -0.0051
| Samsung P3 Foundry S5 | 2023-12-01 | NCM driving% with AC160TS cantilever | 12.76
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | NCM driving% with AC160TS cantilever | 22.38
| Samsung P3 Foundry S5 | 2023-12-01 | Min. Vibration Transmissbility (~100Hz) | 0.031
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Min. Vibration Transmissbility (~100Hz) | 0.061
| Samsung P3 Foundry S5 | 2023-12-01 | Mask Roughness Repeatability (10Repeats, σ ) | 0.007
| Samsung P3 Foundry S5 | 2023-12-01 | Mask Roughness Repeatability (10Repeats, RSD%) | 2.4
| Samsung P3 Foundry S5 | 2023-12-01 | Mask exchange time | 4
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | LowerLimit | 7019
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Lateral(C-D) sensitivity | 14
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Lateral(C-D) noise | 0.228
| Samsung P3 Foundry S5 | 2023-12-01 | Laser intensity with AC160 cantilever | 2.81
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from -Y to +Y (Y scan) | 78.12
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from -Y to +Y (Y scan) | 61.52
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from -Y to +Y (X scan) | 29.29
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from -Y to +Y (X scan) | 17.58
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from -X to +X (Y scan) | 48.82
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from -X to +X (Y scan) | 21.97
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from -X to +X (X scan) | 97.65
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from -X to +X (X scan) | 68.85
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from +Y to -Y (Y scan) | 97.65
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from +Y to -Y (Y scan) | 32.96
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from +Y to -Y (X scan) | 29.29
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from +Y to -Y (X scan) | 6.59
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from +X to -X (Y scan) | 29.29
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from +X to -X (Y scan) | 92.29
| Samsung P3 Foundry S5 | 2023-12-01 | Image offset after moving from +X to -X (X scan) | 19.53
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Image offset after moving from +X to -X (X scan) | 64.45
| Samsung P3 Foundry S5 | 2023-12-01 | Full range of Z Height during Z Scan sweep | 17.2
| Samsung P3 Foundry S5 | 2023-12-01 | Full range of Z Height during Z Scan sweep | 58.4
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of Z Height during Z Scan sweep | 15.9
| Samsung P3 Foundry S5 | 2023-12-01 | Full range of Z Detector during Z Scan sweep | 13.4
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of Z Detector during Z Scan sweep | 12.3
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.39
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of Y2 Detector during Y Scan sweep | 202.9
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.62
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of Y1 Detector during Y Scan sweep | 202.2
| Samsung P3 Foundry S5 | 2023-12-01 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| Samsung P3 Foundry S5 | 2023-12-01 | Full range of Y Detector during Y Scan sweep | 60.7
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of X2 Detector Raw during X Scan sweep | 14.51
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of X2 Detector during X Scan sweep | 196
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of X1 Detector Raw during X Scan sweep | 13.89
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Full range of X1 Detector during X Scan sweep | 193.9
| Samsung P3 Foundry S5 | 2023-12-01 | Full range of X Detector Raw during X Scan sweep | 13.3
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.2
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| Samsung P3 Foundry S5 | 2023-12-01 | A－B difference during 3 times run test | 0.86
| Samsung P3 Foundry S5 | 2023-12-01 | A-B after Auto PSPD Alignment | -0.167
| Samsung P3 Foundry S5 | 2023-12-01 | A＋B difference during 3 times run test | 0.1
| Samsung P3 Foundry S5 | 2023-12-01 | 940nm Step Height Sample (10 Repeats) | 950.948
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | 940nm Step Height Sample (10 Repeats) | 947.83
| Samsung P3 Foundry S5 | 2023-12-01 | 50~200nm Step Height Sample (10 Repeats) | 184.332
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | 50~200nm Step Height Sample (10 Repeats) | 177.83
| Samsung P3 Foundry S5 | 2023-12-01 | 200~500nm Step Height Sample (10 Repeats) | 462.46
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | 200~500nm Step Height Sample (10 Repeats) | 459.12
| Samsung P3 Foundry S5 | 2023-12-01 | * 1800nm Step Height Sample (10 Repeats) | 1806.686
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 0
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.8
| Guangzhou Cansemi IC Manufacturing(Phase III)Co.,Ltd. #1 | 2023-12-01 | %GRR w/ roughness of non-patterened wafer | 5.8
| 부산대학교 물리학과 | 2023-11-30 | Z Scanner position at -3μm | 0.0014
| Sincere Internation | 2023-11-30 | Z Scanner position at -3μm | 0.0023
| 부산대학교 물리학과 | 2023-11-30 | Z Scanner position at +3μm | 0.0148
| Sincere Internation | 2023-11-30 | Z Scanner position at +3μm | 0.0177
| 부산대학교 물리학과 | 2023-11-30 | Z Linearity by Z Scanner position
| Sincere Internation | 2023-11-30 | Z Linearity by Z Scanner position
| 부산대학교 물리학과 | 2023-11-30 | Z Linearity by 3 different sample height | 1.008
| Sincere Internation | 2023-11-30 | Z Linearity by 3 different sample height | 0.999
| 부산대학교 물리학과 | 2023-11-30 | Z Height noise | 0.03
| Sincere Internation | 2023-11-30 | Z Height noise | 0.027
| 부산대학교 물리학과 | 2023-11-30 | Y Scan Stroke | 55
| Sincere Internation | 2023-11-30 | Y Scan Stroke | 137.1
| Sincere Internation | 2023-11-30 | -Y Direction software Limit | -75998.8
| 부산대학교 물리학과 | 2023-11-30 | Y Detector noise (XY Servo on) | 0.37
| Sincere Internation | 2023-11-30 | Y Detector noise (XY Servo on) | 0.3
| 부산대학교 물리학과 | 2023-11-30 | Y Detector noise (XY Servo off) | 0.39
| Sincere Internation | 2023-11-30 | Y Detector noise (XY Servo off) | 0.3
| 부산대학교 물리학과 | 2023-11-30 | Y axis OPM | 1.36
| Sincere Internation | 2023-11-30 | Y axis OPM | 1.24
| 부산대학교 물리학과 | 2023-11-30 | X Scan Stroke | 55
| Sincere Internation | 2023-11-30 | X Scan Stroke | 134.6
| Sincere Internation | 2023-11-30 | -X Direction software Limit | -75831.9
| 부산대학교 물리학과 | 2023-11-30 | X Detector noise (XY Servo on) | 0.35
| Sincere Internation | 2023-11-30 | X Detector noise (XY Servo on) | 0.27
| 부산대학교 물리학과 | 2023-11-30 | X Detector noise (XY Servo off) | 0.36
| Sincere Internation | 2023-11-30 | X Detector noise (XY Servo off) | 0.26
| 부산대학교 물리학과 | 2023-11-30 | X axis OPM | 1.41
| Sincere Internation | 2023-11-30 | X axis OPM | 1.08
| 부산대학교 물리학과 | 2023-11-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Sincere Internation | 2023-11-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 부산대학교 물리학과 | 2023-11-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Sincere Internation | 2023-11-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Sincere Internation | 2023-11-30 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.27
| Sincere Internation | 2023-11-30 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.27
| Sincere Internation | 2023-11-30 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.28
| Sincere Internation | 2023-11-30 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.26
| Sincere Internation | 2023-11-30 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| 부산대학교 물리학과 | 2023-11-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.1
| Sincere Internation | 2023-11-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 부산대학교 물리학과 | 2023-11-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Sincere Internation | 2023-11-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 부산대학교 물리학과 | 2023-11-30 | NCM driving% with AC160TS cantilever | 0.44
| Sincere Internation | 2023-11-30 | NCM driving% with AC160TS cantilever | 0.17
| 부산대학교 물리학과 | 2023-11-30 | NCM Amplitude noise | 13.65
| Sincere Internation | 2023-11-30 | NCM Amplitude noise | 13.77
| 부산대학교 물리학과 | 2023-11-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.198
| Sincere Internation | 2023-11-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.46
| 부산대학교 물리학과 | 2023-11-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.072
| Sincere Internation | 2023-11-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.801
| 부산대학교 물리학과 | 2023-11-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.3
| Sincere Internation | 2023-11-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 201.6
| 부산대학교 물리학과 | 2023-11-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 92.9
| Sincere Internation | 2023-11-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 198
| Sincere Internation | 2023-11-30 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| 부산대학교 물리학과 | 2023-11-30 | Lateral(C-D) sensitivity | 53.7
| Sincere Internation | 2023-11-30 | Lateral(C-D) sensitivity | 8
| 부산대학교 물리학과 | 2023-11-30 | Lateral(C-D) noise | 0.052
| Sincere Internation | 2023-11-30 | Lateral(C-D) noise | 0.114
| 부산대학교 물리학과 | 2023-11-30 | Full range of Z Height during Z Scan sweep | 17.1
| Sincere Internation | 2023-11-30 | Full range of Z Height during Z Scan sweep | 16.6
| 부산대학교 물리학과 | 2023-11-30 | Full range of Z Detector during Z Scan sweep | 11.7
| Sincere Internation | 2023-11-30 | Full range of Z Detector during Z Scan sweep | 12.4
| 부산대학교 물리학과 | 2023-11-30 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Sincere Internation | 2023-11-30 | Full range of Y Detector Raw during Y Scan sweep | 13.6
| 부산대학교 물리학과 | 2023-11-30 | Full range of Y Detector during Y Scan sweep | 67.5
| Sincere Internation | 2023-11-30 | Full range of Y Detector during Y Scan sweep | 137.4
| 부산대학교 물리학과 | 2023-11-30 | Full range of X Detector Raw during X Scan sweep | 15.1
| Sincere Internation | 2023-11-30 | Full range of X Detector Raw during X Scan sweep | 13.6
| 부산대학교 물리학과 | 2023-11-30 | Full range of X Detector during X Scan sweep | 70
| Sincere Internation | 2023-11-30 | Full range of X Detector during X Scan sweep | 134.2
| 부산대학교 물리학과 | 2023-11-30 | A-B noise | 17.13
| Sincere Internation | 2023-11-30 | A-B noise | 7.59
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Z Scanner position at -3μm | 0.0004
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Z Scanner position at +3μm | 0.0128
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Z Linearity by Z Scanner position
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Z Linearity by 3 different sample height | 0.996
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Z Height noise | 0.027
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Y Detector noise (XY Servo on) | 0.28
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Y Detector noise (XY Servo off) | 0.28
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Y axis OPM | 0.59
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | X Detector noise (XY Servo on) | 0.29
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | X Detector noise (XY Servo off) | 0.29
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | X axis OPM | 0.64
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | NCM driving% with AC160TS cantilever | 0.56
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | NCM Amplitude noise | 10.4
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.311
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.862
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.5
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.2
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Lateral(C-D) sensitivity | 5.6
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Lateral(C-D) noise | 0.097
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Full range of Z Height during Z Scan sweep | 16.6
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Full range of Z Detector during Z Scan sweep | 12.8
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Full range of Y Detector Raw during Y Scan sweep | 14
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Full range of Y Detector during Y Scan sweep | 112.5
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Full range of X Detector Raw during X Scan sweep | 14
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | Full range of X Detector during X Scan sweep | 113
| University of Puerto Rico(Lisandro Cunci) | 2023-11-29 | A-B noise | 7.67
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Linearity by Z scanner position | 0.98
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Linearity by Z scanner position | 0.74
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Linearity by 4 different sample height | 1
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Linearity by 4 different sample height | 1.002
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Height noise | 0.036
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Height noise | 0.087
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.9
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 65.6767
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.9
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 65.643
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.5185
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.7
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.5504
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Z Drift (Short Traveling) | 0.16
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y2 Detector noise (XY Servo on) | 0.41
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 254.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y1 Detector noise (XY Servo on) | 0.4
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 269
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 179.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y Ringing(Rpv) | 0.8
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y Non-linearity error | 1.53
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y Detector noise (XY Servo on) | 0.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Repeatability @1st position | 0.04
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Repeatability @ 4th position | 0.03
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Repeatability @ 3rd position | 0.03
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Repeatability @ 2nd position | 0.03
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Accuracy @1st position | 0.38
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Accuracy @ 4th position | 0.35
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Accuracy @ 3rd position | 0.45
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis, Accuracy @ 2nd position | 0.41
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Y axis OPM | 0.87
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | XY Orthogonality | 90.1
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X2 Detector noise (XY Servo on) | 0.45
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 263.6
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X1 Detector noise (XY Servo on) | 0.41
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 272.8
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 181.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X Ringing(Rpv) | 0.79
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X Non-linearity error | 1.7
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X Detector noise (XY Servo on) | 0.33
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Repeatability @1st position | 0.03
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Repeatability @ 4th position | 0.04
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Repeatability @ 3rd position | 0.03
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Repeatability @ 2nd position | 0.03
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Accuracy @1st position | 0.2
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Accuracy @ 4th position | 0.2
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Accuracy @ 3rd position | 0.21
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis, Accuracy @ 2nd position | 0.2
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | X axis OPM | 1.69
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Wafer placement(Loading) error on the load port2 (Y offset) | 44.01
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Wafer placement(Loading) error on the load port2 (X offset) | 22.65
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Wafer placement(Loading) error on the load port1 (Y offset) | 16.85
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Wafer placement(Loading) error on the load port1 (X offset) | 34.25
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Under 50nm Step Height Sample (10 Repeats) | 18.01
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Under 50nm Step Height Sample (10 Repeats) | 18.13
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | StrokeYMin | -153450
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | StrokeYMax | 150439.8
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | StrokeY | 303889.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | StrokeXMin | -204117
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | StrokeXMax | 203707.6
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | StrokeX | 407824.5
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.61
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | RMS of ZeroScan | 0.034
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4e-05
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | OpticalAxisShiftPerFocusY | 0
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00023
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | OpticalAxisShiftPerFocusX | 0.0029
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | NCM driving% with AC160TS cantilever | 5.39
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | NCM driving% with AC160TS cantilever | 22.26
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Min. Vibration Transmissbility (~100Hz) | 0.018
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | LowerLimit | 6400
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Lateral(C-D) sensitivity | 12.9
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Lateral(C-D) noise | 0.182
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from -Y to +Y (Y scan) | 75.44
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from -Y to +Y (X scan) | 16.11
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from -X to +X (Y scan) | 30.76
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from -X to +X (X scan) | 54.93
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from +Y to -Y (Y scan) | 5.13
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from +Y to -Y (X scan) | 4.39
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from +X to -X (Y scan) | 18.31
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Image offset after moving from +X to -X (X scan) | 101.81
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Z Height during Z Scan sweep | 16.2
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Z Height during Z Scan sweep | 42.39
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Z Detector during Z Scan sweep | 11.6
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Z Detector during Z Scan sweep | 13
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Y2 Detector during Y Scan sweep | 187.8
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.53
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of Y1 Detector during Y Scan sweep | 185.4
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of X2 Detector Raw during X Scan sweep | 14.16
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of X2 Detector during X Scan sweep | 187.5
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of X1 Detector Raw during X Scan sweep | 14.37
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Full range of X1 Detector during X Scan sweep | 192.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.32
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.27
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | 940nm Step Height Sample (10 Repeats) | 947.56
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | 940nm Step Height Sample (10 Repeats) | 949.42
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | 50~200nm Step Height Sample (10 Repeats) | 183.26
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | 50~200nm Step Height Sample (10 Repeats) | 183.42
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | 200~500nm Step Height Sample (10 Repeats) | 461.95
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | 200~500nm Step Height Sample (10 Repeats) | 463.9
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | * 1800nm Step Height Sample (10 Repeats) | 1793.53
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | * 1800nm Step Height Sample (10 Repeats) | 1792.42
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 3.3
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 7.5
| Sidtek(Semiconductor Integrated Display Technology) | 2023-11-24 | %GRR w/ roughness of non-patterened wafer | 8.2
| Khalifa University | 2023-11-24 | Z Linearity by Z Scanner position
| Khalifa University | 2023-11-24 | Z Linearity by 3 different sample height | 0.994
| Khalifa University | 2023-11-24 | Z Height noise | 0.022
| Khalifa University | 2023-11-24 | Y Detector noise (XY Servo on) | 0.27
| Khalifa University | 2023-11-24 | Y Detector noise (XY Servo off) | 0.27
| Khalifa University | 2023-11-24 | Y axis OPM | 0.83
| Khalifa University | 2023-11-24 | X Detector noise (XY Servo on) | 0.26
| Khalifa University | 2023-11-24 | X Detector noise (XY Servo off) | 0.26
| Khalifa University | 2023-11-24 | X axis OPM | 1.64
| Khalifa University | 2023-11-24 | SLD beam power at the end of the lens | 0.2
| Khalifa University | 2023-11-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Khalifa University | 2023-11-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Khalifa University | 2023-11-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Khalifa University | 2023-11-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Khalifa University | 2023-11-24 | NCM driving% with AC160TS cantilever | 1.78
| Khalifa University | 2023-11-24 | NCM Amplitude noise | 17.43
| Khalifa University | 2023-11-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.936
| Khalifa University | 2023-11-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.463
| Khalifa University | 2023-11-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.2
| Khalifa University | 2023-11-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 150.8
| Khalifa University | 2023-11-24 | Full range of Z Height during Z Scan sweep | 15.8
| Khalifa University | 2023-11-24 | LD beam power at the end of the lens | 3.21
| Khalifa University | 2023-11-24 | Lateral(C-D) sensitivity | 4
| Khalifa University | 2023-11-24 | Lateral(C-D) noise | 0.11
| Khalifa University | 2023-11-24 | Full range of Z Detector during Z Scan sweep | 13.4
| Khalifa University | 2023-11-24 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| Khalifa University | 2023-11-24 | Full range of Y Detector during Y Scan sweep | 115.7
| Khalifa University | 2023-11-24 | Full range of X Detector Raw during X Scan sweep | 14.7
| Khalifa University | 2023-11-24 | Full range of X Detector during X Scan sweep | 110.7
| Khalifa University | 2023-11-24 | EFM Phase : For the positive bias | 90.1
| Khalifa University | 2023-11-24 | EFM Phase : For the negative bias | -90.1
| Khalifa University | 2023-11-24 | Approach position at -3μm | 0.0034
| Khalifa University | 2023-11-24 | Approach position at +3μm | 0.0042
| Khalifa University | 2023-11-24 | A-B noise | 14.83
| PrimeNano(Eduard) | 2023-11-22 | Z Scanner position at -3μm | 0.0032
| PrimeNano(Eduard) | 2023-11-22 | Z Scanner position at +3μm | 0.0063
| PrimeNano(Eduard) | 2023-11-22 | Z Linearity by Z Scanner position
| PrimeNano(Eduard) | 2023-11-22 | Z Linearity by 3 different sample height | 1.004
| PrimeNano(Eduard) | 2023-11-22 | Z Height noise | 0.028
| PrimeNano(Eduard) | 2023-11-22 | Y2 Detector noise (XY Servo on) | 0.23
| PrimeNano(Eduard) | 2023-11-22 | Y2 Detector noise (XY Servo off) | 0.22
| PrimeNano(Eduard) | 2023-11-22 | Y1 Detector noise (XY Servo on) | 0.23
| PrimeNano(Eduard) | 2023-11-22 | Y1 Detector noise (XY Servo off) | 0.23
| PrimeNano(Eduard) | 2023-11-22 | Y Scan Stroke | 179.32
| PrimeNano(Eduard) | 2023-11-22 | -Y Direction software Limit | -102893
| PrimeNano(Eduard) | 2023-11-22 | Y Detector noise (XY Servo on) | 0.16
| PrimeNano(Eduard) | 2023-11-22 | Y Detector noise (XY Servo off) | 0.16
| PrimeNano(Eduard) | 2023-11-22 | Y axis OPM | 1.53
| PrimeNano(Eduard) | 2023-11-22 | X2 Detector noise (XY Servo on) | 0.24
| PrimeNano(Eduard) | 2023-11-22 | X2 Detector noise (XY Servo off) | 0.22
| PrimeNano(Eduard) | 2023-11-22 | X1 Detector noise (XY Servo on) | 0.25
| PrimeNano(Eduard) | 2023-11-22 | X1 Detector noise (XY Servo off) | 0.24
| PrimeNano(Eduard) | 2023-11-22 | X Scan Stroke | 175.03
| PrimeNano(Eduard) | 2023-11-22 | -X Direction software Limit | -102200
| PrimeNano(Eduard) | 2023-11-22 | X Detector noise (XY Servo on) | 0.19
| PrimeNano(Eduard) | 2023-11-22 | X Detector noise (XY Servo off) | 0.17
| PrimeNano(Eduard) | 2023-11-22 | X axis OPM | 1.35
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.32
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.3
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.3
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| PrimeNano(Eduard) | 2023-11-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| PrimeNano(Eduard) | 2023-11-22 | NCM driving% with AC160TS cantilever | 0.78
| PrimeNano(Eduard) | 2023-11-22 | NCM Amplitude noise | 14.91
| PrimeNano(Eduard) | 2023-11-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.044
| PrimeNano(Eduard) | 2023-11-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.01
| PrimeNano(Eduard) | 2023-11-22 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 241.8
| PrimeNano(Eduard) | 2023-11-22 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 249.2
| PrimeNano(Eduard) | 2023-11-22 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 241.6
| PrimeNano(Eduard) | 2023-11-22 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 246.3
| PrimeNano(Eduard) | 2023-11-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| PrimeNano(Eduard) | 2023-11-22 | Lateral(C-D) sensitivity | 6.3
| PrimeNano(Eduard) | 2023-11-22 | Lateral(C-D) noise | 0.132
| PrimeNano(Eduard) | 2023-11-22 | Full range of Z Height during Z Scan sweep | 17.1
| PrimeNano(Eduard) | 2023-11-22 | Full range of Z Detector during Z Scan sweep | 12.4
| PrimeNano(Eduard) | 2023-11-22 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| PrimeNano(Eduard) | 2023-11-22 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| PrimeNano(Eduard) | 2023-11-22 | Full range of Y Detector during Y Scan sweep | 178.8
| PrimeNano(Eduard) | 2023-11-22 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| PrimeNano(Eduard) | 2023-11-22 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| PrimeNano(Eduard) | 2023-11-22 | Full range of X Detector during X Scan sweep | 174.9
| PrimeNano(Eduard) | 2023-11-22 | A-B noise | 10.14
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Z Scanner position at -3μm | 0.0019
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Z Scanner position at +3μm | 0.0027
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Z Linearity by Z Scanner position
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Z Linearity by 3 different sample height | 0.997
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Z Height noise | 0.029
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Y Scan Stroke | 55
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Y Detector noise (XY Servo on) | 0.27
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Y Detector noise (XY Servo off) | 0.26
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Y axis OPM | 0.64
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | X Scan Stroke | 55
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | X Detector noise (XY Servo on) | 0.27
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | X Detector noise (XY Servo off) | 0.26
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | X axis OPM | 0.56
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.09
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | NCM driving% with AC160TS cantilever | 0.93
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | NCM Amplitude noise | 12.44
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.172
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.847
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.5
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.3
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Lateral(C-D) sensitivity | 58.9
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Lateral(C-D) noise | 0.069
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Full range of Z Height during Z Scan sweep | 17.9
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Full range of Z Detector during Z Scan sweep | 13.3
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Full range of Y Detector during Y Scan sweep | 74
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Full range of X Detector Raw during X Scan sweep | 14.2
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | Full range of X Detector during X Scan sweep | 72
| 강원대학교 화학과 (한솔켐바이오텍) | 2023-11-20 | A-B noise | 6.17
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Z Scanner position at -3μm | 0.006
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Z Scanner position at -3μm | 0.0009
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Z Scanner position at +3μm | 0.0096
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Z Scanner position at +3μm | 0.0115
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Z Linearity by Z Scanner position
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Z Linearity by Z Scanner position
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Z Linearity by 3 different sample height | 1.012
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Z Linearity by 3 different sample height | 1.006
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Z Height noise | 0.027
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Z Height noise | 0.026
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Y Scan Stroke | 136.2
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Y Scan Stroke | 130
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | -Y Direction software Limit | -76155.6
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | -Y Direction software Limit | -76336.9
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Y Detector noise (XY Servo on) | 0.29
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Y Detector noise (XY Servo on) | 0.27
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Y Detector noise (XY Servo off) | 0.29
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Y Detector noise (XY Servo off) | 0.27
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Y axis OPM | 1.06
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Y axis OPM | 0.54
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | X Scan Stroke | 132.5
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | X Scan Stroke | 128.7
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | -X Direction software Limit | -76021.2
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | -X Direction software Limit | -76056.9
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | X Detector noise (XY Servo on) | 0.29
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | X Detector noise (XY Servo on) | 0.27
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | X Detector noise (XY Servo off) | 0.29
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | X Detector noise (XY Servo off) | 0.26
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | X axis OPM | 0.76
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | X axis OPM | 0.85
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.3
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.24
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.29
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.27
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.24
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.24
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.26
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.24
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.24
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.24
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | NCM driving% with AC160TS cantilever | 1.03
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | NCM driving% with AC160TS cantilever | 0.27
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | NCM Amplitude noise | 12.06
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | NCM Amplitude noise | 13.91
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.561
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.561
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.177
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.177
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 178
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 191.4
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 181.5
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 182.7
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Lateral(C-D) sensitivity | 6.9
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Lateral(C-D) sensitivity | 19.6
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Lateral(C-D) noise | 0.148
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Lateral(C-D) noise | 0.114
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Full range of Z Height during Z Scan sweep | 16.7
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Full range of Z Height during Z Scan sweep | 16.6
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Full range of Z Detector during Z Scan sweep | 13.3
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Full range of Z Detector during Z Scan sweep | 13.2
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Full range of Y Detector during Y Scan sweep | 135.9
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Full range of Y Detector during Y Scan sweep | 121.8
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Full range of X Detector Raw during X Scan sweep | 14.6
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Full range of X Detector Raw during X Scan sweep | 13.9
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | Full range of X Detector during X Scan sweep | 132.4
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | Full range of X Detector during X Scan sweep | 125.8
| Zhuhai Gree Electronic Componets#1 | 2023-11-17 | A-B noise | 8.1
| Zhuhai Gree Electronic Componets#2 | 2023-11-17 | A-B noise | 7.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Z Scanner position at -3μm | 0.0054
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Z Scanner position at +3μm | 0.0073
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Z Linearity by Z Scanner position
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Z Linearity by 3 different sample height | 0.998
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Z Height noise | 0.037
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y2 Detector noise (XY Servo on) | 0.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y2 Detector noise (XY Servo off) | 0.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y1 Detector noise (XY Servo on) | 0.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y1 Detector noise (XY Servo off) | 0.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y Scan Stroke | 184
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | -Y Direction software Limit | -103308
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y Detector noise (XY Servo on) | 0.17
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y Detector noise (XY Servo off) | 0.17
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Y axis OPM | 0.89
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X2 Detector noise (XY Servo on) | 0.23
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X2 Detector noise (XY Servo off) | 0.23
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X1 Detector noise (XY Servo on) | 0.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X1 Detector noise (XY Servo off) | 0.23
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X Scan Stroke | 176
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | -X Direction software Limit | -101925
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X Detector noise (XY Servo on) | 0.18
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X Detector noise (XY Servo off) | 0.17
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | X axis OPM | 0.75
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.39
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.38
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.36
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.37
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.36
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | NCM driving% with AC160TS cantilever | 0.3
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | NCM Amplitude noise | 15.42
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.381
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.293
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 256.1
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 254.5
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 234.4
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 252.4
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Lateral(C-D) sensitivity | 12.6
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Lateral(C-D) noise | 0.111
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of Z Height during Z Scan sweep | 16.2
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of Z Detector during Z Scan sweep | 12.8
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of Y Detector during Y Scan sweep | 183.3
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of X1 Detector Raw during X Scan sweep | 14.5
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | Full range of X Detector during X Scan sweep | 175.7
| Fraunhofer Institute for Silicon Technology (ISIT) | 2023-11-16 | A-B noise | 9.24
| Anhui University | 2023-11-15 | Z Scanner position at -3μm | 0.0016
| Anhui University | 2023-11-15 | Z Scanner position at +3μm | 0.0136
| Anhui University | 2023-11-15 | Z Linearity by Z Scanner position
| Anhui University | 2023-11-15 | Z Linearity by 3 different sample height | 0.997
| Anhui University | 2023-11-15 | Z Height noise | 0.028
| Anhui University | 2023-11-15 | Y2 Detector noise (XY Servo on) | 0.26
| Anhui University | 2023-11-15 | Y2 Detector noise (XY Servo off) | 0.27
| Anhui University | 2023-11-15 | Y1 Detector noise (XY Servo on) | 0.28
| Anhui University | 2023-11-15 | Y1 Detector noise (XY Servo off) | 0.27
| Anhui University | 2023-11-15 | Y Scan Stroke | 175.2
| Anhui University | 2023-11-15 | -Y Direction software Limit | -102585
| Anhui University | 2023-11-15 | Y Detector noise (XY Servo on) | 0.2
| Anhui University | 2023-11-15 | Y Detector noise (XY Servo off) | 0.2
| Anhui University | 2023-11-15 | Y axis OPM | 0.75
| Anhui University | 2023-11-15 | X2 Detector noise (XY Servo on) | 0.25
| Anhui University | 2023-11-15 | X2 Detector noise (XY Servo off) | 0.25
| Anhui University | 2023-11-15 | X1 Detector noise (XY Servo on) | 0.26
| Anhui University | 2023-11-15 | X1 Detector noise (XY Servo off) | 0.26
| Anhui University | 2023-11-15 | X Scan Stroke | 180.7
| Anhui University | 2023-11-15 | -X Direction software Limit | -101420
| Anhui University | 2023-11-15 | X Detector noise (XY Servo on) | 0.18
| Anhui University | 2023-11-15 | X Detector noise (XY Servo off) | 0.18
| Anhui University | 2023-11-15 | X axis OPM | 0.93
| Anhui University | 2023-11-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Anhui University | 2023-11-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Anhui University | 2023-11-15 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.4
| Anhui University | 2023-11-15 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.37
| Anhui University | 2023-11-15 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.36
| Anhui University | 2023-11-15 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.32
| Anhui University | 2023-11-15 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Anhui University | 2023-11-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Anhui University | 2023-11-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Anhui University | 2023-11-15 | NCM driving% with AC160TS cantilever | 1.16
| Anhui University | 2023-11-15 | NCM Amplitude noise | 15.44
| Anhui University | 2023-11-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.733
| Anhui University | 2023-11-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.936
| Anhui University | 2023-11-15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 260.4
| Anhui University | 2023-11-15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.7
| Anhui University | 2023-11-15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 261.3
| Anhui University | 2023-11-15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 266.3
| Anhui University | 2023-11-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.09
| Anhui University | 2023-11-15 | Lateral(C-D) sensitivity | 6.1
| Anhui University | 2023-11-15 | Lateral(C-D) noise | 0.128
| Anhui University | 2023-11-15 | Full range of Z Height during Z Scan sweep | 16.9
| Anhui University | 2023-11-15 | Full range of Z Detector during Z Scan sweep | 13
| Anhui University | 2023-11-15 | Full range of Y2 Detector Raw during Y Scan sweep | 13.7
| Anhui University | 2023-11-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| Anhui University | 2023-11-15 | Full range of Y Detector during Y Scan sweep | 174.5
| Anhui University | 2023-11-15 | Full range of X2 Detector Raw during X Scan sweep | 13.6
| Anhui University | 2023-11-15 | Full range of X1 Detector Raw during X Scan sweep | 13.6
| Anhui University | 2023-11-15 | Full range of X Detector during X Scan sweep | 179.1
| Anhui University | 2023-11-15 | A-B noise | 8.4
| University of Wuerzburg | 2023-11-13 | Z Scanner position at -3μm | 0.0039
| DESY (Dr.Chumakov) | 2023-11-13 | Z Scanner position at -3μm | 0.0006
| University of Wuerzburg | 2023-11-13 | Z Scanner position at +3μm | 0.002
| DESY (Dr.Chumakov) | 2023-11-13 | Z Scanner position at +3μm | 0.0121
| University of Wuerzburg | 2023-11-13 | Z Linearity by Z Scanner position
| DESY (Dr.Chumakov) | 2023-11-13 | Z Linearity by Z Scanner position
| University of Wuerzburg | 2023-11-13 | Z Linearity by 3 different sample height | 0.996
| DESY (Dr.Chumakov) | 2023-11-13 | Z Linearity by 3 different sample height | 0.999
| University of Wuerzburg | 2023-11-13 | Z Height noise | 0.027
| DESY (Dr.Chumakov) | 2023-11-13 | Z Height noise | 0.027
| University of Wuerzburg | 2023-11-13 | Y Scan Stroke | 135.5
| University of Wuerzburg | 2023-11-13 | Y Detector noise (XY Servo on) | 0.39
| DESY (Dr.Chumakov) | 2023-11-13 | Y Detector noise (XY Servo on) | 0.26
| University of Wuerzburg | 2023-11-13 | Y Detector noise (XY Servo off) | 0.4
| DESY (Dr.Chumakov) | 2023-11-13 | Y Detector noise (XY Servo off) | 0.25
| University of Wuerzburg | 2023-11-13 | Y axis OPM | 1.31
| DESY (Dr.Chumakov) | 2023-11-13 | Y axis OPM | 1.56
| University of Wuerzburg | 2023-11-13 | X Scan Stroke | 131.7
| University of Wuerzburg | 2023-11-13 | X Detector noise (XY Servo on) | 0.41
| DESY (Dr.Chumakov) | 2023-11-13 | X Detector noise (XY Servo on) | 0.25
| University of Wuerzburg | 2023-11-13 | X Detector noise (XY Servo off) | 0.41
| DESY (Dr.Chumakov) | 2023-11-13 | X Detector noise (XY Servo off) | 0.25
| University of Wuerzburg | 2023-11-13 | X axis OPM | 0.93
| DESY (Dr.Chumakov) | 2023-11-13 | X axis OPM | 1.15
| University of Wuerzburg | 2023-11-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| DESY (Dr.Chumakov) | 2023-11-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| University of Wuerzburg | 2023-11-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| DESY (Dr.Chumakov) | 2023-11-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| University of Wuerzburg | 2023-11-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| DESY (Dr.Chumakov) | 2023-11-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| University of Wuerzburg | 2023-11-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| DESY (Dr.Chumakov) | 2023-11-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.14
| University of Wuerzburg | 2023-11-13 | NCM driving% with AC160TS cantilever | 0.39
| DESY (Dr.Chumakov) | 2023-11-13 | NCM driving% with AC160TS cantilever | 0.9
| University of Wuerzburg | 2023-11-13 | NCM Amplitude noise | 12.29
| DESY (Dr.Chumakov) | 2023-11-13 | NCM Amplitude noise | 12.13
| University of Wuerzburg | 2023-11-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.171
| DESY (Dr.Chumakov) | 2023-11-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.148
| University of Wuerzburg | 2023-11-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.906
| DESY (Dr.Chumakov) | 2023-11-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.301
| University of Wuerzburg | 2023-11-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 185.9
| DESY (Dr.Chumakov) | 2023-11-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.3
| University of Wuerzburg | 2023-11-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 186.5
| DESY (Dr.Chumakov) | 2023-11-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.9
| University of Wuerzburg | 2023-11-13 | Lateral(C-D) sensitivity | 12.2
| DESY (Dr.Chumakov) | 2023-11-13 | Lateral(C-D) sensitivity | 12.93
| University of Wuerzburg | 2023-11-13 | Lateral(C-D) noise | 0.106
| DESY (Dr.Chumakov) | 2023-11-13 | Lateral(C-D) noise | 0.076
| University of Wuerzburg | 2023-11-13 | Full range of Z Height during Z Scan sweep | 17.4
| DESY (Dr.Chumakov) | 2023-11-13 | Full range of Z Height during Z Scan sweep | 17.8
| University of Wuerzburg | 2023-11-13 | Full range of Z Detector during Z Scan sweep | 14.6
| DESY (Dr.Chumakov) | 2023-11-13 | Full range of Z Detector during Z Scan sweep | 14.7
| University of Wuerzburg | 2023-11-13 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| DESY (Dr.Chumakov) | 2023-11-13 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| University of Wuerzburg | 2023-11-13 | Full range of Y Detector during Y Scan sweep | 135.8
| DESY (Dr.Chumakov) | 2023-11-13 | Full range of Y Detector during Y Scan sweep | 109.8
| University of Wuerzburg | 2023-11-13 | Full range of X Detector Raw during X Scan sweep | 14.1
| DESY (Dr.Chumakov) | 2023-11-13 | Full range of X Detector Raw during X Scan sweep | 14.7
| University of Wuerzburg | 2023-11-13 | Full range of X Detector during X Scan sweep | 131.1
| DESY (Dr.Chumakov) | 2023-11-13 | Full range of X Detector during X Scan sweep | 114.2
| University of Wuerzburg | 2023-11-13 | A-B noise | 12.25
| DESY (Dr.Chumakov) | 2023-11-13 | A-B noise | 13.18
| Fukoku | 2023-11-10 | Z Scanner position at -3μm | 0.0036
| Exquisite-Automotive Systems | 2023-11-10 | Z Scanner position at -3μm | 0.0111
| Fukoku | 2023-11-10 | Z Scanner position at +3μm | 0.0127
| Exquisite-Automotive Systems | 2023-11-10 | Z Scanner position at +3μm | 0.0127
| Fukoku | 2023-11-10 | Z Linearity by Z Scanner position
| Exquisite-Automotive Systems | 2023-11-10 | Z Linearity by Z Scanner position
| 인하대학교 물리학과 | 2023-11-10 | Z Linearity by Z Scanner position
| Fukoku | 2023-11-10 | Z Linearity by 3 different sample height | 0.999
| Exquisite-Automotive Systems | 2023-11-10 | Z Linearity by 3 different sample height | 0.996
| 인하대학교 물리학과 | 2023-11-10 | Z Linearity by 3 different sample height | 0.993
| Fukoku | 2023-11-10 | Z Height noise | 0.031
| Exquisite-Automotive Systems | 2023-11-10 | Z Height noise | 0.025
| 인하대학교 물리학과 | 2023-11-10 | Z Height noise | 0.022
| Exquisite-Automotive Systems | 2023-11-10 | Y2 Detector noise (XY Servo on) | 0.32
| Exquisite-Automotive Systems | 2023-11-10 | Y2 Detector noise (XY Servo off) | 0.32
| Exquisite-Automotive Systems | 2023-11-10 | Y1 Detector noise (XY Servo on) | 0.25
| Exquisite-Automotive Systems | 2023-11-10 | Y1 Detector noise (XY Servo off) | 0.25
| Fukoku | 2023-11-10 | Y Scan Stroke | 55
| Exquisite-Automotive Systems | 2023-11-10 | Y Scan Stroke | 176.5
| Exquisite-Automotive Systems | 2023-11-10 | -Y Direction software Limit | -103459
| Fukoku | 2023-11-10 | Y Detector noise (XY Servo on) | 0.28
| Exquisite-Automotive Systems | 2023-11-10 | Y Detector noise (XY Servo on) | 0.21
| 인하대학교 물리학과 | 2023-11-10 | Y Detector noise (XY Servo on) | 0.28
| Fukoku | 2023-11-10 | Y Detector noise (XY Servo off) | 0.3
| Exquisite-Automotive Systems | 2023-11-10 | Y Detector noise (XY Servo off) | 0.21
| 인하대학교 물리학과 | 2023-11-10 | Y Detector noise (XY Servo off) | 0.27
| Fukoku | 2023-11-10 | Y axis OPM | 0.44
| Exquisite-Automotive Systems | 2023-11-10 | Y axis OPM | 0.88
| 인하대학교 물리학과 | 2023-11-10 | Y axis OPM | 0.6
| Exquisite-Automotive Systems | 2023-11-10 | X2 Detector noise (XY Servo on) | 0.26
| Exquisite-Automotive Systems | 2023-11-10 | X2 Detector noise (XY Servo off) | 0.26
| Exquisite-Automotive Systems | 2023-11-10 | X1 Detector noise (XY Servo on) | 0.29
| Exquisite-Automotive Systems | 2023-11-10 | X1 Detector noise (XY Servo off) | 0.28
| Fukoku | 2023-11-10 | X Scan Stroke | 55
| Exquisite-Automotive Systems | 2023-11-10 | X Scan Stroke | 178.3
| Exquisite-Automotive Systems | 2023-11-10 | -X Direction software Limit | -102528
| Fukoku | 2023-11-10 | X Detector noise (XY Servo on) | 0.28
| Exquisite-Automotive Systems | 2023-11-10 | X Detector noise (XY Servo on) | 0.2
| 인하대학교 물리학과 | 2023-11-10 | X Detector noise (XY Servo on) | 0.28
| Fukoku | 2023-11-10 | X Detector noise (XY Servo off) | 0.3
| Exquisite-Automotive Systems | 2023-11-10 | X Detector noise (XY Servo off) | 0.19
| 인하대학교 물리학과 | 2023-11-10 | X Detector noise (XY Servo off) | 0.27
| Fukoku | 2023-11-10 | X axis OPM | 0.93
| Exquisite-Automotive Systems | 2023-11-10 | X axis OPM | 1.2
| 인하대학교 물리학과 | 2023-11-10 | X axis OPM | 0.89
| 인하대학교 물리학과 | 2023-11-10 | SLD beam power at the end of the lens | 0.23
| Fukoku | 2023-11-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 인하대학교 물리학과 | 2023-11-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Fukoku | 2023-11-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 인하대학교 물리학과 | 2023-11-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.28
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.33
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.28
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.28
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| Fukoku | 2023-11-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 인하대학교 물리학과 | 2023-11-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Fukoku | 2023-11-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Exquisite-Automotive Systems | 2023-11-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| 인하대학교 물리학과 | 2023-11-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.18
| Fukoku | 2023-11-10 | NCM driving% with AC160TS cantilever | 0.41
| Exquisite-Automotive Systems | 2023-11-10 | NCM driving% with AC160TS cantilever | 0.15
| 인하대학교 물리학과 | 2023-11-10 | NCM driving% with AC160TS cantilever | 2.97
| Fukoku | 2023-11-10 | NCM Amplitude noise | 12.76
| Exquisite-Automotive Systems | 2023-11-10 | NCM Amplitude noise | 11.76
| 인하대학교 물리학과 | 2023-11-10 | NCM Amplitude noise | 18.03
| Fukoku | 2023-11-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 11.004
| Exquisite-Automotive Systems | 2023-11-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.079
| 인하대학교 물리학과 | 2023-11-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.459
| Fukoku | 2023-11-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.209
| Exquisite-Automotive Systems | 2023-11-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.353
| 인하대학교 물리학과 | 2023-11-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.648
| Exquisite-Automotive Systems | 2023-11-10 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 243.7
| Exquisite-Automotive Systems | 2023-11-10 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 253.6
| Fukoku | 2023-11-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.7
| 인하대학교 물리학과 | 2023-11-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 146.8
| Exquisite-Automotive Systems | 2023-11-10 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 253.8
| Exquisite-Automotive Systems | 2023-11-10 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 262.9
| Fukoku | 2023-11-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.2
| 인하대학교 물리학과 | 2023-11-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 152.6
| 인하대학교 물리학과 | 2023-11-10 | Full range of Z Height during Z Scan sweep | 16.4
| Exquisite-Automotive Systems | 2023-11-10 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| 인하대학교 물리학과 | 2023-11-10 | LD beam power at the end of the lens | 1.7
| Fukoku | 2023-11-10 | Lateral(C-D) sensitivity | 19.2
| Exquisite-Automotive Systems | 2023-11-10 | Lateral(C-D) sensitivity | 20.4
| 인하대학교 물리학과 | 2023-11-10 | Lateral(C-D) sensitivity | 3.2
| Fukoku | 2023-11-10 | Lateral(C-D) noise | 0.069
| Exquisite-Automotive Systems | 2023-11-10 | Lateral(C-D) noise | 0.113
| 인하대학교 물리학과 | 2023-11-10 | Lateral(C-D) noise | 0.139
| Fukoku | 2023-11-10 | Full range of Z Height during Z Scan sweep | 17.4
| Exquisite-Automotive Systems | 2023-11-10 | Full range of Z Height during Z Scan sweep | 15.9
| Fukoku | 2023-11-10 | Full range of Z Detector during Z Scan sweep | 15
| Exquisite-Automotive Systems | 2023-11-10 | Full range of Z Detector during Z Scan sweep | 12.6
| 인하대학교 물리학과 | 2023-11-10 | Full range of Z Detector during Z Scan sweep | 13.3
| Exquisite-Automotive Systems | 2023-11-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Exquisite-Automotive Systems | 2023-11-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.3
| Fukoku | 2023-11-10 | Full range of Y Detector Raw during Y Scan sweep | 14
| 인하대학교 물리학과 | 2023-11-10 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| Fukoku | 2023-11-10 | Full range of Y Detector during Y Scan sweep | 71.3
| Exquisite-Automotive Systems | 2023-11-10 | Full range of Y Detector during Y Scan sweep | 178.3
| 인하대학교 물리학과 | 2023-11-10 | Full range of Y Detector during Y Scan sweep | 114.5
| Exquisite-Automotive Systems | 2023-11-10 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| Exquisite-Automotive Systems | 2023-11-10 | Full range of X1 Detector Raw during X Scan sweep | 13.8
| Fukoku | 2023-11-10 | Full range of X Detector Raw during X Scan sweep | 13.7
| 인하대학교 물리학과 | 2023-11-10 | Full range of X Detector Raw during X Scan sweep | 14.6
| Fukoku | 2023-11-10 | Full range of X Detector during X Scan sweep | 72.7
| Exquisite-Automotive Systems | 2023-11-10 | Full range of X Detector during X Scan sweep | 177.3
| 인하대학교 물리학과 | 2023-11-10 | Full range of X Detector during X Scan sweep | 112
| 인하대학교 물리학과 | 2023-11-10 | EFM Phase : For the positive bias | 90
| 인하대학교 물리학과 | 2023-11-10 | EFM Phase : For the negative bias | -90.5
| 인하대학교 물리학과 | 2023-11-10 | Approach position at -3μm | 0.0052
| 인하대학교 물리학과 | 2023-11-10 | Approach position at +3μm | 0.0104
| Fukoku | 2023-11-10 | A-B noise | 9.74
| Exquisite-Automotive Systems | 2023-11-10 | A-B noise | 13.23
| 인하대학교 물리학과 | 2023-11-10 | A-B noise | 18.52
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Z Linearity by Z Scanner position
| 전북대학교 공동실험실습관 | 2023-11-08 | Z Linearity by Z Scanner position
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Z Linearity by 3 different sample height | 0.997
| 전북대학교 공동실험실습관 | 2023-11-08 | Z Linearity by 3 different sample height | 0.989
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Z Height noise | 0.02
| 전북대학교 공동실험실습관 | 2023-11-08 | Z Height noise | 0.021
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Y Detector noise (XY Servo on) | 0.28
| 전북대학교 공동실험실습관 | 2023-11-08 | Y Detector noise (XY Servo on) | 0.26
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Y Detector noise (XY Servo off) | 0.27
| 전북대학교 공동실험실습관 | 2023-11-08 | Y Detector noise (XY Servo off) | 0.26
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Y axis OPM | 0.92
| 전북대학교 공동실험실습관 | 2023-11-08 | Y axis OPM | 1.42
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | X Detector noise (XY Servo on) | 0.29
| 전북대학교 공동실험실습관 | 2023-11-08 | X Detector noise (XY Servo on) | 0.28
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | X Detector noise (XY Servo off) | 0.28
| 전북대학교 공동실험실습관 | 2023-11-08 | X Detector noise (XY Servo off) | 0.27
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | X axis OPM | 0.72
| 전북대학교 공동실험실습관 | 2023-11-08 | X axis OPM | 1.03
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | SLD beam power at the end of the lens | 0.2
| 전북대학교 공동실험실습관 | 2023-11-08 | SLD beam power at the end of the lens | 0.21
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 전북대학교 공동실험실습관 | 2023-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| 전북대학교 공동실험실습관 | 2023-11-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 전북대학교 공동실험실습관 | 2023-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.16
| 전북대학교 공동실험실습관 | 2023-11-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | NCM driving% with AC160TS cantilever | 3.13
| 전북대학교 공동실험실습관 | 2023-11-08 | NCM driving% with AC160TS cantilever | 9.049
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | NCM Amplitude noise | 17.33
| 전북대학교 공동실험실습관 | 2023-11-08 | NCM Amplitude noise | 15.95
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.77
| 전북대학교 공동실험실습관 | 2023-11-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.494
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 20.192
| 전북대학교 공동실험실습관 | 2023-11-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.706
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 159.6
| 전북대학교 공동실험실습관 | 2023-11-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 156.1
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156.9
| 전북대학교 공동실험실습관 | 2023-11-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Full range of Z Height during Z Scan sweep | 15.4
| 전북대학교 공동실험실습관 | 2023-11-08 | Full range of Z Height during Z Scan sweep | 16.7
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | LD beam power at the end of the lens | 3.8
| 전북대학교 공동실험실습관 | 2023-11-08 | LD beam power at the end of the lens | 2.2
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Lateral(C-D) sensitivity | 3.9
| 전북대학교 공동실험실습관 | 2023-11-08 | Lateral(C-D) sensitivity | 39.1
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Lateral(C-D) noise | 0.112
| 전북대학교 공동실험실습관 | 2023-11-08 | Lateral(C-D) noise | 0.148
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Full range of Z Detector during Z Scan sweep | 15.3
| 전북대학교 공동실험실습관 | 2023-11-08 | Full range of Z Detector during Z Scan sweep | 14.1
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Full range of Y Detector Raw during Y Scan sweep | 15
| 전북대학교 공동실험실습관 | 2023-11-08 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Full range of Y Detector during Y Scan sweep | 120.1
| 전북대학교 공동실험실습관 | 2023-11-08 | Full range of Y Detector during Y Scan sweep | 115
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Full range of X Detector Raw during X Scan sweep | 15
| 전북대학교 공동실험실습관 | 2023-11-08 | Full range of X Detector Raw during X Scan sweep | 14.3
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Full range of X Detector during X Scan sweep | 117.4
| 전북대학교 공동실험실습관 | 2023-11-08 | Full range of X Detector during X Scan sweep | 114.6
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | EFM Phase : For the positive bias | 89.8
| 전북대학교 공동실험실습관 | 2023-11-08 | EFM Phase : For the positive bias | 89.3
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | EFM Phase : For the negative bias | -90.3
| 전북대학교 공동실험실습관 | 2023-11-08 | EFM Phase : For the negative bias | -90.6
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Approach position at -3μm | 0.0015
| 전북대학교 공동실험실습관 | 2023-11-08 | Approach position at -3μm | 0.0004
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | Approach position at +3μm | 0.0155
| 전북대학교 공동실험실습관 | 2023-11-08 | Approach position at +3μm | 0.0069
| Shanghai Chisin Instruments (Demo) | 2023-11-08 | A-B noise | 14.43
| 전북대학교 공동실험실습관 | 2023-11-08 | A-B noise | 12.15
| Trumpf Photonics (PSE) | 2023-11-07 | Z Scanner position at -3μm | 0.011
| Trumpf Photonics (PSE) | 2023-11-07 | Z Scanner position at +3μm | 0.0081
| Trumpf Photonics (PSE) | 2023-11-07 | Z Linearity by Z Scanner position
| Trumpf Photonics (PSE) | 2023-11-07 | Z Linearity by 3 different sample height | 1.016
| Trumpf Photonics (PSE) | 2023-11-07 | Z Height noise | 0.027
| Trumpf Photonics (PSE) | 2023-11-07 | Y2 Detector noise (XY Servo on) | 0.242
| Trumpf Photonics (PSE) | 2023-11-07 | Y2 Detector noise (XY Servo off) | 0.234
| Trumpf Photonics (PSE) | 2023-11-07 | Y1 Detector noise (XY Servo on) | 0.231
| Trumpf Photonics (PSE) | 2023-11-07 | Y1 Detector noise (XY Servo off) | 0.224
| Trumpf Photonics (PSE) | 2023-11-07 | Y Scan Stroke | 180.4
| Trumpf Photonics (PSE) | 2023-11-07 | -Y Direction software Limit | -103051
| Trumpf Photonics (PSE) | 2023-11-07 | Y Detector noise (XY Servo on) | 0.171
| Trumpf Photonics (PSE) | 2023-11-07 | Y Detector noise (XY Servo off) | 0.167
| Trumpf Photonics (PSE) | 2023-11-07 | Y axis OPM | 0.9
| Trumpf Photonics (PSE) | 2023-11-07 | X2 Detector noise (XY Servo on) | 0.255
| Trumpf Photonics (PSE) | 2023-11-07 | X2 Detector noise (XY Servo off) | 0.243
| Trumpf Photonics (PSE) | 2023-11-07 | X1 Detector noise (XY Servo on) | 0.236
| Trumpf Photonics (PSE) | 2023-11-07 | X1 Detector noise (XY Servo off) | 0.222
| Trumpf Photonics (PSE) | 2023-11-07 | X Scan Stroke | 177.8
| Trumpf Photonics (PSE) | 2023-11-07 | -X Direction software Limit | -102038
| Trumpf Photonics (PSE) | 2023-11-07 | X Detector noise (XY Servo on) | 0.187
| Trumpf Photonics (PSE) | 2023-11-07 | X Detector noise (XY Servo off) | 0.175
| Trumpf Photonics (PSE) | 2023-11-07 | X axis OPM | 1.36
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.28
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.27
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Trumpf Photonics (PSE) | 2023-11-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| Trumpf Photonics (PSE) | 2023-11-07 | NCM driving% with AC160TS cantilever | 0.769
| Trumpf Photonics (PSE) | 2023-11-07 | NCM Amplitude noise | 15.62
| Trumpf Photonics (PSE) | 2023-11-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.054
| Trumpf Photonics (PSE) | 2023-11-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.601
| Trumpf Photonics (PSE) | 2023-11-07 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 254.7
| Trumpf Photonics (PSE) | 2023-11-07 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 249.9
| Trumpf Photonics (PSE) | 2023-11-07 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 251.3
| Trumpf Photonics (PSE) | 2023-11-07 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 247.9
| Trumpf Photonics (PSE) | 2023-11-07 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Trumpf Photonics (PSE) | 2023-11-07 | Lateral(C-D) sensitivity | 16.3
| Trumpf Photonics (PSE) | 2023-11-07 | Lateral(C-D) noise | 0.144
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of Z Height during Z Scan sweep | 15.8
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of Z Detector during Z Scan sweep | 12.3
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of Y1 Detector Raw during Y Scan sweep | 14.3
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of Y Detector during Y Scan sweep | 179.75
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of X1 Detector Raw during X Scan sweep | 14.25
| Trumpf Photonics (PSE) | 2023-11-07 | Full range of X Detector during X Scan sweep | 177.65
| Trumpf Photonics (PSE) | 2023-11-07 | A-B noise | 16.55
| Universidad de Cuenca | 2023-11-06 | Z Scanner position at -3μm | 0.0058
| Universidad de Cuenca | 2023-11-06 | Z Scanner position at +3μm | 0.0161
| Universidad de Cuenca | 2023-11-06 | Z Linearity by Z Scanner position
| Universidad de Cuenca | 2023-11-06 | Z Linearity by 3 different sample height | 1.001
| Universidad de Cuenca | 2023-11-06 | Z Height noise | 0.029
| Universidad de Cuenca | 2023-11-06 | Y Detector noise (XY Servo on) | 0.26
| Universidad de Cuenca | 2023-11-06 | Y Detector noise (XY Servo off) | 0.26
| Universidad de Cuenca | 2023-11-06 | Y axis OPM | 1.17
| Universidad de Cuenca | 2023-11-06 | X Detector noise (XY Servo on) | 0.27
| Universidad de Cuenca | 2023-11-06 | X Detector noise (XY Servo off) | 0.27
| Universidad de Cuenca | 2023-11-06 | X axis OPM | 0.81
| Universidad de Cuenca | 2023-11-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Universidad de Cuenca | 2023-11-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Universidad de Cuenca | 2023-11-06 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Universidad de Cuenca | 2023-11-06 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| Universidad de Cuenca | 2023-11-06 | NCM driving% with AC160TS cantilever | 1.12
| Universidad de Cuenca | 2023-11-06 | NCM Amplitude noise | 17.42
| Universidad de Cuenca | 2023-11-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.727
| Universidad de Cuenca | 2023-11-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.352
| Universidad de Cuenca | 2023-11-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 166.5
| Universidad de Cuenca | 2023-11-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 161
| Universidad de Cuenca | 2023-11-06 | Lateral(C-D) sensitivity | 15.833
| Universidad de Cuenca | 2023-11-06 | Lateral(C-D) noise | 0.118
| Universidad de Cuenca | 2023-11-06 | Full range of Z Height during Z Scan sweep | 16.7
| Universidad de Cuenca | 2023-11-06 | Full range of Z Detector during Z Scan sweep | 12.2
| Universidad de Cuenca | 2023-11-06 | Full range of Y Detector Raw during Y Scan sweep | 13.7
| Universidad de Cuenca | 2023-11-06 | Full range of Y Detector during Y Scan sweep | 113.8
| Universidad de Cuenca | 2023-11-06 | Full range of X Detector Raw during X Scan sweep | 14.5
| Universidad de Cuenca | 2023-11-06 | Full range of X Detector during X Scan sweep | 116.6
| Universidad de Cuenca | 2023-11-06 | A-B noise | 10.61
| Longzhong Laboratory | 2023-11-03 | Z Scanner position at -3μm | 0.0039
| Longzhong Laboratory | 2023-11-03 | Z Scanner position at +3μm | 0.0109
| Longzhong Laboratory | 2023-11-03 | Z Linearity by Z Scanner position
| Longzhong Laboratory | 2023-11-03 | Z Linearity by 3 different sample height | 0.994
| Longzhong Laboratory | 2023-11-03 | Z Height noise | 0.027
| Longzhong Laboratory | 2023-11-03 | Y2 Detector noise (XY Servo on) | 0.27
| Longzhong Laboratory | 2023-11-03 | Y2 Detector noise (XY Servo off) | 0.26
| Longzhong Laboratory | 2023-11-03 | Y1 Detector noise (XY Servo on) | 0.23
| Longzhong Laboratory | 2023-11-03 | Y1 Detector noise (XY Servo off) | 0.23
| Longzhong Laboratory | 2023-11-03 | Y Scan Stroke | 174.5
| Longzhong Laboratory | 2023-11-03 | Y Detector noise (XY Servo on) | 0.18
| Longzhong Laboratory | 2023-11-03 | Y Detector noise (XY Servo off) | 0.18
| Longzhong Laboratory | 2023-11-03 | Y axis OPM | 1.13
| Longzhong Laboratory | 2023-11-03 | X2 Detector noise (XY Servo on) | 0.25
| Longzhong Laboratory | 2023-11-03 | X2 Detector noise (XY Servo off) | 0.25
| Longzhong Laboratory | 2023-11-03 | X1 Detector noise (XY Servo on) | 0.28
| Longzhong Laboratory | 2023-11-03 | X1 Detector noise (XY Servo off) | 0.28
| Longzhong Laboratory | 2023-11-03 | X Scan Stroke | 183.4
| Longzhong Laboratory | 2023-11-03 | -X Direction software Limit | -102124
| Longzhong Laboratory | 2023-11-03 | -X Direction software Limit | -103291
| Longzhong Laboratory | 2023-11-03 | X Detector noise (XY Servo on) | 0.2
| Longzhong Laboratory | 2023-11-03 | X Detector noise (XY Servo off) | 0.19
| Longzhong Laboratory | 2023-11-03 | X axis OPM | 1.43
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.38
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.37
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.35
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.35
| Longzhong Laboratory | 2023-11-03 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| Longzhong Laboratory | 2023-11-03 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Longzhong Laboratory | 2023-11-03 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| Longzhong Laboratory | 2023-11-03 | NCM driving% with AC160TS cantilever | 0.31
| Longzhong Laboratory | 2023-11-03 | NCM Amplitude noise | 13.6
| Longzhong Laboratory | 2023-11-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.432
| Longzhong Laboratory | 2023-11-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.355
| Longzhong Laboratory | 2023-11-03 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 251.9
| Longzhong Laboratory | 2023-11-03 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 246.6
| Longzhong Laboratory | 2023-11-03 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 261
| Longzhong Laboratory | 2023-11-03 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 255.7
| Longzhong Laboratory | 2023-11-03 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Longzhong Laboratory | 2023-11-03 | Lateral(C-D) sensitivity | 19.8
| Longzhong Laboratory | 2023-11-03 | Lateral(C-D) noise | 0.152
| Longzhong Laboratory | 2023-11-03 | Full range of Z Height during Z Scan sweep | 16.9
| Longzhong Laboratory | 2023-11-03 | Full range of Z Detector during Z Scan sweep | 13.3
| Longzhong Laboratory | 2023-11-03 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Longzhong Laboratory | 2023-11-03 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Longzhong Laboratory | 2023-11-03 | Full range of Y Detector during Y Scan sweep | 178.1
| Longzhong Laboratory | 2023-11-03 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Longzhong Laboratory | 2023-11-03 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| Longzhong Laboratory | 2023-11-03 | Full range of X Detector during X Scan sweep | 183.6
| Longzhong Laboratory | 2023-11-03 | A-B noise | 6.7
| CXT #4 | 2023-11-01 | Z Linearity by Z scanner position | 0.506
| CXT #4 | 2023-11-01 | Z Linearity by Z scanner position | 0.28131
| CXT #4 | 2023-11-01 | Z Linearity by 3 different sample height | 1.004
| CXT #4 | 2023-11-01 | Z Height noise | 0.029
| CXT #4 | 2023-11-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.1
| CXT #4 | 2023-11-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.9
| CXT #4 | 2023-11-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1
| CXT #4 | 2023-11-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11
| CXT #4 | 2023-11-01 | Z Drift (Short Traveling) | 0.29
| CXT #4 | 2023-11-01 | Y2 Detector noise (XY Servo on) | 0.272
| CXT #4 | 2023-11-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 267.3
| CXT #4 | 2023-11-01 | Y1 Detector noise (XY Servo on) | 0.277
| CXT #4 | 2023-11-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 272.3
| CXT #4 | 2023-11-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 197.2
| CXT #4 | 2023-11-01 | Y Ringing(Rpv) | 0.214
| CXT #4 | 2023-11-01 | Y Non-linearity error | -3.01
| CXT #4 | 2023-11-01 | Y Detector noise (XY Servo on) | 0.204
| CXT #4 | 2023-11-01 | Y axis, Repeatability @1st position | 0.07
| CXT #4 | 2023-11-01 | Y axis, Repeatability @ 4th position | 0.07
| CXT #4 | 2023-11-01 | Y axis, Repeatability @ 3rd position | 0.08
| CXT #4 | 2023-11-01 | Y axis, Repeatability @ 2nd position | 0.08
| CXT #4 | 2023-11-01 | Y axis, Accuracy @1st position | 0.78
| CXT #4 | 2023-11-01 | Y axis, Accuracy @ 4th position | 0.72
| CXT #4 | 2023-11-01 | Y axis, Accuracy @ 3rd position | 0.75
| CXT #4 | 2023-11-01 | Y axis, Accuracy @ 2nd position | 0.79
| CXT #4 | 2023-11-01 | Y axis OPM | 0.682
| CXT #4 | 2023-11-01 | XY Orthogonality | 89.9
| CXT #4 | 2023-11-01 | X2 Detector noise (XY Servo on) | 0.295
| CXT #4 | 2023-11-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 278.3
| CXT #4 | 2023-11-01 | X1 Detector noise (XY Servo on) | 0.295
| CXT #4 | 2023-11-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 274.6
| CXT #4 | 2023-11-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.7
| CXT #4 | 2023-11-01 | X Ringing(Rpv) | 0.307
| CXT #4 | 2023-11-01 | X Non-linearity error | -2.75
| CXT #4 | 2023-11-01 | X Detector noise (XY Servo on) | 0.23
| CXT #4 | 2023-11-01 | X axis, Repeatability @1st position | 0.02
| CXT #4 | 2023-11-01 | X axis, Repeatability @ 4th position | 0.03
| CXT #4 | 2023-11-01 | X axis, Repeatability @ 3rd position | 0.03
| CXT #4 | 2023-11-01 | X axis, Repeatability @ 2nd position | 0.03
| CXT #4 | 2023-11-01 | X axis, Accuracy @1st position | 0.12
| CXT #4 | 2023-11-01 | X axis, Accuracy @ 4th position | 0.26
| CXT #4 | 2023-11-01 | X axis, Accuracy @ 3rd position | 0.18
| CXT #4 | 2023-11-01 | X axis, Accuracy @ 2nd position | 0.2
| CXT #4 | 2023-11-01 | X axis OPM | 0.814
| CXT #4 | 2023-11-01 | Under 50nm Step Height Sample (10 Repeats) | 18.09
| CXT #4 | 2023-11-01 | Torque Profile (50mm, 20um/s) | 2.48
| CXT #4 | 2023-11-01 | Torque Profile (50mm, 200um/s) | 2.28
| CXT #4 | 2023-11-01 | StrokeYMin | -153666
| CXT #4 | 2023-11-01 | StrokeYMax | 150152.6
| CXT #4 | 2023-11-01 | StrokeY | 303818.1
| CXT #4 | 2023-11-01 | StrokeXMin | -203978
| CXT #4 | 2023-11-01 | StrokeXMax | 203758.1
| CXT #4 | 2023-11-01 | StrokeX | 407735.6
| CXT #4 | 2023-11-01 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.72
| CXT #4 | 2023-11-01 | Roll motion to X+ full scan | 20
| CXT #4 | 2023-11-01 | RMS of ZeroScan | 0.057
| CXT #4 | 2023-11-01 | Profiler noise @200um | 0.21
| CXT #4 | 2023-11-01 | Position repeatability(10repeats, Position Based) | 0.37
| CXT #4 | 2023-11-01 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.0001
| CXT #4 | 2023-11-01 | OpticalAxisShiftPerFocusY | -0.0026
| CXT #4 | 2023-11-01 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| CXT #4 | 2023-11-01 | OpticalAxisShiftPerFocusX | -0.0027
| CXT #4 | 2023-11-01 | OPM Repeatability @5mm | 1.6
| CXT #4 | 2023-11-01 | OPM Repeatability @25mm | 7.72
| CXT #4 | 2023-11-01 | OPM Repeatability @1mm | 0.84
| CXT #4 | 2023-11-01 | OPM Repeatability @10mm | 4.97
| CXT #4 | 2023-11-01 | NCM driving% with AC160TS cantilever | 19.66
| CXT #4 | 2023-11-01 | NCM driving% with AC160TS cantilever | 31.22
| CXT #4 | 2023-11-01 | Min. Vibration Transmissbility (~100Hz) | 0.045
| CXT #4 | 2023-11-01 | Max OPM @5mm | 29.31
| CXT #4 | 2023-11-01 | Max OPM @25mm | 136.65
| CXT #4 | 2023-11-01 | Max OPM @1mm | 8.48
| CXT #4 | 2023-11-01 | Max OPM @10mm | 50.3
| CXT #4 | 2023-11-01 | LowerLimit | 4737
| CXT #4 | 2023-11-01 | Lateral(C-D) sensitivity | 22.6
| CXT #4 | 2023-11-01 | Lateral(C-D) noise | 0.248
| CXT #4 | 2023-11-01 | Image offset after moving from -Y to +Y (Y scan) | 142.82
| CXT #4 | 2023-11-01 | Image offset after moving from -Y to +Y (X scan) | 35.16
| CXT #4 | 2023-11-01 | Image offset after moving from -X to +X (Y scan) | 180.91
| CXT #4 | 2023-11-01 | Image offset after moving from -X to +X (X scan) | 155.27
| CXT #4 | 2023-11-01 | Image offset after moving from +Y to -Y (Y scan) | 52.73
| CXT #4 | 2023-11-01 | Image offset after moving from +Y to -Y (X scan) | 98.88
| CXT #4 | 2023-11-01 | Image offset after moving from +X to -X (Y scan) | 111.33
| CXT #4 | 2023-11-01 | Image offset after moving from +X to -X (X scan) | 109.13
| CXT #4 | 2023-11-01 | Full range of Z Height during Z Scan sweep | 17
| CXT #4 | 2023-11-01 | Full range of Z Height during Z Scan sweep | 47.1
| CXT #4 | 2023-11-01 | Full range of Z Detector during Z Scan sweep | 13
| CXT #4 | 2023-11-01 | Full range of Z Detector during Z Scan sweep | 14.1
| CXT #4 | 2023-11-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.21
| CXT #4 | 2023-11-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.26
| CXT #4 | 2023-11-01 | Full range of X2 Detector Raw during X Scan sweep | 13.81
| CXT #4 | 2023-11-01 | Full range of X1 Detector Raw during X Scan sweep | 13.66
| CXT #4 | 2023-11-01 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| CXT #4 | 2023-11-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| CXT #4 | 2023-11-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.59
| CXT #4 | 2023-11-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.67
| CXT #4 | 2023-11-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| CXT #4 | 2023-11-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| CXT #4 | 2023-11-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| CXT #4 | 2023-11-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22
| CXT #4 | 2023-11-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| CXT #4 | 2023-11-01 | 940nm Step Height Sample (10 Repeats) | 183.29
| CXT #4 | 2023-11-01 | 50~200nm Step Height Sample (10 Repeats) | 176.02
| CXT #4 | 2023-11-01 | 200~500nm Step Height Sample (10 Repeats) | 463.15
| CXT #4 | 2023-11-01 | 200~500nm Step Height Sample (10 Repeats) | 18.08
| CXT #4 | 2023-11-01 | * 1800nm Step Height Sample (10 Repeats) | 462.75
| CXT #4 | 2023-11-01 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.4
| CXT #4 | 2023-11-01 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 12.4
| CXT #4 | 2023-11-01 | %GRR w/ roughness of non-patterened wafer | 13.2
| AI IMAGE TECHNOLOGY | 2023-10-31 | Z Scanner position at -3μm | 0.0049
| Qingfang (Qingdao) Technology | 2023-10-31 | Z Scanner position at -3μm | 0.009
| AI IMAGE TECHNOLOGY | 2023-10-31 | Z Scanner position at +3μm | 0.0125
| Qingfang (Qingdao) Technology | 2023-10-31 | Z Scanner position at +3μm | 0.0024
| AI IMAGE TECHNOLOGY | 2023-10-31 | Z Linearity by Z Scanner position
| Qingfang (Qingdao) Technology | 2023-10-31 | Z Linearity by Z Scanner position
| AI IMAGE TECHNOLOGY | 2023-10-31 | Z Linearity by 3 different sample height | 1.003
| Qingfang (Qingdao) Technology | 2023-10-31 | Z Linearity by 3 different sample height | 1.001
| AI IMAGE TECHNOLOGY | 2023-10-31 | Z Height noise | 0.025
| Qingfang (Qingdao) Technology | 2023-10-31 | Z Height noise | 0.027
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y2 Detector noise (XY Servo on) | 0.22
| Qingfang (Qingdao) Technology | 2023-10-31 | Y2 Detector noise (XY Servo on) | 0.2
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y2 Detector noise (XY Servo off) | 0.22
| Qingfang (Qingdao) Technology | 2023-10-31 | Y2 Detector noise (XY Servo off) | 0.21
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y1 Detector noise (XY Servo on) | 0.22
| Qingfang (Qingdao) Technology | 2023-10-31 | Y1 Detector noise (XY Servo on) | 0.22
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y1 Detector noise (XY Servo off) | 0.22
| Qingfang (Qingdao) Technology | 2023-10-31 | Y1 Detector noise (XY Servo off) | 0.22
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y Scan Stroke | 172.2
| Qingfang (Qingdao) Technology | 2023-10-31 | Y Scan Stroke | 175.23
| AI IMAGE TECHNOLOGY | 2023-10-31 | -Y Direction software Limit | -103241
| Qingfang (Qingdao) Technology | 2023-10-31 | -Y Direction software Limit | -75556.2
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y Detector noise (XY Servo on) | 0.16
| Qingfang (Qingdao) Technology | 2023-10-31 | Y Detector noise (XY Servo on) | 0.15
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y Detector noise (XY Servo off) | 0.16
| Qingfang (Qingdao) Technology | 2023-10-31 | Y Detector noise (XY Servo off) | 0.16
| AI IMAGE TECHNOLOGY | 2023-10-31 | Y axis OPM | 0.62
| Qingfang (Qingdao) Technology | 2023-10-31 | Y axis OPM | 1.1
| AI IMAGE TECHNOLOGY | 2023-10-31 | X2 Detector noise (XY Servo on) | 0.23
| Qingfang (Qingdao) Technology | 2023-10-31 | X2 Detector noise (XY Servo on) | 0.22
| AI IMAGE TECHNOLOGY | 2023-10-31 | X2 Detector noise (XY Servo off) | 0.23
| Qingfang (Qingdao) Technology | 2023-10-31 | X2 Detector noise (XY Servo off) | 0.22
| AI IMAGE TECHNOLOGY | 2023-10-31 | X1 Detector noise (XY Servo on) | 0.23
| Qingfang (Qingdao) Technology | 2023-10-31 | X1 Detector noise (XY Servo on) | 0.22
| AI IMAGE TECHNOLOGY | 2023-10-31 | X1 Detector noise (XY Servo off) | 0.23
| Qingfang (Qingdao) Technology | 2023-10-31 | X1 Detector noise (XY Servo off) | 0.22
| AI IMAGE TECHNOLOGY | 2023-10-31 | X Scan Stroke | 171.3
| Qingfang (Qingdao) Technology | 2023-10-31 | X Scan Stroke | 173.26
| AI IMAGE TECHNOLOGY | 2023-10-31 | -X Direction software Limit | -101697
| Qingfang (Qingdao) Technology | 2023-10-31 | -X Direction software Limit | -75306.2
| AI IMAGE TECHNOLOGY | 2023-10-31 | X Detector noise (XY Servo on) | 0.17
| Qingfang (Qingdao) Technology | 2023-10-31 | X Detector noise (XY Servo on) | 0.16
| AI IMAGE TECHNOLOGY | 2023-10-31 | X Detector noise (XY Servo off) | 0.16
| Qingfang (Qingdao) Technology | 2023-10-31 | X Detector noise (XY Servo off) | 0.16
| AI IMAGE TECHNOLOGY | 2023-10-31 | X axis OPM | 1.25
| Qingfang (Qingdao) Technology | 2023-10-31 | X axis OPM | 0.72
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.38
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.37
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.39
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.32
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.32
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.31
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.33
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| AI IMAGE TECHNOLOGY | 2023-10-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Qingfang (Qingdao) Technology | 2023-10-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| AI IMAGE TECHNOLOGY | 2023-10-31 | NCM driving% with AC160TS cantilever | 0.25
| Qingfang (Qingdao) Technology | 2023-10-31 | NCM driving% with AC160TS cantilever | 1.37
| AI IMAGE TECHNOLOGY | 2023-10-31 | NCM Amplitude noise | 17.8
| Qingfang (Qingdao) Technology | 2023-10-31 | NCM Amplitude noise | 14.62
| AI IMAGE TECHNOLOGY | 2023-10-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.225
| Qingfang (Qingdao) Technology | 2023-10-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.415
| AI IMAGE TECHNOLOGY | 2023-10-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.57
| Qingfang (Qingdao) Technology | 2023-10-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.949
| AI IMAGE TECHNOLOGY | 2023-10-31 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 244.7
| Qingfang (Qingdao) Technology | 2023-10-31 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 240
| AI IMAGE TECHNOLOGY | 2023-10-31 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 236.8
| Qingfang (Qingdao) Technology | 2023-10-31 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 246.7
| AI IMAGE TECHNOLOGY | 2023-10-31 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 240.8
| Qingfang (Qingdao) Technology | 2023-10-31 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 240.5
| AI IMAGE TECHNOLOGY | 2023-10-31 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 235.8
| Qingfang (Qingdao) Technology | 2023-10-31 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 245.9
| AI IMAGE TECHNOLOGY | 2023-10-31 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Qingfang (Qingdao) Technology | 2023-10-31 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| AI IMAGE TECHNOLOGY | 2023-10-31 | Lateral(C-D) sensitivity | 12.8
| Qingfang (Qingdao) Technology | 2023-10-31 | Lateral(C-D) sensitivity | 9.1
| AI IMAGE TECHNOLOGY | 2023-10-31 | Lateral(C-D) noise | 0.123
| Qingfang (Qingdao) Technology | 2023-10-31 | Lateral(C-D) noise | 0.148
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of Z Height during Z Scan sweep | 16
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of Z Height during Z Scan sweep | 16.4
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of Z Detector during Z Scan sweep | 13.5
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of Z Detector during Z Scan sweep | 13.1
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of Y Detector during Y Scan sweep | 173.5
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of Y Detector during Y Scan sweep | 174.9
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| AI IMAGE TECHNOLOGY | 2023-10-31 | Full range of X Detector during X Scan sweep | 173.9
| Qingfang (Qingdao) Technology | 2023-10-31 | Full range of X Detector during X Scan sweep | 173.1
| AI IMAGE TECHNOLOGY | 2023-10-31 | A-B noise | 10.01
| Qingfang (Qingdao) Technology | 2023-10-31 | A-B noise | 9.64
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Z Linearity by Z Scanner position
| University of Aveiro (PSE) | 2023-10-30 | Z Linearity by Z Scanner position
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Z Linearity by 3 different sample height | 1.003
| University of Aveiro (PSE) | 2023-10-30 | Z Linearity by 3 different sample height | 1.002
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Z Height noise | 0.019
| University of Aveiro (PSE) | 2023-10-30 | Z Height noise | 0.022
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Y Detector noise (XY Servo on) | 0.29
| University of Aveiro (PSE) | 2023-10-30 | Y Detector noise (XY Servo on) | 0.26
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Y Detector noise (XY Servo off) | 0.24
| University of Aveiro (PSE) | 2023-10-30 | Y Detector noise (XY Servo off) | 0.25
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Y axis OPM | 1.04
| University of Aveiro (PSE) | 2023-10-30 | Y axis OPM | 1.33
| 한국교통대학교 화학산업연구소 | 2023-10-30 | X Detector noise (XY Servo on) | 0.31
| University of Aveiro (PSE) | 2023-10-30 | X Detector noise (XY Servo on) | 0.26
| 한국교통대학교 화학산업연구소 | 2023-10-30 | X Detector noise (XY Servo off) | 0.25
| University of Aveiro (PSE) | 2023-10-30 | X Detector noise (XY Servo off) | 0.25
| 한국교통대학교 화학산업연구소 | 2023-10-30 | X axis OPM | 1.67
| University of Aveiro (PSE) | 2023-10-30 | X axis OPM | 1.1
| 한국교통대학교 화학산업연구소 | 2023-10-30 | SLD beam power at the end of the lens | 0.23
| University of Aveiro (PSE) | 2023-10-30 | SLD beam power at the end of the lens | 0.22
| 한국교통대학교 화학산업연구소 | 2023-10-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| University of Aveiro (PSE) | 2023-10-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한국교통대학교 화학산업연구소 | 2023-10-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| University of Aveiro (PSE) | 2023-10-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 한국교통대학교 화학산업연구소 | 2023-10-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| University of Aveiro (PSE) | 2023-10-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 한국교통대학교 화학산업연구소 | 2023-10-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| University of Aveiro (PSE) | 2023-10-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| 한국교통대학교 화학산업연구소 | 2023-10-30 | NCM driving% with AC160TS cantilever | 1.14
| University of Aveiro (PSE) | 2023-10-30 | NCM driving% with AC160TS cantilever | 12.24
| 한국교통대학교 화학산업연구소 | 2023-10-30 | NCM Amplitude noise | 17.35
| University of Aveiro (PSE) | 2023-10-30 | NCM Amplitude noise | 17.46
| 한국교통대학교 화학산업연구소 | 2023-10-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.866
| University of Aveiro (PSE) | 2023-10-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.587
| 한국교통대학교 화학산업연구소 | 2023-10-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.496
| University of Aveiro (PSE) | 2023-10-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.952
| 한국교통대학교 화학산업연구소 | 2023-10-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 155.5
| University of Aveiro (PSE) | 2023-10-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 158.4
| 한국교통대학교 화학산업연구소 | 2023-10-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 153.8
| University of Aveiro (PSE) | 2023-10-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.9
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Full range of Z Height during Z Scan sweep | 15.4
| University of Aveiro (PSE) | 2023-10-30 | Full range of Z Height during Z Scan sweep | 16.5
| 한국교통대학교 화학산업연구소 | 2023-10-30 | LD beam power at the end of the lens | 1.54
| University of Aveiro (PSE) | 2023-10-30 | LD beam power at the end of the lens | 1.88
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Lateral(C-D) sensitivity | 6.8
| University of Aveiro (PSE) | 2023-10-30 | Lateral(C-D) sensitivity | 2.8
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Lateral(C-D) noise | 0.12
| University of Aveiro (PSE) | 2023-10-30 | Lateral(C-D) noise | 0.132
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Full range of Z Detector during Z Scan sweep | 13.2
| University of Aveiro (PSE) | 2023-10-30 | Full range of Z Detector during Z Scan sweep | 14.4
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| University of Aveiro (PSE) | 2023-10-30 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Full range of Y Detector during Y Scan sweep | 117.2
| University of Aveiro (PSE) | 2023-10-30 | Full range of Y Detector during Y Scan sweep | 119.7
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Full range of X Detector Raw during X Scan sweep | 15.1
| University of Aveiro (PSE) | 2023-10-30 | Full range of X Detector Raw during X Scan sweep | 15.1
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Full range of X Detector during X Scan sweep | 115.8
| University of Aveiro (PSE) | 2023-10-30 | Full range of X Detector during X Scan sweep | 118.9
| 한국교통대학교 화학산업연구소 | 2023-10-30 | EFM Phase : For the positive bias | 89.5
| University of Aveiro (PSE) | 2023-10-30 | EFM Phase : For the positive bias | 90.3
| 한국교통대학교 화학산업연구소 | 2023-10-30 | EFM Phase : For the negative bias | -89.1
| University of Aveiro (PSE) | 2023-10-30 | EFM Phase : For the negative bias | -89.9
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Approach position at -3μm | 0.0108
| University of Aveiro (PSE) | 2023-10-30 | Approach position at -3μm | 0.004
| 한국교통대학교 화학산업연구소 | 2023-10-30 | Approach position at +3μm | 0.0172
| University of Aveiro (PSE) | 2023-10-30 | Approach position at +3μm | 0.0002
| 한국교통대학교 화학산업연구소 | 2023-10-30 | A-B noise | 5.2
| University of Aveiro (PSE) | 2023-10-30 | A-B noise | 6.73
| 경상국립대학교 화학과 | 2023-10-27 | Z Scanner position at -3μm | 0.0001
| 부경대학교 고분자공학과 | 2023-10-27 | Z Scanner position at -3μm | 0.0044
| 경상국립대학교 화학과 | 2023-10-27 | Z Scanner position at +3μm | 0.0071
| 부경대학교 고분자공학과 | 2023-10-27 | Z Scanner position at +3μm | 0.0126
| 경상국립대학교 화학과 | 2023-10-27 | Z Linearity by Z Scanner position
| 부경대학교 고분자공학과 | 2023-10-27 | Z Linearity by Z Scanner position
| 경상국립대학교 화학과 | 2023-10-27 | Z Linearity by 3 different sample height | 0.996
| 부경대학교 고분자공학과 | 2023-10-27 | Z Linearity by 3 different sample height | 0.998
| 경상국립대학교 화학과 | 2023-10-27 | Z Height noise | 0.026
| 부경대학교 고분자공학과 | 2023-10-27 | Z Height noise | 0.026
| 경상국립대학교 화학과 | 2023-10-27 | Y Scan Stroke | 55
| 부경대학교 고분자공학과 | 2023-10-27 | Y Scan Stroke | 55
| 경상국립대학교 화학과 | 2023-10-27 | Y Detector noise (XY Servo on) | 0.35
| 부경대학교 고분자공학과 | 2023-10-27 | Y Detector noise (XY Servo on) | 0.33
| 경상국립대학교 화학과 | 2023-10-27 | Y Detector noise (XY Servo off) | 0.35
| 부경대학교 고분자공학과 | 2023-10-27 | Y Detector noise (XY Servo off) | 0.32
| 경상국립대학교 화학과 | 2023-10-27 | Y axis OPM | 0.59
| 부경대학교 고분자공학과 | 2023-10-27 | Y axis OPM | 0.88
| 경상국립대학교 화학과 | 2023-10-27 | X Scan Stroke | 55
| 부경대학교 고분자공학과 | 2023-10-27 | X Scan Stroke | 55
| 경상국립대학교 화학과 | 2023-10-27 | X Detector noise (XY Servo on) | 0.37
| 부경대학교 고분자공학과 | 2023-10-27 | X Detector noise (XY Servo on) | 0.33
| 경상국립대학교 화학과 | 2023-10-27 | X Detector noise (XY Servo off) | 0.38
| 부경대학교 고분자공학과 | 2023-10-27 | X Detector noise (XY Servo off) | 0.33
| 경상국립대학교 화학과 | 2023-10-27 | X axis OPM | 0.49
| 부경대학교 고분자공학과 | 2023-10-27 | X axis OPM | 0.92
| 경상국립대학교 화학과 | 2023-10-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 부경대학교 고분자공학과 | 2023-10-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 경상국립대학교 화학과 | 2023-10-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| 부경대학교 고분자공학과 | 2023-10-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 경상국립대학교 화학과 | 2023-10-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 부경대학교 고분자공학과 | 2023-10-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| 경상국립대학교 화학과 | 2023-10-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| 부경대학교 고분자공학과 | 2023-10-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| 경상국립대학교 화학과 | 2023-10-27 | NCM driving% with AC160TS cantilever | 1.77
| 부경대학교 고분자공학과 | 2023-10-27 | NCM driving% with AC160TS cantilever | 1.7
| 경상국립대학교 화학과 | 2023-10-27 | NCM Amplitude noise | 9.14
| 부경대학교 고분자공학과 | 2023-10-27 | NCM Amplitude noise | 11.69
| 경상국립대학교 화학과 | 2023-10-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.701
| 부경대학교 고분자공학과 | 2023-10-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.787
| 경상국립대학교 화학과 | 2023-10-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.405
| 부경대학교 고분자공학과 | 2023-10-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.1
| 경상국립대학교 화학과 | 2023-10-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.6
| 부경대학교 고분자공학과 | 2023-10-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 94.6
| 경상국립대학교 화학과 | 2023-10-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.5
| 부경대학교 고분자공학과 | 2023-10-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.2
| 경상국립대학교 화학과 | 2023-10-27 | Lateral(C-D) sensitivity | 17.8
| 부경대학교 고분자공학과 | 2023-10-27 | Lateral(C-D) sensitivity | 26.8
| 경상국립대학교 화학과 | 2023-10-27 | Lateral(C-D) noise | 0.078
| 부경대학교 고분자공학과 | 2023-10-27 | Lateral(C-D) noise | 0.043
| 경상국립대학교 화학과 | 2023-10-27 | Full range of Z Height during Z Scan sweep | 16.9
| 부경대학교 고분자공학과 | 2023-10-27 | Full range of Z Height during Z Scan sweep | 16.7
| 경상국립대학교 화학과 | 2023-10-27 | Full range of Z Detector during Z Scan sweep | 16.5
| 부경대학교 고분자공학과 | 2023-10-27 | Full range of Z Detector during Z Scan sweep | 13.3
| 경상국립대학교 화학과 | 2023-10-27 | Full range of Y Detector Raw during Y Scan sweep | 15.4
| 부경대학교 고분자공학과 | 2023-10-27 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| 경상국립대학교 화학과 | 2023-10-27 | Full range of Y Detector during Y Scan sweep | 71.2
| 부경대학교 고분자공학과 | 2023-10-27 | Full range of Y Detector during Y Scan sweep | 69.7
| 경상국립대학교 화학과 | 2023-10-27 | Full range of X Detector Raw during X Scan sweep | 14.5
| 부경대학교 고분자공학과 | 2023-10-27 | Full range of X Detector Raw during X Scan sweep | 14.8
| 경상국립대학교 화학과 | 2023-10-27 | Full range of X Detector during X Scan sweep | 70.3
| 부경대학교 고분자공학과 | 2023-10-27 | Full range of X Detector during X Scan sweep | 70.5
| 경상국립대학교 화학과 | 2023-10-27 | A-B noise | 11.93
| 부경대학교 고분자공학과 | 2023-10-27 | A-B noise | 14.49
| Samsung P3D | 2023-10-26 | Z Linearity by Z scanner position | 1.13
| Samsung P3D | 2023-10-26 | Z Linearity by 5 different sample height | 1.006
| Samsung P3D | 2023-10-26 | Z Height noise | 0.038
| Samsung P3D | 2023-10-26 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.8
| Samsung P3D | 2023-10-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.8
| Samsung P3D | 2023-10-26 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.3
| Samsung P3D | 2023-10-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3
| Samsung P3D | 2023-10-26 | Z Drift (Short Traveling) | 0.16
| Samsung P3D | 2023-10-26 | Y2 Detector noise (XY Servo on) | 0.76
| Samsung P3D | 2023-10-26 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 272.4
| Samsung P3D | 2023-10-26 | Y1 Detector noise (XY Servo on) | 0.75
| Samsung P3D | 2023-10-26 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 264.1
| Samsung P3D | 2023-10-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 189.8
| Samsung P3D | 2023-10-26 | Y Ringing(Rpv) | 0.41
| Samsung P3D | 2023-10-26 | Y Non-linearity error | 3.59
| Samsung P3D | 2023-10-26 | Y Detector noise (XY Servo on) | 0.71
| Samsung P3D | 2023-10-26 | Y axis, Repeatability @1st position | 0.04
| Samsung P3D | 2023-10-26 | Y axis, Repeatability @ 4th position | 0.04
| Samsung P3D | 2023-10-26 | Y axis, Repeatability @ 3rd position | 0.03
| Samsung P3D | 2023-10-26 | Y axis, Repeatability @ 2nd position | 0.03
| Samsung P3D | 2023-10-26 | Y axis, Accuracy @1st position | 0.25
| Samsung P3D | 2023-10-26 | Y axis, Accuracy @ 4th position | 0.35
| Samsung P3D | 2023-10-26 | Y axis, Accuracy @ 3rd position | 0.29
| Samsung P3D | 2023-10-26 | Y axis, Accuracy @ 2nd position | 0.21
| Samsung P3D | 2023-10-26 | Y axis OPM | 2.21
| Samsung P3D | 2023-10-26 | XY Orthogonality | 90
| Samsung P3D | 2023-10-26 | X2 Detector noise (XY Servo on) | 0.4
| Samsung P3D | 2023-10-26 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 285
| Samsung P3D | 2023-10-26 | X1 Detector noise (XY Servo on) | 0.39
| Samsung P3D | 2023-10-26 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 278.5
| Samsung P3D | 2023-10-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.9
| Samsung P3D | 2023-10-26 | X Ringing(Rpv) | 0.22
| Samsung P3D | 2023-10-26 | X Non-linearity error | 2.78
| Samsung P3D | 2023-10-26 | X Detector noise (XY Servo on) | 0.33
| Samsung P3D | 2023-10-26 | X axis, Repeatability @1st position | 0.03
| Samsung P3D | 2023-10-26 | X axis, Repeatability @ 4th position | 0.04
| Samsung P3D | 2023-10-26 | X axis, Repeatability @ 3rd position | 0.03
| Samsung P3D | 2023-10-26 | X axis, Repeatability @ 2nd position | 0.03
| Samsung P3D | 2023-10-26 | X axis, Accuracy @1st position | 0.14
| Samsung P3D | 2023-10-26 | X axis, Accuracy @ 4th position | 0.21
| Samsung P3D | 2023-10-26 | X axis, Accuracy @ 3rd position | 0.15
| Samsung P3D | 2023-10-26 | X axis, Accuracy @ 2nd position | 0.21
| Samsung P3D | 2023-10-26 | X axis OPM | 1.34
| Samsung P3D | 2023-10-26 | Under 50nm Step Height Sample (10 Repeats) | 18.14
| Samsung P3D | 2023-10-26 | Torque Profile (50mm, 20um/s) | 4.634
| Samsung P3D | 2023-10-26 | Torque Profile (50mm, 200um/s) | 3.853
| Samsung P3D | 2023-10-26 | StrokeYMin | -154526
| Samsung P3D | 2023-10-26 | StrokeYMax | 151682
| Samsung P3D | 2023-10-26 | StrokeY | 306208
| Samsung P3D | 2023-10-26 | StrokeXMin | -204406
| Samsung P3D | 2023-10-26 | StrokeXMax | 203359
| Samsung P3D | 2023-10-26 | StrokeX | 407765
| Samsung P3D | 2023-10-26 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.83
| Samsung P3D | 2023-10-26 | Roll motion to X+ full scan | 25.4
| Samsung P3D | 2023-10-26 | RMS of ZeroScan | 0.03
| Samsung P3D | 2023-10-26 | Profiler noise @200um | 0.15
| Samsung P3D | 2023-10-26 | Position repeatability(10repeats, Position Based) | 0.52
| Samsung P3D | 2023-10-26 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00012
| Samsung P3D | 2023-10-26 | OpticalAxisShiftPerFocusY | -0.0004
| Samsung P3D | 2023-10-26 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00024
| Samsung P3D | 2023-10-26 | OpticalAxisShiftPerFocusX | 0.0035
| Samsung P3D | 2023-10-26 | OPM Repeatability @5mm | 1.01
| Samsung P3D | 2023-10-26 | OPM Repeatability @25mm | 3.28
| Samsung P3D | 2023-10-26 | OPM Repeatability @1mm | 0.77
| Samsung P3D | 2023-10-26 | OPM Repeatability @10mm | 4.89
| Samsung P3D | 2023-10-26 | NCM driving% with AC160TS cantilever | 33.3
| Samsung P3D | 2023-10-26 | Min. Vibration Transmissbility (~100Hz) | 0.025
| Samsung P3D | 2023-10-26 | Max OPM @5mm | 25.87
| Samsung P3D | 2023-10-26 | Max OPM @25mm | 125.7
| Samsung P3D | 2023-10-26 | Max OPM @1mm | 9.93
| Samsung P3D | 2023-10-26 | Max OPM @10mm | 38.07
| Samsung P3D | 2023-10-26 | LowerLimit | 4719
| Samsung P3D | 2023-10-26 | Lateral(C-D) sensitivity | 22.2
| Samsung P3D | 2023-10-26 | Lateral(C-D) noise | 0.214
| Samsung P3D | 2023-10-26 | Image offset after moving from -Y to +Y (Y scan) | 12.45
| Samsung P3D | 2023-10-26 | Image offset after moving from -Y to +Y (X scan) | 16.85
| Samsung P3D | 2023-10-26 | Image offset after moving from -X to +X (Y scan) | 6.59
| Samsung P3D | 2023-10-26 | Image offset after moving from -X to +X (X scan) | 102.54
| Samsung P3D | 2023-10-26 | Image offset after moving from +Y to -Y (Y scan) | 98.14
| Samsung P3D | 2023-10-26 | Image offset after moving from +Y to -Y (X scan) | 5.86
| Samsung P3D | 2023-10-26 | Image offset after moving from +X to -X (Y scan) | 16.85
| Samsung P3D | 2023-10-26 | Image offset after moving from +X to -X (X scan) | 117.19
| Samsung P3D | 2023-10-26 | Full range of Z Height during Z Scan sweep | 16.5
| Samsung P3D | 2023-10-26 | Full range of Z Detector during Z Scan sweep | 11.6
| Samsung P3D | 2023-10-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.05
| Samsung P3D | 2023-10-26 | Full range of Y2 Detector during Y Scan sweep | 205.4
| Samsung P3D | 2023-10-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.69
| Samsung P3D | 2023-10-26 | Full range of Y1 Detector during Y Scan sweep | 210.8
| Samsung P3D | 2023-10-26 | Full range of X2 Detector Raw during X Scan sweep | 13.84
| Samsung P3D | 2023-10-26 | Full range of X2 Detector during X Scan sweep | 212.4
| Samsung P3D | 2023-10-26 | Full range of X1 Detector Raw during X Scan sweep | 14.08
| Samsung P3D | 2023-10-26 | Full range of X1 Detector during X Scan sweep | 194.9
| Samsung P3D | 2023-10-26 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Samsung P3D | 2023-10-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Samsung P3D | 2023-10-26 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.45
| Samsung P3D | 2023-10-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| Samsung P3D | 2023-10-26 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Samsung P3D | 2023-10-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Samsung P3D | 2023-10-26 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.34
| Samsung P3D | 2023-10-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22
| Samsung P3D | 2023-10-26 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| Samsung P3D | 2023-10-26 | 940nm Step Height Sample (10 Repeats) | 949.67
| Samsung P3D | 2023-10-26 | 50~200nm Step Height Sample (10 Repeats) | 183.56
| Samsung P3D | 2023-10-26 | 200~500nm Step Height Sample (10 Repeats) | 463.61
| Samsung P3D | 2023-10-26 | * 1800nm Step Height Sample (10 Repeats) | 1792.72
| Samsung P3D | 2023-10-26 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.3
| Samsung P3D | 2023-10-26 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.1
| Samsung P3D | 2023-10-26 | %GRR w/ roughness of non-patterened wafer | 7.4
| Guangzhou University | 2023-10-20 | Z Scanner position at -3μm | 0.0014
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Z Scanner position at -3μm | 0.0003
| TYSiC#8 | 2023-10-20 | Z Scanner position at -3μm | 0.0025
| TYSiC#9 | 2023-10-20 | Z Scanner position at -3μm | 0.0048
| Guangzhou University | 2023-10-20 | Z Scanner position at +3μm | 0.0036
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Z Scanner position at +3μm | 0.0066
| TYSiC#8 | 2023-10-20 | Z Scanner position at +3μm | 0.0145
| TYSiC#9 | 2023-10-20 | Z Scanner position at +3μm | 0.01
| Guangzhou University | 2023-10-20 | Z Linearity by Z Scanner position
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Z Linearity by Z Scanner position
| TYSiC#8 | 2023-10-20 | Z Linearity by Z Scanner position
| TYSiC#9 | 2023-10-20 | Z Linearity by Z Scanner position
| Guangzhou University | 2023-10-20 | Z Linearity by 3 different sample height | 1.003
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Z Linearity by 3 different sample height | 1.01
| TYSiC#8 | 2023-10-20 | Z Linearity by 3 different sample height | 1.008
| TYSiC#9 | 2023-10-20 | Z Linearity by 3 different sample height | 1.003
| Guangzhou University | 2023-10-20 | Z Height noise | 0.039
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Z Height noise | 0.027
| TYSiC#8 | 2023-10-20 | Z Height noise | 0.025
| TYSiC#9 | 2023-10-20 | Z Height noise | 0.027
| Guangzhou University | 2023-10-20 | Y Scan Stroke | 135
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Y Scan Stroke | 55
| TYSiC#8 | 2023-10-20 | Y Scan Stroke | 134.9
| TYSiC#9 | 2023-10-20 | Y Scan Stroke | 127.6
| TYSiC#8 | 2023-10-20 | -Y Direction software Limit | -100964
| TYSiC#9 | 2023-10-20 | -Y Direction software Limit | -100932
| Guangzhou University | 2023-10-20 | Y Detector noise (XY Servo on) | 0.54
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Y Detector noise (XY Servo on) | 0.3
| TYSiC#8 | 2023-10-20 | Y Detector noise (XY Servo on) | 0.28
| TYSiC#9 | 2023-10-20 | Y Detector noise (XY Servo on) | 0.32
| Guangzhou University | 2023-10-20 | Y Detector noise (XY Servo off) | 0.49
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Y Detector noise (XY Servo off) | 0.3
| TYSiC#8 | 2023-10-20 | Y Detector noise (XY Servo off) | 0.28
| TYSiC#9 | 2023-10-20 | Y Detector noise (XY Servo off) | 0.29
| Guangzhou University | 2023-10-20 | Y axis OPM | 1.41
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Y axis OPM | 0.97
| TYSiC#8 | 2023-10-20 | Y axis OPM | 1.35
| TYSiC#9 | 2023-10-20 | Y axis OPM | 1.03
| Guangzhou University | 2023-10-20 | X Scan Stroke | 130.6
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | X Scan Stroke | 55
| TYSiC#8 | 2023-10-20 | X Scan Stroke | 128.7
| TYSiC#9 | 2023-10-20 | X Scan Stroke | 129.9
| TYSiC#8 | 2023-10-20 | -X Direction software Limit | -102747
| TYSiC#9 | 2023-10-20 | -X Direction software Limit | -100932
| Guangzhou University | 2023-10-20 | X Detector noise (XY Servo on) | 0.52
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | X Detector noise (XY Servo on) | 0.32
| TYSiC#8 | 2023-10-20 | X Detector noise (XY Servo on) | 0.29
| TYSiC#9 | 2023-10-20 | X Detector noise (XY Servo on) | 0.29
| Guangzhou University | 2023-10-20 | X Detector noise (XY Servo off) | 0.48
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | X Detector noise (XY Servo off) | 0.33
| TYSiC#8 | 2023-10-20 | X Detector noise (XY Servo off) | 0.29
| TYSiC#9 | 2023-10-20 | X Detector noise (XY Servo off) | 0.27
| Guangzhou University | 2023-10-20 | X axis OPM | 1.21
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | X axis OPM | 0.85
| TYSiC#8 | 2023-10-20 | X axis OPM | 0.81
| TYSiC#9 | 2023-10-20 | X axis OPM | 0.73
| Guangzhou University | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| TYSiC#8 | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| TYSiC#9 | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Guangzhou University | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| TYSiC#8 | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| TYSiC#9 | 2023-10-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| TYSiC#8 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.28
| TYSiC#9 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.31
| TYSiC#8 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.31
| TYSiC#9 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.32
| TYSiC#8 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| TYSiC#9 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| TYSiC#8 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| TYSiC#9 | 2023-10-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| TYSiC#8 | 2023-10-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| TYSiC#9 | 2023-10-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| Guangzhou University | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| TYSiC#8 | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| TYSiC#9 | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Guangzhou University | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.16
| TYSiC#8 | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| TYSiC#9 | 2023-10-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Guangzhou University | 2023-10-20 | NCM driving% with AC160TS cantilever | 0.6
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | NCM driving% with AC160TS cantilever | 0.23
| TYSiC#8 | 2023-10-20 | NCM driving% with AC160TS cantilever | 2.39
| TYSiC#9 | 2023-10-20 | NCM driving% with AC160TS cantilever | 0.22
| Guangzhou University | 2023-10-20 | NCM Amplitude noise | 13.92
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | NCM Amplitude noise | 15.27
| TYSiC#8 | 2023-10-20 | NCM Amplitude noise | 16.1
| TYSiC#9 | 2023-10-20 | NCM Amplitude noise | 15.5
| Guangzhou University | 2023-10-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.026
| TYSiC#8 | 2023-10-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.58
| TYSiC#9 | 2023-10-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.106
| Guangzhou University | 2023-10-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.402
| TYSiC#8 | 2023-10-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.68
| TYSiC#9 | 2023-10-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.239
| Guangzhou University | 2023-10-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 190.1
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 89
| TYSiC#8 | 2023-10-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 187.4
| TYSiC#9 | 2023-10-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 177.3
| Guangzhou University | 2023-10-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 183.9
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.5
| TYSiC#8 | 2023-10-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 186.5
| TYSiC#9 | 2023-10-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 177.5
| TYSiC#8 | 2023-10-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| TYSiC#9 | 2023-10-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Guangzhou University | 2023-10-20 | Lateral(C-D) sensitivity | 14.8
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Lateral(C-D) sensitivity | 15.1
| TYSiC#8 | 2023-10-20 | Lateral(C-D) sensitivity | 12.7
| TYSiC#9 | 2023-10-20 | Lateral(C-D) sensitivity | 10.8
| Guangzhou University | 2023-10-20 | Lateral(C-D) noise | 0.097
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Lateral(C-D) noise | 0.053
| TYSiC#8 | 2023-10-20 | Lateral(C-D) noise | 0.102
| TYSiC#9 | 2023-10-20 | Lateral(C-D) noise | 0.123
| Guangzhou University | 2023-10-20 | Full range of Z Height during Z Scan sweep | 15.9
| TYSiC#8 | 2023-10-20 | Full range of Z Height during Z Scan sweep | 16
| TYSiC#9 | 2023-10-20 | Full range of Z Height during Z Scan sweep | 15.8
| Guangzhou University | 2023-10-20 | Full range of Z Detector during Z Scan sweep | 12.5
| TYSiC#8 | 2023-10-20 | Full range of Z Detector during Z Scan sweep | 14
| TYSiC#9 | 2023-10-20 | Full range of Z Detector during Z Scan sweep | 12.1
| Guangzhou University | 2023-10-20 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| TYSiC#8 | 2023-10-20 | Full range of Y Detector Raw during Y Scan sweep | 14
| TYSiC#9 | 2023-10-20 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Guangzhou University | 2023-10-20 | Full range of Y Detector during Y Scan sweep | 134.9
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Full range of Y Detector during Y Scan sweep | 66.4
| TYSiC#8 | 2023-10-20 | Full range of Y Detector during Y Scan sweep | 135.9
| TYSiC#9 | 2023-10-20 | Full range of Y Detector during Y Scan sweep | 128
| Guangzhou University | 2023-10-20 | Full range of X Detector Raw during X Scan sweep | 14.2
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Full range of X Detector Raw during X Scan sweep | 14.3
| TYSiC#8 | 2023-10-20 | Full range of X Detector Raw during X Scan sweep | 13.4
| TYSiC#9 | 2023-10-20 | Full range of X Detector Raw during X Scan sweep | 14.6
| Guangzhou University | 2023-10-20 | Full range of X Detector during X Scan sweep | 130.5
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | Full range of X Detector during X Scan sweep | 68.5
| TYSiC#8 | 2023-10-20 | Full range of X Detector during X Scan sweep | 130.8
| TYSiC#9 | 2023-10-20 | Full range of X Detector during X Scan sweep | 127.6
| Guangzhou University | 2023-10-20 | A-B noise | 15.7
| Nagaoka Univ. of Technology(Prof. Syouji) | 2023-10-20 | A-B noise | 10.01
| TYSiC#8 | 2023-10-20 | A-B noise | 9.42
| TYSiC#9 | 2023-10-20 | A-B noise | 18.18
| TYSiC#10 | 2023-10-19 | Z Scanner position at -3μm | 0.0195
| TYSiC#11 | 2023-10-19 | Z Scanner position at -3μm | 0.0086
| TYSiC#12 | 2023-10-19 | Z Scanner position at -3μm | 0.0055
| TYSiC#10 | 2023-10-19 | Z Scanner position at +3μm | 0.0093
| TYSiC#11 | 2023-10-19 | Z Scanner position at +3μm | 0.0102
| TYSiC#12 | 2023-10-19 | Z Scanner position at +3μm | 0.0138
| TYSiC#10 | 2023-10-19 | Z Linearity by Z Scanner position
| TYSiC#11 | 2023-10-19 | Z Linearity by Z Scanner position
| TYSiC#12 | 2023-10-19 | Z Linearity by Z Scanner position
| TYSiC#10 | 2023-10-19 | Z Linearity by 3 different sample height | 1.011
| TYSiC#11 | 2023-10-19 | Z Linearity by 3 different sample height | 1.002
| TYSiC#12 | 2023-10-19 | Z Linearity by 3 different sample height | 0.989
| TYSiC#10 | 2023-10-19 | Z Height noise | 0.03
| TYSiC#11 | 2023-10-19 | Z Height noise | 0.029
| TYSiC#12 | 2023-10-19 | Z Height noise | 0.026
| TYSiC#10 | 2023-10-19 | Y Scan Stroke | 133.9
| TYSiC#11 | 2023-10-19 | Y Scan Stroke | 128.6
| TYSiC#12 | 2023-10-19 | Y Scan Stroke | 131
| TYSiC#10 | 2023-10-19 | -Y Direction software Limit | -100779
| TYSiC#11 | 2023-10-19 | -Y Direction software Limit | -100834
| TYSiC#12 | 2023-10-19 | -Y Direction software Limit | -100928
| TYSiC#10 | 2023-10-19 | Y Detector noise (XY Servo on) | 0.284
| TYSiC#11 | 2023-10-19 | Y Detector noise (XY Servo on) | 0.27
| TYSiC#12 | 2023-10-19 | Y Detector noise (XY Servo on) | 0.3
| TYSiC#10 | 2023-10-19 | Y Detector noise (XY Servo off) | 0.3
| TYSiC#11 | 2023-10-19 | Y Detector noise (XY Servo off) | 0.28
| TYSiC#12 | 2023-10-19 | Y Detector noise (XY Servo off) | 0.3
| TYSiC#10 | 2023-10-19 | Y axis OPM | 1.65
| TYSiC#11 | 2023-10-19 | Y axis OPM | 0.83
| TYSiC#12 | 2023-10-19 | Y axis OPM | 0.71
| TYSiC#10 | 2023-10-19 | X Scan Stroke | 135.6
| TYSiC#11 | 2023-10-19 | X Scan Stroke | 124.1
| TYSiC#12 | 2023-10-19 | X Scan Stroke | 131.7
| TYSiC#10 | 2023-10-19 | -X Direction software Limit | -100391
| TYSiC#11 | 2023-10-19 | -X Direction software Limit | -100901
| TYSiC#12 | 2023-10-19 | -X Direction software Limit | -100796
| TYSiC#10 | 2023-10-19 | X Detector noise (XY Servo on) | 0.25
| TYSiC#11 | 2023-10-19 | X Detector noise (XY Servo on) | 0.3
| TYSiC#12 | 2023-10-19 | X Detector noise (XY Servo on) | 0.29
| TYSiC#10 | 2023-10-19 | X Detector noise (XY Servo off) | 0.25
| TYSiC#11 | 2023-10-19 | X Detector noise (XY Servo off) | 0.28
| TYSiC#12 | 2023-10-19 | X Detector noise (XY Servo off) | 0.29
| TYSiC#10 | 2023-10-19 | X axis OPM | 0.67
| TYSiC#11 | 2023-10-19 | X axis OPM | 0.89
| TYSiC#12 | 2023-10-19 | X axis OPM | 0.75
| TYSiC#10 | 2023-10-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| TYSiC#11 | 2023-10-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| TYSiC#12 | 2023-10-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| TYSiC#10 | 2023-10-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| TYSiC#11 | 2023-10-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| TYSiC#12 | 2023-10-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| TYSiC#10 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.34
| TYSiC#11 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.38
| TYSiC#12 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.27
| TYSiC#10 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| TYSiC#11 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| TYSiC#12 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.28
| TYSiC#10 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.33
| TYSiC#11 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.34
| TYSiC#12 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.27
| TYSiC#10 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.32
| TYSiC#11 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.36
| TYSiC#12 | 2023-10-19 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.27
| TYSiC#10 | 2023-10-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| TYSiC#11 | 2023-10-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.34
| TYSiC#12 | 2023-10-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| TYSiC#10 | 2023-10-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| TYSiC#11 | 2023-10-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| TYSiC#12 | 2023-10-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| TYSiC#10 | 2023-10-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| TYSiC#11 | 2023-10-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| TYSiC#12 | 2023-10-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| TYSiC#10 | 2023-10-19 | NCM driving% with AC160TS cantilever | 2.15
| TYSiC#11 | 2023-10-19 | NCM driving% with AC160TS cantilever | 1.27
| TYSiC#12 | 2023-10-19 | NCM driving% with AC160TS cantilever | 1.93
| TYSiC#10 | 2023-10-19 | NCM Amplitude noise | 16.22
| TYSiC#11 | 2023-10-19 | NCM Amplitude noise | 13.33
| TYSiC#12 | 2023-10-19 | NCM Amplitude noise | 16.1
| TYSiC#10 | 2023-10-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.43
| TYSiC#11 | 2023-10-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.271
| TYSiC#12 | 2023-10-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.423
| TYSiC#10 | 2023-10-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.12
| TYSiC#11 | 2023-10-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.784
| TYSiC#12 | 2023-10-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.791
| TYSiC#10 | 2023-10-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 191.3
| TYSiC#11 | 2023-10-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 182
| TYSiC#12 | 2023-10-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 192.9
| TYSiC#10 | 2023-10-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 193.8
| TYSiC#11 | 2023-10-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 186.6
| TYSiC#12 | 2023-10-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 193.9
| TYSiC#10 | 2023-10-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| TYSiC#11 | 2023-10-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| TYSiC#12 | 2023-10-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| TYSiC#10 | 2023-10-19 | Lateral(C-D) sensitivity | 6.5
| TYSiC#11 | 2023-10-19 | Lateral(C-D) sensitivity | 9.9
| TYSiC#12 | 2023-10-19 | Lateral(C-D) sensitivity | 12.6
| TYSiC#10 | 2023-10-19 | Lateral(C-D) noise | 0.115
| TYSiC#11 | 2023-10-19 | Lateral(C-D) noise | 0.112
| TYSiC#12 | 2023-10-19 | Lateral(C-D) noise | 0.17
| TYSiC#10 | 2023-10-19 | Full range of Z Height during Z Scan sweep | 16.3
| TYSiC#11 | 2023-10-19 | Full range of Z Height during Z Scan sweep | 16.1
| TYSiC#12 | 2023-10-19 | Full range of Z Height during Z Scan sweep | 16.3
| TYSiC#10 | 2023-10-19 | Full range of Z Detector during Z Scan sweep | 11.6
| TYSiC#11 | 2023-10-19 | Full range of Z Detector during Z Scan sweep | 13
| TYSiC#12 | 2023-10-19 | Full range of Z Detector during Z Scan sweep | 12.7
| TYSiC#10 | 2023-10-19 | Full range of Y Detector Raw during Y Scan sweep | 14
| TYSiC#11 | 2023-10-19 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| TYSiC#12 | 2023-10-19 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| TYSiC#10 | 2023-10-19 | Full range of Y Detector during Y Scan sweep | 132.8
| TYSiC#11 | 2023-10-19 | Full range of Y Detector during Y Scan sweep | 129.2
| TYSiC#12 | 2023-10-19 | Full range of Y Detector during Y Scan sweep | 129
| TYSiC#10 | 2023-10-19 | Full range of X Detector Raw during X Scan sweep | 14
| TYSiC#11 | 2023-10-19 | Full range of X Detector Raw during X Scan sweep | 13.5
| TYSiC#12 | 2023-10-19 | Full range of X Detector Raw during X Scan sweep | 13.5
| TYSiC#10 | 2023-10-19 | Full range of X Detector during X Scan sweep | 135.4
| TYSiC#11 | 2023-10-19 | Full range of X Detector during X Scan sweep | 125.5
| TYSiC#12 | 2023-10-19 | Full range of X Detector during X Scan sweep | 129.9
| TYSiC#10 | 2023-10-19 | A-B noise | 15
| TYSiC#11 | 2023-10-19 | A-B noise | 10.56
| TYSiC#12 | 2023-10-19 | A-B noise | 12.8
| RMUTR University | 2023-10-17 | Z Scanner position at -3μm | 0.0032
| RMUTR University | 2023-10-17 | Z Scanner position at +3μm | 0.0098
| RMUTR University | 2023-10-17 | Z Linearity by Z Scanner position
| RMUTR University | 2023-10-17 | Z Linearity by 3 different sample height | 1.007
| RMUTR University | 2023-10-17 | Z Height noise | 0.018
| RMUTR University | 2023-10-17 | Y Scan Stroke | 55
| RMUTR University | 2023-10-17 | Y Detector noise (XY Servo on) | 0.26
| RMUTR University | 2023-10-17 | Y Detector noise (XY Servo off) | 0.26
| RMUTR University | 2023-10-17 | Y axis OPM | 1.12
| RMUTR University | 2023-10-17 | X Scan Stroke | 55
| RMUTR University | 2023-10-17 | X Detector noise (XY Servo on) | 0.26
| RMUTR University | 2023-10-17 | X Detector noise (XY Servo off) | 0.26
| RMUTR University | 2023-10-17 | X axis OPM | 0.59
| RMUTR University | 2023-10-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| RMUTR University | 2023-10-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| RMUTR University | 2023-10-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| RMUTR University | 2023-10-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| RMUTR University | 2023-10-17 | NCM driving% with AC160TS cantilever | 0.14
| RMUTR University | 2023-10-17 | NCM Amplitude noise | 18.4
| RMUTR University | 2023-10-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.418
| RMUTR University | 2023-10-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.604
| RMUTR University | 2023-10-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.4
| RMUTR University | 2023-10-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99.4
| RMUTR University | 2023-10-17 | Lateral(C-D) sensitivity | 4.9
| RMUTR University | 2023-10-17 | Lateral(C-D) noise | 0.093
| RMUTR University | 2023-10-17 | Full range of Z Height during Z Scan sweep | 16.2
| RMUTR University | 2023-10-17 | Full range of Z Detector during Z Scan sweep | 12.6
| RMUTR University | 2023-10-17 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| RMUTR University | 2023-10-17 | Full range of Y Detector during Y Scan sweep | 71.5
| RMUTR University | 2023-10-17 | Full range of X Detector Raw during X Scan sweep | 14.1
| RMUTR University | 2023-10-17 | Full range of X Detector during X Scan sweep | 70.3
| RMUTR University | 2023-10-17 | A-B noise | 12.71
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Z Scanner position at -3μm | 0.0001
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Z Scanner position at +3μm | 0.0036
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Z Linearity by Z Scanner position
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Z Linearity by 3 different sample height | 1.016
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Z Height noise | 0.027
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Y Scan Stroke | 110
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Y Detector noise (XY Servo on) | 0.28
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Y Detector noise (XY Servo off) | 0.29
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Y axis OPM | 0.89
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | X Scan Stroke | 110
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | X Detector noise (XY Servo on) | 0.29
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | X Detector noise (XY Servo off) | 0.29
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | X axis OPM | 0.59
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.49
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.58
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.47
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | NCM driving% with AC160TS cantilever | 0.56
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | NCM Amplitude noise | 15.77
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.3
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.772
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 247.7
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 252.3
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Lateral(C-D) sensitivity | 6.9
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Lateral(C-D) noise | 0.1
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Full range of Z Height during Z Scan sweep | 16.1
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Full range of Z Detector during Z Scan sweep | 11.6
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Full range of Y Detector during Y Scan sweep | 177.2
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Full range of X Detector Raw during X Scan sweep | 14.5
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | Full range of X Detector during X Scan sweep | 182.3
| University of Mississippi(Vig Sundaresan) | 2023-10-16 | A-B noise | 10.29
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Z Scanner position at -3μm | 0.0012
| Cornell University(Liu) | 2023-10-13 | Z Scanner position at -3μm | 0.0056
| NovaSic (PSE) | 2023-10-13 | Z Scanner position at -3μm | 0.0078
| IDEADED (PSE) | 2023-10-13 | Z Scanner position at -3μm | 0.009
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Z Scanner position at +3μm | 0.0012
| Cornell University(Liu) | 2023-10-13 | Z Scanner position at +3μm | 0.0168
| NovaSic (PSE) | 2023-10-13 | Z Scanner position at +3μm | 0.0068
| IDEADED (PSE) | 2023-10-13 | Z Scanner position at +3μm | 0.0157
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Z Linearity by Z Scanner position
| Cornell University(Liu) | 2023-10-13 | Z Linearity by Z Scanner position
| NovaSic (PSE) | 2023-10-13 | Z Linearity by Z Scanner position
| IDEADED (PSE) | 2023-10-13 | Z Linearity by Z Scanner position
| 삼성전자 (PE) | 2023-10-13 | Z Linearity by Z Scanner position
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Z Linearity by 3 different sample height | 1.01
| Cornell University(Liu) | 2023-10-13 | Z Linearity by 3 different sample height | 1.009
| NovaSic (PSE) | 2023-10-13 | Z Linearity by 3 different sample height | 1.001
| IDEADED (PSE) | 2023-10-13 | Z Linearity by 3 different sample height | 0.999
| 삼성전자 (PE) | 2023-10-13 | Z Linearity by 3 different sample height | 0.998
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Z Height noise | 0.025
| Cornell University(Liu) | 2023-10-13 | Z Height noise | 0.027
| NovaSic (PSE) | 2023-10-13 | Z Height noise | 0.022
| IDEADED (PSE) | 2023-10-13 | Z Height noise | 0.027
| 삼성전자 (PE) | 2023-10-13 | Z Height noise | 0.03
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y2 Detector noise (XY Servo on) | 0.32
| NovaSic (PSE) | 2023-10-13 | Y2 Detector noise (XY Servo on) | 0.33
| IDEADED (PSE) | 2023-10-13 | Y2 Detector noise (XY Servo on) | 0.26
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y2 Detector noise (XY Servo off) | 0.22
| NovaSic (PSE) | 2023-10-13 | Y2 Detector noise (XY Servo off) | 0.33
| IDEADED (PSE) | 2023-10-13 | Y2 Detector noise (XY Servo off) | 0.23
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y1 Detector noise (XY Servo on) | 0.33
| NovaSic (PSE) | 2023-10-13 | Y1 Detector noise (XY Servo on) | 0.24
| IDEADED (PSE) | 2023-10-13 | Y1 Detector noise (XY Servo on) | 0.27
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y1 Detector noise (XY Servo off) | 0.25
| NovaSic (PSE) | 2023-10-13 | Y1 Detector noise (XY Servo off) | 0.24
| IDEADED (PSE) | 2023-10-13 | Y1 Detector noise (XY Servo off) | 0.23
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y Scan Stroke | 172.6
| Cornell University(Liu) | 2023-10-13 | Y Scan Stroke | 55
| NovaSic (PSE) | 2023-10-13 | Y Scan Stroke | 184.9
| IDEADED (PSE) | 2023-10-13 | Y Scan Stroke | 180
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | -Y Direction software Limit | -75407.5
| NovaSic (PSE) | 2023-10-13 | -Y Direction software Limit | -102705
| IDEADED (PSE) | 2023-10-13 | -Y Direction software Limit | -102594
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y Detector noise (XY Servo on) | 0.26
| Cornell University(Liu) | 2023-10-13 | Y Detector noise (XY Servo on) | 0.27
| NovaSic (PSE) | 2023-10-13 | Y Detector noise (XY Servo on) | 0.21
| IDEADED (PSE) | 2023-10-13 | Y Detector noise (XY Servo on) | 0.21
| 삼성전자 (PE) | 2023-10-13 | Y Detector noise (XY Servo on) | 0.27
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y Detector noise (XY Servo off) | 0.17
| Cornell University(Liu) | 2023-10-13 | Y Detector noise (XY Servo off) | 0.26
| NovaSic (PSE) | 2023-10-13 | Y Detector noise (XY Servo off) | 0.21
| IDEADED (PSE) | 2023-10-13 | Y Detector noise (XY Servo off) | 0.16
| 삼성전자 (PE) | 2023-10-13 | Y Detector noise (XY Servo off) | 0.26
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Y axis OPM | 0.63
| Cornell University(Liu) | 2023-10-13 | Y axis OPM | 0.34
| NovaSic (PSE) | 2023-10-13 | Y axis OPM | 0.88
| IDEADED (PSE) | 2023-10-13 | Y axis OPM | 1.55
| 삼성전자 (PE) | 2023-10-13 | Y axis OPM | 1.71
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X2 Detector noise (XY Servo on) | 0.35
| NovaSic (PSE) | 2023-10-13 | X2 Detector noise (XY Servo on) | 0.24
| IDEADED (PSE) | 2023-10-13 | X2 Detector noise (XY Servo on) | 0.28
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X2 Detector noise (XY Servo off) | 0.23
| NovaSic (PSE) | 2023-10-13 | X2 Detector noise (XY Servo off) | 0.24
| IDEADED (PSE) | 2023-10-13 | X2 Detector noise (XY Servo off) | 0.26
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X1 Detector noise (XY Servo on) | 0.35
| NovaSic (PSE) | 2023-10-13 | X1 Detector noise (XY Servo on) | 0.22
| IDEADED (PSE) | 2023-10-13 | X1 Detector noise (XY Servo on) | 0.28
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X1 Detector noise (XY Servo off) | 0.23
| NovaSic (PSE) | 2023-10-13 | X1 Detector noise (XY Servo off) | 0.22
| IDEADED (PSE) | 2023-10-13 | X1 Detector noise (XY Servo off) | 0.25
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X Scan Stroke | 164.2
| Cornell University(Liu) | 2023-10-13 | X Scan Stroke | 55
| NovaSic (PSE) | 2023-10-13 | X Scan Stroke | 179
| IDEADED (PSE) | 2023-10-13 | X Scan Stroke | 176.6
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | -X Direction software Limit | -75410.6
| NovaSic (PSE) | 2023-10-13 | -X Direction software Limit | -101918
| IDEADED (PSE) | 2023-10-13 | -X Direction software Limit | -102324
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X Detector noise (XY Servo on) | 0.29
| Cornell University(Liu) | 2023-10-13 | X Detector noise (XY Servo on) | 0.27
| NovaSic (PSE) | 2023-10-13 | X Detector noise (XY Servo on) | 0.17
| IDEADED (PSE) | 2023-10-13 | X Detector noise (XY Servo on) | 0.23
| 삼성전자 (PE) | 2023-10-13 | X Detector noise (XY Servo on) | 0.27
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X Detector noise (XY Servo off) | 0.17
| Cornell University(Liu) | 2023-10-13 | X Detector noise (XY Servo off) | 0.26
| NovaSic (PSE) | 2023-10-13 | X Detector noise (XY Servo off) | 0.17
| IDEADED (PSE) | 2023-10-13 | X Detector noise (XY Servo off) | 0.18
| 삼성전자 (PE) | 2023-10-13 | X Detector noise (XY Servo off) | 0.26
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | X axis OPM | 0.86
| Cornell University(Liu) | 2023-10-13 | X axis OPM | 0.52
| NovaSic (PSE) | 2023-10-13 | X axis OPM | 0.77
| IDEADED (PSE) | 2023-10-13 | X axis OPM | 0.88
| 삼성전자 (PE) | 2023-10-13 | X axis OPM | 1.06
| 삼성전자 (PE) | 2023-10-13 | SLD beam power at the end of the lens | 0.26
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Cornell University(Liu) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 삼성전자 (PE) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| Cornell University(Liu) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| 삼성전자 (PE) | 2023-10-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.32
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.34
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.3
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.26
| NovaSic (PSE) | 2023-10-13 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| IDEADED (PSE) | 2023-10-13 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.27
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Cornell University(Liu) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| NovaSic (PSE) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| IDEADED (PSE) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 삼성전자 (PE) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Cornell University(Liu) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| NovaSic (PSE) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| IDEADED (PSE) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 삼성전자 (PE) | 2023-10-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.17
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of Wafer Zeroscan(-60000,0) | 0.28
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of Wafer Zeroscan(60000,0) | 0.28
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of Wafer Zeroscan(0,-60000) | 0.27
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of Wafer Zeroscan(0,60000) | 0.27
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | RMS of Z Drive of Wafer ZeroScan (0,0) | 0.28
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | NCM driving% with AC160TS cantilever | 0.83
| Cornell University(Liu) | 2023-10-13 | NCM driving% with AC160TS cantilever | 0.75
| NovaSic (PSE) | 2023-10-13 | NCM driving% with AC160TS cantilever | 0.58
| IDEADED (PSE) | 2023-10-13 | NCM driving% with AC160TS cantilever | 0.37
| 삼성전자 (PE) | 2023-10-13 | NCM driving% with AC160TS cantilever | 7.33
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | NCM Amplitude noise | 11.29
| Cornell University(Liu) | 2023-10-13 | NCM Amplitude noise | 17.93
| NovaSic (PSE) | 2023-10-13 | NCM Amplitude noise | 15.31
| IDEADED (PSE) | 2023-10-13 | NCM Amplitude noise | 17.41
| 삼성전자 (PE) | 2023-10-13 | NCM Amplitude noise | 15.57
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| Cornell University(Liu) | 2023-10-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.565
| NovaSic (PSE) | 2023-10-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.544
| IDEADED (PSE) | 2023-10-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.179
| 삼성전자 (PE) | 2023-10-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.218
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 27
| Cornell University(Liu) | 2023-10-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.563
| NovaSic (PSE) | 2023-10-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.588
| IDEADED (PSE) | 2023-10-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.453
| 삼성전자 (PE) | 2023-10-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.373
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247
| NovaSic (PSE) | 2023-10-13 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 259.5
| IDEADED (PSE) | 2023-10-13 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 255.3
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 260.9
| NovaSic (PSE) | 2023-10-13 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 255.8
| IDEADED (PSE) | 2023-10-13 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 252.1
| Cornell University(Liu) | 2023-10-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106
| 삼성전자 (PE) | 2023-10-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.4
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.7
| NovaSic (PSE) | 2023-10-13 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 258.9
| IDEADED (PSE) | 2023-10-13 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 247.1
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 243.2
| NovaSic (PSE) | 2023-10-13 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 254.4
| IDEADED (PSE) | 2023-10-13 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249.6
| Cornell University(Liu) | 2023-10-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.8
| 삼성전자 (PE) | 2023-10-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.2
| 삼성전자 (PE) | 2023-10-13 | Full range of Z Height during Z Scan sweep | 16.1
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| NovaSic (PSE) | 2023-10-13 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| IDEADED (PSE) | 2023-10-13 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| 삼성전자 (PE) | 2023-10-13 | LD beam power at the end of the lens | 2.3
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Lateral(C-D) sensitivity | 37.904
| Cornell University(Liu) | 2023-10-13 | Lateral(C-D) sensitivity | 5.7
| NovaSic (PSE) | 2023-10-13 | Lateral(C-D) sensitivity | 5.5
| IDEADED (PSE) | 2023-10-13 | Lateral(C-D) sensitivity | 8.3
| 삼성전자 (PE) | 2023-10-13 | Lateral(C-D) sensitivity | 8
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Lateral(C-D) noise | 0.107
| Cornell University(Liu) | 2023-10-13 | Lateral(C-D) noise | 0.105
| NovaSic (PSE) | 2023-10-13 | Lateral(C-D) noise | 0.129
| IDEADED (PSE) | 2023-10-13 | Lateral(C-D) noise | 0.131
| 삼성전자 (PE) | 2023-10-13 | Lateral(C-D) noise | 0.132
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of Z Height during Z Scan sweep | 16
| Cornell University(Liu) | 2023-10-13 | Full range of Z Height during Z Scan sweep | 15.8
| NovaSic (PSE) | 2023-10-13 | Full range of Z Height during Z Scan sweep | 16.5
| IDEADED (PSE) | 2023-10-13 | Full range of Z Height during Z Scan sweep | 15.9
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of Z Detector during Z Scan sweep | 11.8
| Cornell University(Liu) | 2023-10-13 | Full range of Z Detector during Z Scan sweep | 12.8
| NovaSic (PSE) | 2023-10-13 | Full range of Z Detector during Z Scan sweep | 14
| IDEADED (PSE) | 2023-10-13 | Full range of Z Detector during Z Scan sweep | 12
| 삼성전자 (PE) | 2023-10-13 | Full range of Z Detector during Z Scan sweep | 15.1
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of Y2 Detector Raw during Y Scan sweep | 13.5
| NovaSic (PSE) | 2023-10-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| IDEADED (PSE) | 2023-10-13 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of Y1 Detector Raw during Y Scan sweep | 13.7
| NovaSic (PSE) | 2023-10-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| IDEADED (PSE) | 2023-10-13 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Cornell University(Liu) | 2023-10-13 | Full range of Y Detector Raw during Y Scan sweep | 14
| 삼성전자 (PE) | 2023-10-13 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of Y Detector during Y Scan sweep | 179.5
| Cornell University(Liu) | 2023-10-13 | Full range of Y Detector during Y Scan sweep | 74.3
| NovaSic (PSE) | 2023-10-13 | Full range of Y Detector during Y Scan sweep | 178.2
| IDEADED (PSE) | 2023-10-13 | Full range of Y Detector during Y Scan sweep | 179.8
| 삼성전자 (PE) | 2023-10-13 | Full range of Y Detector during Y Scan sweep | 119
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of X2 Detector Raw during X Scan sweep | 13.5
| NovaSic (PSE) | 2023-10-13 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| IDEADED (PSE) | 2023-10-13 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of X1 Detector Raw during X Scan sweep | 13.5
| NovaSic (PSE) | 2023-10-13 | Full range of X1 Detector Raw during X Scan sweep | 14
| IDEADED (PSE) | 2023-10-13 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Cornell University(Liu) | 2023-10-13 | Full range of X Detector Raw during X Scan sweep | 14.2
| 삼성전자 (PE) | 2023-10-13 | Full range of X Detector Raw during X Scan sweep | 15.2
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | Full range of X Detector during X Scan sweep | 169.1
| Cornell University(Liu) | 2023-10-13 | Full range of X Detector during X Scan sweep | 73.7
| NovaSic (PSE) | 2023-10-13 | Full range of X Detector during X Scan sweep | 178.3
| IDEADED (PSE) | 2023-10-13 | Full range of X Detector during X Scan sweep | 176.6
| 삼성전자 (PE) | 2023-10-13 | Full range of X Detector during X Scan sweep | 119.6
| 삼성전자 (PE) | 2023-10-13 | EFM Phase : For the positive bias | 89.7
| 삼성전자 (PE) | 2023-10-13 | EFM Phase : For the negative bias | -90.7
| 삼성전자 (PE) | 2023-10-13 | Approach position at -3μm | 0.0119
| 삼성전자 (PE) | 2023-10-13 | Approach position at +3μm | 0.0103
| IIT Bombay(Indian Institute of Technology Bombay | 2023-10-13 | A-B noise | 7.93
| Cornell University(Liu) | 2023-10-13 | A-B noise | 11.69
| NovaSic (PSE) | 2023-10-13 | A-B noise | 10.09
| IDEADED (PSE) | 2023-10-13 | A-B noise | 9.21
| 삼성전자 (PE) | 2023-10-13 | A-B noise | 9.04
| TDK | 2023-10-12 | Z Scanner position at -3μm | 0.0057
| National Chung Hsing University | 2023-10-12 | Z Scanner position at -3μm | 0.0027
| TDK | 2023-10-12 | Z Scanner position at +3μm | 0.0124
| National Chung Hsing University | 2023-10-12 | Z Scanner position at +3μm | 0.0015
| TDK | 2023-10-12 | Z Linearity by Z Scanner position
| National Chung Hsing University | 2023-10-12 | Z Linearity by Z Scanner position
| TDK | 2023-10-12 | Z Linearity by 3 different sample height | 1.002
| National Chung Hsing University | 2023-10-12 | Z Linearity by 3 different sample height | 0.996
| TDK | 2023-10-12 | Z Height noise | 0.032
| National Chung Hsing University | 2023-10-12 | Z Height noise | 0.026
| TDK | 2023-10-12 | Y Detector noise (XY Servo on) | 0.26
| National Chung Hsing University | 2023-10-12 | Y Detector noise (XY Servo on) | 0.29
| TDK | 2023-10-12 | Y Detector noise (XY Servo off) | 0.27
| National Chung Hsing University | 2023-10-12 | Y Detector noise (XY Servo off) | 0.29
| TDK | 2023-10-12 | Y axis OPM | 1.57
| National Chung Hsing University | 2023-10-12 | Y axis OPM | 1.67
| TDK | 2023-10-12 | X Detector noise (XY Servo on) | 0.26
| National Chung Hsing University | 2023-10-12 | X Detector noise (XY Servo on) | 0.28
| TDK | 2023-10-12 | X Detector noise (XY Servo off) | 0.26
| National Chung Hsing University | 2023-10-12 | X Detector noise (XY Servo off) | 0.28
| TDK | 2023-10-12 | X axis OPM | 0.79
| National Chung Hsing University | 2023-10-12 | X axis OPM | 1.37
| TDK | 2023-10-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| National Chung Hsing University | 2023-10-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| TDK | 2023-10-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| National Chung Hsing University | 2023-10-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| TDK | 2023-10-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| National Chung Hsing University | 2023-10-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| TDK | 2023-10-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.18
| National Chung Hsing University | 2023-10-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| TDK | 2023-10-12 | NCM driving% with AC160TS cantilever | 0.64
| National Chung Hsing University | 2023-10-12 | NCM driving% with AC160TS cantilever | 0.99
| TDK | 2023-10-12 | NCM Amplitude noise | 15.27
| National Chung Hsing University | 2023-10-12 | NCM Amplitude noise | 14.49
| TDK | 2023-10-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.519
| National Chung Hsing University | 2023-10-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.066
| TDK | 2023-10-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.902
| National Chung Hsing University | 2023-10-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.822
| TDK | 2023-10-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 155.3
| National Chung Hsing University | 2023-10-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 174.4
| TDK | 2023-10-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 155
| National Chung Hsing University | 2023-10-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 169.1
| TDK | 2023-10-12 | Lateral(C-D) sensitivity | 10.8
| National Chung Hsing University | 2023-10-12 | Lateral(C-D) sensitivity | 11.5
| TDK | 2023-10-12 | Lateral(C-D) noise | 0.101
| National Chung Hsing University | 2023-10-12 | Lateral(C-D) noise | 0.11
| TDK | 2023-10-12 | Full range of Z Height during Z Scan sweep | 16.3
| National Chung Hsing University | 2023-10-12 | Full range of Z Height during Z Scan sweep | 15.6
| TDK | 2023-10-12 | Full range of Z Detector during Z Scan sweep | 13.1
| National Chung Hsing University | 2023-10-12 | Full range of Z Detector during Z Scan sweep | 12.6
| TDK | 2023-10-12 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| National Chung Hsing University | 2023-10-12 | Full range of Y Detector Raw during Y Scan sweep | 14
| TDK | 2023-10-12 | Full range of Y Detector during Y Scan sweep | 105.2
| National Chung Hsing University | 2023-10-12 | Full range of Y Detector during Y Scan sweep | 121.4
| TDK | 2023-10-12 | Full range of X Detector Raw during X Scan sweep | 13.9
| National Chung Hsing University | 2023-10-12 | Full range of X Detector Raw during X Scan sweep | 14.3
| TDK | 2023-10-12 | Full range of X Detector during X Scan sweep | 107.5
| National Chung Hsing University | 2023-10-12 | Full range of X Detector during X Scan sweep | 120.2
| TDK | 2023-10-12 | A-B noise | 10.24
| National Chung Hsing University | 2023-10-12 | A-B noise | 10.14
| SZMTC | 2023-10-10 | Z Linearity by Z scanner position | 0.94
| SZMTC | 2023-10-10 | Z Linearity by Z scanner position | 0.35
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Linearity by Z scanner position | 1.19
| Shenzhen JWKW | 2023-10-10 | Z Linearity by Z scanner position | 0.66
| SZMTC | 2023-10-10 | Z Linearity by 3 different sample height | 1.002
| SZMTC | 2023-10-10 | Z Linearity by 3 different sample height | 1
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Linearity by 3 different sample height | 1.001
| Shenzhen JWKW | 2023-10-10 | Z Linearity by 3 different sample height | 0.999
| SZMTC | 2023-10-10 | Z Height noise | 0.036
| SZMTC | 2023-10-10 | Z Height noise | 0.047
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Height noise | 0.026
| Shenzhen JWKW | 2023-10-10 | Z Height noise | 0.026
| SZMTC | 2023-10-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.9
| SZMTC | 2023-10-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.3
| Shenzhen JWKW | 2023-10-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.3
| SZMTC | 2023-10-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.8
| SZMTC | 2023-10-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.5
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.3
| Shenzhen JWKW | 2023-10-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.4
| SZMTC | 2023-10-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.6
| SZMTC | 2023-10-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.9
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| Shenzhen JWKW | 2023-10-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.7
| SZMTC | 2023-10-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.6
| SZMTC | 2023-10-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.8
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.6
| Shenzhen JWKW | 2023-10-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.2
| SZMTC | 2023-10-10 | Z Drift (Short Traveling) | 0.06
| SZMTC | 2023-10-10 | Z Drift (Short Traveling) | 0.16
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Z Drift (Short Traveling) | 3.48
| Shenzhen JWKW | 2023-10-10 | Z Drift (Short Traveling) | 3.2
| SZMTC | 2023-10-10 | Y2 Detector noise (XY Servo on) | 0.28
| SZMTC | 2023-10-10 | Y2 Detector noise (XY Servo on) | 0.27
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y2 Detector noise (XY Servo on) | 0.44
| Shenzhen JWKW | 2023-10-10 | Y2 Detector noise (XY Servo on) | 0.34
| SZMTC | 2023-10-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 259.9
| SZMTC | 2023-10-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 255.7
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 242.7
| Shenzhen JWKW | 2023-10-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 252
| SZMTC | 2023-10-10 | Y1 Detector noise (XY Servo on) | 0.22
| SZMTC | 2023-10-10 | Y1 Detector noise (XY Servo on) | 0.3
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y1 Detector noise (XY Servo on) | 0.4
| Shenzhen JWKW | 2023-10-10 | Y1 Detector noise (XY Servo on) | 0.37
| SZMTC | 2023-10-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 248.2
| SZMTC | 2023-10-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 256.8
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 255.2
| Shenzhen JWKW | 2023-10-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 242.5
| SZMTC | 2023-10-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 163.6
| SZMTC | 2023-10-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.9
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 163.6
| Shenzhen JWKW | 2023-10-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 152.2
| SZMTC | 2023-10-10 | Y Ringing(Rpv) | 0.1
| SZMTC | 2023-10-10 | Y Ringing(Rpv) | 0.23
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y Ringing(Rpv) | 0.2
| Shenzhen JWKW | 2023-10-10 | Y Ringing(Rpv) | 0.18
| SZMTC | 2023-10-10 | Y Detector noise (XY Servo on) | 0.19
| SZMTC | 2023-10-10 | Y Detector noise (XY Servo on) | 0.21
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y Detector noise (XY Servo on) | 0.3
| Shenzhen JWKW | 2023-10-10 | Y Detector noise (XY Servo on) | 0.27
| SZMTC | 2023-10-10 | Y axis OPM | 0.94
| SZMTC | 2023-10-10 | Y axis OPM | 0.88
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Y axis OPM | 1.98
| Shenzhen JWKW | 2023-10-10 | Y axis OPM | 0.48
| SZMTC | 2023-10-10 | XY Orthogonality | 89.9
| SZMTC | 2023-10-10 | XY Orthogonality | 89.9
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | XY Orthogonality | 90.04
| Shenzhen JWKW | 2023-10-10 | XY Orthogonality | 90
| SZMTC | 2023-10-10 | X2 Detector noise (XY Servo on) | 0.34
| SZMTC | 2023-10-10 | X2 Detector noise (XY Servo on) | 0.34
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X2 Detector noise (XY Servo on) | 0.41
| Shenzhen JWKW | 2023-10-10 | X2 Detector noise (XY Servo on) | 0.4
| SZMTC | 2023-10-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 254.5
| SZMTC | 2023-10-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 263.8
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 233.4
| Shenzhen JWKW | 2023-10-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 245.9
| SZMTC | 2023-10-10 | X1 Detector noise (XY Servo on) | 0.34
| SZMTC | 2023-10-10 | X1 Detector noise (XY Servo on) | 0.36
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X1 Detector noise (XY Servo on) | 0.38
| Shenzhen JWKW | 2023-10-10 | X1 Detector noise (XY Servo on) | 0.42
| SZMTC | 2023-10-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 253
| SZMTC | 2023-10-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 257.8
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 230.7
| Shenzhen JWKW | 2023-10-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 248.2
| SZMTC | 2023-10-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 164.5
| SZMTC | 2023-10-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 170.8
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 164.5
| Shenzhen JWKW | 2023-10-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 158
| SZMTC | 2023-10-10 | X Ringing(Rpv) | 0.37
| SZMTC | 2023-10-10 | X Detector noise (XY Servo on) | 0.24
| SZMTC | 2023-10-10 | X Detector noise (XY Servo on) | 0.25
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X Detector noise (XY Servo on) | 0.28
| Shenzhen JWKW | 2023-10-10 | X Detector noise (XY Servo on) | 0.35
| SZMTC | 2023-10-10 | X axis OPM | 0.8
| SZMTC | 2023-10-10 | X axis OPM | 0.65
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | X axis OPM | 0.69
| Shenzhen JWKW | 2023-10-10 | X axis OPM | 0.85
| SZMTC | 2023-10-10 | StrokeYMin | -102188
| SZMTC | 2023-10-10 | StrokeYMin | -102108
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | StrokeYMin | -102219
| Shenzhen JWKW | 2023-10-10 | StrokeYMin | -101236
| SZMTC | 2023-10-10 | StrokeYMax | 103747.5
| SZMTC | 2023-10-10 | StrokeYMax | 103459.4
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | StrokeYMax | 103945
| Shenzhen JWKW | 2023-10-10 | StrokeYMax | 103723.5
| SZMTC | 2023-10-10 | StrokeY | 205935
| SZMTC | 2023-10-10 | StrokeY | 205567.5
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | StrokeY | 206163.8
| Shenzhen JWKW | 2023-10-10 | StrokeY | 204959.2
| SZMTC | 2023-10-10 | StrokeXMin | -116274
| SZMTC | 2023-10-10 | StrokeXMin | -116145
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | StrokeXMin | -116161
| Shenzhen JWKW | 2023-10-10 | StrokeXMin | -100396
| SZMTC | 2023-10-10 | StrokeXMax | 115285
| SZMTC | 2023-10-10 | StrokeXMax | 115696.3
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | StrokeXMax | 115811.3
| Shenzhen JWKW | 2023-10-10 | StrokeXMax | 106058.5
| SZMTC | 2023-10-10 | StrokeX | 231558.8
| SZMTC | 2023-10-10 | StrokeX | 231841.5
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | StrokeX | 231972.6
| Shenzhen JWKW | 2023-10-10 | StrokeX | 206454.6
| SZMTC | 2023-10-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.6
| SZMTC | 2023-10-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.4
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.4
| Shenzhen JWKW | 2023-10-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.76
| SZMTC | 2023-10-10 | RMS of ZeroScan | 0.036
| SZMTC | 2023-10-10 | RMS of ZeroScan | 0.048
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | RMS of ZeroScan | 0.03
| Shenzhen JWKW | 2023-10-10 | RMS of ZeroScan | 0.037
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| Shenzhen JWKW | 2023-10-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 3e-05
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusY | -0.0006
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusY | -0.0003
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | OpticalAxisShiftPerFocusY | 0.001
| Shenzhen JWKW | 2023-10-10 | OpticalAxisShiftPerFocusY | 0.0013
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.0002
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00025
| Shenzhen JWKW | 2023-10-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7e-05
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusX | 0.0026
| SZMTC | 2023-10-10 | OpticalAxisShiftPerFocusX | 0.0022
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | OpticalAxisShiftPerFocusX | 0.0021
| Shenzhen JWKW | 2023-10-10 | OpticalAxisShiftPerFocusX | 0.0023
| SZMTC | 2023-10-10 | NCM driving% with AC160TS cantilever | 12.7
| SZMTC | 2023-10-10 | NCM driving% with AC160TS cantilever | 17.03
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | NCM driving% with AC160TS cantilever | 10.86
| Shenzhen JWKW | 2023-10-10 | NCM driving% with AC160TS cantilever | 12.4
| SZMTC | 2023-10-10 | LowerLimit | 4840
| SZMTC | 2023-10-10 | LowerLimit | 4680
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | LowerLimit | 3955.6
| Shenzhen JWKW | 2023-10-10 | LowerLimit | 5908.8
| SZMTC | 2023-10-10 | Lateral(C-D) sensitivity | 13.9
| SZMTC | 2023-10-10 | Lateral(C-D) sensitivity | 14.4
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Lateral(C-D) sensitivity | 47.9
| Shenzhen JWKW | 2023-10-10 | Lateral(C-D) sensitivity | 47
| SZMTC | 2023-10-10 | Lateral(C-D) noise | 0.124
| SZMTC | 2023-10-10 | Lateral(C-D) noise | 0.15
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Lateral(C-D) noise | 0.141
| Shenzhen JWKW | 2023-10-10 | Lateral(C-D) noise | 0.23
| SZMTC | 2023-10-10 | Full range of Z Height during Z Scan sweep | 15.8
| SZMTC | 2023-10-10 | Full range of Z Height during Z Scan sweep | 16.9
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of Z Height during Z Scan sweep | 17.2
| Shenzhen JWKW | 2023-10-10 | Full range of Z Height during Z Scan sweep | 17.8
| SZMTC | 2023-10-10 | Full range of Z Detector during Z Scan sweep | 12.6
| SZMTC | 2023-10-10 | Full range of Z Detector during Z Scan sweep | 12.2
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of Z Detector during Z Scan sweep | 12.5
| Shenzhen JWKW | 2023-10-10 | Full range of Z Detector during Z Scan sweep | 13.1
| SZMTC | 2023-10-10 | Full range of Y2 Detector Raw during Y Scan sweep | 13.73
| SZMTC | 2023-10-10 | Full range of Y2 Detector Raw during Y Scan sweep | 13.7
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14.59
| Shenzhen JWKW | 2023-10-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14.76
| SZMTC | 2023-10-10 | Full range of Y2 Detector during Y Scan sweep | 164.95
| SZMTC | 2023-10-10 | Full range of Y2 Detector during Y Scan sweep | 162.71
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of Y2 Detector during Y Scan sweep | 191.1
| Shenzhen JWKW | 2023-10-10 | Full range of Y2 Detector during Y Scan sweep | 178.9
| SZMTC | 2023-10-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.17
| SZMTC | 2023-10-10 | Full range of Y1 Detector Raw during Y Scan sweep | 13.86
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.46
| Shenzhen JWKW | 2023-10-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.58
| SZMTC | 2023-10-10 | Full range of Y1 Detector during Y Scan sweep | 170.26
| SZMTC | 2023-10-10 | Full range of Y1 Detector during Y Scan sweep | 172.36
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of Y1 Detector during Y Scan sweep | 180.7
| Shenzhen JWKW | 2023-10-10 | Full range of Y1 Detector during Y Scan sweep | 185.5
| SZMTC | 2023-10-10 | Full range of X2 Detector Raw during X Scan sweep | 13.67
| SZMTC | 2023-10-10 | Full range of X2 Detector Raw during X Scan sweep | 13.84
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of X2 Detector Raw during X Scan sweep | 15.03
| Shenzhen JWKW | 2023-10-10 | Full range of X2 Detector Raw during X Scan sweep | 14.54
| SZMTC | 2023-10-10 | Full range of X2 Detector during X Scan sweep | 178.1
| SZMTC | 2023-10-10 | Full range of X2 Detector during X Scan sweep | 172.22
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of X2 Detector during X Scan sweep | 187.5
| Shenzhen JWKW | 2023-10-10 | Full range of X2 Detector during X Scan sweep | 169.6
| SZMTC | 2023-10-10 | Full range of X1 Detector Raw during X Scan sweep | 13.84
| SZMTC | 2023-10-10 | Full range of X1 Detector Raw during X Scan sweep | 13.92
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of X1 Detector Raw during X Scan sweep | 15.49
| Shenzhen JWKW | 2023-10-10 | Full range of X1 Detector Raw during X Scan sweep | 14.53
| SZMTC | 2023-10-10 | Full range of X1 Detector during X Scan sweep | 187.3
| SZMTC | 2023-10-10 | Full range of X1 Detector during X Scan sweep | 171.89
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Full range of X1 Detector during X Scan sweep | 189.6
| Shenzhen JWKW | 2023-10-10 | Full range of X1 Detector during X Scan sweep | 167.8
| SZMTC | 2023-10-10 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| SZMTC | 2023-10-10 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Shenzhen JWKW | 2023-10-10 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| SZMTC | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| SZMTC | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| Shenzhen JWKW | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| SZMTC | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.41
| SZMTC | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.5
| Shenzhen JWKW | 2023-10-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.15
| SZMTC | 2023-10-10 | 940nm Step Height Sample (10 Repeats) | 949.24
| SZMTC | 2023-10-10 | 940nm Step Height Sample (10 Repeats) | 948.42
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | 940nm Step Height Sample (10 Repeats) | 950.3
| Shenzhen JWKW | 2023-10-10 | 940nm Step Height Sample (10 Repeats) | 947.39
| SZMTC | 2023-10-10 | 50~200nm Step Height Sample (10 Repeats) | 183.02
| SZMTC | 2023-10-10 | 50~200nm Step Height Sample (10 Repeats) | 183.37
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | 50~200nm Step Height Sample (10 Repeats) | 185.1
| Shenzhen JWKW | 2023-10-10 | 50~200nm Step Height Sample (10 Repeats) | 183.42
| SZMTC | 2023-10-10 | 200~500nm Step Height Sample (10 Repeats) | 461.56
| SZMTC | 2023-10-10 | 200~500nm Step Height Sample (10 Repeats) | 462.29
| Quanzhou San'an Semiconductor Tech #2 | 2023-10-10 | 200~500nm Step Height Sample (10 Repeats) | 461.79
| Shenzhen JWKW | 2023-10-10 | 200~500nm Step Height Sample (10 Repeats) | 457.39
| SZMTC | 2023-10-10 | * 1800nm Step Height Sample (10 Repeats) | 1784.57
| SZMTC | 2023-10-10 | * 1800nm Step Height Sample (10 Repeats) | 1780.67
| SZMTC | 2023-10-10 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 1.5
| SZMTC | 2023-10-10 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.2
| SZMTC | 2023-10-10 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 1.8
| SZMTC | 2023-10-10 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.3
| SZMTC | 2023-10-10 | %GRR w/ roughness of non-patterened wafer | 2.3
| SZMTC | 2023-10-10 | %GRR w/ roughness of non-patterened wafer | 3.2
| Brookhaven National Lab(BNL) | 2023-10-06 | Z Linearity by Z Scanner position
| Brookhaven National Lab(BNL) | 2023-10-06 | Z Linearity by 3 different sample height | 1.002
| Brookhaven National Lab(BNL) | 2023-10-06 | Z Height noise | 0.021
| Brookhaven National Lab(BNL) | 2023-10-06 | Y Detector noise (XY Servo on) | 0.29
| Brookhaven National Lab(BNL) | 2023-10-06 | Y Detector noise (XY Servo off) | 0.28
| Brookhaven National Lab(BNL) | 2023-10-06 | Y axis OPM | 0.65
| Brookhaven National Lab(BNL) | 2023-10-06 | X Detector noise (XY Servo on) | 0.28
| Brookhaven National Lab(BNL) | 2023-10-06 | X Detector noise (XY Servo off) | 0.27
| Brookhaven National Lab(BNL) | 2023-10-06 | X axis OPM | 0.78
| Brookhaven National Lab(BNL) | 2023-10-06 | SLD beam power at the end of the lens | 0.26
| Brookhaven National Lab(BNL) | 2023-10-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Brookhaven National Lab(BNL) | 2023-10-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Brookhaven National Lab(BNL) | 2023-10-06 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Brookhaven National Lab(BNL) | 2023-10-06 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Brookhaven National Lab(BNL) | 2023-10-06 | NCM driving% with AC160TS cantilever | 1.322
| Brookhaven National Lab(BNL) | 2023-10-06 | NCM Amplitude noise | 16.47
| Brookhaven National Lab(BNL) | 2023-10-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.312
| Brookhaven National Lab(BNL) | 2023-10-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.473
| Brookhaven National Lab(BNL) | 2023-10-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160.6
| Brookhaven National Lab(BNL) | 2023-10-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156.5
| Brookhaven National Lab(BNL) | 2023-10-06 | Full range of Z Height during Z Scan sweep | 16.4
| Brookhaven National Lab(BNL) | 2023-10-06 | LD beam power at the end of the lens | 2.6
| Brookhaven National Lab(BNL) | 2023-10-06 | Lateral(C-D) sensitivity | 14.9
| Brookhaven National Lab(BNL) | 2023-10-06 | Lateral(C-D) noise | 0.155
| Brookhaven National Lab(BNL) | 2023-10-06 | Full range of Z Detector during Z Scan sweep | 13.4
| Brookhaven National Lab(BNL) | 2023-10-06 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| Brookhaven National Lab(BNL) | 2023-10-06 | Full range of Y Detector during Y Scan sweep | 114.6
| Brookhaven National Lab(BNL) | 2023-10-06 | Full range of X Detector Raw during X Scan sweep | 14.3
| Brookhaven National Lab(BNL) | 2023-10-06 | Full range of X Detector during X Scan sweep | 112
| Brookhaven National Lab(BNL) | 2023-10-06 | EFM Phase : For the positive bias | 90.4
| Brookhaven National Lab(BNL) | 2023-10-06 | EFM Phase : For the negative bias | -89.5
| Brookhaven National Lab(BNL) | 2023-10-06 | Approach position at -3μm | 0.007
| Brookhaven National Lab(BNL) | 2023-10-06 | Approach position at +3μm | 0.0072
| Brookhaven National Lab(BNL) | 2023-10-06 | A-B noise | 15.49
| SUMCO | 2023-10-05 | Z Linearity by Z scanner position | 1.12
| SUMCO | 2023-10-05 | Z Linearity by 5 different sample height | 1.007
| SUMCO | 2023-10-05 | Z Height noise | 0.028
| SUMCO | 2023-10-05 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.2
| SUMCO | 2023-10-05 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.3
| SUMCO | 2023-10-05 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2
| SUMCO | 2023-10-05 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3
| SUMCO | 2023-10-05 | Z Drift (Short Traveling) | 0.17
| SUMCO | 2023-10-05 | Y2 Detector noise (XY Servo on) | 0.44
| SUMCO | 2023-10-05 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 282.1
| SUMCO | 2023-10-05 | Y1 Detector noise (XY Servo on) | 0.43
| SUMCO | 2023-10-05 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 285.8
| SUMCO | 2023-10-05 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 188.8
| SUMCO | 2023-10-05 | Y Ringing(Rpv) | 0.1
| SUMCO | 2023-10-05 | Y Non-linearity error | 3.11
| SUMCO | 2023-10-05 | Y Detector noise (XY Servo on) | 0.33
| SUMCO | 2023-10-05 | Y axis, Repeatability @1st position | 0.05
| SUMCO | 2023-10-05 | Y axis, Repeatability @ 4th position | 0.03
| SUMCO | 2023-10-05 | Y axis, Repeatability @ 3rd position | 0.03
| SUMCO | 2023-10-05 | Y axis, Repeatability @ 2nd position | 0.07
| SUMCO | 2023-10-05 | Y axis, Accuracy @1st position | 0.57
| SUMCO | 2023-10-05 | Y axis, Accuracy @ 4th position | 0.45
| SUMCO | 2023-10-05 | Y axis, Accuracy @ 3rd position | 0.44
| SUMCO | 2023-10-05 | Y axis, Accuracy @ 2nd position | 0.69
| SUMCO | 2023-10-05 | Y axis OPM | 0.58
| SUMCO | 2023-10-05 | XY Orthogonality | 89.9
| SUMCO | 2023-10-05 | X2 Detector noise (XY Servo on) | 0.4
| SUMCO | 2023-10-05 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 284.1
| SUMCO | 2023-10-05 | X1 Detector noise (XY Servo on) | 0.4
| SUMCO | 2023-10-05 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 277.7
| SUMCO | 2023-10-05 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.1
| SUMCO | 2023-10-05 | X Ringing(Rpv) | 0.1
| SUMCO | 2023-10-05 | X Non-linearity error | 3.24
| SUMCO | 2023-10-05 | X Detector noise (XY Servo on) | 0.31
| SUMCO | 2023-10-05 | X axis, Repeatability @1st position | 0.03
| SUMCO | 2023-10-05 | X axis, Repeatability @ 4th position | 0.03
| SUMCO | 2023-10-05 | X axis, Repeatability @ 3rd position | 0.07
| SUMCO | 2023-10-05 | X axis, Repeatability @ 2nd position | 0.07
| SUMCO | 2023-10-05 | X axis, Accuracy @1st position | 0.22
| SUMCO | 2023-10-05 | X axis, Accuracy @ 4th position | 0.15
| SUMCO | 2023-10-05 | X axis, Accuracy @ 3rd position | 0.37
| SUMCO | 2023-10-05 | X axis, Accuracy @ 2nd position | 0.31
| SUMCO | 2023-10-05 | X axis OPM | 0.81
| SUMCO | 2023-10-05 | Under 50nm Step Height Sample (10 Repeats) | 18.08
| SUMCO | 2023-10-05 | StrokeYMin | -153105
| SUMCO | 2023-10-05 | StrokeYMax | 150717
| SUMCO | 2023-10-05 | StrokeY | 303821.9
| SUMCO | 2023-10-05 | StrokeXMin | -204111
| SUMCO | 2023-10-05 | StrokeXMax | 203569.9
| SUMCO | 2023-10-05 | StrokeX | 407681.2
| SUMCO | 2023-10-05 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.99
| SUMCO | 2023-10-05 | RMS of ZeroScan | 0.032
| SUMCO | 2023-10-05 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00013
| SUMCO | 2023-10-05 | OpticalAxisShiftPerFocusY | -0.0031
| SUMCO | 2023-10-05 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00056
| SUMCO | 2023-10-05 | OpticalAxisShiftPerFocusX | 0.006
| SUMCO | 2023-10-05 | NCM driving% with AC160TS cantilever | 5.58
| SUMCO | 2023-10-05 | Min. Vibration Transmissbility (~100Hz) | 0.0002
| SUMCO | 2023-10-05 | Lateral(C-D) sensitivity | 117.336
| SUMCO | 2023-10-05 | Lateral(C-D) noise | 0.195
| SUMCO | 2023-10-05 | Image offset after moving from -Y to +Y (Y scan) | 52.73
| SUMCO | 2023-10-05 | Image offset after moving from -Y to +Y (X scan) | 50.54
| SUMCO | 2023-10-05 | Image offset after moving from -X to +X (Y scan) | 13.92
| SUMCO | 2023-10-05 | Image offset after moving from -X to +X (X scan) | 84.23
| SUMCO | 2023-10-05 | Image offset after moving from +Y to -Y (Y scan) | 93.75
| SUMCO | 2023-10-05 | Image offset after moving from +Y to -Y (X scan) | 24.9
| SUMCO | 2023-10-05 | Image offset after moving from +X to -X (Y scan) | 30.76
| SUMCO | 2023-10-05 | Image offset after moving from +X to -X (X scan) | 73.97
| SUMCO | 2023-10-05 | Full range of Z Height during Z Scan sweep | 16.3
| SUMCO | 2023-10-05 | Full range of Z Detector during Z Scan sweep | 13.5
| SUMCO | 2023-10-05 | Full range of Y2 Detector Raw during Y Scan sweep | 13.89
| SUMCO | 2023-10-05 | Full range of Y2 Detector during Y Scan sweep | 210.8541
| SUMCO | 2023-10-05 | Full range of Y1 Detector Raw during Y Scan sweep | 14.12
| SUMCO | 2023-10-05 | Full range of Y1 Detector during Y Scan sweep | 201.3093
| SUMCO | 2023-10-05 | Full range of X2 Detector Raw during X Scan sweep | 14.01
| SUMCO | 2023-10-05 | Full range of X2 Detector during X Scan sweep | 208.3893
| SUMCO | 2023-10-05 | Full range of X1 Detector Raw during X Scan sweep | 14.02
| SUMCO | 2023-10-05 | Full range of X1 Detector during X Scan sweep | 211.0628
| SUMCO | 2023-10-05 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| SUMCO | 2023-10-05 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| SUMCO | 2023-10-05 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.2
| SUMCO | 2023-10-05 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24
| SUMCO | 2023-10-05 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| SUMCO | 2023-10-05 | 940nm Step Height Sample (10 Repeats) | 948.44
| SUMCO | 2023-10-05 | 50~200nm Step Height Sample (10 Repeats) | 183.17
| SUMCO | 2023-10-05 | 200~500nm Step Height Sample (10 Repeats) | 462.44
| SUMCO | 2023-10-05 | * 1800nm Step Height Sample (10 Repeats) | 1794.3
| SUMCO | 2023-10-05 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.5
| SUMCO | 2023-10-05 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4
| SUMCO | 2023-10-05 | %GRR w/ roughness of non-patterened wafer | 4.7
| SUMCO #2 | 2023-10-04 | Z Linearity by Z scanner position | 1.12
| SUMCO #2 | 2023-10-04 | Z Linearity by 5 different sample height | 1.007
| SUMCO #2 | 2023-10-04 | Z Height noise | 0.028
| SUMCO #2 | 2023-10-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.1804
| SUMCO #2 | 2023-10-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.2552
| SUMCO #2 | 2023-10-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2496
| SUMCO #2 | 2023-10-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2577
| SUMCO #2 | 2023-10-04 | Y2 Detector noise (XY Servo on) | 0.437
| SUMCO #2 | 2023-10-04 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 282.08
| SUMCO #2 | 2023-10-04 | Y1 Detector noise (XY Servo on) | 0.433
| SUMCO #2 | 2023-10-04 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 285.78
| SUMCO #2 | 2023-10-04 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 188.786
| SUMCO #2 | 2023-10-04 | Y Ringing(Rpv) | 0.099
| SUMCO #2 | 2023-10-04 | Y Non-linearity error | 3.11
| SUMCO #2 | 2023-10-04 | Y Detector noise (XY Servo on) | 0.327
| SUMCO #2 | 2023-10-04 | Y axis, Repeatability @1st position | 0.05
| SUMCO #2 | 2023-10-04 | Y axis, Repeatability @ 4th position | 0.03
| SUMCO #2 | 2023-10-04 | Y axis, Repeatability @ 3rd position | 0.03
| SUMCO #2 | 2023-10-04 | Y axis, Repeatability @ 2nd position | 0.07
| SUMCO #2 | 2023-10-04 | Y axis, Accuracy @1st position | 0.57
| SUMCO #2 | 2023-10-04 | Y axis, Accuracy @ 4th position | 0.45
| SUMCO #2 | 2023-10-04 | Y axis, Accuracy @ 3rd position | 0.44
| SUMCO #2 | 2023-10-04 | Y axis, Accuracy @ 2nd position | 0.69
| SUMCO #2 | 2023-10-04 | Y axis OPM | 0.578
| SUMCO #2 | 2023-10-04 | X2 Detector noise (XY Servo on) | 0.4
| SUMCO #2 | 2023-10-04 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 284.15
| SUMCO #2 | 2023-10-04 | X1 Detector noise (XY Servo on) | 0.397
| SUMCO #2 | 2023-10-04 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 277.7
| SUMCO #2 | 2023-10-04 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.057
| SUMCO #2 | 2023-10-04 | X Ringing(Rpv) | 0.102
| SUMCO #2 | 2023-10-04 | X Non-linearity error | 3.24
| SUMCO #2 | 2023-10-04 | X Detector noise (XY Servo on) | 0.314
| SUMCO #2 | 2023-10-04 | X axis, Repeatability @1st position | 0.03
| SUMCO #2 | 2023-10-04 | X axis, Repeatability @ 4th position | 0.03
| SUMCO #2 | 2023-10-04 | X axis, Repeatability @ 3rd position | 0.07
| SUMCO #2 | 2023-10-04 | X axis, Repeatability @ 2nd position | 0.07
| SUMCO #2 | 2023-10-04 | X axis, Accuracy @1st position | 0.22
| SUMCO #2 | 2023-10-04 | X axis, Accuracy @ 4th position | 0.15
| SUMCO #2 | 2023-10-04 | X axis, Accuracy @ 3rd position | 0.37
| SUMCO #2 | 2023-10-04 | X axis, Accuracy @ 2nd position | 0.31
| SUMCO #2 | 2023-10-04 | X axis OPM | 0.81
| SUMCO #2 | 2023-10-04 | Under 50nm Step Height Sample (10 Repeats) | 18.079
| SUMCO #2 | 2023-10-04 | StrokeYMin | -153105
| SUMCO #2 | 2023-10-04 | StrokeYMax | 150717
| SUMCO #2 | 2023-10-04 | StrokeY | 303821.9
| SUMCO #2 | 2023-10-04 | StrokeXMin | -204111
| SUMCO #2 | 2023-10-04 | StrokeXMax | 203569.9
| SUMCO #2 | 2023-10-04 | StrokeX | 407681.2
| SUMCO #2 | 2023-10-04 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.987
| SUMCO #2 | 2023-10-04 | RMS of ZeroScan | 0.032
| SUMCO #2 | 2023-10-04 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.000125
| SUMCO #2 | 2023-10-04 | OpticalAxisShiftPerFocusY | -0.0031
| SUMCO #2 | 2023-10-04 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000556
| SUMCO #2 | 2023-10-04 | OpticalAxisShiftPerFocusX | 0.006
| SUMCO #2 | 2023-10-04 | NCM driving% with AC160TS cantilever | 5.58
| SUMCO #2 | 2023-10-04 | Min. Vibration Transmissbility (~100Hz) | 0.0002
| SUMCO #2 | 2023-10-04 | Lateral(C-D) sensitivity | 117.336
| SUMCO #2 | 2023-10-04 | Lateral(C-D) noise | 0.195
| SUMCO #2 | 2023-10-04 | Laser intensity with AC160 cantilever | 2.911
| SUMCO #2 | 2023-10-04 | Image offset after moving from -Y to +Y (Y scan) | 52.73
| SUMCO #2 | 2023-10-04 | Image offset after moving from -Y to +Y (X scan) | 50.54
| SUMCO #2 | 2023-10-04 | Image offset after moving from -X to +X (Y scan) | 13.92
| SUMCO #2 | 2023-10-04 | Image offset after moving from -X to +X (X scan) | 84.23
| SUMCO #2 | 2023-10-04 | Image offset after moving from +Y to -Y (Y scan) | 93.75
| SUMCO #2 | 2023-10-04 | Image offset after moving from +Y to -Y (X scan) | 24.9
| SUMCO #2 | 2023-10-04 | Image offset after moving from +X to -X (Y scan) | 30.76
| SUMCO #2 | 2023-10-04 | Image offset after moving from +X to -X (X scan) | 73.97
| SUMCO #2 | 2023-10-04 | Full range of Z Height during Z Scan sweep | 16.3472
| SUMCO #2 | 2023-10-04 | Full range of Z Detector during Z Scan sweep | 13.47
| SUMCO #2 | 2023-10-04 | Full range of Y2 Detector Raw during Y Scan sweep | 13.89
| SUMCO #2 | 2023-10-04 | Full range of Y2 Detector during Y Scan sweep | 210.8541
| SUMCO #2 | 2023-10-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.12
| SUMCO #2 | 2023-10-04 | Full range of Y1 Detector during Y Scan sweep | 201.3093
| SUMCO #2 | 2023-10-04 | Full range of X2 Detector Raw during X Scan sweep | 14.01
| SUMCO #2 | 2023-10-04 | Full range of X2 Detector during X Scan sweep | 208.3893
| SUMCO #2 | 2023-10-04 | Full range of X1 Detector Raw during X Scan sweep | 14.02
| SUMCO #2 | 2023-10-04 | Full range of X1 Detector during X Scan sweep | 211.0628
| SUMCO #2 | 2023-10-04 | Allowable max warpage, w/ SiC sample chuck | 300
| SUMCO #2 | 2023-10-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| SUMCO #2 | 2023-10-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.2
| SUMCO #2 | 2023-10-04 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24
| SUMCO #2 | 2023-10-04 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| SUMCO #2 | 2023-10-04 | A－B difference during 3 times run test | 0.19
| SUMCO #2 | 2023-10-04 | A＋B difference during 3 times run test | 0.02
| SUMCO #2 | 2023-10-04 | 940nm Step Height Sample (10 Repeats) | 948.441
| SUMCO #2 | 2023-10-04 | 50~200nm Step Height Sample (10 Repeats) | 183.168
| SUMCO #2 | 2023-10-04 | 200~500nm Step Height Sample (10 Repeats) | 462.442
| SUMCO #2 | 2023-10-04 | * 1800nm Step Height Sample (10 Repeats) | 1794.299
| SUMCO #2 | 2023-10-04 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.5
| SUMCO #2 | 2023-10-04 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4
| SUMCO #2 | 2023-10-04 | %GRR w/ roughness of non-patterened wafer | 4.7
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Z Scanner position at -3μm | 0.012
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Z Scanner position at -3μm | 0.0033
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Z Scanner position at -3μm | 0.0014
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Z Scanner position at +3μm | 0.0064
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Z Scanner position at +3μm | 0.0084
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Z Scanner position at +3μm | 0.0068
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Z Linearity by Z Scanner position
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Z Linearity by Z Scanner position
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Z Linearity by Z Scanner position
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Z Linearity by 3 different sample height | 1.008
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Z Linearity by 3 different sample height | 0.999
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Z Linearity by 3 different sample height | 0.996
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Z Height noise | 0.027
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Z Height noise | 0.027
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Z Height noise | 0.029
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y2 Detector noise (XY Servo on) | 0.31
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y2 Detector noise (XY Servo on) | 0.34
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y2 Detector noise (XY Servo off) | 0.3
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y2 Detector noise (XY Servo off) | 0.22
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y1 Detector noise (XY Servo on) | 0.3
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y1 Detector noise (XY Servo on) | 0.32
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y1 Detector noise (XY Servo off) | 0.29
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y1 Detector noise (XY Servo off) | 0.22
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y Scan Stroke | 171.4
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y Scan Stroke | 181.6
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Y Scan Stroke | 55
| Hebei Synlight Semiconductor #7 | 2023-10-04 | -Y Direction software Limit | -102953
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y Detector noise (XY Servo on) | 0.22
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y Detector noise (XY Servo on) | 0.26
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Y Detector noise (XY Servo on) | 0.28
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y Detector noise (XY Servo off) | 0.21
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y Detector noise (XY Servo off) | 0.16
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Y Detector noise (XY Servo off) | 0.27
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Y axis OPM | 1.11
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Y axis OPM | 1.14
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Y axis OPM | 1.26
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X2 Detector noise (XY Servo on) | 0.31
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X2 Detector noise (XY Servo on) | 0.34
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X2 Detector noise (XY Servo off) | 0.3
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X2 Detector noise (XY Servo off) | 0.21
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X1 Detector noise (XY Servo on) | 0.32
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X1 Detector noise (XY Servo on) | 0.35
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X1 Detector noise (XY Servo off) | 0.31
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X1 Detector noise (XY Servo off) | 0.22
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X Scan Stroke | 173
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X Scan Stroke | 172.7
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | X Scan Stroke | 55
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | -X Direction software Limit | -102246
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | -X Direction software Limit | -103054
| Hebei Synlight Semiconductor #7 | 2023-10-04 | -X Direction software Limit | -102140
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X Detector noise (XY Servo on) | 0.23
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X Detector noise (XY Servo on) | 0.28
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | X Detector noise (XY Servo on) | 0.29
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X Detector noise (XY Servo off) | 0.22
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X Detector noise (XY Servo off) | 0.16
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | X Detector noise (XY Servo off) | 0.28
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | X axis OPM | 1.25
| Hebei Synlight Semiconductor #7 | 2023-10-04 | X axis OPM | 1.06
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | X axis OPM | 1.44
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.33
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.36
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.31
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.3
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.29
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Hebei Synlight Semiconductor #7 | 2023-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | NCM driving% with AC160TS cantilever | 0.96
| Hebei Synlight Semiconductor #7 | 2023-10-04 | NCM driving% with AC160TS cantilever | 0.19
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | NCM driving% with AC160TS cantilever | 0.24
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | NCM Amplitude noise | 13.78
| Hebei Synlight Semiconductor #7 | 2023-10-04 | NCM Amplitude noise | 14.74
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | NCM Amplitude noise | 17.82
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.124
| Hebei Synlight Semiconductor #7 | 2023-10-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.651
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.281
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.459
| Hebei Synlight Semiconductor #7 | 2023-10-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.648
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.297
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.2
| Hebei Synlight Semiconductor #7 | 2023-10-04 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 252.2
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 239
| Hebei Synlight Semiconductor #7 | 2023-10-04 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 246.4
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.6
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 247.6
| Hebei Synlight Semiconductor #7 | 2023-10-04 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.1
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 243.1
| Hebei Synlight Semiconductor #7 | 2023-10-04 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 244.9
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.3
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Lateral(C-D) sensitivity | 3.4
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Lateral(C-D) sensitivity | 3.8
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Lateral(C-D) sensitivity | 22.8
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Lateral(C-D) noise | 0.159
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Lateral(C-D) noise | 0.123
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Lateral(C-D) noise | 0.062
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of Z Height during Z Scan sweep | 15.9
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of Z Height during Z Scan sweep | 16.8
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Full range of Z Height during Z Scan sweep | 16
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of Z Detector during Z Scan sweep | 11.6
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of Z Detector during Z Scan sweep | 12.2
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Full range of Z Detector during Z Scan sweep | 11.7
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of Y2 Detector Raw during Y Scan sweep | 13.9
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.8
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of Y Detector during Y Scan sweep | 169.7
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of Y Detector during Y Scan sweep | 182.8
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Full range of Y Detector during Y Scan sweep | 72.2
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of X2 Detector Raw during X Scan sweep | 14
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Full range of X Detector Raw during X Scan sweep | 14.2
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | Full range of X Detector during X Scan sweep | 172.4
| Hebei Synlight Semiconductor #7 | 2023-10-04 | Full range of X Detector during X Scan sweep | 172.5
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | Full range of X Detector during X Scan sweep | 71.8
| Fraunhofer Research Institution for Energy Infrastructures and Geothermal Systems IEG (EMFT) | 2023-10-04 | A-B noise | 11.36
| Hebei Synlight Semiconductor #7 | 2023-10-04 | A-B noise | 10.37
| 경북하이브리드연구원 (다민사이언스) | 2023-10-04 | A-B noise | 10.8
| Zhejiang University_Quzhou | 2023-09-26 | Z Scanner position at -3μm | 0.0039
| Zhejiang University_Quzhou | 2023-09-26 | Z Scanner position at +3μm | 0.0157
| Zhejiang University_Quzhou | 2023-09-26 | Z Linearity by Z Scanner position
| Zhejiang University_Quzhou | 2023-09-26 | Z Linearity by 3 different sample height | 0.992
| Zhejiang University_Quzhou | 2023-09-26 | Z Height noise | 0.027
| Zhejiang University_Quzhou | 2023-09-26 | Y Detector noise (XY Servo on) | 0.27
| Zhejiang University_Quzhou | 2023-09-26 | Y Detector noise (XY Servo off) | 0.27
| Zhejiang University_Quzhou | 2023-09-26 | Y axis OPM | 0.46
| Zhejiang University_Quzhou | 2023-09-26 | X Detector noise (XY Servo on) | 0.26
| Zhejiang University_Quzhou | 2023-09-26 | X Detector noise (XY Servo off) | 0.26
| Zhejiang University_Quzhou | 2023-09-26 | X axis OPM | 0.94
| Zhejiang University_Quzhou | 2023-09-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Zhejiang University_Quzhou | 2023-09-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Zhejiang University_Quzhou | 2023-09-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Zhejiang University_Quzhou | 2023-09-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Zhejiang University_Quzhou | 2023-09-26 | NCM driving% with AC160TS cantilever | 0.93
| Zhejiang University_Quzhou | 2023-09-26 | NCM Amplitude noise | 17.14
| Zhejiang University_Quzhou | 2023-09-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.963
| Zhejiang University_Quzhou | 2023-09-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.488
| Zhejiang University_Quzhou | 2023-09-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 152.9
| Zhejiang University_Quzhou | 2023-09-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.5
| Zhejiang University_Quzhou | 2023-09-26 | Lateral(C-D) sensitivity | 10.3
| Zhejiang University_Quzhou | 2023-09-26 | Lateral(C-D) noise | 0.119
| Zhejiang University_Quzhou | 2023-09-26 | Full range of Z Height during Z Scan sweep | 15.7
| Zhejiang University_Quzhou | 2023-09-26 | Full range of Z Detector during Z Scan sweep | 12.8
| Zhejiang University_Quzhou | 2023-09-26 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Zhejiang University_Quzhou | 2023-09-26 | Full range of Y Detector during Y Scan sweep | 110.2
| Zhejiang University_Quzhou | 2023-09-26 | Full range of X Detector Raw during X Scan sweep | 14.2
| Zhejiang University_Quzhou | 2023-09-26 | Full range of X Detector during X Scan sweep | 110.1
| Zhejiang University_Quzhou | 2023-09-26 | A-B noise | 9.36
| East China University of Science and Technology | 2023-09-25 | Z Scanner position at -3μm | 0.0064
| East China University of Science and Technology | 2023-09-25 | Z Scanner position at +3μm | 0.0165
| East China University of Science and Technology | 2023-09-25 | Z Linearity by Z Scanner position
| East China University of Science and Technology | 2023-09-25 | Z Linearity by 3 different sample height | 1.006
| East China University of Science and Technology | 2023-09-25 | Z Height noise | 0.028
| East China University of Science and Technology | 2023-09-25 | Y Scan Stroke | 55
| East China University of Science and Technology | 2023-09-25 | Y Detector noise (XY Servo on) | 0.27
| East China University of Science and Technology | 2023-09-25 | Y Detector noise (XY Servo off) | 0.27
| East China University of Science and Technology | 2023-09-25 | Y axis OPM | 0.27
| East China University of Science and Technology | 2023-09-25 | X Scan Stroke | 55
| East China University of Science and Technology | 2023-09-25 | X Detector noise (XY Servo on) | 0.27
| East China University of Science and Technology | 2023-09-25 | X Detector noise (XY Servo off) | 0.27
| East China University of Science and Technology | 2023-09-25 | X axis OPM | 1.32
| East China University of Science and Technology | 2023-09-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| East China University of Science and Technology | 2023-09-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| East China University of Science and Technology | 2023-09-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| East China University of Science and Technology | 2023-09-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| East China University of Science and Technology | 2023-09-25 | NCM driving% with AC160TS cantilever | 1.33
| East China University of Science and Technology | 2023-09-25 | NCM Amplitude noise | 15.31
| East China University of Science and Technology | 2023-09-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.112
| East China University of Science and Technology | 2023-09-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.415
| East China University of Science and Technology | 2023-09-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101
| East China University of Science and Technology | 2023-09-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.2
| East China University of Science and Technology | 2023-09-25 | Lateral(C-D) sensitivity | 5.5
| East China University of Science and Technology | 2023-09-25 | Lateral(C-D) noise | 0.08
| East China University of Science and Technology | 2023-09-25 | Full range of Z Height during Z Scan sweep | 15.9
| East China University of Science and Technology | 2023-09-25 | Full range of Z Detector during Z Scan sweep | 12
| East China University of Science and Technology | 2023-09-25 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| East China University of Science and Technology | 2023-09-25 | Full range of Y Detector during Y Scan sweep | 69.6
| East China University of Science and Technology | 2023-09-25 | Full range of X Detector Raw during X Scan sweep | 13.9
| East China University of Science and Technology | 2023-09-25 | Full range of X Detector during X Scan sweep | 70.8
| East China University of Science and Technology | 2023-09-25 | A-B noise | 10.17
| TYSiC#3 | 2023-09-20 | Z Scanner position at -3μm | 0.002445
| TYSiC#4 | 2023-09-20 | Z Scanner position at -3μm | 0.0005
| TYSiC#5 | 2023-09-20 | Z Scanner position at -3μm | 0.0076
| TYSiC#6 | 2023-09-20 | Z Scanner position at -3μm | 0.0005
| TYSiC#7 | 2023-09-20 | Z Scanner position at -3μm | 0.0056
| 경상국립대학교 반도체공학과 | 2023-09-20 | Z Scanner position at -3μm | 0.0024
| TYSiC#3 | 2023-09-20 | Z Scanner position at +3μm | 0.008226
| TYSiC#4 | 2023-09-20 | Z Scanner position at +3μm | 0.0105
| TYSiC#5 | 2023-09-20 | Z Scanner position at +3μm | 0.0075
| TYSiC#6 | 2023-09-20 | Z Scanner position at +3μm | 0.0194
| TYSiC#7 | 2023-09-20 | Z Scanner position at +3μm | 0.0127
| 경상국립대학교 반도체공학과 | 2023-09-20 | Z Scanner position at +3μm | 0.0138
| TYSiC#3 | 2023-09-20 | Z Linearity by Z Scanner position
| TYSiC#4 | 2023-09-20 | Z Linearity by Z Scanner position
| TYSiC#5 | 2023-09-20 | Z Linearity by Z Scanner position
| TYSiC#6 | 2023-09-20 | Z Linearity by Z Scanner position
| TYSiC#7 | 2023-09-20 | Z Linearity by Z Scanner position
| 경상국립대학교 반도체공학과 | 2023-09-20 | Z Linearity by Z Scanner position
| TYSiC#3 | 2023-09-20 | Z Linearity by 3 different sample height | 1.0038
| TYSiC#4 | 2023-09-20 | Z Linearity by 3 different sample height | 0.992
| TYSiC#5 | 2023-09-20 | Z Linearity by 3 different sample height | 0.991
| TYSiC#6 | 2023-09-20 | Z Linearity by 3 different sample height | 0.991
| TYSiC#7 | 2023-09-20 | Z Linearity by 3 different sample height | 0.983
| 경상국립대학교 반도체공학과 | 2023-09-20 | Z Linearity by 3 different sample height | 1.004
| TYSiC#3 | 2023-09-20 | Z Height noise | 0.025
| TYSiC#4 | 2023-09-20 | Z Height noise | 0.036
| TYSiC#5 | 2023-09-20 | Z Height noise | 0.029
| TYSiC#6 | 2023-09-20 | Z Height noise | 0.036
| TYSiC#7 | 2023-09-20 | Z Height noise | 0.039
| 경상국립대학교 반도체공학과 | 2023-09-20 | Z Height noise | 0.026
| TYSiC#3 | 2023-09-20 | Y Scan Stroke | 132.9
| TYSiC#4 | 2023-09-20 | Y Scan Stroke | 130.6
| TYSiC#5 | 2023-09-20 | Y Scan Stroke | 128.6
| TYSiC#6 | 2023-09-20 | Y Scan Stroke | 131.9
| TYSiC#7 | 2023-09-20 | Y Scan Stroke | 129
| 경상국립대학교 반도체공학과 | 2023-09-20 | Y Scan Stroke | 55
| TYSiC#3 | 2023-09-20 | -Y Direction software Limit | -100349
| TYSiC#5 | 2023-09-20 | -Y Direction software Limit | -100691
| TYSiC#7 | 2023-09-20 | -Y Direction software Limit | -100468
| TYSiC#3 | 2023-09-20 | Y Detector noise (XY Servo on) | 0.263
| TYSiC#4 | 2023-09-20 | Y Detector noise (XY Servo on) | 0.28
| TYSiC#5 | 2023-09-20 | Y Detector noise (XY Servo on) | 0.28
| TYSiC#6 | 2023-09-20 | Y Detector noise (XY Servo on) | 0.25
| TYSiC#7 | 2023-09-20 | Y Detector noise (XY Servo on) | 0.27
| 경상국립대학교 반도체공학과 | 2023-09-20 | Y Detector noise (XY Servo on) | 0.35
| TYSiC#3 | 2023-09-20 | Y Detector noise (XY Servo off) | 0.251
| TYSiC#4 | 2023-09-20 | Y Detector noise (XY Servo off) | 0.28
| TYSiC#5 | 2023-09-20 | Y Detector noise (XY Servo off) | 0.26
| TYSiC#6 | 2023-09-20 | Y Detector noise (XY Servo off) | 0.24
| TYSiC#7 | 2023-09-20 | Y Detector noise (XY Servo off) | 0.26
| 경상국립대학교 반도체공학과 | 2023-09-20 | Y Detector noise (XY Servo off) | 0.35
| TYSiC#3 | 2023-09-20 | Y axis OPM | 0.5564
| TYSiC#4 | 2023-09-20 | Y axis OPM | 0.99
| TYSiC#5 | 2023-09-20 | Y axis OPM | 0.82
| TYSiC#6 | 2023-09-20 | Y axis OPM | 0.92
| TYSiC#7 | 2023-09-20 | Y axis OPM | 0.57
| 경상국립대학교 반도체공학과 | 2023-09-20 | Y axis OPM | 0.8
| TYSiC#3 | 2023-09-20 | X Scan Stroke | 127.3
| TYSiC#4 | 2023-09-20 | X Scan Stroke | 130.7
| TYSiC#5 | 2023-09-20 | X Scan Stroke | 130.7
| TYSiC#6 | 2023-09-20 | X Scan Stroke | 132.8
| TYSiC#7 | 2023-09-20 | X Scan Stroke | 129.2
| 경상국립대학교 반도체공학과 | 2023-09-20 | X Scan Stroke | 55
| TYSiC#3 | 2023-09-20 | -X Direction software Limit | -100629
| TYSiC#4 | 2023-09-20 | -X Direction software Limit | 103518
| TYSiC#4 | 2023-09-20 | -X Direction software Limit | 103135
| TYSiC#5 | 2023-09-20 | -X Direction software Limit | -100440
| TYSiC#6 | 2023-09-20 | -X Direction software Limit | -100605
| TYSiC#6 | 2023-09-20 | -X Direction software Limit | -100844
| TYSiC#7 | 2023-09-20 | -X Direction software Limit | -100556
| TYSiC#3 | 2023-09-20 | X Detector noise (XY Servo on) | 0.277
| TYSiC#4 | 2023-09-20 | X Detector noise (XY Servo on) | 0.29
| TYSiC#5 | 2023-09-20 | X Detector noise (XY Servo on) | 0.27
| TYSiC#6 | 2023-09-20 | X Detector noise (XY Servo on) | 0.26
| TYSiC#7 | 2023-09-20 | X Detector noise (XY Servo on) | 0.28
| 경상국립대학교 반도체공학과 | 2023-09-20 | X Detector noise (XY Servo on) | 0.36
| TYSiC#3 | 2023-09-20 | X Detector noise (XY Servo off) | 0.259
| TYSiC#4 | 2023-09-20 | X Detector noise (XY Servo off) | 0.27
| TYSiC#5 | 2023-09-20 | X Detector noise (XY Servo off) | 0.25
| TYSiC#6 | 2023-09-20 | X Detector noise (XY Servo off) | 0.26
| TYSiC#7 | 2023-09-20 | X Detector noise (XY Servo off) | 0.26
| 경상국립대학교 반도체공학과 | 2023-09-20 | X Detector noise (XY Servo off) | 0.36
| TYSiC#3 | 2023-09-20 | X axis OPM | 0.933
| TYSiC#4 | 2023-09-20 | X axis OPM | 0.86
| TYSiC#5 | 2023-09-20 | X axis OPM | 0.86
| TYSiC#6 | 2023-09-20 | X axis OPM | 1.12
| TYSiC#7 | 2023-09-20 | X axis OPM | 0.66
| 경상국립대학교 반도체공학과 | 2023-09-20 | X axis OPM | 0.41
| TYSiC#3 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.261
| TYSiC#4 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| TYSiC#5 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| TYSiC#6 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| TYSiC#7 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| 경상국립대학교 반도체공학과 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| TYSiC#3 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.417
| TYSiC#4 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| TYSiC#5 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| TYSiC#6 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| TYSiC#7 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| 경상국립대학교 반도체공학과 | 2023-09-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| TYSiC#3 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.337
| TYSiC#4 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.3
| TYSiC#5 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.35
| TYSiC#6 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.29
| TYSiC#7 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(-90000,0) | 0.39
| TYSiC#3 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.355
| TYSiC#4 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.29
| TYSiC#5 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.36
| TYSiC#6 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.31
| TYSiC#7 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(90000,0) | 0.39
| TYSiC#3 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.327
| TYSiC#4 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| TYSiC#5 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.32
| TYSiC#6 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.29
| TYSiC#7 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,-90000) | 0.42
| TYSiC#3 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.312
| TYSiC#4 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.31
| TYSiC#5 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.35
| TYSiC#6 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.28
| TYSiC#7 | 2023-09-20 | RMS of Z Height of Wafer Zeroscan(0,90000) | 0.37
| TYSiC#3 | 2023-09-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.32
| TYSiC#4 | 2023-09-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| TYSiC#5 | 2023-09-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.3
| TYSiC#6 | 2023-09-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.28
| TYSiC#7 | 2023-09-20 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.37
| TYSiC#3 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.216
| TYSiC#4 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| TYSiC#5 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| TYSiC#6 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| TYSiC#7 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| 경상국립대학교 반도체공학과 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| TYSiC#3 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.394
| TYSiC#4 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| TYSiC#5 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| TYSiC#6 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| TYSiC#7 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 경상국립대학교 반도체공학과 | 2023-09-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| TYSiC#3 | 2023-09-20 | NCM driving% with AC160TS cantilever | 0.744
| TYSiC#4 | 2023-09-20 | NCM driving% with AC160TS cantilever | 1.05
| TYSiC#5 | 2023-09-20 | NCM driving% with AC160TS cantilever | 0.27
| TYSiC#6 | 2023-09-20 | NCM driving% with AC160TS cantilever | 1.9
| TYSiC#7 | 2023-09-20 | NCM driving% with AC160TS cantilever | 0.34
| 경상국립대학교 반도체공학과 | 2023-09-20 | NCM driving% with AC160TS cantilever | 0.37
| TYSiC#3 | 2023-09-20 | NCM Amplitude noise | 14.478
| TYSiC#4 | 2023-09-20 | NCM Amplitude noise | 11.3
| TYSiC#5 | 2023-09-20 | NCM Amplitude noise | 15.07
| TYSiC#6 | 2023-09-20 | NCM Amplitude noise | 17.62
| TYSiC#7 | 2023-09-20 | NCM Amplitude noise | 8.2
| 경상국립대학교 반도체공학과 | 2023-09-20 | NCM Amplitude noise | 17.59
| TYSiC#3 | 2023-09-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.455
| TYSiC#4 | 2023-09-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.643
| TYSiC#5 | 2023-09-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.156
| TYSiC#6 | 2023-09-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.463
| TYSiC#7 | 2023-09-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.026
| 경상국립대학교 반도체공학과 | 2023-09-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.421
| TYSiC#3 | 2023-09-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.898
| TYSiC#4 | 2023-09-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.46
| TYSiC#5 | 2023-09-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.963
| TYSiC#6 | 2023-09-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.111
| TYSiC#7 | 2023-09-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.93
| 경상국립대학교 반도체공학과 | 2023-09-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.527
| TYSiC#3 | 2023-09-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 186.5
| TYSiC#4 | 2023-09-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 178.5
| TYSiC#5 | 2023-09-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 181.2
| TYSiC#6 | 2023-09-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 186.7
| TYSiC#7 | 2023-09-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 178.6
| 경상국립대학교 반도체공학과 | 2023-09-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99
| TYSiC#3 | 2023-09-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 179.9
| TYSiC#4 | 2023-09-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 182
| TYSiC#5 | 2023-09-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 182.2
| TYSiC#6 | 2023-09-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 189.8
| TYSiC#7 | 2023-09-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 181.3
| 경상국립대학교 반도체공학과 | 2023-09-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100.3
| TYSiC#3 | 2023-09-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.043
| TYSiC#4 | 2023-09-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| TYSiC#5 | 2023-09-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| TYSiC#6 | 2023-09-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| TYSiC#7 | 2023-09-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.05
| TYSiC#3 | 2023-09-20 | Lateral(C-D) sensitivity | 2.5
| TYSiC#4 | 2023-09-20 | Lateral(C-D) sensitivity | 4
| TYSiC#5 | 2023-09-20 | Lateral(C-D) sensitivity | 6.1
| TYSiC#6 | 2023-09-20 | Lateral(C-D) sensitivity | 2.2
| TYSiC#7 | 2023-09-20 | Lateral(C-D) sensitivity | 3.4
| 경상국립대학교 반도체공학과 | 2023-09-20 | Lateral(C-D) sensitivity | 5.3
| TYSiC#3 | 2023-09-20 | Lateral(C-D) noise | 0.172
| TYSiC#4 | 2023-09-20 | Lateral(C-D) noise | 0.14
| TYSiC#5 | 2023-09-20 | Lateral(C-D) noise | 0.162
| TYSiC#6 | 2023-09-20 | Lateral(C-D) noise | 0.145
| TYSiC#7 | 2023-09-20 | Lateral(C-D) noise | 0.167
| 경상국립대학교 반도체공학과 | 2023-09-20 | Lateral(C-D) noise | 0.073
| TYSiC#3 | 2023-09-20 | Full range of Z Height during Z Scan sweep | 16.45
| TYSiC#4 | 2023-09-20 | Full range of Z Height during Z Scan sweep | 16.7
| TYSiC#5 | 2023-09-20 | Full range of Z Height during Z Scan sweep | 15.7
| TYSiC#6 | 2023-09-20 | Full range of Z Height during Z Scan sweep | 16.2
| TYSiC#7 | 2023-09-20 | Full range of Z Height during Z Scan sweep | 16
| 경상국립대학교 반도체공학과 | 2023-09-20 | Full range of Z Height during Z Scan sweep | 16.1
| TYSiC#3 | 2023-09-20 | Full range of Z Detector during Z Scan sweep | 13.75
| TYSiC#4 | 2023-09-20 | Full range of Z Detector during Z Scan sweep | 11.7
| TYSiC#5 | 2023-09-20 | Full range of Z Detector during Z Scan sweep | 11.2
| TYSiC#6 | 2023-09-20 | Full range of Z Detector during Z Scan sweep | 12.9
| TYSiC#7 | 2023-09-20 | Full range of Z Detector during Z Scan sweep | 11.9
| 경상국립대학교 반도체공학과 | 2023-09-20 | Full range of Z Detector during Z Scan sweep | 13.1
| TYSiC#3 | 2023-09-20 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| TYSiC#4 | 2023-09-20 | Full range of Y Detector Raw during Y Scan sweep | 15
| TYSiC#5 | 2023-09-20 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| TYSiC#6 | 2023-09-20 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| TYSiC#7 | 2023-09-20 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 경상국립대학교 반도체공학과 | 2023-09-20 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| TYSiC#3 | 2023-09-20 | Full range of Y Detector during Y Scan sweep | 133.1
| TYSiC#4 | 2023-09-20 | Full range of Y Detector during Y Scan sweep | 130.3
| TYSiC#5 | 2023-09-20 | Full range of Y Detector during Y Scan sweep | 128.9
| TYSiC#6 | 2023-09-20 | Full range of Y Detector during Y Scan sweep | 131.7
| TYSiC#7 | 2023-09-20 | Full range of Y Detector during Y Scan sweep | 129.3
| 경상국립대학교 반도체공학과 | 2023-09-20 | Full range of Y Detector during Y Scan sweep | 70
| TYSiC#3 | 2023-09-20 | Full range of X Detector Raw during X Scan sweep | 14.1
| TYSiC#4 | 2023-09-20 | Full range of X Detector Raw during X Scan sweep | 14.5
| TYSiC#5 | 2023-09-20 | Full range of X Detector Raw during X Scan sweep | 14.4
| TYSiC#6 | 2023-09-20 | Full range of X Detector Raw during X Scan sweep | 14
| TYSiC#7 | 2023-09-20 | Full range of X Detector Raw during X Scan sweep | 14.3
| 경상국립대학교 반도체공학과 | 2023-09-20 | Full range of X Detector Raw during X Scan sweep | 13.9
| TYSiC#3 | 2023-09-20 | Full range of X Detector during X Scan sweep | 127.3
| TYSiC#4 | 2023-09-20 | Full range of X Detector during X Scan sweep | 131.1
| TYSiC#5 | 2023-09-20 | Full range of X Detector during X Scan sweep | 131
| TYSiC#6 | 2023-09-20 | Full range of X Detector during X Scan sweep | 132.9
| TYSiC#7 | 2023-09-20 | Full range of X Detector during X Scan sweep | 129.3
| 경상국립대학교 반도체공학과 | 2023-09-20 | Full range of X Detector during X Scan sweep | 69.8
| TYSiC#3 | 2023-09-20 | A-B noise | 8.86
| TYSiC#4 | 2023-09-20 | A-B noise | 6.69
| TYSiC#5 | 2023-09-20 | A-B noise | 9.14
| TYSiC#6 | 2023-09-20 | A-B noise | 8.62
| TYSiC#7 | 2023-09-20 | A-B noise | 10.79
| 경상국립대학교 반도체공학과 | 2023-09-20 | A-B noise | 10.7
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Linearity by Z scanner position | 0.7
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Linearity by 3 different sample height | 1.001
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Height noise | 0.028
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.8
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.6
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.5
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4
| TSMC F23 (Kumamoto) | 2023-09-19 | Z Drift (Short Traveling) | 0.12
| TSMC F23 (Kumamoto) | 2023-09-19 | Y2 Detector noise (XY Servo on) | 0.38
| TSMC F23 (Kumamoto) | 2023-09-19 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276
| TSMC F23 (Kumamoto) | 2023-09-19 | Y1 Detector noise (XY Servo on) | 0.4
| TSMC F23 (Kumamoto) | 2023-09-19 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 266.3
| TSMC F23 (Kumamoto) | 2023-09-19 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 191.2
| TSMC F23 (Kumamoto) | 2023-09-19 | Y Ringing(Rpv) | 0.18
| TSMC F23 (Kumamoto) | 2023-09-19 | Y Non-linearity error | -3.37
| TSMC F23 (Kumamoto) | 2023-09-19 | Y Detector noise (XY Servo on) | 0.34
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Repeatability @1st position | 0.04
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Repeatability @ 4th position | 0.03
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Repeatability @ 3rd position | 0.06
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Repeatability @ 2nd position | 0.05
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Accuracy @1st position | 0.56
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Accuracy @ 4th position | 0.41
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Accuracy @ 3rd position | 0.66
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis, Accuracy @ 2nd position | 0.49
| TSMC F23 (Kumamoto) | 2023-09-19 | Y axis OPM | 0.65
| TSMC F23 (Kumamoto) | 2023-09-19 | XY Orthogonality | 90
| TSMC F23 (Kumamoto) | 2023-09-19 | X2 Detector noise (XY Servo on) | 0.34
| TSMC F23 (Kumamoto) | 2023-09-19 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 284.8
| TSMC F23 (Kumamoto) | 2023-09-19 | X1 Detector noise (XY Servo on) | 0.36
| TSMC F23 (Kumamoto) | 2023-09-19 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 271.7
| TSMC F23 (Kumamoto) | 2023-09-19 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.7
| TSMC F23 (Kumamoto) | 2023-09-19 | X Ringing(Rpv) | 0.41
| TSMC F23 (Kumamoto) | 2023-09-19 | X Non-linearity error | 3.03
| TSMC F23 (Kumamoto) | 2023-09-19 | X Detector noise (XY Servo on) | 0.28
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Repeatability @1st position | 0.05
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Repeatability @ 4th position | 0.04
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Repeatability @ 3rd position | 0.04
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Repeatability @ 2nd position | 0.06
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Accuracy @1st position | 0.63
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Accuracy @ 4th position | 0.46
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Accuracy @ 3rd position | 0.54
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis, Accuracy @ 2nd position | 0.47
| TSMC F23 (Kumamoto) | 2023-09-19 | X axis OPM | 0.73
| TSMC F23 (Kumamoto) | 2023-09-19 | Under 50nm Step Height Sample (10 Repeats) | 18.3
| TSMC F23 (Kumamoto) | 2023-09-19 | Torque Profile (50mm, 20um/s) | 2.7
| TSMC F23 (Kumamoto) | 2023-09-19 | Torque Profile (50mm, 200um/s) | 2.3
| TSMC F23 (Kumamoto) | 2023-09-19 | StrokeYMin | -153405
| TSMC F23 (Kumamoto) | 2023-09-19 | StrokeYMax | 150941
| TSMC F23 (Kumamoto) | 2023-09-19 | StrokeY | 304346
| TSMC F23 (Kumamoto) | 2023-09-19 | StrokeXMin | -204117
| TSMC F23 (Kumamoto) | 2023-09-19 | StrokeXMax | 205033
| TSMC F23 (Kumamoto) | 2023-09-19 | StrokeX | 409150
| TSMC F23 (Kumamoto) | 2023-09-19 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.56
| TSMC F23 (Kumamoto) | 2023-09-19 | Roll motion to X+ full scan | 6.7
| TSMC F23 (Kumamoto) | 2023-09-19 | RMS of ZeroScan | 0.041
| TSMC F23 (Kumamoto) | 2023-09-19 | Profiler noise @200um | 0.14
| TSMC F23 (Kumamoto) | 2023-09-19 | Position repeatability(10repeats, Position Based) | 0.32
| TSMC F23 (Kumamoto) | 2023-09-19 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00044
| TSMC F23 (Kumamoto) | 2023-09-19 | OpticalAxisShiftPerFocusY | -0.0033
| TSMC F23 (Kumamoto) | 2023-09-19 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00019
| TSMC F23 (Kumamoto) | 2023-09-19 | OpticalAxisShiftPerFocusX | -0.0042
| TSMC F23 (Kumamoto) | 2023-09-19 | OPM Repeatability @5mm | 2.5
| TSMC F23 (Kumamoto) | 2023-09-19 | OPM Repeatability @25mm | 2.94
| TSMC F23 (Kumamoto) | 2023-09-19 | OPM Repeatability @1mm | 1.05
| TSMC F23 (Kumamoto) | 2023-09-19 | OPM Repeatability @10mm | 1.7
| TSMC F23 (Kumamoto) | 2023-09-19 | NCM driving% with AC160TS cantilever | 7.7
| TSMC F23 (Kumamoto) | 2023-09-19 | Min. Vibration Transmissbility (~100Hz) | 0.038
| TSMC F23 (Kumamoto) | 2023-09-19 | Max OPM @5mm | 34.26
| TSMC F23 (Kumamoto) | 2023-09-19 | Max OPM @25mm | 39.06
| TSMC F23 (Kumamoto) | 2023-09-19 | Max OPM @1mm | 11.73
| TSMC F23 (Kumamoto) | 2023-09-19 | Max OPM @10mm | 19.24
| TSMC F23 (Kumamoto) | 2023-09-19 | LowerLimit | 4640
| TSMC F23 (Kumamoto) | 2023-09-19 | Lateral(C-D) sensitivity | 7.7
| TSMC F23 (Kumamoto) | 2023-09-19 | Lateral(C-D) noise | 0.198
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from -Y to +Y (Y scan) | 90.82
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from -Y to +Y (X scan) | 18.31
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from -X to +X (Y scan) | 43.94
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from -X to +X (X scan) | 37.35
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from +Y to -Y (Y scan) | 12.45
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from +Y to -Y (X scan) | 32.23
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from +X to -X (Y scan) | 30.76
| TSMC F23 (Kumamoto) | 2023-09-19 | Image offset after moving from +X to -X (X scan) | 13.18
| TSMC F23 (Kumamoto) | 2023-09-19 | Full range of Z Height during Z Scan sweep | 17.2
| TSMC F23 (Kumamoto) | 2023-09-19 | Full range of Z Detector during Z Scan sweep | 11.9
| TSMC F23 (Kumamoto) | 2023-09-19 | Full range of Y2 Detector Raw during Y Scan sweep | 14.13
| TSMC F23 (Kumamoto) | 2023-09-19 | Full range of Y1 Detector Raw during Y Scan sweep | 14.45
| TSMC F23 (Kumamoto) | 2023-09-19 | Full range of X2 Detector Raw during X Scan sweep | 14.11
| TSMC F23 (Kumamoto) | 2023-09-19 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| TSMC F23 (Kumamoto) | 2023-09-19 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| TSMC F23 (Kumamoto) | 2023-09-19 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| TSMC F23 (Kumamoto) | 2023-09-19 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.64
| TSMC F23 (Kumamoto) | 2023-09-19 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.65
| TSMC F23 (Kumamoto) | 2023-09-19 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| TSMC F23 (Kumamoto) | 2023-09-19 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| TSMC F23 (Kumamoto) | 2023-09-19 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| TSMC F23 (Kumamoto) | 2023-09-19 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| TSMC F23 (Kumamoto) | 2023-09-19 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| TSMC F23 (Kumamoto) | 2023-09-19 | 50~200nm Step Height Sample (10 Repeats) | 184.22
| TSMC F23 (Kumamoto) | 2023-09-19 | 200~500nm Step Height Sample (10 Repeats) | 459.99
| TSMC F23 (Kumamoto) | 2023-09-19 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.3
| TSMC F23 (Kumamoto) | 2023-09-19 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 12
| TSMC F23 (Kumamoto) | 2023-09-19 | %GRR w/ roughness of non-patterened wafer | 14.5
| Swegan AB (PSE) | 2023-09-19 | Z Scanner position at -3μm | 0.002616
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Z Scanner position at -3μm | 0.000633
| Swegan AB (PSE) | 2023-09-19 | Z Scanner position at +3μm | 0.001106
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Z Scanner position at +3μm | 0.002405
| Swegan AB (PSE) | 2023-09-19 | Z Linearity by Z Scanner position
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Z Linearity by Z Scanner position
| Swegan AB (PSE) | 2023-09-19 | Z Linearity by 3 different sample height | 0.99807
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Z Linearity by 3 different sample height | 0.99345
| Swegan AB (PSE) | 2023-09-19 | Z Height noise | 0.0189
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Z Height noise | 0.0255
| Swegan AB (PSE) | 2023-09-19 | Y2 Detector noise (XY Servo on) | 0.262
| Swegan AB (PSE) | 2023-09-19 | Y2 Detector noise (XY Servo off) | 0.1796
| Swegan AB (PSE) | 2023-09-19 | Y1 Detector noise (XY Servo on) | 0.2491
| Swegan AB (PSE) | 2023-09-19 | Y1 Detector noise (XY Servo off) | 0.1796
| Swegan AB (PSE) | 2023-09-19 | Y Scan Stroke | 169.2095
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Y Scan Stroke | 55
| Swegan AB (PSE) | 2023-09-19 | Y Detector noise (XY Servo on) | 0.2058
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Y Detector noise (XY Servo on) | 0.346
| Swegan AB (PSE) | 2023-09-19 | Y Detector noise (XY Servo off) | 0.13
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Y Detector noise (XY Servo off) | 0.3429
| Swegan AB (PSE) | 2023-09-19 | Y axis OPM | 1.1536
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Y axis OPM | 0.4402
| Swegan AB (PSE) | 2023-09-19 | X2 Detector noise (XY Servo on) | 0.2599
| Swegan AB (PSE) | 2023-09-19 | X2 Detector noise (XY Servo off) | 0.1792
| Swegan AB (PSE) | 2023-09-19 | X1 Detector noise (XY Servo on) | 0.2502
| Swegan AB (PSE) | 2023-09-19 | X1 Detector noise (XY Servo off) | 0.1751
| Swegan AB (PSE) | 2023-09-19 | X Scan Stroke | 171.324
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | X Scan Stroke | 55
| Swegan AB (PSE) | 2023-09-19 | -X Direction software Limit | -75038.5
| Swegan AB (PSE) | 2023-09-19 | -X Direction software Limit | -75288
| Swegan AB (PSE) | 2023-09-19 | X Detector noise (XY Servo on) | 0.209
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | X Detector noise (XY Servo on) | 0.349
| Swegan AB (PSE) | 2023-09-19 | X Detector noise (XY Servo off) | 0.1296
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | X Detector noise (XY Servo off) | 0.3468
| Swegan AB (PSE) | 2023-09-19 | X axis OPM | 1.0092
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | X axis OPM | 0.5199
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.245
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.263
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.27
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.26
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.28
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.28
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.26
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.168
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Swegan AB (PSE) | 2023-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.206
| Swegan AB (PSE) | 2023-09-19 | NCM driving% with AC160TS cantilever | 0.351
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | NCM driving% with AC160TS cantilever | 0.367
| Swegan AB (PSE) | 2023-09-19 | NCM Amplitude noise | 13.5951
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | NCM Amplitude noise | 11.6805
| Swegan AB (PSE) | 2023-09-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.522
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.367
| Swegan AB (PSE) | 2023-09-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.798
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.623
| Swegan AB (PSE) | 2023-09-19 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 232.3
| Swegan AB (PSE) | 2023-09-19 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 236.1
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.8
| Swegan AB (PSE) | 2023-09-19 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 236.1
| Swegan AB (PSE) | 2023-09-19 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 239.8
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.2
| Swegan AB (PSE) | 2023-09-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Swegan AB (PSE) | 2023-09-19 | Lateral(C-D) sensitivity | 4.5
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Lateral(C-D) sensitivity | 4.8627
| Swegan AB (PSE) | 2023-09-19 | Lateral(C-D) noise | 0.088
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Lateral(C-D) noise | 0.052
| Swegan AB (PSE) | 2023-09-19 | Full range of Z Height during Z Scan sweep | 16.4
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Full range of Z Height during Z Scan sweep | 16.4
| Swegan AB (PSE) | 2023-09-19 | Full range of Z Detector during Z Scan sweep | 12.7
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Full range of Z Detector during Z Scan sweep | 13.3
| Swegan AB (PSE) | 2023-09-19 | Full range of Y2 Detector Raw during Y Scan sweep | 14.65
| Swegan AB (PSE) | 2023-09-19 | Full range of Y1 Detector Raw during Y Scan sweep | 14.25
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| Swegan AB (PSE) | 2023-09-19 | Full range of Y Detector during Y Scan sweep | 169.25
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Full range of Y Detector during Y Scan sweep | 71.5
| Swegan AB (PSE) | 2023-09-19 | Full range of X2 Detector Raw during X Scan sweep | 14.45
| Swegan AB (PSE) | 2023-09-19 | Full range of X1 Detector Raw during X Scan sweep | 14.35
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Full range of X Detector Raw during X Scan sweep | 13.6
| Swegan AB (PSE) | 2023-09-19 | Full range of X Detector during X Scan sweep | 171.55
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | Full range of X Detector during X Scan sweep | 69.4
| Swegan AB (PSE) | 2023-09-19 | A-B noise | 7.7224
| Park Systems Taiwan(demo 장비)(PST) | 2023-09-19 | A-B noise | 6.9598
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Z Scanner position at -3μm | 0.0036
| Wuhan Aroptics-tech | 2023-09-18 | Z Scanner position at -3μm | 0.0104
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Z Scanner position at +3μm | 0.0044
| Wuhan Aroptics-tech | 2023-09-18 | Z Scanner position at +3μm | 0.0151
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Z Linearity by Z Scanner position
| Wuhan Aroptics-tech | 2023-09-18 | Z Linearity by Z Scanner position
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Z Linearity by 3 different sample height | 0.999
| Wuhan Aroptics-tech | 2023-09-18 | Z Linearity by 3 different sample height | 0.995
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Z Height noise | 0.03
| Wuhan Aroptics-tech | 2023-09-18 | Z Height noise | 0.018
| Wuhan Aroptics-tech | 2023-09-18 | Y2 Detector noise (XY Servo on) | 0.21
| Wuhan Aroptics-tech | 2023-09-18 | Y2 Detector noise (XY Servo off) | 0.18
| Wuhan Aroptics-tech | 2023-09-18 | Y1 Detector noise (XY Servo on) | 0.2
| Wuhan Aroptics-tech | 2023-09-18 | Y1 Detector noise (XY Servo off) | 0.17
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Y Scan Stroke | 55
| Wuhan Aroptics-tech | 2023-09-18 | Y Scan Stroke | 180.2
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Y Detector noise (XY Servo on) | 0.34
| Wuhan Aroptics-tech | 2023-09-18 | Y Detector noise (XY Servo on) | 0.16
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Y Detector noise (XY Servo off) | 0.35
| Wuhan Aroptics-tech | 2023-09-18 | Y Detector noise (XY Servo off) | 0.13
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Y axis OPM | 0.56
| Wuhan Aroptics-tech | 2023-09-18 | Y axis OPM | 1.12
| Wuhan Aroptics-tech | 2023-09-18 | X2 Detector noise (XY Servo on) | 0.22
| Wuhan Aroptics-tech | 2023-09-18 | X2 Detector noise (XY Servo off) | 0.18
| Wuhan Aroptics-tech | 2023-09-18 | X1 Detector noise (XY Servo on) | 0.22
| Wuhan Aroptics-tech | 2023-09-18 | X1 Detector noise (XY Servo off) | 0.17
| 부산대학교 재료공학부 (세원) | 2023-09-18 | X Scan Stroke | 55
| Wuhan Aroptics-tech | 2023-09-18 | X Scan Stroke | 182.1
| Wuhan Aroptics-tech | 2023-09-18 | -X Direction software Limit | -75204
| Wuhan Aroptics-tech | 2023-09-18 | -X Direction software Limit | -75222
| 부산대학교 재료공학부 (세원) | 2023-09-18 | X Detector noise (XY Servo on) | 0.34
| Wuhan Aroptics-tech | 2023-09-18 | X Detector noise (XY Servo on) | 0.18
| 부산대학교 재료공학부 (세원) | 2023-09-18 | X Detector noise (XY Servo off) | 0.35
| Wuhan Aroptics-tech | 2023-09-18 | X Detector noise (XY Servo off) | 0.13
| 부산대학교 재료공학부 (세원) | 2023-09-18 | X axis OPM | 0.68
| Wuhan Aroptics-tech | 2023-09-18 | X axis OPM | 1.01
| 부산대학교 재료공학부 (세원) | 2023-09-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 부산대학교 재료공학부 (세원) | 2023-09-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.33
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.35
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.32
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.32
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.31
| 부산대학교 재료공학부 (세원) | 2023-09-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 부산대학교 재료공학부 (세원) | 2023-09-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Wuhan Aroptics-tech | 2023-09-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| 부산대학교 재료공학부 (세원) | 2023-09-18 | NCM driving% with AC160TS cantilever | 0.45
| Wuhan Aroptics-tech | 2023-09-18 | NCM driving% with AC160TS cantilever | 1.82
| 부산대학교 재료공학부 (세원) | 2023-09-18 | NCM Amplitude noise | 16.16
| Wuhan Aroptics-tech | 2023-09-18 | NCM Amplitude noise | 15.73
| 부산대학교 재료공학부 (세원) | 2023-09-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.453
| Wuhan Aroptics-tech | 2023-09-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.034
| 부산대학교 재료공학부 (세원) | 2023-09-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.461
| Wuhan Aroptics-tech | 2023-09-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.772
| Wuhan Aroptics-tech | 2023-09-18 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 265
| Wuhan Aroptics-tech | 2023-09-18 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 253.5
| 부산대학교 재료공학부 (세원) | 2023-09-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99
| Wuhan Aroptics-tech | 2023-09-18 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 251.9
| Wuhan Aroptics-tech | 2023-09-18 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 261.1
| 부산대학교 재료공학부 (세원) | 2023-09-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99.5
| Wuhan Aroptics-tech | 2023-09-18 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Lateral(C-D) sensitivity | 6.9
| Wuhan Aroptics-tech | 2023-09-18 | Lateral(C-D) sensitivity | 11.9
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Lateral(C-D) noise | 0.1
| Wuhan Aroptics-tech | 2023-09-18 | Lateral(C-D) noise | 0.098
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Full range of Z Height during Z Scan sweep | 16.7
| Wuhan Aroptics-tech | 2023-09-18 | Full range of Z Height during Z Scan sweep | 15.9
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Full range of Z Detector during Z Scan sweep | 11.7
| Wuhan Aroptics-tech | 2023-09-18 | Full range of Z Detector during Z Scan sweep | 12.8
| Wuhan Aroptics-tech | 2023-09-18 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| Wuhan Aroptics-tech | 2023-09-18 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Full range of Y Detector during Y Scan sweep | 73
| Wuhan Aroptics-tech | 2023-09-18 | Full range of Y Detector during Y Scan sweep | 177.7
| Wuhan Aroptics-tech | 2023-09-18 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Wuhan Aroptics-tech | 2023-09-18 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Full range of X Detector Raw during X Scan sweep | 14.2
| 부산대학교 재료공학부 (세원) | 2023-09-18 | Full range of X Detector during X Scan sweep | 70.7
| Wuhan Aroptics-tech | 2023-09-18 | Full range of X Detector during X Scan sweep | 183.8
| 부산대학교 재료공학부 (세원) | 2023-09-18 | A-B noise | 8.94
| Wuhan Aroptics-tech | 2023-09-18 | A-B noise | 14.4
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Linearity by Z scanner position | 1.08
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Linearity by Z scanner position | 0.62
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Linearity by 3 different sample height | 0.999
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Linearity by 3 different sample height | 1.004
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Height noise | 0.038
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Height noise | 0.036
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.7
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.7
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.3
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.6
| Infineon Villach (PSE) #7 | 2023-09-14 | Z Drift (Short Traveling) | 0.22
| Infineon Villach (PSE) #8 | 2023-09-14 | Z Drift (Short Traveling) | 0.15
| Infineon Villach (PSE) #7 | 2023-09-14 | Y2 Detector noise (XY Servo on) | 0.27
| Infineon Villach (PSE) #8 | 2023-09-14 | Y2 Detector noise (XY Servo on) | 0.26
| Infineon Villach (PSE) #7 | 2023-09-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 238.7
| Infineon Villach (PSE) #8 | 2023-09-14 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 242.4
| Infineon Villach (PSE) #7 | 2023-09-14 | Y1 Detector noise (XY Servo on) | 0.35
| Infineon Villach (PSE) #8 | 2023-09-14 | Y1 Detector noise (XY Servo on) | 0.26
| Infineon Villach (PSE) #7 | 2023-09-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 252
| Infineon Villach (PSE) #8 | 2023-09-14 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 248
| Infineon Villach (PSE) #7 | 2023-09-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.1
| Infineon Villach (PSE) #8 | 2023-09-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 182.2
| Infineon Villach (PSE) #7 | 2023-09-14 | Y Ringing(Rpv) | 0.51
| Infineon Villach (PSE) #8 | 2023-09-14 | Y Ringing(Rpv) | 0.86
| Infineon Villach (PSE) #7 | 2023-09-14 | Y Non-linearity error | 1.76
| Infineon Villach (PSE) #8 | 2023-09-14 | Y Non-linearity error | 2.55
| Infineon Villach (PSE) #7 | 2023-09-14 | Y Detector noise (XY Servo on) | 0.23
| Infineon Villach (PSE) #8 | 2023-09-14 | Y Detector noise (XY Servo on) | 0.19
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Repeatability @1st position | 0.04
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Repeatability @1st position | 0.04
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Repeatability @ 4th position | 0.05
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Repeatability @ 4th position | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Repeatability @ 3rd position | 0.04
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Repeatability @ 3rd position | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Repeatability @ 2nd position | 0.03
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Repeatability @ 2nd position | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Accuracy @1st position | 0.2
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Accuracy @1st position | 0.86
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Accuracy @ 4th position | 0.24
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Accuracy @ 4th position | 0.99
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Accuracy @ 3rd position | 0.18
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Accuracy @ 3rd position | 1.02
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis, Accuracy @ 2nd position | 0.15
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis, Accuracy @ 2nd position | 0.92
| Infineon Villach (PSE) #7 | 2023-09-14 | Y axis OPM | 0.38
| Infineon Villach (PSE) #8 | 2023-09-14 | Y axis OPM | 0.21
| Infineon Villach (PSE) #7 | 2023-09-14 | XY Orthogonality | 90
| Infineon Villach (PSE) #8 | 2023-09-14 | XY Orthogonality | 89.7
| Infineon Villach (PSE) #7 | 2023-09-14 | X2 Detector noise (XY Servo on) | 0.27
| Infineon Villach (PSE) #8 | 2023-09-14 | X2 Detector noise (XY Servo on) | 0.24
| Infineon Villach (PSE) #7 | 2023-09-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 244.5
| Infineon Villach (PSE) #8 | 2023-09-14 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 251.1
| Infineon Villach (PSE) #7 | 2023-09-14 | X1 Detector noise (XY Servo on) | 0.24
| Infineon Villach (PSE) #8 | 2023-09-14 | X1 Detector noise (XY Servo on) | 0.26
| Infineon Villach (PSE) #7 | 2023-09-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 243.5
| Infineon Villach (PSE) #8 | 2023-09-14 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 256.5
| Infineon Villach (PSE) #7 | 2023-09-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 166.1
| Infineon Villach (PSE) #8 | 2023-09-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 182.3
| Infineon Villach (PSE) #7 | 2023-09-14 | X Ringing(Rpv) | 0.42
| Infineon Villach (PSE) #8 | 2023-09-14 | X Ringing(Rpv) | 0.54
| Infineon Villach (PSE) #7 | 2023-09-14 | X Non-linearity error | 2.1
| Infineon Villach (PSE) #8 | 2023-09-14 | X Non-linearity error | 2.86
| Infineon Villach (PSE) #7 | 2023-09-14 | X Detector noise (XY Servo on) | 0.19
| Infineon Villach (PSE) #8 | 2023-09-14 | X Detector noise (XY Servo on) | 0.18
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Repeatability @1st position | 0.04
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Repeatability @1st position | 0.02
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Repeatability @ 4th position | 0.03
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Repeatability @ 4th position | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Repeatability @ 3rd position | 0.03
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Repeatability @ 3rd position | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Repeatability @ 2nd position | 0.04
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Repeatability @ 2nd position | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Accuracy @1st position | 0.21
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Accuracy @1st position | 0.1
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Accuracy @ 4th position | 0.15
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Accuracy @ 4th position | 0.2
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Accuracy @ 3rd position | 0.22
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Accuracy @ 3rd position | 0.14
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis, Accuracy @ 2nd position | 0.21
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis, Accuracy @ 2nd position | 0.19
| Infineon Villach (PSE) #7 | 2023-09-14 | X axis OPM | 0.49
| Infineon Villach (PSE) #8 | 2023-09-14 | X axis OPM | 0.29
| Infineon Villach (PSE) #7 | 2023-09-14 | Under 50nm Step Height Sample (10 Repeats) | 15.95
| Infineon Villach (PSE) #8 | 2023-09-14 | Under 50nm Step Height Sample (10 Repeats) | 16.07
| Infineon Villach (PSE) #7 | 2023-09-14 | StrokeYMin | -109373
| Infineon Villach (PSE) #8 | 2023-09-14 | StrokeYMin | -109072
| Infineon Villach (PSE) #7 | 2023-09-14 | StrokeYMax | 103655.2
| Infineon Villach (PSE) #8 | 2023-09-14 | StrokeYMax | 105427
| Infineon Villach (PSE) #7 | 2023-09-14 | StrokeY | 213027.8
| Infineon Villach (PSE) #8 | 2023-09-14 | StrokeY | 214499
| Infineon Villach (PSE) #7 | 2023-09-14 | StrokeXMin | -141789
| Infineon Villach (PSE) #8 | 2023-09-14 | StrokeXMin | -140216
| Infineon Villach (PSE) #7 | 2023-09-14 | StrokeXMax | 137664.9
| Infineon Villach (PSE) #8 | 2023-09-14 | StrokeXMax | 139352
| Infineon Villach (PSE) #7 | 2023-09-14 | StrokeX | 279454.2
| Infineon Villach (PSE) #8 | 2023-09-14 | StrokeX | 279568
| Infineon Villach (PSE) #7 | 2023-09-14 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.98
| Infineon Villach (PSE) #8 | 2023-09-14 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.74
| Infineon Villach (PSE) #7 | 2023-09-14 | RMS of ZeroScan | 0.035
| Infineon Villach (PSE) #8 | 2023-09-14 | RMS of ZeroScan | 0.033
| Infineon Villach (PSE) #7 | 2023-09-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Infineon Villach (PSE) #8 | 2023-09-14 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| Infineon Villach (PSE) #7 | 2023-09-14 | OpticalAxisShiftPerFocusY | -0.001
| Infineon Villach (PSE) #8 | 2023-09-14 | OpticalAxisShiftPerFocusY | -0.0018
| Infineon Villach (PSE) #7 | 2023-09-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| Infineon Villach (PSE) #8 | 2023-09-14 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00021
| Infineon Villach (PSE) #7 | 2023-09-14 | OpticalAxisShiftPerFocusX | -0.0015
| Infineon Villach (PSE) #8 | 2023-09-14 | OpticalAxisShiftPerFocusX | -0.005
| Infineon Villach (PSE) #7 | 2023-09-14 | NCM driving% with AC160TS cantilever | 5.49
| Infineon Villach (PSE) #8 | 2023-09-14 | NCM driving% with AC160TS cantilever | 6.91
| Infineon Villach (PSE) #7 | 2023-09-14 | Min. Vibration Transmissbility (~100Hz) | 0.002
| Infineon Villach (PSE) #8 | 2023-09-14 | Min. Vibration Transmissbility (~100Hz) | 0.001
| Infineon Villach (PSE) #7 | 2023-09-14 | LowerLimit | 12850
| Infineon Villach (PSE) #8 | 2023-09-14 | LowerLimit | 12598
| Infineon Villach (PSE) #7 | 2023-09-14 | Lateral(C-D) sensitivity | 7.4
| Infineon Villach (PSE) #8 | 2023-09-14 | Lateral(C-D) sensitivity | 10
| Infineon Villach (PSE) #7 | 2023-09-14 | Lateral(C-D) noise | 0.109
| Infineon Villach (PSE) #8 | 2023-09-14 | Lateral(C-D) noise | 0.219
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from -Y to +Y (Y scan) | 84.23
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from -Y to +Y (Y scan) | 45.41
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from -Y to +Y (X scan) | 15.38
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from -Y to +Y (X scan) | 28.56
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from -X to +X (Y scan) | 25.63
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from -X to +X (Y scan) | 12.45
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from -X to +X (X scan) | 112.06
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from -X to +X (X scan) | 22.7
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from +Y to -Y (Y scan) | 152.34
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from +Y to -Y (Y scan) | 68.84
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from +Y to -Y (X scan) | 10.25
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from +Y to -Y (X scan) | 18.31
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from +X to -X (Y scan) | 76.17
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from +X to -X (Y scan) | 54.93
| Infineon Villach (PSE) #7 | 2023-09-14 | Image offset after moving from +X to -X (X scan) | 109.13
| Infineon Villach (PSE) #8 | 2023-09-14 | Image offset after moving from +X to -X (X scan) | 65.91
| Infineon Villach (PSE) #7 | 2023-09-14 | Full range of Z Height during Z Scan sweep | 16.4
| Infineon Villach (PSE) #8 | 2023-09-14 | Full range of Z Height during Z Scan sweep | 16.8
| Infineon Villach (PSE) #7 | 2023-09-14 | Full range of Z Detector during Z Scan sweep | 12.4
| Infineon Villach (PSE) #8 | 2023-09-14 | Full range of Z Detector during Z Scan sweep | 12.4
| Infineon Villach (PSE) #7 | 2023-09-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.18
| Infineon Villach (PSE) #8 | 2023-09-14 | Full range of Y2 Detector Raw during Y Scan sweep | 12.97
| Infineon Villach (PSE) #7 | 2023-09-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.14
| Infineon Villach (PSE) #8 | 2023-09-14 | Full range of Y1 Detector Raw during Y Scan sweep | 13.11
| Infineon Villach (PSE) #7 | 2023-09-14 | Full range of X2 Detector Raw during X Scan sweep | 14.63
| Infineon Villach (PSE) #8 | 2023-09-14 | Full range of X2 Detector Raw during X Scan sweep | 14.13
| Infineon Villach (PSE) #7 | 2023-09-14 | Full range of X1 Detector Raw during X Scan sweep | 14.08
| Infineon Villach (PSE) #8 | 2023-09-14 | Full range of X1 Detector Raw during X Scan sweep | 14.01
| Infineon Villach (PSE) #7 | 2023-09-14 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Infineon Villach (PSE) #8 | 2023-09-14 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) | 101.01
| Infineon Villach (PSE) #7 | 2023-09-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Infineon Villach (PSE) #8 | 2023-09-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Infineon Villach (PSE) #7 | 2023-09-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.4
| Infineon Villach (PSE) #8 | 2023-09-14 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Infineon Villach (PSE) #7 | 2023-09-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| Infineon Villach (PSE) #8 | 2023-09-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| Infineon Villach (PSE) #7 | 2023-09-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Infineon Villach (PSE) #8 | 2023-09-14 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Infineon Villach (PSE) #7 | 2023-09-14 | 50~200nm Step Height Sample (10 Repeats) | 183.17
| Infineon Villach (PSE) #8 | 2023-09-14 | 50~200nm Step Height Sample (10 Repeats) | 177.65
| Infineon Villach (PSE) #7 | 2023-09-14 | 200~500nm Step Height Sample (10 Repeats) | 460.5
| Infineon Villach (PSE) #8 | 2023-09-14 | 200~500nm Step Height Sample (10 Repeats) | 462.69
| Infineon Villach (PSE) #7 | 2023-09-14 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.8
| Infineon Villach (PSE) #8 | 2023-09-14 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.1
| Infineon Villach (PSE) #7 | 2023-09-14 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.1
| Infineon Villach (PSE) #8 | 2023-09-14 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.6
| Infineon Villach (PSE) #7 | 2023-09-14 | %GRR w/ roughness of non-patterened wafer | 4.2
| Infineon Villach (PSE) #8 | 2023-09-14 | %GRR w/ roughness of non-patterened wafer | 6.2
| PSJ DEMO | 2023-09-14 | Z Scanner position at -3μm | 0.0184
| PSJ DEMO | 2023-09-14 | Z Scanner position at +3μm | 0.0038
| PSJ DEMO | 2023-09-14 | Z Linearity by 3 different sample height | 1.001
| PSJ DEMO | 2023-09-14 | Z Height noise | 0.025
| PSJ DEMO | 2023-09-14 | Y Scan Stroke | 55
| PSJ DEMO | 2023-09-14 | Y Detector noise (XY Servo on) | 0.29
| PSJ DEMO | 2023-09-14 | Y Detector noise (XY Servo off) | 0.3
| PSJ DEMO | 2023-09-14 | Y axis OPM | 0.9
| PSJ DEMO | 2023-09-14 | X Scan Stroke | 55
| PSJ DEMO | 2023-09-14 | X Detector noise (XY Servo on) | 0.29
| PSJ DEMO | 2023-09-14 | X Detector noise (XY Servo off) | 0.28
| PSJ DEMO | 2023-09-14 | X axis OPM | 0.8
| PSJ DEMO | 2023-09-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| PSJ DEMO | 2023-09-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| PSJ DEMO | 2023-09-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| PSJ DEMO | 2023-09-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| PSJ DEMO | 2023-09-14 | NCM driving% with AC160TS cantilever | 5.76
| PSJ DEMO | 2023-09-14 | NCM Amplitude noise | 15.07
| PSJ DEMO | 2023-09-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.096
| PSJ DEMO | 2023-09-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.226
| PSJ DEMO | 2023-09-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102
| PSJ DEMO | 2023-09-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100
| PSJ DEMO | 2023-09-14 | Lateral(C-D) sensitivity | 47.7
| PSJ DEMO | 2023-09-14 | Lateral(C-D) noise | 0.071
| PSJ DEMO | 2023-09-14 | Full range of Z Height during Z Scan sweep | 15.9
| PSJ DEMO | 2023-09-14 | Full range of Z Detector during Z Scan sweep | 12.6
| PSJ DEMO | 2023-09-14 | Full range of Y Detector Raw during Y Scan sweep | 14
| PSJ DEMO | 2023-09-14 | Full range of Y Detector during Y Scan sweep | 71.3
| PSJ DEMO | 2023-09-14 | Full range of X Detector Raw during X Scan sweep | 14.1
| PSJ DEMO | 2023-09-14 | Full range of X Detector during X Scan sweep | 70.6
| PSJ DEMO | 2023-09-14 | A-B noise | 14.61
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Z Scanner position at -3μm | 0.001
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Z Scanner position at +3μm | 0.0008
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Z Linearity by 3 different sample height | 0.995
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Z Height noise | 0.026
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Y Scan Stroke | 16
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Y Detector noise (XY Servo on) | 0.036
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Y Detector noise (XY Servo off) | 0.038
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Y axis OPM | 0.593
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | X Scan Stroke | 15.2
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | X Detector noise (XY Servo on) | 0.024
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | X Detector noise (XY Servo off) | 0.024
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | X axis OPM | 1.348
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | NCM driving% with AC160TS cantilever | 0.96
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | NCM Amplitude noise | 17.2
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.867
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.508
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 35.7
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 21.9
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Lateral(C-D) sensitivity | 3.9
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Lateral(C-D) noise | 0.099
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Full range of Z Height during Z Scan sweep | 15.6
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Full range of Z Detector during Z Scan sweep | 12.2
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Full range of Y Detector Raw during Y Scan sweep | 9.9
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Full range of Y Detector during Y Scan sweep | 17.6
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Full range of X Detector Raw during X Scan sweep | 15.8
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | Full range of X Detector during X Scan sweep | 17.3
| Indian Institute of Science Education and Research Thiruvananthapuram | 2023-09-12 | A-B noise | 11.09
| Infineon Kulim #2 | 2023-09-08 | Z Linearity by Z scanner position | 0.55
| Infineon Kulim #2 | 2023-09-08 | Z Linearity by Z scanner position | 0.77
| Infineon Kulim #2 | 2023-09-08 | Z Linearity by 4 different sample height | 1.005
| Infineon Kulim #2_1 | 2023-09-08 | Z Linearity by 4 different sample height | 1.0034
| Infineon Kulim #2 | 2023-09-08 | Z Height noise | 0.028
| Infineon Kulim #2 | 2023-09-08 | Z Height noise | 0.03
| Infineon Kulim #2 | 2023-09-08 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.3688
| Infineon Kulim #2 | 2023-09-08 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.3731
| Infineon Kulim #2 | 2023-09-08 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8696
| Infineon Kulim #2 | 2023-09-08 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.8477
| Infineon Kulim #2 | 2023-09-08 | Y2 Detector noise (XY Servo on) | 0.271
| Infineon Kulim #2 | 2023-09-08 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 243.21
| Infineon Kulim #2 | 2023-09-08 | Y1 Detector noise (XY Servo on) | 0.344
| Infineon Kulim #2 | 2023-09-08 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 238.89
| Infineon Kulim #2 | 2023-09-08 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 170.31
| Infineon Kulim #2 | 2023-09-08 | Y Ringing(Rpv) | 0.097
| Infineon Kulim #2 | 2023-09-08 | Y Non-linearity error | 1.78
| Infineon Kulim #2 | 2023-09-08 | Y Detector noise (XY Servo on) | 0.22
| Infineon Kulim #2 | 2023-09-08 | Y axis, Repeatability @1st position | 0.05
| Infineon Kulim #2 | 2023-09-08 | Y axis, Repeatability @ 4th position | 0.05
| Infineon Kulim #2 | 2023-09-08 | Y axis, Repeatability @ 3rd position | 0.07
| Infineon Kulim #2 | 2023-09-08 | Y axis, Repeatability @ 2nd position | 0.06
| Infineon Kulim #2 | 2023-09-08 | Y axis, Accuracy @1st position | 1.15
| Infineon Kulim #2 | 2023-09-08 | Y axis, Accuracy @ 4th position | 1.06
| Infineon Kulim #2 | 2023-09-08 | Y axis, Accuracy @ 3rd position | 1.83
| Infineon Kulim #2 | 2023-09-08 | Y axis, Accuracy @ 2nd position | 1.79
| Infineon Kulim #2 | 2023-09-08 | Y axis OPM | 0.231
| Infineon Kulim #2 | 2023-09-08 | X2 Detector noise (XY Servo on) | 0.321
| Infineon Kulim #2 | 2023-09-08 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 249.12
| Infineon Kulim #2 | 2023-09-08 | X1 Detector noise (XY Servo on) | 0.299
| Infineon Kulim #2 | 2023-09-08 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 249.36
| Infineon Kulim #2 | 2023-09-08 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 170.776
| Infineon Kulim #2 | 2023-09-08 | X Ringing(Rpv) | 0.355
| Infineon Kulim #2 | 2023-09-08 | X Non-linearity error | 1.88
| Infineon Kulim #2 | 2023-09-08 | X Detector noise (XY Servo on) | 0.227
| Infineon Kulim #2 | 2023-09-08 | X axis, Repeatability @1st position | 0.02
| Infineon Kulim #2 | 2023-09-08 | X axis, Repeatability @ 4th position | 0.02
| Infineon Kulim #2 | 2023-09-08 | X axis, Repeatability @ 3rd position | 0.03
| Infineon Kulim #2 | 2023-09-08 | X axis, Repeatability @ 2nd position | 0.05
| Infineon Kulim #2 | 2023-09-08 | X axis, Accuracy @1st position | 0.9
| Infineon Kulim #2 | 2023-09-08 | X axis, Accuracy @ 4th position | 0.79
| Infineon Kulim #2 | 2023-09-08 | X axis, Accuracy @ 3rd position | 0.9
| Infineon Kulim #2 | 2023-09-08 | X axis, Accuracy @ 2nd position | 1.1
| Infineon Kulim #2 | 2023-09-08 | X axis OPM | 0.662
| Infineon Kulim #2 | 2023-09-08 | StrokeYMin | -109238
| Infineon Kulim #2 | 2023-09-08 | StrokeYMax | 103270.8
| Infineon Kulim #2 | 2023-09-08 | StrokeY | 212508.3
| Infineon Kulim #2 | 2023-09-08 | StrokeXMin | -140980
| Infineon Kulim #2 | 2023-09-08 | StrokeXMax | 136519.4
| Infineon Kulim #2 | 2023-09-08 | StrokeX | 277499.2
| Infineon Kulim #2 | 2023-09-08 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.966
| Infineon Kulim #2 | 2023-09-08 | RMS of ZeroScan | 0.033
| Infineon Kulim #2 | 2023-09-08 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6.7e-05
| Infineon Kulim #2 | 2023-09-08 | OpticalAxisShiftPerFocusY | 0.0009
| Infineon Kulim #2 | 2023-09-08 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.000864
| Infineon Kulim #2 | 2023-09-08 | OpticalAxisShiftPerFocusX | 0.002
| Infineon Kulim #2 | 2023-09-08 | NCM driving% with AC160TS cantilever | 26.82
| Infineon Kulim #2 | 2023-09-08 | NCM driving% with AC160TS cantilever | 37.37
| Infineon Kulim #2 | 2023-09-08 | Min. Vibration Transmissbility (~100Hz) | 0.0002
| Infineon Kulim #2 | 2023-09-08 | Lateral(C-D) sensitivity | 91.787
| Infineon Kulim #2 | 2023-09-08 | Lateral(C-D) noise | 0.134
| Infineon Kulim #2 | 2023-09-08 | Laser intensity with AC160 cantilever | 3.066
| Infineon Kulim #2 | 2023-09-08 | Laser intensity with AC160 cantilever | 2.995
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from -Y to +Y (Y scan) | 120.85
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from -Y to +Y (X scan) | 28.56
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from -X to +X (Y scan) | 41.75
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from -X to +X (X scan) | 19.04
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from +Y to -Y (Y scan) | 122.31
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from +Y to -Y (X scan) | 15.38
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from +X to -X (Y scan) | 23.44
| Infineon Kulim #2 | 2023-09-08 | Image offset after moving from +X to -X (X scan) | 70.31
| Infineon Kulim #2 | 2023-09-08 | Full range of Z Height during Z Scan sweep | 15.168
| Infineon Kulim #2 | 2023-09-08 | Full range of Z Height during Z Scan sweep | 16.4835
| Infineon Kulim #2 | 2023-09-08 | Full range of Z Detector during Z Scan sweep | 11.42
| Infineon Kulim #2 | 2023-09-08 | Full range of Z Detector during Z Scan sweep | 11.04
| Infineon Kulim #2 | 2023-09-08 | Full range of Y2 Detector Raw during Y Scan sweep | 13.15
| Infineon Kulim #2 | 2023-09-08 | Full range of Y2 Detector during Y Scan sweep | 180.2772
| Infineon Kulim #2 | 2023-09-08 | Full range of Y1 Detector Raw during Y Scan sweep | 13.31
| Infineon Kulim #2 | 2023-09-08 | Full range of Y1 Detector during Y Scan sweep | 185.7076
| Infineon Kulim #2 | 2023-09-08 | Full range of X2 Detector Raw during X Scan sweep | 13.17
| Infineon Kulim #2 | 2023-09-08 | Full range of X2 Detector during X Scan sweep | 183.1662
| Infineon Kulim #2 | 2023-09-08 | Full range of X1 Detector Raw during X Scan sweep | 12.87
| Infineon Kulim #2 | 2023-09-08 | Full range of X1 Detector during X Scan sweep | 191.4485
| Infineon Kulim #2 | 2023-09-08 | Allowable max warpage, w/ SiC sample chuck | 300
| Infineon Kulim #2 | 2023-09-08 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Infineon Kulim #2 | 2023-09-08 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22
| Infineon Kulim #2 | 2023-09-08 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| Infineon Kulim #2 | 2023-09-08 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| Infineon Kulim #2 | 2023-09-08 | A－B difference during 3 times run test | 0.76
| Infineon Kulim #2 | 2023-09-08 | A－B difference during 3 times run test | 0.13
| Infineon Kulim #2 | 2023-09-08 | A＋B difference during 3 times run test | 0.12
| Infineon Kulim #2 | 2023-09-08 | A＋B difference during 3 times run test | 0.16
| Infineon Kulim #2 | 2023-09-08 | 940nm Step Height Sample (10 Repeats) | 950.616
| Infineon Kulim #2 | 2023-09-08 | 50~200nm Step Height Sample (10 Repeats) | 183.996
| Infineon Kulim #2 | 2023-09-08 | 50~200nm Step Height Sample (10 Repeats) | 176.691
| Infineon Kulim #2 | 2023-09-08 | 200~500nm Step Height Sample (10 Repeats) | 462.813
| Infineon Kulim #2 | 2023-09-08 | 200~500nm Step Height Sample (10 Repeats) | 462.714
| Infineon Kulim #2 | 2023-09-08 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 6.8
| Infineon Kulim #2 | 2023-09-08 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.2
| Infineon Kulim #2 | 2023-09-08 | %GRR w/ roughness of non-patterened wafer | 7.5
| Xavier University | 2023-09-08 | Z Scanner position at -3μm | 0.0096
| Xavier University | 2023-09-08 | Z Scanner position at +3μm | 0.0173
| Xavier University | 2023-09-08 | Z Linearity by 3 different sample height | 0.998
| Rebeca Ribeiro | 2023-09-08 | Z Linearity by 3 different sample height | 1.009
| Xavier University | 2023-09-08 | Z Height noise | 0.028
| Rebeca Ribeiro | 2023-09-08 | Z Height noise | 0.021
| Xavier University | 2023-09-08 | Y Scan Stroke | 55
| Xavier University | 2023-09-08 | Y Detector noise (XY Servo on) | 0.28
| Rebeca Ribeiro | 2023-09-08 | Y Detector noise (XY Servo on) | 0.26
| Xavier University | 2023-09-08 | Y Detector noise (XY Servo off) | 0.27
| Rebeca Ribeiro | 2023-09-08 | Y Detector noise (XY Servo off) | 0.25
| Xavier University | 2023-09-08 | Y axis OPM | 0.53
| Rebeca Ribeiro | 2023-09-08 | Y axis OPM | 1.4
| Xavier University | 2023-09-08 | X Scan Stroke | 55
| Xavier University | 2023-09-08 | X Detector noise (XY Servo on) | 0.28
| Rebeca Ribeiro | 2023-09-08 | X Detector noise (XY Servo on) | 0.26
| Xavier University | 2023-09-08 | X Detector noise (XY Servo off) | 0.27
| Rebeca Ribeiro | 2023-09-08 | X Detector noise (XY Servo off) | 0.25
| Xavier University | 2023-09-08 | X axis OPM | 0.32
| Rebeca Ribeiro | 2023-09-08 | X axis OPM | 0.46
| Xavier University | 2023-09-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Rebeca Ribeiro | 2023-09-08 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Xavier University | 2023-09-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Rebeca Ribeiro | 2023-09-08 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Xavier University | 2023-09-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Rebeca Ribeiro | 2023-09-08 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Xavier University | 2023-09-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Rebeca Ribeiro | 2023-09-08 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Xavier University | 2023-09-08 | NCM driving% with AC160TS cantilever | 1.95
| Rebeca Ribeiro | 2023-09-08 | NCM driving% with AC160TS cantilever | 0.742
| Xavier University | 2023-09-08 | NCM Amplitude noise | 13.6
| Rebeca Ribeiro | 2023-09-08 | NCM Amplitude noise | 17.95
| Xavier University | 2023-09-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.411
| Rebeca Ribeiro | 2023-09-08 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.379
| Xavier University | 2023-09-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.625
| Rebeca Ribeiro | 2023-09-08 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.694
| Xavier University | 2023-09-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.7
| Rebeca Ribeiro | 2023-09-08 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.7
| Xavier University | 2023-09-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.5
| Rebeca Ribeiro | 2023-09-08 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 151.9
| Rebeca Ribeiro | 2023-09-08 | Full range of Z Height during Z Scan sweep | 15.9
| Xavier University | 2023-09-08 | Lateral(C-D) sensitivity | 4.2
| Rebeca Ribeiro | 2023-09-08 | Lateral(C-D) sensitivity | 24
| Xavier University | 2023-09-08 | Lateral(C-D) noise | 0.114
| Rebeca Ribeiro | 2023-09-08 | Lateral(C-D) noise | 0.088
| Xavier University | 2023-09-08 | Full range of Z Height during Z Scan sweep | 16.3
| Xavier University | 2023-09-08 | Full range of Z Detector during Z Scan sweep | 12.8
| Rebeca Ribeiro | 2023-09-08 | Full range of Z Detector during Z Scan sweep | 12.9
| Xavier University | 2023-09-08 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| Rebeca Ribeiro | 2023-09-08 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Xavier University | 2023-09-08 | Full range of Y Detector during Y Scan sweep | 71.8
| Rebeca Ribeiro | 2023-09-08 | Full range of Y Detector during Y Scan sweep | 112.6
| Xavier University | 2023-09-08 | Full range of X Detector Raw during X Scan sweep | 13.8
| Rebeca Ribeiro | 2023-09-08 | Full range of X Detector Raw during X Scan sweep | 14.8
| Xavier University | 2023-09-08 | Full range of X Detector during X Scan sweep | 70.2
| Rebeca Ribeiro | 2023-09-08 | Full range of X Detector during X Scan sweep | 113.3
| Rebeca Ribeiro | 2023-09-08 | Approach position at -3μm | 0.0033
| Rebeca Ribeiro | 2023-09-08 | Approach position at +3μm | 0.0033
| Xavier University | 2023-09-08 | A-B noise | 7.76
| Rebeca Ribeiro | 2023-09-08 | A-B noise | 9.89
| SweGan, Sweden | 2023-09-07 | Z Scanner position at -3μm | 0.0026
| SweGan, Sweden | 2023-09-07 | Z Scanner position at +3μm | 0.0011
| SweGan, Sweden | 2023-09-07 | Z Linearity by 3 different sample height | 0.998
| SweGan, Sweden | 2023-09-07 | Z Height noise | 0.019
| SweGan, Sweden | 2023-09-07 | Y2 Detector noise (XY Servo on) | 0.26
| SweGan, Sweden | 2023-09-07 | Y2 Detector noise (XY Servo off) | 0.18
| SweGan, Sweden | 2023-09-07 | Y1 Detector noise (XY Servo on) | 0.25
| SweGan, Sweden | 2023-09-07 | Y1 Detector noise (XY Servo off) | 0.18
| SweGan, Sweden | 2023-09-07 | Y Scan Stroke | 169.2095
| SweGan, Sweden | 2023-09-07 | Y Detector noise (XY Servo on) | 0.21
| SweGan, Sweden | 2023-09-07 | Y Detector noise (XY Servo off) | 0.13
| SweGan, Sweden | 2023-09-07 | Y axis OPM | 1.15
| SweGan, Sweden | 2023-09-07 | X2 Detector noise (XY Servo on) | 0.26
| SweGan, Sweden | 2023-09-07 | X2 Detector noise (XY Servo off) | 0.18
| SweGan, Sweden | 2023-09-07 | X1 Detector noise (XY Servo on) | 0.25
| SweGan, Sweden | 2023-09-07 | X1 Detector noise (XY Servo off) | 0.18
| SweGan, Sweden | 2023-09-07 | X Scan Stroke | 171.324
| SweGan, Sweden | 2023-09-07 | -X Direction software Limit | -75038.5
| SweGan, Sweden | 2023-09-07 | -X Direction software Limit | -75288
| SweGan, Sweden | 2023-09-07 | X Detector noise (XY Servo on) | 0.21
| SweGan, Sweden | 2023-09-07 | X Detector noise (XY Servo off) | 0.13
| SweGan, Sweden | 2023-09-07 | X axis OPM | 1.01
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of Wafer Zeroscan (Top) | 0.26
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of Wafer Zeroscan (Right) | 0.28
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of Wafer Zeroscan (Left) | 0.28
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of Wafer ZeroScan (Center) | 0.26
| SweGan, Sweden | 2023-09-07 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.27
| SweGan, Sweden | 2023-09-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| SweGan, Sweden | 2023-09-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| SweGan, Sweden | 2023-09-07 | NCM driving% with AC160TS cantilever | 0.35
| SweGan, Sweden | 2023-09-07 | NCM Amplitude noise | 13.6
| SweGan, Sweden | 2023-09-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.522
| SweGan, Sweden | 2023-09-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.798
| SweGan, Sweden | 2023-09-07 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 232.3
| SweGan, Sweden | 2023-09-07 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 236.1
| SweGan, Sweden | 2023-09-07 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 236.1
| SweGan, Sweden | 2023-09-07 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 239.8
| SweGan, Sweden | 2023-09-07 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| SweGan, Sweden | 2023-09-07 | Lateral(C-D) sensitivity | 4.5
| SweGan, Sweden | 2023-09-07 | Lateral(C-D) noise | 0.088
| SweGan, Sweden | 2023-09-07 | Full range of Z Height during Z Scan sweep | 16.4
| SweGan, Sweden | 2023-09-07 | Full range of Z Detector during Z Scan sweep | 12.7
| SweGan, Sweden | 2023-09-07 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| SweGan, Sweden | 2023-09-07 | Full range of Y1 Detector Raw during Y Scan sweep | 14.3
| SweGan, Sweden | 2023-09-07 | Full range of Y Detector during Y Scan sweep | 169.3
| SweGan, Sweden | 2023-09-07 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| SweGan, Sweden | 2023-09-07 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| SweGan, Sweden | 2023-09-07 | Full range of X Detector during X Scan sweep | 171.6
| SweGan, Sweden | 2023-09-07 | A-B noise | 7.72
| 부산대학교_공과대학 | 2023-08-31 | Z Scanner position at -3μm | 0.005
| Sincere International Trading Company Limited #2 | 2023-08-31 | Z Scanner position at -3μm | 0.0113
| University of Texas at Arlington | 2023-08-31 | Z Scanner position at -3μm | 0.0001
| 부산대학교_공과대학 | 2023-08-31 | Z Scanner position at +3μm | 0.0043
| Sincere International Trading Company Limited #2 | 2023-08-31 | Z Scanner position at +3μm | 0.0058
| University of Texas at Arlington | 2023-08-31 | Z Scanner position at +3μm | 0.0015
| 부산대학교_공과대학 | 2023-08-31 | Z Linearity by 3 different sample height | 1.006
| Sincere International Trading Company Limited #2 | 2023-08-31 | Z Linearity by 3 different sample height | 1.002
| University of Texas at Arlington | 2023-08-31 | Z Linearity by 3 different sample height | 1.003
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Z Linearity by 3 different sample height | 1.007
| 부산대학교_공과대학 | 2023-08-31 | Z Height noise | 0.029
| Sincere International Trading Company Limited #2 | 2023-08-31 | Z Height noise | 0.03
| University of Texas at Arlington | 2023-08-31 | Z Height noise | 0.034
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Z Height noise | 0.036
| 부산대학교_공과대학 | 2023-08-31 | Y2 Detector noise (XY Servo on) | 0.26
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y2 Detector noise (XY Servo on) | 0.4
| 부산대학교_공과대학 | 2023-08-31 | Y2 Detector noise (XY Servo off) | 0.26
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y2 Detector noise (XY Servo off) | 0.39
| 부산대학교_공과대학 | 2023-08-31 | Y1 Detector noise (XY Servo on) | 0.29
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y1 Detector noise (XY Servo on) | 0.35
| 부산대학교_공과대학 | 2023-08-31 | Y1 Detector noise (XY Servo off) | 0.28
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y1 Detector noise (XY Servo off) | 0.33
| 부산대학교_공과대학 | 2023-08-31 | Y Scan Stroke | 183.2
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y Scan Stroke | 194.6
| 부산대학교_공과대학 | 2023-08-31 | Y Detector noise (XY Servo on) | 0.2
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y Detector noise (XY Servo on) | 0.27
| University of Texas at Arlington | 2023-08-31 | Y Detector noise (XY Servo on) | 0.29
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Y Detector noise (XY Servo on) | 0.29
| 부산대학교_공과대학 | 2023-08-31 | Y Detector noise (XY Servo off) | 0.19
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y Detector noise (XY Servo off) | 0.26
| University of Texas at Arlington | 2023-08-31 | Y Detector noise (XY Servo off) | 0.28
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Y Detector noise (XY Servo off) | 0.28
| 부산대학교_공과대학 | 2023-08-31 | Y axis OPM | 1.32
| Sincere International Trading Company Limited #2 | 2023-08-31 | Y axis OPM | 1.25
| University of Texas at Arlington | 2023-08-31 | Y axis OPM | 1.02
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Y axis OPM | 0.42
| 부산대학교_공과대학 | 2023-08-31 | X2 Detector noise (XY Servo on) | 0.3
| Sincere International Trading Company Limited #2 | 2023-08-31 | X2 Detector noise (XY Servo on) | 0.37
| 부산대학교_공과대학 | 2023-08-31 | X2 Detector noise (XY Servo off) | 0.3
| Sincere International Trading Company Limited #2 | 2023-08-31 | X2 Detector noise (XY Servo off) | 0.35
| 부산대학교_공과대학 | 2023-08-31 | X1 Detector noise (XY Servo on) | 0.27
| Sincere International Trading Company Limited #2 | 2023-08-31 | X1 Detector noise (XY Servo on) | 0.41
| 부산대학교_공과대학 | 2023-08-31 | X1 Detector noise (XY Servo off) | 0.27
| Sincere International Trading Company Limited #2 | 2023-08-31 | X1 Detector noise (XY Servo off) | 0.4
| 부산대학교_공과대학 | 2023-08-31 | X Scan Stroke | 183.1
| Sincere International Trading Company Limited #2 | 2023-08-31 | X Scan Stroke | 198.9
| 부산대학교_공과대학 | 2023-08-31 | -X Direction software Limit | -75407.5
| 부산대학교_공과대학 | 2023-08-31 | -X Direction software Limit | -75546.5
| Sincere International Trading Company Limited #2 | 2023-08-31 | -X Direction software Limit | -150886
| Sincere International Trading Company Limited #2 | 2023-08-31 | -X Direction software Limit | -150017
| 부산대학교_공과대학 | 2023-08-31 | X Detector noise (XY Servo on) | 0.21
| Sincere International Trading Company Limited #2 | 2023-08-31 | X Detector noise (XY Servo on) | 0.29
| University of Texas at Arlington | 2023-08-31 | X Detector noise (XY Servo on) | 0.3
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | X Detector noise (XY Servo on) | 0.31
| 부산대학교_공과대학 | 2023-08-31 | X Detector noise (XY Servo off) | 0.21
| Sincere International Trading Company Limited #2 | 2023-08-31 | X Detector noise (XY Servo off) | 0.28
| University of Texas at Arlington | 2023-08-31 | X Detector noise (XY Servo off) | 0.3
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | X Detector noise (XY Servo off) | 0.3
| 부산대학교_공과대학 | 2023-08-31 | X axis OPM | 0.51
| Sincere International Trading Company Limited #2 | 2023-08-31 | X axis OPM | 2.37
| University of Texas at Arlington | 2023-08-31 | X axis OPM | 1
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | X axis OPM | 0.67
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| University of Texas at Arlington | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| University of Texas at Arlington | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Top) | 0.33
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Top) | 0.36
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Right) | 0.31
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Right) | 0.4
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Left) | 0.32
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Left) | 0.39
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of Wafer ZeroScan (Center) | 0.31
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of Wafer ZeroScan (Center) | 0.37
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.33
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.36
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| University of Texas at Arlington | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 부산대학교_공과대학 | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Sincere International Trading Company Limited #2 | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| University of Texas at Arlington | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 부산대학교_공과대학 | 2023-08-31 | NCM driving% with AC160TS cantilever | 0.48
| Sincere International Trading Company Limited #2 | 2023-08-31 | NCM driving% with AC160TS cantilever | 1.11
| University of Texas at Arlington | 2023-08-31 | NCM driving% with AC160TS cantilever | 0.24
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | NCM driving% with AC160TS cantilever | 5.26
| 부산대학교_공과대학 | 2023-08-31 | NCM Amplitude noise | 15.12
| Sincere International Trading Company Limited #2 | 2023-08-31 | NCM Amplitude noise | 14.47
| University of Texas at Arlington | 2023-08-31 | NCM Amplitude noise | 16.39
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | NCM Amplitude noise | 17.24
| 부산대학교_공과대학 | 2023-08-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.635
| Sincere International Trading Company Limited #2 | 2023-08-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.48
| University of Texas at Arlington | 2023-08-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.92
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.802
| 부산대학교_공과대학 | 2023-08-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.562
| Sincere International Trading Company Limited #2 | 2023-08-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.795
| University of Texas at Arlington | 2023-08-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.824
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.554
| 부산대학교_공과대학 | 2023-08-31 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 266.1
| Sincere International Trading Company Limited #2 | 2023-08-31 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 276.8
| 부산대학교_공과대학 | 2023-08-31 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 270.7
| Sincere International Trading Company Limited #2 | 2023-08-31 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 270.2
| University of Texas at Arlington | 2023-08-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 164.3
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.9
| 부산대학교_공과대학 | 2023-08-31 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 261.4
| Sincere International Trading Company Limited #2 | 2023-08-31 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 275
| 부산대학교_공과대학 | 2023-08-31 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 275.1
| Sincere International Trading Company Limited #2 | 2023-08-31 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 281
| University of Texas at Arlington | 2023-08-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 164.5
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.5
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Full range of Z Height during Z Scan sweep | 15.4
| 부산대학교_공과대학 | 2023-08-31 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Sincere International Trading Company Limited #2 | 2023-08-31 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| 부산대학교_공과대학 | 2023-08-31 | Lateral(C-D) sensitivity | 13.9
| Sincere International Trading Company Limited #2 | 2023-08-31 | Lateral(C-D) sensitivity | 7.9
| University of Texas at Arlington | 2023-08-31 | Lateral(C-D) sensitivity | 16.6
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Lateral(C-D) sensitivity | 4
| 부산대학교_공과대학 | 2023-08-31 | Lateral(C-D) noise | 0.162
| Sincere International Trading Company Limited #2 | 2023-08-31 | Lateral(C-D) noise | 0.117
| University of Texas at Arlington | 2023-08-31 | Lateral(C-D) noise | 0.108
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Lateral(C-D) noise | 0.099
| 부산대학교_공과대학 | 2023-08-31 | Full range of Z Height during Z Scan sweep | 16.5
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of Z Height during Z Scan sweep | 16.1
| University of Texas at Arlington | 2023-08-31 | Full range of Z Height during Z Scan sweep | 17.2
| 부산대학교_공과대학 | 2023-08-31 | Full range of Z Detector during Z Scan sweep | 11.2
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of Z Detector during Z Scan sweep | 12
| University of Texas at Arlington | 2023-08-31 | Full range of Z Detector during Z Scan sweep | 12.3
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Full range of Z Detector during Z Scan sweep | 12.6
| 부산대학교_공과대학 | 2023-08-31 | Full range of Y2 Detector Raw during Y Scan sweep | 13.6
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| 부산대학교_공과대학 | 2023-08-31 | Full range of Y1 Detector Raw during Y Scan sweep | 13.7
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| University of Texas at Arlington | 2023-08-31 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Full range of Y Detector Raw during Y Scan sweep | 14
| 부산대학교_공과대학 | 2023-08-31 | Full range of Y Detector during Y Scan sweep | 183.5
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of Y Detector during Y Scan sweep | 194
| University of Texas at Arlington | 2023-08-31 | Full range of Y Detector during Y Scan sweep | 113.2
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Full range of Y Detector during Y Scan sweep | 109.8
| 부산대학교_공과대학 | 2023-08-31 | Full range of X2 Detector Raw during X Scan sweep | 13.6
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| 부산대학교_공과대학 | 2023-08-31 | Full range of X1 Detector Raw during X Scan sweep | 13.6
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| University of Texas at Arlington | 2023-08-31 | Full range of X Detector Raw during X Scan sweep | 13.6
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Full range of X Detector Raw during X Scan sweep | 15
| 부산대학교_공과대학 | 2023-08-31 | Full range of X Detector during X Scan sweep | 182.9
| Sincere International Trading Company Limited #2 | 2023-08-31 | Full range of X Detector during X Scan sweep | 197
| University of Texas at Arlington | 2023-08-31 | Full range of X Detector during X Scan sweep | 112.2
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Full range of X Detector during X Scan sweep | 111.2
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Approach position at -3μm | 0.0133
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | Approach position at +3μm | 0.0162
| 부산대학교_공과대학 | 2023-08-31 | A-B noise | 15.17
| Sincere International Trading Company Limited #2 | 2023-08-31 | A-B noise | 10.78
| University of Texas at Arlington | 2023-08-31 | A-B noise | 11.84
| Westlake University_Jia Xing Huang Lab | 2023-08-31 | A-B noise | 17.68
| TSMC F23(Kumamoto) | 2023-08-30 | Z Linearity by Z scanner position | 0.7
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Linearity by Z scanner position | 0.57
| TSMC F23(Kumamoto) | 2023-08-30 | Z Linearity by 5 different sample height | 1.001
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Linearity by 4 different sample height | 1.013
| TSMC F23(Kumamoto) | 2023-08-30 | Z Height noise | 0.028
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Height noise | 0.04
| TSMC F23(Kumamoto) | 2023-08-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.8056
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.3
| TSMC F23(Kumamoto) | 2023-08-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.55
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.31
| TSMC F23(Kumamoto) | 2023-08-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.4902
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.18
| TSMC F23(Kumamoto) | 2023-08-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3915
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.16
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Drift (Short Traveling) | 0.12
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Z Detector offset | 23.5
| TSMC F23(Kumamoto) | 2023-08-30 | Y2 Detector noise (XY Servo on) | 0.38
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y2 Detector noise (XY Servo on) | 0.31
| TSMC F23(Kumamoto) | 2023-08-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 244.7
| TSMC F23(Kumamoto) | 2023-08-30 | Y1 Detector noise (XY Servo on) | 0.401
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y1 Detector noise (XY Servo on) | 0.3
| TSMC F23(Kumamoto) | 2023-08-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 266.3
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 244.57
| TSMC F23(Kumamoto) | 2023-08-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 191.2
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.83
| TSMC F23(Kumamoto) | 2023-08-30 | Y Ringing(Rpv) | 0.177
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y Ringing(Rpv) | 0.11
| TSMC F23(Kumamoto) | 2023-08-30 | Y Non-linearity error | -3.37
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y Non-linearity error | -1.73
| TSMC F23(Kumamoto) | 2023-08-30 | Y Detector noise (XY Servo on) | 0.34
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y Detector noise (XY Servo on) | 0.21
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Repeatability @1st position | 0.04
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y axis, Repeatability @1st position | 0.41
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Repeatability @ 4th position | 0.03
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y axis, Repeatability @ 4th position | 0.49
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Repeatability @ 3rd position | 0.06
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y axis, Repeatability @ 3rd position | 0.41
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Repeatability @ 2nd position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y axis, Repeatability @ 2nd position | 0.45
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Accuracy @1st position | 0.56
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Accuracy @ 4th position | 0.41
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Accuracy @ 3rd position | 0.66
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis, Accuracy @ 2nd position | 0.49
| TSMC F23(Kumamoto) | 2023-08-30 | Y axis OPM | 0.653
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Y axis OPM | 0.22
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | XY Orthogonality | 89.9
| TSMC F23(Kumamoto) | 2023-08-30 | X2 Detector noise (XY Servo on) | 0.34
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X2 Detector noise (XY Servo on) | 0.26
| TSMC F23(Kumamoto) | 2023-08-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 284.8
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 258.02
| TSMC F23(Kumamoto) | 2023-08-30 | X1 Detector noise (XY Servo on) | 0.355
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X1 Detector noise (XY Servo on) | 0.26
| TSMC F23(Kumamoto) | 2023-08-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 271.7
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 260.68
| TSMC F23(Kumamoto) | 2023-08-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.7
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.3
| TSMC F23(Kumamoto) | 2023-08-30 | X Ringing(Rpv) | 0.41
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X Ringing(Rpv) | 0.15
| TSMC F23(Kumamoto) | 2023-08-30 | X Non-linearity error | 3.03
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X Non-linearity error | 2.36
| TSMC F23(Kumamoto) | 2023-08-30 | X Detector noise (XY Servo on) | 0.28
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X Detector noise (XY Servo on) | 0.19
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Repeatability @1st position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X axis, Repeatability @1st position | 0.2
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Repeatability @ 4th position | 0.04
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X axis, Repeatability @ 4th position | 0.23
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Repeatability @ 3rd position | 0.04
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X axis, Repeatability @ 3rd position | 0.27
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Repeatability @ 2nd position | 0.06
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X axis, Repeatability @ 2nd position | 0.35
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Accuracy @1st position | 0.63
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Accuracy @ 4th position | 0.46
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Accuracy @ 3rd position | 0.54
| TSMC F23(Kumamoto) | 2023-08-30 | X axis, Accuracy @ 2nd position | 0.47
| TSMC F23(Kumamoto) | 2023-08-30 | X axis OPM | 0.728
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | X axis OPM | 0.49
| TSMC F23(Kumamoto) | 2023-08-30 | Under 50nm Step Height Sample (10 Repeats) | 18.3
| TSMC F23(Kumamoto) | 2023-08-30 | Torque Profile (50mm, 20um/s) | 2.67
| TSMC F23(Kumamoto) | 2023-08-30 | Torque Profile (50mm, 200um/s) | 2.27
| TSMC F23(Kumamoto) | 2023-08-30 | StrokeYMin | -153405
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | StrokeYMin | -102001
| TSMC F23(Kumamoto) | 2023-08-30 | StrokeYMax | 150941
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | StrokeYMax | 102522
| TSMC F23(Kumamoto) | 2023-08-30 | StrokeY | 304346
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | StrokeY | 204523
| TSMC F23(Kumamoto) | 2023-08-30 | StrokeXMin | -204117
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | StrokeXMin | -102600
| TSMC F23(Kumamoto) | 2023-08-30 | StrokeXMax | 205033
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | StrokeXMax | 104950
| TSMC F23(Kumamoto) | 2023-08-30 | StrokeX | 409150
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | StrokeX | 207550
| TSMC F23(Kumamoto) | 2023-08-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.557
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.528
| TSMC F23(Kumamoto) | 2023-08-30 | Roll motion to X+ full scan | 6.7
| TSMC F23(Kumamoto) | 2023-08-30 | RMS of ZeroScan | 0.041
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | RMS of ZeroScan | 0.021
| TSMC F23(Kumamoto) | 2023-08-30 | Profiler noise @200um | 0.14
| TSMC F23(Kumamoto) | 2023-08-30 | PosMin | -26.074
| TSMC F23(Kumamoto) | 2023-08-30 | PosMax | 25.796
| TSMC F23(Kumamoto) | 2023-08-30 | Position repeatability(10repeats, Position Based) | 0.32
| TSMC F23(Kumamoto) | 2023-08-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00044
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 2e-05
| TSMC F23(Kumamoto) | 2023-08-30 | OpticalAxisShiftPerFocusY | -0.0033
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | OpticalAxisShiftPerFocusY | 0.0021
| TSMC F23(Kumamoto) | 2023-08-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00019
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 2e-05
| TSMC F23(Kumamoto) | 2023-08-30 | OpticalAxisShiftPerFocusX | -0.0042
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | OpticalAxisShiftPerFocusX | 0.0044
| TSMC F23(Kumamoto) | 2023-08-30 | OPM Repeatability @5mm | 2.5
| TSMC F23(Kumamoto) | 2023-08-30 | OPM Repeatability @25mm | 2.94
| TSMC F23(Kumamoto) | 2023-08-30 | OPM Repeatability @1mm | 1.05
| TSMC F23(Kumamoto) | 2023-08-30 | OPM Repeatability @10mm | 1.7
| TSMC F23(Kumamoto) | 2023-08-30 | NCM driving% with AC160TS cantilever | 7.7
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | NCM driving% with AC160TS cantilever | 11.95
| TSMC F23(Kumamoto) | 2023-08-30 | Min. Vibration Transmissbility (~100Hz) | 0.038
| TSMC F23(Kumamoto) | 2023-08-30 | MaxLength | 50
| TSMC F23(Kumamoto) | 2023-08-30 | Max OPM @5mm | 34.26
| TSMC F23(Kumamoto) | 2023-08-30 | Max OPM @25mm | 39.06
| TSMC F23(Kumamoto) | 2023-08-30 | Max OPM @1mm | 11.73
| TSMC F23(Kumamoto) | 2023-08-30 | Max OPM @10mm | 19.24
| TSMC F23(Kumamoto) | 2023-08-30 | Lateral(C-D) sensitivity | 7.7
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Lateral(C-D) sensitivity | 31.2
| TSMC F23(Kumamoto) | 2023-08-30 | Lateral(C-D) noise | 0.198
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Lateral(C-D) noise | 0.096
| TSMC F23(Kumamoto) | 2023-08-30 | Laser intensity with AC160 cantilever | 3.3
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Laser intensity with AC160 cantilever | 2.8
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from -Y to +Y (Y scan) | 90.82
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from -Y to +Y (X scan) | 18.31
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from -X to +X (Y scan) | 43.94
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from -X to +X (X scan) | 37.35
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from +Y to -Y (Y scan) | 12.45
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from +Y to -Y (X scan) | 32.23
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from +X to -X (Y scan) | 30.76
| TSMC F23(Kumamoto) | 2023-08-30 | Image offset after moving from +X to -X (X scan) | 13.18
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of Z Height during Z Scan sweep | 17.2
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of Z Height during Z Scan sweep | 17.1
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of Z Detector during Z Scan sweep | 11.9
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of Z Detector during Z Scan sweep | 12.38
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.13
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.12
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of Y2 Detector during Y Scan sweep | 204.2195
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of Y2 Detector during Y Scan sweep | 168.85
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.45
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of Y1 Detector Raw during Y Scan sweep | 13.89
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of Y1 Detector during Y Scan sweep | 208.7762
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of Y1 Detector during Y Scan sweep | 172.63
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of X2 Detector Raw during X Scan sweep | 14.11
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of X2 Detector during X Scan sweep | 212.0797
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of X2 Detector during X Scan sweep | 177.66
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of X1 Detector Raw during X Scan sweep | 13.87
| TSMC F23(Kumamoto) | 2023-08-30 | Full range of X1 Detector during X Scan sweep | 213.2958
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | Full range of X1 Detector during X Scan sweep | 170.99
| TSMC F23(Kumamoto) | 2023-08-30 | Allowable max warpage, w/ SiC sample chuck | 300
| TSMC F23(Kumamoto) | 2023-08-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| TSMC F23(Kumamoto) | 2023-08-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.64
| TSMC F23(Kumamoto) | 2023-08-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.65
| TSMC F23(Kumamoto) | 2023-08-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| TSMC F23(Kumamoto) | 2023-08-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| TSMC F23(Kumamoto) | 2023-08-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.62
| TSMC F23(Kumamoto) | 2023-08-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| TSMC F23(Kumamoto) | 2023-08-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| TSMC F23(Kumamoto) | 2023-08-30 | A－B difference during 3 times run test | 0.58
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | A-B after Auto PSPD Alignment | -0.17
| TSMC F23(Kumamoto) | 2023-08-30 | A＋B difference during 3 times run test | 0.01
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | 940nm Step Height Sample (10 Repeats) | 948.22
| TSMC F23(Kumamoto) | 2023-08-30 | 50~200nm Step Height Sample (10 Repeats) | 184.22
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | 50~200nm Step Height Sample (10 Repeats) | 183.56
| TSMC F23(Kumamoto) | 2023-08-30 | 200~500nm Step Height Sample (10 Repeats) | 459.99
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | 200~500nm Step Height Sample (10 Repeats) | 461.968
| TSMC F23(Kumamoto) | 2023-08-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.3
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.1
| TSMC F23(Kumamoto) | 2023-08-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 12
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.9
| TSMC F23(Kumamoto) | 2023-08-30 | %GRR w/ roughness of non-patterened wafer | 14.5
| SINCERE INTERNATION TRADE CO.,LIMITED #2 | 2023-08-30 | %GRR w/ roughness of non-patterened wafer | 10
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Linearity by Z scanner position | 0.8
| CXT #4 | 2023-08-28 | Z Linearity by Z scanner position | 0.506
| CXT #4 | 2023-08-28 | Z Linearity by Z scanner position | 0.28131
| Nexchip #3 | 2023-08-28 | Z Linearity by Z scanner position | 0.65
| CXT #4 | 2023-08-28 | Z Linearity by 5 different sample height | 1.004
| Nexchip #3 | 2023-08-28 | Z Linearity by 5 different sample height | 1.002
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Linearity by 4 different sample height | 1.001
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Height noise | 0.039
| CXT #4 | 2023-08-28 | Z Height noise | 0.03
| Nexchip #3 | 2023-08-28 | Z Height noise | 0.03
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 29.7616
| CXT #4 | 2023-08-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.1
| Nexchip #3 | 2023-08-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 29.28
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 29.8249
| CXT #4 | 2023-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.9
| Nexchip #3 | 2023-08-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 29.2
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.856
| CXT #4 | 2023-08-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1
| Nexchip #3 | 2023-08-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.7217
| CXT #4 | 2023-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11
| Nexchip #3 | 2023-08-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.71
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Drift (Short Traveling) | 0.19
| CXT #4 | 2023-08-28 | Z Drift (Short Traveling) | 0.29
| Nexchip #3 | 2023-08-28 | Z Drift (Short Traveling) | 0.29
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Z Detector offset | 28.7
| CXT #4 | 2023-08-28 | Z Detector offset | 37.7
| CXT #4 | 2023-08-28 | Z Detector offset | 2.2
| Nexchip #3 | 2023-08-28 | Z Detector offset | 2.85
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y2 Detector noise (XY Servo on) | 0.28
| CXT #4 | 2023-08-28 | Y2 Detector noise (XY Servo on) | 0.272
| Nexchip #3 | 2023-08-28 | Y2 Detector noise (XY Servo on) | 0.38
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 241
| CXT #4 | 2023-08-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 267.3
| Nexchip #3 | 2023-08-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 263.56
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y1 Detector noise (XY Servo on) | 0.34
| CXT #4 | 2023-08-28 | Y1 Detector noise (XY Servo on) | 0.204
| Nexchip #3 | 2023-08-28 | Y1 Detector noise (XY Servo on) | 0.29
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 254.8
| CXT #4 | 2023-08-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 272.3
| Nexchip #3 | 2023-08-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 246.07
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.8
| CXT #4 | 2023-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 197.2
| Nexchip #3 | 2023-08-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 186.89
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y Ringing(Rpv) | 0.204
| CXT #4 | 2023-08-28 | Y Ringing(Rpv) | 0.473
| Nexchip #3 | 2023-08-28 | Y Ringing(Rpv) | 0.265
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y Non-linearity error | 2.93
| CXT #4 | 2023-08-28 | Y Non-linearity error | -3.01
| Nexchip #3 | 2023-08-28 | Y Non-linearity error | 1.39
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y Detector noise (XY Servo on) | 0.23
| CXT #4 | 2023-08-28 | Y Detector noise (XY Servo on) | 0.295
| Nexchip #3 | 2023-08-28 | Y Detector noise (XY Servo on) | 0.25
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y axis, Repeatability @1st position | 0.08
| CXT #4 | 2023-08-28 | Y axis, Repeatability @1st position | 0.07
| Nexchip #3 | 2023-08-28 | Y axis, Repeatability @1st position | 0.09
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y axis, Repeatability @ 4th position | 0.22
| CXT #4 | 2023-08-28 | Y axis, Repeatability @ 4th position | 0.07
| Nexchip #3 | 2023-08-28 | Y axis, Repeatability @ 4th position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y axis, Repeatability @ 3rd position | 0.22
| CXT #4 | 2023-08-28 | Y axis, Repeatability @ 3rd position | 0.08
| Nexchip #3 | 2023-08-28 | Y axis, Repeatability @ 3rd position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y axis, Repeatability @ 2nd position | 0.27
| CXT #4 | 2023-08-28 | Y axis, Repeatability @ 2nd position | 0.08
| Nexchip #3 | 2023-08-28 | Y axis, Repeatability @ 2nd position | 0.08
| CXT #4 | 2023-08-28 | Y axis, Accuracy @1st position | 0.78
| Nexchip #3 | 2023-08-28 | Y axis, Accuracy @1st position | 0.69
| CXT #4 | 2023-08-28 | Y axis, Accuracy @ 4th position | 0.72
| Nexchip #3 | 2023-08-28 | Y axis, Accuracy @ 4th position | 0.46
| CXT #4 | 2023-08-28 | Y axis, Accuracy @ 3rd position | 0.75
| Nexchip #3 | 2023-08-28 | Y axis, Accuracy @ 3rd position | 0.61
| CXT #4 | 2023-08-28 | Y axis, Accuracy @ 2nd position | 0.79
| Nexchip #3 | 2023-08-28 | Y axis, Accuracy @ 2nd position | 0.73
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Y axis OPM | 0.78
| CXT #4 | 2023-08-28 | Y axis OPM | 0.873
| Nexchip #3 | 2023-08-28 | Y axis OPM | 0.652
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | XY Orthogonality | 90
| CXT #4 | 2023-08-28 | XY Orthogonality | 89.9
| Nexchip #3 | 2023-08-28 | XY Orthogonality | 90.2
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X2 Detector noise (XY Servo on) | 0.29
| CXT #4 | 2023-08-28 | X2 Detector noise (XY Servo on) | 0.277
| Nexchip #3 | 2023-08-28 | X2 Detector noise (XY Servo on) | 0.38
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 252.6
| CXT #4 | 2023-08-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 278.3
| Nexchip #3 | 2023-08-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 266.93
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X1 Detector noise (XY Servo on) | 0.3
| CXT #4 | 2023-08-28 | X1 Detector noise (XY Servo on) | 0.295
| Nexchip #3 | 2023-08-28 | X1 Detector noise (XY Servo on) | 0.35
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 246.4
| CXT #4 | 2023-08-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 274.6
| Nexchip #3 | 2023-08-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 241.76
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 195.5
| CXT #4 | 2023-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 186.7
| Nexchip #3 | 2023-08-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 187.14
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X Ringing(Rpv) | 0.101
| CXT #4 | 2023-08-28 | X Ringing(Rpv) | 0.214
| Nexchip #3 | 2023-08-28 | X Ringing(Rpv) | 0.211
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X Non-linearity error | 1.62
| CXT #4 | 2023-08-28 | X Non-linearity error | -2.75
| Nexchip #3 | 2023-08-28 | X Non-linearity error | 2.32
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X Detector noise (XY Servo on) | 0.22
| CXT #4 | 2023-08-28 | X Detector noise (XY Servo on) | 0.23
| Nexchip #3 | 2023-08-28 | X Detector noise (XY Servo on) | 0.29
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X axis, Repeatability @1st position | 0.09
| CXT #4 | 2023-08-28 | X axis, Repeatability @1st position | 0.02
| Nexchip #3 | 2023-08-28 | X axis, Repeatability @1st position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X axis, Repeatability @ 4th position | 0.1
| CXT #4 | 2023-08-28 | X axis, Repeatability @ 4th position | 0.03
| Nexchip #3 | 2023-08-28 | X axis, Repeatability @ 4th position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X axis, Repeatability @ 3rd position | 0.08
| CXT #4 | 2023-08-28 | X axis, Repeatability @ 3rd position | 0.03
| Nexchip #3 | 2023-08-28 | X axis, Repeatability @ 3rd position | 0.06
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X axis, Repeatability @ 2nd position | 0.07
| CXT #4 | 2023-08-28 | X axis, Repeatability @ 2nd position | 0.03
| Nexchip #3 | 2023-08-28 | X axis, Repeatability @ 2nd position | 0.06
| CXT #4 | 2023-08-28 | X axis, Accuracy @1st position | 0.12
| Nexchip #3 | 2023-08-28 | X axis, Accuracy @1st position | 0.32
| CXT #4 | 2023-08-28 | X axis, Accuracy @ 4th position | 0.26
| Nexchip #3 | 2023-08-28 | X axis, Accuracy @ 4th position | 0.39
| CXT #4 | 2023-08-28 | X axis, Accuracy @ 3rd position | 0.18
| Nexchip #3 | 2023-08-28 | X axis, Accuracy @ 3rd position | 0.37
| CXT #4 | 2023-08-28 | X axis, Accuracy @ 2nd position | 0.2
| Nexchip #3 | 2023-08-28 | X axis, Accuracy @ 2nd position | 0.3
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | X axis OPM | 0.43
| CXT #4 | 2023-08-28 | X axis OPM | 0.837
| Nexchip #3 | 2023-08-28 | X axis OPM | 0.582
| CXT #4 | 2023-08-28 | Under 50nm Step Height Sample (10 Repeats) | 18.09
| CXT #4 | 2023-08-28 | Under 50nm Step Height Sample (10 Repeats) | 18.079
| Nexchip #3 | 2023-08-28 | Under 50nm Step Height Sample (10 Repeats) | 18.168
| CXT #4 | 2023-08-28 | Torque Profile (50mm, 20um/s) | 2.486
| CXT #4 | 2023-08-28 | Torque Profile (50mm, 200um/s) | 2.285
| Nexchip #3 | 2023-08-28 | Torque Profile (10mm, 20um/s) | 2.631
| Nexchip #3 | 2023-08-28 | Torque Profile (10mm, 200um/s) | 2.691
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | StrokeYMin | -101952
| CXT #4 | 2023-08-28 | StrokeYMin | -153665
| Nexchip #3 | 2023-08-28 | StrokeYMin | -153770
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | StrokeYMax | 103766
| CXT #4 | 2023-08-28 | StrokeYMax | 150152
| Nexchip #3 | 2023-08-28 | StrokeYMax | 150477
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | StrokeY | 205718
| CXT #4 | 2023-08-28 | StrokeY | 303817
| Nexchip #3 | 2023-08-28 | StrokeY | 304247
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | StrokeXMin | -102485
| CXT #4 | 2023-08-28 | StrokeXMin | -203977
| Nexchip #3 | 2023-08-28 | StrokeXMin | -203814
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | StrokeXMax | 104400
| CXT #4 | 2023-08-28 | StrokeXMax | 203758
| Nexchip #3 | 2023-08-28 | StrokeXMax | 203963
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | StrokeX | 206885
| CXT #4 | 2023-08-28 | StrokeX | 407735
| Nexchip #3 | 2023-08-28 | StrokeX | 407777
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.762
| CXT #4 | 2023-08-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.715
| Nexchip #3 | 2023-08-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.77
| CXT #4 | 2023-08-28 | Roll motion to X+ full scan | 0.2
| Nexchip #3 | 2023-08-28 | Roll motion to X+ full scan | 10.7
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | RMS of ZeroScan | 0.033
| CXT #4 | 2023-08-28 | RMS of ZeroScan | 0.057
| Nexchip #3 | 2023-08-28 | RMS of ZeroScan | 0.06
| CXT #4 | 2023-08-28 | Profiler noise @200um | 0.211
| Nexchip #3 | 2023-08-28 | Profiler noise @200um | 0.219
| CXT #4 | 2023-08-28 | PosMin | -26.097
| Nexchip #3 | 2023-08-28 | PosMin | -5.2
| CXT #4 | 2023-08-28 | PosMax | 25.852
| Nexchip #3 | 2023-08-28 | PosMax | 5.2
| CXT #4 | 2023-08-28 | Position repeatability(10repeats, Position Based) | 0.37
| Nexchip #3 | 2023-08-28 | Position repeatability(10repeats, Position Based) | 0.58
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 4.22e-05
| CXT #4 | 2023-08-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.0001
| Nexchip #3 | 2023-08-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | OpticalAxisShiftPerFocusY | -0.0013
| CXT #4 | 2023-08-28 | OpticalAxisShiftPerFocusY | -0.0026
| Nexchip #3 | 2023-08-28 | OpticalAxisShiftPerFocusY | 0.0006
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 6.99e-05
| CXT #4 | 2023-08-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| Nexchip #3 | 2023-08-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 6e-05
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | OpticalAxisShiftPerFocusX | -0.0012
| CXT #4 | 2023-08-28 | OpticalAxisShiftPerFocusX | -0.0027
| Nexchip #3 | 2023-08-28 | OpticalAxisShiftPerFocusX | -0.0031
| CXT #4 | 2023-08-28 | OPM Repeatability @5mm | 1.596
| Nexchip #3 | 2023-08-28 | OPM Repeatability @5mm | 1.3
| CXT #4 | 2023-08-28 | OPM Repeatability @25mm | 7.721
| CXT #4 | 2023-08-28 | OPM Repeatability @1mm | 0.844
| Nexchip #3 | 2023-08-28 | OPM Repeatability @1mm | 0.438
| CXT #4 | 2023-08-28 | OPM Repeatability @10mm | 4.972
| Nexchip #3 | 2023-08-28 | OPM Repeatability @10mm | 1.313
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | NCM driving% with AC160TS cantilever | 35
| CXT #4 | 2023-08-28 | NCM driving% with AC160TS cantilever | 9.8
| CXT #4 | 2023-08-28 | NCM driving% with AC160TS cantilever | 31.22
| Nexchip #3 | 2023-08-28 | NCM driving% with AC160TS cantilever | 27.3
| CXT #4 | 2023-08-28 | Min. Vibration Transmissbility (~100Hz) | 0.005
| Nexchip #3 | 2023-08-28 | Min. Vibration Transmissbility (~100Hz) | 0.02
| CXT #4 | 2023-08-28 | MaxLength | 51
| Nexchip #3 | 2023-08-28 | MaxLength | 20
| CXT #4 | 2023-08-28 | Max OPM @5mm | 29.31
| Nexchip #3 | 2023-08-28 | Max OPM @5mm | 12.698
| CXT #4 | 2023-08-28 | Max OPM @25mm | 136.646
| CXT #4 | 2023-08-28 | Max OPM @1mm | 8.484
| Nexchip #3 | 2023-08-28 | Max OPM @1mm | 5.599
| CXT #4 | 2023-08-28 | Max OPM @10mm | 50.296
| Nexchip #3 | 2023-08-28 | Max OPM @10mm | 18.873
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Lateral(C-D) sensitivity | 30.1
| CXT #4 | 2023-08-28 | Lateral(C-D) sensitivity | 67.678
| Nexchip #3 | 2023-08-28 | Lateral(C-D) sensitivity | 33.8
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Lateral(C-D) noise | 0.13
| CXT #4 | 2023-08-28 | Lateral(C-D) noise | 0.248
| Nexchip #3 | 2023-08-28 | Lateral(C-D) noise | 0.233
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Laser intensity with AC160 cantilever | 2.9
| CXT #4 | 2023-08-28 | Laser intensity with AC160 cantilever | 3.03
| CXT #4 | 2023-08-28 | Laser intensity with AC160 cantilever | 2.971
| Nexchip #3 | 2023-08-28 | Laser intensity with AC160 cantilever | 3.02
| CXT #4 | 2023-08-28 | Image offset after moving from -Y to +Y (Y scan) | 142.82
| Nexchip #3 | 2023-08-28 | Image offset after moving from -Y to +Y (Y scan) | 43.21
| CXT #4 | 2023-08-28 | Image offset after moving from -Y to +Y (X scan) | 35.16
| Nexchip #3 | 2023-08-28 | Image offset after moving from -Y to +Y (X scan) | 5.86
| CXT #4 | 2023-08-28 | Image offset after moving from -X to +X (Y scan) | 180.91
| Nexchip #3 | 2023-08-28 | Image offset after moving from -X to +X (Y scan) | 13.18
| CXT #4 | 2023-08-28 | Image offset after moving from -X to +X (X scan) | 155.27
| Nexchip #3 | 2023-08-28 | Image offset after moving from -X to +X (X scan) | 79.83
| CXT #4 | 2023-08-28 | Image offset after moving from +Y to -Y (Y scan) | 52.73
| Nexchip #3 | 2023-08-28 | Image offset after moving from +Y to -Y (Y scan) | 9.52
| CXT #4 | 2023-08-28 | Image offset after moving from +Y to -Y (X scan) | 98.88
| Nexchip #3 | 2023-08-28 | Image offset after moving from +Y to -Y (X scan) | 4.39
| CXT #4 | 2023-08-28 | Image offset after moving from +X to -X (Y scan) | 111.33
| Nexchip #3 | 2023-08-28 | Image offset after moving from +X to -X (Y scan) | 42.48
| CXT #4 | 2023-08-28 | Image offset after moving from +X to -X (X scan) | 109.13
| Nexchip #3 | 2023-08-28 | Image offset after moving from +X to -X (X scan) | 106.93
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of Z Height during Z Scan sweep | 16.1
| CXT #4 | 2023-08-28 | Full range of Z Height during Z Scan sweep | 17
| CXT #4 | 2023-08-28 | Full range of Z Height during Z Scan sweep | 47.1
| Nexchip #3 | 2023-08-28 | Full range of Z Height during Z Scan sweep | 16.1
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of Z Detector during Z Scan sweep | 10.9
| CXT #4 | 2023-08-28 | Full range of Z Detector during Z Scan sweep | 13
| CXT #4 | 2023-08-28 | Full range of Z Detector during Z Scan sweep | 14.1
| Nexchip #3 | 2023-08-28 | Full range of Z Detector during Z Scan sweep | 10.98
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 13.4
| CXT #4 | 2023-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 13.2
| Nexchip #3 | 2023-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of Y2 Detector during Y Scan sweep | 183.8798
| CXT #4 | 2023-08-28 | Full range of Y2 Detector during Y Scan sweep | 193.239
| Nexchip #3 | 2023-08-28 | Full range of Y2 Detector during Y Scan sweep | 190.8135
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 13.49
| CXT #4 | 2023-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 13.3
| Nexchip #3 | 2023-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of Y1 Detector during Y Scan sweep | 179.2352
| CXT #4 | 2023-08-28 | Full range of Y1 Detector during Y Scan sweep | 194.1475
| Nexchip #3 | 2023-08-28 | Full range of Y1 Detector during Y Scan sweep | 191.5977
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of X2 Detector Raw during X Scan sweep | 12.73
| CXT #4 | 2023-08-28 | Full range of X2 Detector Raw during X Scan sweep | 12.8
| Nexchip #3 | 2023-08-28 | Full range of X2 Detector Raw during X Scan sweep | 14.02
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of X2 Detector during X Scan sweep | 191.2303
| CXT #4 | 2023-08-28 | Full range of X2 Detector during X Scan sweep | 203.3629
| Nexchip #3 | 2023-08-28 | Full range of X2 Detector during X Scan sweep | 190.9822
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of X1 Detector Raw during X Scan sweep | 12.66
| CXT #4 | 2023-08-28 | Full range of X1 Detector Raw during X Scan sweep | 12.7
| Nexchip #3 | 2023-08-28 | Full range of X1 Detector Raw during X Scan sweep | 14.95
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | Full range of X1 Detector during X Scan sweep | 192.0903
| CXT #4 | 2023-08-28 | Full range of X1 Detector during X Scan sweep | 198.5324
| Nexchip #3 | 2023-08-28 | Full range of X1 Detector during X Scan sweep | 198.1771
| CXT #4 | 2023-08-28 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 900
| Nexchip #3 | 2023-08-28 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 782.26
| CXT #4 | 2023-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Nexchip #3 | 2023-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| CXT #4 | 2023-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.79
| Nexchip #3 | 2023-08-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.7
| CXT #4 | 2023-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| Nexchip #3 | 2023-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.71
| CXT #4 | 2023-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Nexchip #3 | 2023-08-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| CXT #4 | 2023-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| Nexchip #3 | 2023-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| CXT #4 | 2023-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| Nexchip #3 | 2023-08-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.19
| CXT #4 | 2023-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.26
| Nexchip #3 | 2023-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.23
| CXT #4 | 2023-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| Nexchip #3 | 2023-08-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.04
| CXT #4 | 2023-08-28 | A－B difference during 3 times run test | 0.47
| Nexchip #3 | 2023-08-28 | A－B difference during 3 times run test | 0.71
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | A-B after Auto PSPD Alignment | 0.076
| Nexchip #3 | 2023-08-28 | A-B after Auto PSPD Alignment | -0.26
| CXT #4 | 2023-08-28 | A＋B difference during 3 times run test | 0.03
| Nexchip #3 | 2023-08-28 | A＋B difference during 3 times run test | 0.09
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | 940nm Step Height Sample (10 Repeats) | 951.98
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | 50~200nm Step Height Sample (10 Repeats) | 184.16
| CXT #4 | 2023-08-28 | 50~200nm Step Height Sample (10 Repeats) | 176.02
| CXT #4 | 2023-08-28 | 50~200nm Step Height Sample (10 Repeats) | 183.292
| Nexchip #3 | 2023-08-28 | 50~200nm Step Height Sample (10 Repeats) | 185.029
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | 200~500nm Step Height Sample (10 Repeats) | 461.34
| CXT #4 | 2023-08-28 | 200~500nm Step Height Sample (10 Repeats) | 463.15
| CXT #4 | 2023-08-28 | 200~500nm Step Height Sample (10 Repeats) | 462.754
| Nexchip #3 | 2023-08-28 | 200~500nm Step Height Sample (10 Repeats) | 462.545
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.3
| CXT #4 | 2023-08-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.4
| Nexchip #3 | 2023-08-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 13.4
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 9.4
| CXT #4 | 2023-08-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 12.4
| Nexchip #3 | 2023-08-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 6
| SINCERE INTERNATION TRADE CO.,LIMITED #3 | 2023-08-28 | %GRR w/ roughness of non-patterened wafer | 9.7
| CXT #4 | 2023-08-28 | %GRR w/ roughness of non-patterened wafer | 13.2
| Nexchip #3 | 2023-08-28 | %GRR w/ roughness of non-patterened wafer | 14.7
| Hisense Broadband,Inc. | 2023-08-28 | Z Scanner position at -3μm | 0.0037
| Park Systems Singapore(demo 장비) | 2023-08-28 | Z Scanner position at -3μm | 0.0062
| Hisense Broadband,Inc. | 2023-08-28 | Z Scanner position at +3μm | 0.0173
| Park Systems Singapore(demo 장비) | 2023-08-28 | Z Scanner position at +3μm | 0.0008
| Hisense Broadband,Inc. | 2023-08-28 | Z Linearity by Z Scanner position
| Park Systems Singapore(demo 장비) | 2023-08-28 | Z Linearity by Z Scanner position
| Hisense Broadband,Inc. | 2023-08-28 | Z Linearity by 3 different sample height | 1.004
| Park Systems Singapore(demo 장비) | 2023-08-28 | Z Linearity by 3 different sample height | 0.996
| Hisense Broadband,Inc. | 2023-08-28 | Z Height noise | 0.029
| Park Systems Singapore(demo 장비) | 2023-08-28 | Z Height noise | 0.029
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y2 Detector noise (XY Servo on) | 0.41
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y2 Detector noise (XY Servo off) | 0.36
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y1 Detector noise (XY Servo on) | 0.4
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y1 Detector noise (XY Servo off) | 0.36
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y Scan Stroke | 169.533
| Park Systems Singapore(demo 장비) | 2023-08-28 | -Y Direction software Limit | -75138
| Hisense Broadband,Inc. | 2023-08-28 | Y Detector noise (XY Servo on) | 0.27
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y Detector noise (XY Servo on) | 0.3
| Hisense Broadband,Inc. | 2023-08-28 | Y Detector noise (XY Servo off) | 0.27
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y Detector noise (XY Servo off) | 0.26
| Hisense Broadband,Inc. | 2023-08-28 | Y axis OPM | 0.67
| Park Systems Singapore(demo 장비) | 2023-08-28 | Y axis OPM | 0.76
| Park Systems Singapore(demo 장비) | 2023-08-28 | X2 Detector noise (XY Servo on) | 0.41
| Park Systems Singapore(demo 장비) | 2023-08-28 | X2 Detector noise (XY Servo off) | 0.36
| Park Systems Singapore(demo 장비) | 2023-08-28 | X1 Detector noise (XY Servo on) | 0.4
| Park Systems Singapore(demo 장비) | 2023-08-28 | X1 Detector noise (XY Servo off) | 0.37
| Park Systems Singapore(demo 장비) | 2023-08-28 | X Scan Stroke | 170.25
| Park Systems Singapore(demo 장비) | 2023-08-28 | -X Direction software Limit | -75546
| Hisense Broadband,Inc. | 2023-08-28 | X Detector noise (XY Servo on) | 0.26
| Park Systems Singapore(demo 장비) | 2023-08-28 | X Detector noise (XY Servo on) | 0.31
| Hisense Broadband,Inc. | 2023-08-28 | X Detector noise (XY Servo off) | 0.26
| Park Systems Singapore(demo 장비) | 2023-08-28 | X Detector noise (XY Servo off) | 0.26
| Hisense Broadband,Inc. | 2023-08-28 | X axis OPM | 1.14
| Park Systems Singapore(demo 장비) | 2023-08-28 | X axis OPM | 0.85
| Hisense Broadband,Inc. | 2023-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Hisense Broadband,Inc. | 2023-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of Wafer Zeroscan(-60000,0) | 0.29
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of Wafer Zeroscan(60000,0) | 0.3
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of Wafer Zeroscan(0,-60000) | 0.3
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of Wafer Zeroscan(0,60000) | 0.29
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Height of Wafer ZeroScan (0,0) | 0.29
| Hisense Broadband,Inc. | 2023-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Hisense Broadband,Inc. | 2023-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Park Systems Singapore(demo 장비) | 2023-08-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| Hisense Broadband,Inc. | 2023-08-28 | NCM driving% with AC160TS cantilever | 1.29
| Park Systems Singapore(demo 장비) | 2023-08-28 | NCM driving% with AC160TS cantilever | 1.595
| Hisense Broadband,Inc. | 2023-08-28 | NCM Amplitude noise | 15.46
| Park Systems Singapore(demo 장비) | 2023-08-28 | NCM Amplitude noise | 15.14
| Hisense Broadband,Inc. | 2023-08-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.174
| Park Systems Singapore(demo 장비) | 2023-08-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.331
| Hisense Broadband,Inc. | 2023-08-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.939
| Park Systems Singapore(demo 장비) | 2023-08-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.702
| Park Systems Singapore(demo 장비) | 2023-08-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 252.3
| Park Systems Singapore(demo 장비) | 2023-08-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 244.5
| Hisense Broadband,Inc. | 2023-08-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 154.7
| Park Systems Singapore(demo 장비) | 2023-08-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 248.5
| Park Systems Singapore(demo 장비) | 2023-08-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 250.4
| Hisense Broadband,Inc. | 2023-08-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156.1
| Park Systems Singapore(demo 장비) | 2023-08-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Hisense Broadband,Inc. | 2023-08-28 | Lateral(C-D) sensitivity | 5.8
| Park Systems Singapore(demo 장비) | 2023-08-28 | Lateral(C-D) sensitivity | 6
| Hisense Broadband,Inc. | 2023-08-28 | Lateral(C-D) noise | 0.109
| Park Systems Singapore(demo 장비) | 2023-08-28 | Lateral(C-D) noise | 0.042
| Hisense Broadband,Inc. | 2023-08-28 | Full range of Z Height during Z Scan sweep | 15.6
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of Z Height during Z Scan sweep | 16
| Hisense Broadband,Inc. | 2023-08-28 | Full range of Z Detector during Z Scan sweep | 12.5
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of Z Detector during Z Scan sweep | 11.6
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of Y2 Detector Raw during Y Scan sweep | 13.5
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of Y1 Detector Raw during Y Scan sweep | 13.7
| Hisense Broadband,Inc. | 2023-08-28 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Hisense Broadband,Inc. | 2023-08-28 | Full range of Y Detector during Y Scan sweep | 115.8
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of Y Detector during Y Scan sweep | 169.2
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of X2 Detector Raw during X Scan sweep | 13.7
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of X1 Detector Raw during X Scan sweep | 13.6
| Hisense Broadband,Inc. | 2023-08-28 | Full range of X Detector Raw during X Scan sweep | 13.9
| Hisense Broadband,Inc. | 2023-08-28 | Full range of X Detector during X Scan sweep | 110.4
| Park Systems Singapore(demo 장비) | 2023-08-28 | Full range of X Detector during X Scan sweep | 170.15
| Hisense Broadband,Inc. | 2023-08-28 | A-B noise | 9.96
| Park Systems Singapore(demo 장비) | 2023-08-28 | A-B noise | 8.33
| Columbia Univ | 2023-08-25 | Z Linearity by 3 different sample height | 0.998
| Columbia Univ | 2023-08-25 | Z Height noise | 0.022
| Columbia Univ | 2023-08-25 | Y Detector noise (XY Servo on) | 0.28
| Columbia Univ | 2023-08-25 | Y Detector noise (XY Servo off) | 0.27
| Columbia Univ | 2023-08-25 | Y axis OPM | 1.71
| Columbia Univ | 2023-08-25 | X Detector noise (XY Servo on) | 0.29
| Columbia Univ | 2023-08-25 | X Detector noise (XY Servo off) | 0.28
| Columbia Univ | 2023-08-25 | X axis OPM | 0.76
| Columbia Univ | 2023-08-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Columbia Univ | 2023-08-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Columbia Univ | 2023-08-25 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Columbia Univ | 2023-08-25 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Columbia Univ | 2023-08-25 | NCM driving% with AC160TS cantilever | 10.78
| Columbia Univ | 2023-08-25 | NCM Amplitude noise | 8.98
| Columbia Univ | 2023-08-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.603
| Columbia Univ | 2023-08-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.304
| Columbia Univ | 2023-08-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 157.7
| Columbia Univ | 2023-08-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.4
| Columbia Univ | 2023-08-25 | Full range of Z Height during Z Scan sweep | 16.4
| Columbia Univ | 2023-08-25 | Lateral(C-D) sensitivity | 3.1
| Columbia Univ | 2023-08-25 | Lateral(C-D) noise | 0.114
| Columbia Univ | 2023-08-25 | Full range of Z Detector during Z Scan sweep | 14.7
| Columbia Univ | 2023-08-25 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| Columbia Univ | 2023-08-25 | Full range of Y Detector during Y Scan sweep | 117.1
| Columbia Univ | 2023-08-25 | Full range of X Detector Raw during X Scan sweep | 14.3
| Columbia Univ | 2023-08-25 | Full range of X Detector during X Scan sweep | 117.1
| Columbia Univ | 2023-08-25 | Approach position at -3μm | 0.0104
| Columbia Univ | 2023-08-25 | Approach position at +3μm | 0.0107
| Columbia Univ | 2023-08-25 | A-B noise | 6.64
| 주식회사 에이치앤에스 | 2023-08-22 | Z Scanner position at -3μm | 0.0037
| 주식회사 에이치앤에스 | 2023-08-22 | Z Scanner position at +3μm | 0.0057
| 주식회사 에이치앤에스 | 2023-08-22 | Z Linearity by 3 different sample height | 1
| PSE demo lab | 2023-08-22 | Z Linearity by 3 different sample height | 1.006
| 주식회사 에이치앤에스 | 2023-08-22 | Z Height noise | 0.028
| PSE demo lab | 2023-08-22 | Z Height noise | 0.033
| 주식회사 에이치앤에스 | 2023-08-22 | Y Scan Stroke | 55
| 주식회사 에이치앤에스 | 2023-08-22 | Y Detector noise (XY Servo on) | 0.27
| PSE demo lab | 2023-08-22 | Y Detector noise (XY Servo on) | 0.27
| 주식회사 에이치앤에스 | 2023-08-22 | Y Detector noise (XY Servo off) | 0.26
| PSE demo lab | 2023-08-22 | Y Detector noise (XY Servo off) | 0.26
| 주식회사 에이치앤에스 | 2023-08-22 | Y axis OPM | 0.35
| PSE demo lab | 2023-08-22 | Y axis OPM | 0.64
| 주식회사 에이치앤에스 | 2023-08-22 | X Scan Stroke | 55
| 주식회사 에이치앤에스 | 2023-08-22 | X Detector noise (XY Servo on) | 0.27
| PSE demo lab | 2023-08-22 | X Detector noise (XY Servo on) | 0.28
| 주식회사 에이치앤에스 | 2023-08-22 | X Detector noise (XY Servo off) | 0.26
| PSE demo lab | 2023-08-22 | X Detector noise (XY Servo off) | 0.27
| 주식회사 에이치앤에스 | 2023-08-22 | X axis OPM | 0.88
| PSE demo lab | 2023-08-22 | X axis OPM | 0.37
| 주식회사 에이치앤에스 | 2023-08-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| PSE demo lab | 2023-08-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 주식회사 에이치앤에스 | 2023-08-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| PSE demo lab | 2023-08-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 주식회사 에이치앤에스 | 2023-08-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| PSE demo lab | 2023-08-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 주식회사 에이치앤에스 | 2023-08-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| PSE demo lab | 2023-08-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 주식회사 에이치앤에스 | 2023-08-22 | NCM driving% with AC160TS cantilever | 0.35
| PSE demo lab | 2023-08-22 | NCM driving% with AC160TS cantilever | 0.46
| 주식회사 에이치앤에스 | 2023-08-22 | NCM Amplitude noise | 11.09
| PSE demo lab | 2023-08-22 | NCM Amplitude noise | 14.77
| 주식회사 에이치앤에스 | 2023-08-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.257
| PSE demo lab | 2023-08-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| 주식회사 에이치앤에스 | 2023-08-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.59
| PSE demo lab | 2023-08-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.803
| 주식회사 에이치앤에스 | 2023-08-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102
| PSE demo lab | 2023-08-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 166.9
| 주식회사 에이치앤에스 | 2023-08-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 100.9
| PSE demo lab | 2023-08-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 163.9
| PSE demo lab | 2023-08-22 | Full range of Z Height during Z Scan sweep | 16.3
| 주식회사 에이치앤에스 | 2023-08-22 | Lateral(C-D) sensitivity | 8.8
| PSE demo lab | 2023-08-22 | Lateral(C-D) sensitivity | 3.1
| 주식회사 에이치앤에스 | 2023-08-22 | Lateral(C-D) noise | 0.047
| PSE demo lab | 2023-08-22 | Lateral(C-D) noise | 0.149
| 주식회사 에이치앤에스 | 2023-08-22 | Full range of Z Height during Z Scan sweep | 16.1
| 주식회사 에이치앤에스 | 2023-08-22 | Full range of Z Detector during Z Scan sweep | 11.3
| PSE demo lab | 2023-08-22 | Full range of Z Detector during Z Scan sweep | 13.7
| 주식회사 에이치앤에스 | 2023-08-22 | Full range of Y Detector Raw during Y Scan sweep | 14
| PSE demo lab | 2023-08-22 | Full range of Y Detector Raw during Y Scan sweep | 14
| 주식회사 에이치앤에스 | 2023-08-22 | Full range of Y Detector during Y Scan sweep | 71.1
| PSE demo lab | 2023-08-22 | Full range of Y Detector during Y Scan sweep | 116.5
| 주식회사 에이치앤에스 | 2023-08-22 | Full range of X Detector Raw during X Scan sweep | 14.2
| PSE demo lab | 2023-08-22 | Full range of X Detector Raw during X Scan sweep | 14.1
| 주식회사 에이치앤에스 | 2023-08-22 | Full range of X Detector during X Scan sweep | 71.6
| PSE demo lab | 2023-08-22 | Full range of X Detector during X Scan sweep | 115.2
| PSE demo lab | 2023-08-22 | Approach position at -3μm | 0.0051
| PSE demo lab | 2023-08-22 | Approach position at +3μm | 0.0133
| 주식회사 에이치앤에스 | 2023-08-22 | A-B noise | 4.38
| PSE demo lab | 2023-08-22 | A-B noise | 18.46
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Linearity by Z scanner position | 1.38
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Linearity by Z scanner position | 0.05
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Linearity by 4 different sample height | 1.017
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Linearity by 4 different sample height | 1.012
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Height noise | 0.03
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Height noise | 0.038
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.1501
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.42
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.8995
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.05
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.0557
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.98
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.0022
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Drift (Short Traveling) | 0.23
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Drift (Short Traveling) | 0.22
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Z Detector offset | 58
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Z Detector offset | 8.3
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y2 Detector noise (XY Servo on) | 0.307
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y2 Detector noise (XY Servo on) | 0.346
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 251.15
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 235.467
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y1 Detector noise (XY Servo on) | 0.294
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y1 Detector noise (XY Servo on) | 0.329
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 235.8
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 243.821
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.9
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.899
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y Ringing(Rpv) | 0.126
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y Ringing(Rpv) | 0.11
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y Non-linearity error | 2.23
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y Non-linearity error | 1.82
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y Detector noise (XY Servo on) | 0.214
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y Detector noise (XY Servo on) | 0.241
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y axis, Repeatability @1st position | 0.45
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y axis, Repeatability @1st position | 0.33
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y axis, Repeatability @ 4th position | 0.48
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y axis, Repeatability @ 4th position | 0.27
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y axis, Repeatability @ 3rd position | 0.44
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y axis, Repeatability @ 3rd position | 0.22
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y axis, Repeatability @ 2nd position | 0.42
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y axis, Repeatability @ 2nd position | 0.33
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Y axis OPM | 1.208
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Y axis OPM | 0.67
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | XY Orthogonality | 90
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | XY Orthogonality | 90.1
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X2 Detector noise (XY Servo on) | 0.325
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X2 Detector noise (XY Servo on) | 0.342
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 238.24
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 243.35
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X1 Detector noise (XY Servo on) | 0.317
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X1 Detector noise (XY Servo on) | 0.352
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 250.24
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 242.94
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 160.76
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.272
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X Ringing(Rpv) | 0.175
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X Ringing(Rpv) | 0.098
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X Non-linearity error | 2.33
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X Non-linearity error | 1.57
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X Detector noise (XY Servo on) | 0.23
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X Detector noise (XY Servo on) | 0.25
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X axis, Repeatability @1st position | 0.2
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X axis, Repeatability @1st position | 0.2
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X axis, Repeatability @ 4th position | 0.23
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X axis, Repeatability @ 4th position | 0.16
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X axis, Repeatability @ 3rd position | 0.21
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X axis, Repeatability @ 3rd position | 0.18
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X axis, Repeatability @ 2nd position | 0.2
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X axis, Repeatability @ 2nd position | 0.14
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | X axis OPM | 0.679
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | X axis OPM | 1.012
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | StrokeYMin | -102001
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | StrokeYMin | -102004
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | StrokeYMax | 102522
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | StrokeYMax | 103617
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | StrokeY | 204523
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | StrokeY | 205621
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | StrokeXMin | -102800
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | StrokeXMin | -102732
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | StrokeXMax | 104950
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | StrokeXMax | 103921
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | StrokeX | 207750
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | StrokeX | 206653
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.835
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.515
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | RMS of ZeroScan | 0.032
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | RMS of ZeroScan | 0.034
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6e-05
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8e-05
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | OpticalAxisShiftPerFocusY | -0.001
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | OpticalAxisShiftPerFocusY | -0.0004
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00022
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | OpticalAxisShiftPerFocusX | -0.0018
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | OpticalAxisShiftPerFocusX | 0.0011
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | NCM driving% with AC160TS cantilever | 3.55
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | NCM driving% with AC160TS cantilever | 8.15
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Lateral(C-D) sensitivity | 32
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Lateral(C-D) sensitivity | 20.6
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Lateral(C-D) noise | 0.117
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Lateral(C-D) noise | 0.11
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Laser intensity with AC160 cantilever | 2.823
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Laser intensity with AC160 cantilever | 2.9
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of Z Height during Z Scan sweep | 16.8
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of Z Height during Z Scan sweep | 16.29
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of Z Detector during Z Scan sweep | 12.19
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of Z Detector during Z Scan sweep | 11.79
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of Y2 Detector Raw during Y Scan sweep | 14.02
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of Y2 Detector Raw during Y Scan sweep | 14.32
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of Y2 Detector during Y Scan sweep | 166.173
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of Y2 Detector during Y Scan sweep | 170.3298
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of Y1 Detector Raw during Y Scan sweep | 14.23
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of Y1 Detector during Y Scan sweep | 174.3079
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of Y1 Detector during Y Scan sweep | 180.7926
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of X2 Detector Raw during X Scan sweep | 14.05
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of X2 Detector Raw during X Scan sweep | 14.72
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of X2 Detector during X Scan sweep | 178.6979
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of X2 Detector during X Scan sweep | 189.7594
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of X1 Detector Raw during X Scan sweep | 14.65
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | Full range of X1 Detector during X Scan sweep | 174.8713
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | Full range of X1 Detector during X Scan sweep | 182.4511
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.36
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.22
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | A-B after Auto PSPD Alignment | 0.15
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | A-B after Auto PSPD Alignment | 0.214
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | 940nm Step Height Sample (10 Repeats) | 951
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | 940nm Step Height Sample (10 Repeats) | 949.81
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | 50~200nm Step Height Sample (10 Repeats) | 184.56
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | 50~200nm Step Height Sample (10 Repeats) | 184.18
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | 200~500nm Step Height Sample (10 Repeats) | 461.87
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | 200~500nm Step Height Sample (10 Repeats) | 460.78
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.3
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 7.1
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 6.3
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 6
| SINCERE INTERNATION TRADE CO.,LIMITED #1 | 2023-08-21 | %GRR w/ roughness of non-patterened wafer | 6.7
| SINCERE INTERNATION TRADE CO.,LIMITED #5 | 2023-08-21 | %GRR w/ roughness of non-patterened wafer | 9.3
| Dyson UK | 2023-08-21 | Z Scanner position at -3μm | 0.0017
| Dyson UK | 2023-08-21 | Z Scanner position at +3μm | 0.0108
| Dyson UK | 2023-08-21 | Z Linearity by 3 different sample height | 1.002
| Dyson UK | 2023-08-21 | Z Height noise | 0.026
| Dyson UK | 2023-08-21 | Y Scan Stroke | 55
| Dyson UK | 2023-08-21 | Y Detector noise (XY Servo on) | 0.28
| Dyson UK | 2023-08-21 | Y Detector noise (XY Servo off) | 0.26
| Dyson UK | 2023-08-21 | Y axis OPM | 0.88
| Dyson UK | 2023-08-21 | X Scan Stroke | 55
| Dyson UK | 2023-08-21 | X Detector noise (XY Servo on) | 0.29
| Dyson UK | 2023-08-21 | X Detector noise (XY Servo off) | 0.27
| Dyson UK | 2023-08-21 | X axis OPM | 1.17
| Dyson UK | 2023-08-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Dyson UK | 2023-08-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Dyson UK | 2023-08-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.12
| Dyson UK | 2023-08-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Dyson UK | 2023-08-21 | NCM driving% with AC160TS cantilever | 1.32
| Dyson UK | 2023-08-21 | NCM Amplitude noise | 16.1
| Dyson UK | 2023-08-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.27
| Dyson UK | 2023-08-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.416
| Dyson UK | 2023-08-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.6
| Dyson UK | 2023-08-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103.1
| Dyson UK | 2023-08-21 | Lateral(C-D) sensitivity | 4.7
| Dyson UK | 2023-08-21 | Lateral(C-D) noise | 0.091
| Dyson UK | 2023-08-21 | Full range of Z Height during Z Scan sweep | 15.9
| Dyson UK | 2023-08-21 | Full range of Z Detector during Z Scan sweep | 10.8
| Dyson UK | 2023-08-21 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| Dyson UK | 2023-08-21 | Full range of Y Detector during Y Scan sweep | 72.2
| Dyson UK | 2023-08-21 | Full range of X Detector Raw during X Scan sweep | 13.7
| Dyson UK | 2023-08-21 | Full range of X Detector during X Scan sweep | 75.4
| Dyson UK | 2023-08-21 | A-B noise | 11.79
| GRANDIT | 2023-08-18 | Z Scanner position at -3μm | 0.0046
| GRANDIT | 2023-08-18 | Z Scanner position at +3μm | 0.0116
| GRANDIT | 2023-08-18 | Z Linearity by 3 different sample height | 1.005
| GRANDIT | 2023-08-18 | Z Height noise | 0.023
| GRANDIT | 2023-08-18 | Y Detector noise (XY Servo on) | 0.41
| GRANDIT | 2023-08-18 | Y Detector noise (XY Servo off) | 0.36
| GRANDIT | 2023-08-18 | Y axis OPM | 1.13
| GRANDIT | 2023-08-18 | X Detector noise (XY Servo on) | 0.41
| GRANDIT | 2023-08-18 | X Detector noise (XY Servo off) | 0.36
| GRANDIT | 2023-08-18 | X axis OPM | 1.28
| GRANDIT | 2023-08-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| GRANDIT | 2023-08-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| GRANDIT | 2023-08-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.09
| GRANDIT | 2023-08-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.14
| GRANDIT | 2023-08-18 | NCM driving% with AC160TS cantilever | 0.89
| GRANDIT | 2023-08-18 | NCM Amplitude noise | 14.7
| GRANDIT | 2023-08-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.439
| GRANDIT | 2023-08-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.405
| GRANDIT | 2023-08-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 152.3
| GRANDIT | 2023-08-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 156.4
| GRANDIT | 2023-08-18 | Lateral(C-D) sensitivity | 6.8
| GRANDIT | 2023-08-18 | Lateral(C-D) noise | 0.086
| GRANDIT | 2023-08-18 | Full range of Z Height during Z Scan sweep | 16.3
| GRANDIT | 2023-08-18 | Full range of Z Detector during Z Scan sweep | 12.4
| GRANDIT | 2023-08-18 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| GRANDIT | 2023-08-18 | Full range of Y Detector during Y Scan sweep | 112.3
| GRANDIT | 2023-08-18 | Full range of X Detector Raw during X Scan sweep | 13.9
| GRANDIT | 2023-08-18 | Full range of X Detector during X Scan sweep | 108.9
| GRANDIT | 2023-08-18 | A-B noise | 10.3
| Infineon Villach #7 | 2023-08-08 | Z Linearity by Z scanner position | 1.08
| Infineon Villach #8 | 2023-08-08 | Z Linearity by Z scanner position | 0.62
| Infineon Villach #7 | 2023-08-08 | Z Linearity by 4 different sample height | 0.999
| Infineon Villach #8 | 2023-08-08 | Z Linearity by 4 different sample height | 1.004
| Infineon Villach #7 | 2023-08-08 | Z Height noise | 0.038
| Infineon Villach #8 | 2023-08-08 | Z Height noise | 0.036
| Infineon Villach #7 | 2023-08-08 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.718
| Infineon Villach #8 | 2023-08-08 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.5
| Infineon Villach #7 | 2023-08-08 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.5907
| Infineon Villach #8 | 2023-08-08 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.3
| Infineon Villach #7 | 2023-08-08 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.371
| Infineon Villach #8 | 2023-08-08 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.01
| Infineon Villach #7 | 2023-08-08 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3827
| Infineon Villach #8 | 2023-08-08 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.6
| Infineon Villach #7 | 2023-08-08 | Y2 Detector noise (XY Servo on) | 0.268
| Infineon Villach #8 | 2023-08-08 | Y2 Detector noise (XY Servo on) | 0.26
| Infineon Villach #7 | 2023-08-08 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 238.71
| Infineon Villach #8 | 2023-08-08 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 242.365
| Infineon Villach #7 | 2023-08-08 | Y1 Detector noise (XY Servo on) | 0.349
| Infineon Villach #8 | 2023-08-08 | Y1 Detector noise (XY Servo on) | 0.26
| Infineon Villach #7 | 2023-08-08 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 251.95
| Infineon Villach #8 | 2023-08-08 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 247.992
| Infineon Villach #7 | 2023-08-08 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.15
| Infineon Villach #8 | 2023-08-08 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 182.155
| Infineon Villach #7 | 2023-08-08 | Y Ringing(Rpv) | 0.509
| Infineon Villach #8 | 2023-08-08 | Y Ringing(Rpv) | 0.86
| Infineon Villach #7 | 2023-08-08 | Y Non-linearity error | 1.76
| Infineon Villach #8 | 2023-08-08 | Y Non-linearity error | 2.55
| Infineon Villach #7 | 2023-08-08 | Y Detector noise (XY Servo on) | 0.227
| Infineon Villach #8 | 2023-08-08 | Y Detector noise (XY Servo on) | 0.19
| Infineon Villach #7 | 2023-08-08 | Y axis, Repeatability @1st position | 0.04
| Infineon Villach #8 | 2023-08-08 | Y axis, Repeatability @1st position | 0.04
| Infineon Villach #7 | 2023-08-08 | Y axis, Repeatability @ 4th position | 0.05
| Infineon Villach #8 | 2023-08-08 | Y axis, Repeatability @ 4th position | 0.03
| Infineon Villach #7 | 2023-08-08 | Y axis, Repeatability @ 3rd position | 0.04
| Infineon Villach #8 | 2023-08-08 | Y axis, Repeatability @ 3rd position | 0.03
| Infineon Villach #7 | 2023-08-08 | Y axis, Repeatability @ 2nd position | 0.03
| Infineon Villach #8 | 2023-08-08 | Y axis, Repeatability @ 2nd position | 0.03
| Infineon Villach #7 | 2023-08-08 | Y axis, Accuracy @1st position | 0.2
| Infineon Villach #8 | 2023-08-08 | Y axis, Accuracy @1st position | 0.86
| Infineon Villach #7 | 2023-08-08 | Y axis, Accuracy @ 4th position | 0.24
| Infineon Villach #8 | 2023-08-08 | Y axis, Accuracy @ 4th position | 0.99
| Infineon Villach #7 | 2023-08-08 | Y axis, Accuracy @ 3rd position | 0.18
| Infineon Villach #8 | 2023-08-08 | Y axis, Accuracy @ 3rd position | 1.02
| Infineon Villach #7 | 2023-08-08 | Y axis, Accuracy @ 2nd position | 0.15
| Infineon Villach #8 | 2023-08-08 | Y axis, Accuracy @ 2nd position | 0.92
| Infineon Villach #7 | 2023-08-08 | Y axis OPM | 0.38
| Infineon Villach #8 | 2023-08-08 | Y axis OPM | 0.205
| Infineon Villach #7 | 2023-08-08 | X2 Detector noise (XY Servo on) | 0.265
| Infineon Villach #8 | 2023-08-08 | X2 Detector noise (XY Servo on) | 0.24
| Infineon Villach #7 | 2023-08-08 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 244.53
| Infineon Villach #8 | 2023-08-08 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 251.071
| Infineon Villach #7 | 2023-08-08 | X1 Detector noise (XY Servo on) | 0.243
| Infineon Villach #8 | 2023-08-08 | X1 Detector noise (XY Servo on) | 0.26
| Infineon Villach #7 | 2023-08-08 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 243.5
| Infineon Villach #8 | 2023-08-08 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 256.493
| Infineon Villach #7 | 2023-08-08 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 166.07
| Infineon Villach #8 | 2023-08-08 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 182.292
| Infineon Villach #7 | 2023-08-08 | X Ringing(Rpv) | 0.42
| Infineon Villach #8 | 2023-08-08 | X Ringing(Rpv) | 0.54
| Infineon Villach #7 | 2023-08-08 | X Non-linearity error | 2.1
| Infineon Villach #8 | 2023-08-08 | X Non-linearity error | 2.86
| Infineon Villach #7 | 2023-08-08 | X Detector noise (XY Servo on) | 0.187
| Infineon Villach #8 | 2023-08-08 | X Detector noise (XY Servo on) | 0.18
| Infineon Villach #7 | 2023-08-08 | X axis, Repeatability @1st position | 0.04
| Infineon Villach #8 | 2023-08-08 | X axis, Repeatability @1st position | 0.02
| Infineon Villach #7 | 2023-08-08 | X axis, Repeatability @ 4th position | 0.03
| Infineon Villach #8 | 2023-08-08 | X axis, Repeatability @ 4th position | 0.03
| Infineon Villach #7 | 2023-08-08 | X axis, Repeatability @ 3rd position | 0.03
| Infineon Villach #8 | 2023-08-08 | X axis, Repeatability @ 3rd position | 0.03
| Infineon Villach #7 | 2023-08-08 | X axis, Repeatability @ 2nd position | 0.04
| Infineon Villach #8 | 2023-08-08 | X axis, Repeatability @ 2nd position | 0.03
| Infineon Villach #7 | 2023-08-08 | X axis, Accuracy @1st position | 0.21
| Infineon Villach #8 | 2023-08-08 | X axis, Accuracy @1st position | 0.1
| Infineon Villach #7 | 2023-08-08 | X axis, Accuracy @ 4th position | 0.15
| Infineon Villach #8 | 2023-08-08 | X axis, Accuracy @ 4th position | 0.2
| Infineon Villach #7 | 2023-08-08 | X axis, Accuracy @ 3rd position | 0.22
| Infineon Villach #8 | 2023-08-08 | X axis, Accuracy @ 3rd position | 0.14
| Infineon Villach #7 | 2023-08-08 | X axis, Accuracy @ 2nd position | 0.21
| Infineon Villach #8 | 2023-08-08 | X axis, Accuracy @ 2nd position | 0.19
| Infineon Villach #7 | 2023-08-08 | X axis OPM | 0.493
| Infineon Villach #8 | 2023-08-08 | X axis OPM | 0.287
| Infineon Villach #7 | 2023-08-08 | StrokeYMin | -109373
| Infineon Villach #8 | 2023-08-08 | StrokeYMin | -109072
| Infineon Villach #7 | 2023-08-08 | StrokeYMax | 103655.2
| Infineon Villach #8 | 2023-08-08 | StrokeYMax | 105427
| Infineon Villach #7 | 2023-08-08 | StrokeY | 213027.8
| Infineon Villach #8 | 2023-08-08 | StrokeY | 214499
| Infineon Villach #7 | 2023-08-08 | StrokeXMin | -141789
| Infineon Villach #8 | 2023-08-08 | StrokeXMin | -140216
| Infineon Villach #7 | 2023-08-08 | StrokeXMax | 137664.9
| Infineon Villach #8 | 2023-08-08 | StrokeXMax | 139352
| Infineon Villach #7 | 2023-08-08 | StrokeX | 279454.2
| Infineon Villach #8 | 2023-08-08 | StrokeX | 279568
| Infineon Villach #7 | 2023-08-08 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.975
| Infineon Villach #8 | 2023-08-08 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.74
| Infineon Villach #7 | 2023-08-08 | RMS of ZeroScan | 0.035
| Infineon Villach #8 | 2023-08-08 | RMS of ZeroScan | 0.033
| Infineon Villach #7 | 2023-08-08 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Infineon Villach #8 | 2023-08-08 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| Infineon Villach #7 | 2023-08-08 | OpticalAxisShiftPerFocusY | -0.001
| Infineon Villach #8 | 2023-08-08 | OpticalAxisShiftPerFocusY | -0.0018
| Infineon Villach #7 | 2023-08-08 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| Infineon Villach #8 | 2023-08-08 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00021
| Infineon Villach #7 | 2023-08-08 | OpticalAxisShiftPerFocusX | -0.0015
| Infineon Villach #8 | 2023-08-08 | OpticalAxisShiftPerFocusX | -0.005
| Infineon Villach #7 | 2023-08-08 | NCM driving% with AC160TS cantilever | 5.49
| Infineon Villach #8 | 2023-08-08 | NCM driving% with AC160TS cantilever | 6.91
| Infineon Villach #7 | 2023-08-08 | Min. Vibration Transmissbility (~100Hz) | 0.002
| Infineon Villach #8 | 2023-08-08 | Min. Vibration Transmissbility (~100Hz) | 0.001
| Infineon Villach #7 | 2023-08-08 | Lateral(C-D) sensitivity | 7.4
| Infineon Villach #8 | 2023-08-08 | Lateral(C-D) sensitivity | 10
| Infineon Villach #7 | 2023-08-08 | Lateral(C-D) noise | 0.109
| Infineon Villach #8 | 2023-08-08 | Lateral(C-D) noise | 0.219
| Infineon Villach #7 | 2023-08-08 | Laser intensity with AC160 cantilever | 3.1
| Infineon Villach #8 | 2023-08-08 | Laser intensity with AC160 cantilever | 2.94
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from -Y to +Y (Y scan) | 84.23
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from -Y to +Y (Y scan) | 45.41
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from -Y to +Y (X scan) | 15.38
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from -Y to +Y (X scan) | 28.56
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from -X to +X (Y scan) | 25.63
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from -X to +X (Y scan) | 12.45
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from -X to +X (X scan) | 112.06
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from -X to +X (X scan) | 22.7
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from +Y to -Y (Y scan) | 152.34
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from +Y to -Y (Y scan) | 68.84
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from +Y to -Y (X scan) | 10.25
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from +Y to -Y (X scan) | 18.31
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from +X to -X (Y scan) | 76.17
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from +X to -X (Y scan) | 54.93
| Infineon Villach #7 | 2023-08-08 | Image offset after moving from +X to -X (X scan) | 109.13
| Infineon Villach #8 | 2023-08-08 | Image offset after moving from +X to -X (X scan) | 65.91
| Infineon Villach #7 | 2023-08-08 | Full range of Z Height during Z Scan sweep | 16.39
| Infineon Villach #8 | 2023-08-08 | Full range of Z Height during Z Scan sweep | 16.8
| Infineon Villach #7 | 2023-08-08 | Full range of Z Detector during Z Scan sweep | 12.4
| Infineon Villach #8 | 2023-08-08 | Full range of Z Detector during Z Scan sweep | 12.44
| Infineon Villach #7 | 2023-08-08 | Full range of Y2 Detector Raw during Y Scan sweep | 14.18
| Infineon Villach #8 | 2023-08-08 | Full range of Y2 Detector Raw during Y Scan sweep | 12.97
| Infineon Villach #7 | 2023-08-08 | Full range of Y2 Detector during Y Scan sweep | 188.2289
| Infineon Villach #8 | 2023-08-08 | Full range of Y2 Detector during Y Scan sweep | 187.1382
| Infineon Villach #7 | 2023-08-08 | Full range of Y1 Detector Raw during Y Scan sweep | 14.14
| Infineon Villach #8 | 2023-08-08 | Full range of Y1 Detector Raw during Y Scan sweep | 13.11
| Infineon Villach #7 | 2023-08-08 | Full range of Y1 Detector during Y Scan sweep | 193.9383
| Infineon Villach #8 | 2023-08-08 | Full range of Y1 Detector during Y Scan sweep | 188.87
| Infineon Villach #7 | 2023-08-08 | Full range of X2 Detector Raw during X Scan sweep | 14.63
| Infineon Villach #8 | 2023-08-08 | Full range of X2 Detector Raw during X Scan sweep | 14.13
| Infineon Villach #7 | 2023-08-08 | Full range of X2 Detector during X Scan sweep | 183.9668
| Infineon Villach #8 | 2023-08-08 | Full range of X2 Detector during X Scan sweep | 181.0211
| Infineon Villach #7 | 2023-08-08 | Full range of X1 Detector Raw during X Scan sweep | 14.08
| Infineon Villach #8 | 2023-08-08 | Full range of X1 Detector Raw during X Scan sweep | 14.01
| Infineon Villach #7 | 2023-08-08 | Full range of X1 Detector during X Scan sweep | 187.3771
| Infineon Villach #8 | 2023-08-08 | Full range of X1 Detector during X Scan sweep | 194.902
| Infineon Villach #7 | 2023-08-08 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| Infineon Villach #8 | 2023-08-08 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Infineon Villach #7 | 2023-08-08 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.4
| Infineon Villach #8 | 2023-08-08 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Infineon Villach #7 | 2023-08-08 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| Infineon Villach #8 | 2023-08-08 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| Infineon Villach #7 | 2023-08-08 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.08
| Infineon Villach #8 | 2023-08-08 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Infineon Villach #7 | 2023-08-08 | A－B difference during 3 times run test | 0.5
| Infineon Villach #8 | 2023-08-08 | A－B difference during 3 times run test | 0.39
| Infineon Villach #7 | 2023-08-08 | A＋B difference during 3 times run test | 0.19
| Infineon Villach #8 | 2023-08-08 | A＋B difference during 3 times run test | 0.02
| Infineon Villach #7 | 2023-08-08 | 50~200nm Step Height Sample (10 Repeats) | 183.17
| Infineon Villach #8 | 2023-08-08 | 50~200nm Step Height Sample (10 Repeats) | 177.65
| Infineon Villach #7 | 2023-08-08 | 200~500nm Step Height Sample (10 Repeats) | 460.501
| Infineon Villach #8 | 2023-08-08 | 200~500nm Step Height Sample (10 Repeats) | 462.69
| Infineon Villach #7 | 2023-08-08 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.8
| Infineon Villach #8 | 2023-08-08 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.1
| Infineon Villach #7 | 2023-08-08 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.1
| Infineon Villach #8 | 2023-08-08 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.6
| Infineon Villach #7 | 2023-08-08 | %GRR w/ roughness of non-patterened wafer | 4.2
| Infineon Villach #8 | 2023-08-08 | %GRR w/ roughness of non-patterened wafer | 6.2
| George Mason University | 2023-08-07 | Z Scanner position at -3μm | 0.0089
| Northwest University | 2023-08-07 | Z Scanner position at -3μm | 0.0043
| ICMAB-CSIC | 2023-08-07 | Z Scanner position at -3μm | 0.0049
| George Mason University | 2023-08-07 | Z Scanner position at +3μm | 0.0018
| Northwest University | 2023-08-07 | Z Scanner position at +3μm | 0.0003
| ICMAB-CSIC | 2023-08-07 | Z Scanner position at +3μm | 0.0101
| George Mason University | 2023-08-07 | Z Linearity by 3 different sample height | 1.006
| Northwest University | 2023-08-07 | Z Linearity by 3 different sample height | 0.999
| ICMAB-CSIC | 2023-08-07 | Z Linearity by 3 different sample height | 1
| George Mason University | 2023-08-07 | Z Height noise | 0.031
| Northwest University | 2023-08-07 | Z Height noise | 0.019
| George Mason University | 2023-08-07 | Y Scan Stroke | 110
| George Mason University | 2023-08-07 | Y Detector noise (XY Servo on) | 0.32
| Northwest University | 2023-08-07 | Y Detector noise (XY Servo on) | 0.37
| ICMAB-CSIC | 2023-08-07 | Y Detector noise (XY Servo on) | 0.29
| George Mason University | 2023-08-07 | Y Detector noise (XY Servo off) | 0.31
| Northwest University | 2023-08-07 | Y Detector noise (XY Servo off) | 0.35
| ICMAB-CSIC | 2023-08-07 | Y Detector noise (XY Servo off) | 0.31
| George Mason University | 2023-08-07 | Y axis OPM | 0.94
| Northwest University | 2023-08-07 | Y axis OPM | 0.25
| ICMAB-CSIC | 2023-08-07 | Y axis OPM | 1.31
| George Mason University | 2023-08-07 | X Scan Stroke | 110
| George Mason University | 2023-08-07 | X Detector noise (XY Servo on) | 0.31
| Northwest University | 2023-08-07 | X Detector noise (XY Servo on) | 0.36
| ICMAB-CSIC | 2023-08-07 | X Detector noise (XY Servo on) | 0.29
| George Mason University | 2023-08-07 | X Detector noise (XY Servo off) | 0.3
| Northwest University | 2023-08-07 | X Detector noise (XY Servo off) | 0.35
| ICMAB-CSIC | 2023-08-07 | X Detector noise (XY Servo off) | 0.32
| George Mason University | 2023-08-07 | X axis OPM | 1.01
| Northwest University | 2023-08-07 | X axis OPM | 0.76
| ICMAB-CSIC | 2023-08-07 | X axis OPM | 0.76
| George Mason University | 2023-08-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.41
| Northwest University | 2023-08-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| ICMAB-CSIC | 2023-08-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| George Mason University | 2023-08-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| Northwest University | 2023-08-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| ICMAB-CSIC | 2023-08-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| George Mason University | 2023-08-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| Northwest University | 2023-08-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| ICMAB-CSIC | 2023-08-07 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| George Mason University | 2023-08-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| Northwest University | 2023-08-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| ICMAB-CSIC | 2023-08-07 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| George Mason University | 2023-08-07 | NCM driving% with AC160TS cantilever | 0.15
| Northwest University | 2023-08-07 | NCM driving% with AC160TS cantilever | 1.27
| ICMAB-CSIC | 2023-08-07 | NCM driving% with AC160TS cantilever | 0.24
| Northwest University | 2023-08-07 | NCM Amplitude noise | 13.9
| ICMAB-CSIC | 2023-08-07 | NCM Amplitude noise | 16.56
| George Mason University | 2023-08-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.093
| Northwest University | 2023-08-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.439
| ICMAB-CSIC | 2023-08-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.233
| George Mason University | 2023-08-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.657
| Northwest University | 2023-08-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.993
| ICMAB-CSIC | 2023-08-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.075
| George Mason University | 2023-08-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 260.3
| Northwest University | 2023-08-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 161.2
| ICMAB-CSIC | 2023-08-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.2
| George Mason University | 2023-08-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 255.8
| Northwest University | 2023-08-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.3
| ICMAB-CSIC | 2023-08-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162
| George Mason University | 2023-08-07 | Lateral(C-D) sensitivity | 12.4
| Northwest University | 2023-08-07 | Lateral(C-D) sensitivity | 11.3
| ICMAB-CSIC | 2023-08-07 | Lateral(C-D) sensitivity | 14.3
| George Mason University | 2023-08-07 | Lateral(C-D) noise | 0.093
| Northwest University | 2023-08-07 | Lateral(C-D) noise | 0.095
| ICMAB-CSIC | 2023-08-07 | Lateral(C-D) noise | 0.086
| George Mason University | 2023-08-07 | Full range of Z Height during Z Scan sweep | 15.9
| Northwest University | 2023-08-07 | Full range of Z Height during Z Scan sweep | 16.3
| ICMAB-CSIC | 2023-08-07 | Full range of Z Height during Z Scan sweep | 15.8
| George Mason University | 2023-08-07 | Full range of Z Detector during Z Scan sweep | 11.1
| Northwest University | 2023-08-07 | Full range of Z Detector during Z Scan sweep | 12.1
| ICMAB-CSIC | 2023-08-07 | Full range of Z Detector during Z Scan sweep | 11.3
| George Mason University | 2023-08-07 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Northwest University | 2023-08-07 | Full range of Y Detector Raw during Y Scan sweep | 14
| ICMAB-CSIC | 2023-08-07 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| George Mason University | 2023-08-07 | Full range of Y Detector during Y Scan sweep | 187.5
| Northwest University | 2023-08-07 | Full range of Y Detector during Y Scan sweep | 113
| ICMAB-CSIC | 2023-08-07 | Full range of Y Detector during Y Scan sweep | 113.6
| George Mason University | 2023-08-07 | Full range of X Detector Raw during X Scan sweep | 14.5
| Northwest University | 2023-08-07 | Full range of X Detector Raw during X Scan sweep | 14.1
| ICMAB-CSIC | 2023-08-07 | Full range of X Detector Raw during X Scan sweep | 15
| George Mason University | 2023-08-07 | Full range of X Detector during X Scan sweep | 184.9
| Northwest University | 2023-08-07 | Full range of X Detector during X Scan sweep | 114.3
| ICMAB-CSIC | 2023-08-07 | Full range of X Detector during X Scan sweep | 118.5
| George Mason University | 2023-08-07 | A-B noise | 8.48
| Northwest University | 2023-08-07 | A-B noise | 13.91
| ICMAB-CSIC | 2023-08-07 | A-B noise | 8.78
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Linearity by Z scanner position | 0.616
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Linearity by 4 different sample height | 1.0156
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Height noise | 0.04
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.7
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.7
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.7
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y2 Detector noise (XY Servo on) | 0.43
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 225.9
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y1 Detector noise (XY Servo on) | 0.32
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 228
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.9
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y Ringing(Rpv) | 0.11
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y Non-linearity error | 1.9
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y Detector noise (XY Servo on) | 0.22
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y axis, Repeatability @1st position | 0.3
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y axis, Repeatability @ 4th position | 0.44
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y axis, Repeatability @ 3rd position | 0.39
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y axis, Repeatability @ 2nd position | 0.46
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Y axis OPM | 0.97
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X2 Detector noise (XY Servo on) | 0.38
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 239.6
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X1 Detector noise (XY Servo on) | 0.44
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 228.1
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.3
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X Ringing(Rpv) | 0.16
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X Non-linearity error | 1.7
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X Detector noise (XY Servo on) | 0.32
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X axis, Repeatability @1st position | 0.29
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X axis, Repeatability @ 4th position | 0.18
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X axis, Repeatability @ 3rd position | 0.22
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X axis, Repeatability @ 2nd position | 0.26
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | X axis OPM | 1.41
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | StrokeYMin | -102001
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | StrokeYMax | 102522
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | StrokeY | 204523
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | StrokeXMin | -102800
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | StrokeXMax | 104950
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | StrokeX | 207750
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.665
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | RMS of ZeroScan | 0.03
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | NCM driving% with AC160TS cantilever | 13.4
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Lateral(C-D) sensitivity | 16.7
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Lateral(C-D) noise | 0.173
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Laser intensity with AC160 cantilever | 2.98
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of Z Height during Z Scan sweep | 16.6
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of Z Detector during Z Scan sweep | 12.02
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of Y2 Detector during Y Scan sweep | 174.8839
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of Y1 Detector Raw during Y Scan sweep | 14.27
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of Y1 Detector during Y Scan sweep | 185.0495
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of X2 Detector Raw during X Scan sweep | 13.89
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of X2 Detector during X Scan sweep | 174.103
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of X1 Detector Raw during X Scan sweep | 14.22
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | Full range of X1 Detector during X Scan sweep | 179.701
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | 940nm Step Height Sample (10 Repeats) | 953.46
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | 50~200nm Step Height Sample (10 Repeats) | 184.38
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | 200~500nm Step Height Sample (10 Repeats) | 460.2
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.7
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.3
| SINCERE INTERNATION TRADE CO.,LIMITED #4 | 2023-08-04 | %GRR w/ roughness of non-patterened wafer | 5.9
| 성균관대학교 신소재공학부 | 2023-08-04 | Z Linearity by 3 different sample height | 1.009
| 성균관대학교 신소재공학부 | 2023-08-04 | Z Height noise | 0.021
| 성균관대학교 신소재공학부 | 2023-08-04 | Y Detector noise (XY Servo on) | 0.28
| 성균관대학교 신소재공학부 | 2023-08-04 | Y Detector noise (XY Servo off) | 0.27
| 성균관대학교 신소재공학부 | 2023-08-04 | Y axis OPM | 0.72
| 성균관대학교 신소재공학부 | 2023-08-04 | X Detector noise (XY Servo on) | 0.29
| 성균관대학교 신소재공학부 | 2023-08-04 | X Detector noise (XY Servo off) | 0.28
| 성균관대학교 신소재공학부 | 2023-08-04 | X axis OPM | 0.93
| 성균관대학교 신소재공학부 | 2023-08-04 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| 성균관대학교 신소재공학부 | 2023-08-04 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 성균관대학교 신소재공학부 | 2023-08-04 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.11
| 성균관대학교 신소재공학부 | 2023-08-04 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 성균관대학교 신소재공학부 | 2023-08-04 | NCM driving% with AC160TS cantilever | 0.978
| 성균관대학교 신소재공학부 | 2023-08-04 | NCM Amplitude noise | 7.83
| 성균관대학교 신소재공학부 | 2023-08-04 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.243
| 성균관대학교 신소재공학부 | 2023-08-04 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.081
| 성균관대학교 신소재공학부 | 2023-08-04 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.9
| 성균관대학교 신소재공학부 | 2023-08-04 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 153.6
| 성균관대학교 신소재공학부 | 2023-08-04 | Full range of Z Height during Z Scan sweep | 16.1
| 성균관대학교 신소재공학부 | 2023-08-04 | Lateral(C-D) sensitivity | 5.9
| 성균관대학교 신소재공학부 | 2023-08-04 | Lateral(C-D) noise | 0.145
| 성균관대학교 신소재공학부 | 2023-08-04 | Full range of Z Detector during Z Scan sweep | 13.4
| 성균관대학교 신소재공학부 | 2023-08-04 | Full range of Y Detector Raw during Y Scan sweep | 15
| 성균관대학교 신소재공학부 | 2023-08-04 | Full range of Y Detector during Y Scan sweep | 115.6
| 성균관대학교 신소재공학부 | 2023-08-04 | Full range of X Detector Raw during X Scan sweep | 14.9
| 성균관대학교 신소재공학부 | 2023-08-04 | Full range of X Detector during X Scan sweep | 114.5
| 성균관대학교 신소재공학부 | 2023-08-04 | Approach position at -3μm | 0.0122
| 성균관대학교 신소재공학부 | 2023-08-04 | Approach position at +3μm | 0.0159
| 성균관대학교 신소재공학부 | 2023-08-04 | A-B noise | 14.95
| ICFO | 2023-08-01 | Z Linearity by 3 different sample height | 1.001
| ICFO | 2023-08-01 | Z Height noise | 0.022
| ICFO | 2023-08-01 | Y Detector noise (XY Servo on) | 0.28
| ICFO | 2023-08-01 | Y Detector noise (XY Servo off) | 0.27
| ICFO | 2023-08-01 | Y axis OPM | 0.63
| ICFO | 2023-08-01 | X Detector noise (XY Servo on) | 0.29
| ICFO | 2023-08-01 | X Detector noise (XY Servo off) | 0.28
| ICFO | 2023-08-01 | X axis OPM | 1.06
| ICFO | 2023-08-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| ICFO | 2023-08-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| ICFO | 2023-08-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| ICFO | 2023-08-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| ICFO | 2023-08-01 | NCM driving% with AC160TS cantilever | 3.01
| ICFO | 2023-08-01 | NCM Amplitude noise | 8
| ICFO | 2023-08-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.526
| ICFO | 2023-08-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.631
| ICFO | 2023-08-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 149
| ICFO | 2023-08-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 155.3
| ICFO | 2023-08-01 | Full range of Z Height during Z Scan sweep | 16.7
| ICFO | 2023-08-01 | Lateral(C-D) sensitivity | 5.7
| ICFO | 2023-08-01 | Lateral(C-D) noise | 0.086
| ICFO | 2023-08-01 | Full range of Z Detector during Z Scan sweep | 14.7
| ICFO | 2023-08-01 | Full range of Y Detector Raw during Y Scan sweep | 15.3
| ICFO | 2023-08-01 | Full range of Y Detector during Y Scan sweep | 114
| ICFO | 2023-08-01 | Full range of X Detector Raw during X Scan sweep | 14.6
| ICFO | 2023-08-01 | Full range of X Detector during X Scan sweep | 113.4
| ICFO | 2023-08-01 | Approach position at -3μm | 0.003
| ICFO | 2023-08-01 | Approach position at +3μm | 0.0035
| ICFO | 2023-08-01 | A-B noise | 8.48
| School of Physical and Mathematical Sciences | 2023-07-31 | Z Scanner position at -3μm | 0.0069
| School of Physical and Mathematical Sciences | 2023-07-31 | Z Scanner position at +3μm | 0.0026
| School of Physical and Mathematical Sciences | 2023-07-31 | Z Linearity by 3 different sample height | 0.991
| School of Physical and Mathematical Sciences | 2023-07-31 | Z Height noise | 0.021
| School of Physical and Mathematical Sciences | 2023-07-31 | Y Scan Stroke | 55
| School of Physical and Mathematical Sciences | 2023-07-31 | Y Detector noise (XY Servo on) | 0.28
| School of Physical and Mathematical Sciences | 2023-07-31 | Y Detector noise (XY Servo off) | 0.27
| School of Physical and Mathematical Sciences | 2023-07-31 | Y axis OPM | 0.9
| School of Physical and Mathematical Sciences | 2023-07-31 | X Scan Stroke | 55
| School of Physical and Mathematical Sciences | 2023-07-31 | X Detector noise (XY Servo on) | 0.29
| School of Physical and Mathematical Sciences | 2023-07-31 | X Detector noise (XY Servo off) | 0.27
| School of Physical and Mathematical Sciences | 2023-07-31 | X axis OPM | 1.39
| School of Physical and Mathematical Sciences | 2023-07-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| School of Physical and Mathematical Sciences | 2023-07-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| School of Physical and Mathematical Sciences | 2023-07-31 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| School of Physical and Mathematical Sciences | 2023-07-31 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| School of Physical and Mathematical Sciences | 2023-07-31 | NCM driving% with AC160TS cantilever | 2.12
| School of Physical and Mathematical Sciences | 2023-07-31 | NCM Amplitude noise | 14.23
| School of Physical and Mathematical Sciences | 2023-07-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.139
| School of Physical and Mathematical Sciences | 2023-07-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.493
| School of Physical and Mathematical Sciences | 2023-07-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.6
| School of Physical and Mathematical Sciences | 2023-07-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103
| School of Physical and Mathematical Sciences | 2023-07-31 | Lateral(C-D) sensitivity | 12.6
| School of Physical and Mathematical Sciences | 2023-07-31 | Lateral(C-D) noise | 0.063
| School of Physical and Mathematical Sciences | 2023-07-31 | Full range of Z Height during Z Scan sweep | 15.6
| School of Physical and Mathematical Sciences | 2023-07-31 | Full range of Z Detector during Z Scan sweep | 10.6
| School of Physical and Mathematical Sciences | 2023-07-31 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| School of Physical and Mathematical Sciences | 2023-07-31 | Full range of Y Detector during Y Scan sweep | 75.9
| School of Physical and Mathematical Sciences | 2023-07-31 | Full range of X Detector Raw during X Scan sweep | 14.2
| School of Physical and Mathematical Sciences | 2023-07-31 | Full range of X Detector during X Scan sweep | 74.3
| School of Physical and Mathematical Sciences | 2023-07-31 | A-B noise | 10.25
| Shanghai GTX Semiconductor | 2023-07-28 | Z Linearity by Z scanner position | 0.575
| Shanghai GTX Semiconductor | 2023-07-28 | Z Linearity by Z scanner position | 0.56782
| Shanghai GTX Semiconductor | 2023-07-28 | Z Linearity by 5 different sample height | 1.003
| Shanghai GTX Semiconductor_1 | 2023-07-28 | Z Linearity by 5 different sample height | 1.0047
| Shanghai GTX Semiconductor | 2023-07-28 | Z Height noise | 0.04
| Shanghai GTX Semiconductor | 2023-07-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7182
| Shanghai GTX Semiconductor | 2023-07-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.4613
| Shanghai GTX Semiconductor | 2023-07-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.26
| Shanghai GTX Semiconductor | 2023-07-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2446
| Shanghai GTX Semiconductor | 2023-07-28 | Y2 Detector noise (XY Servo on) | 0.505
| Shanghai GTX Semiconductor | 2023-07-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.48
| Shanghai GTX Semiconductor | 2023-07-28 | Y1 Detector noise (XY Servo on) | 0.485
| Shanghai GTX Semiconductor | 2023-07-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 269.83
| Shanghai GTX Semiconductor | 2023-07-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 190.25
| Shanghai GTX Semiconductor | 2023-07-28 | Y Ringing(Rpv) | 0.292
| Shanghai GTX Semiconductor | 2023-07-28 | Y Non-linearity error | 2.49
| Shanghai GTX Semiconductor | 2023-07-28 | Y Detector noise (XY Servo on) | 0.451
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Repeatability @1st position | 0.07
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Repeatability @ 4th position | 0.07
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Repeatability @ 3rd position | 0.08
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Repeatability @ 2nd position | 0.1
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Accuracy @1st position | 1.13
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Accuracy @ 4th position | 1.09
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Accuracy @ 3rd position | 1.27
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis, Accuracy @ 2nd position | 1.77
| Shanghai GTX Semiconductor | 2023-07-28 | Y axis OPM | 0.7921
| Shanghai GTX Semiconductor | 2023-07-28 | X2 Detector noise (XY Servo on) | 0.405
| Shanghai GTX Semiconductor | 2023-07-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 283.4
| Shanghai GTX Semiconductor | 2023-07-28 | X1 Detector noise (XY Servo on) | 0.443
| Shanghai GTX Semiconductor | 2023-07-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 271.21
| Shanghai GTX Semiconductor | 2023-07-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 190.36
| Shanghai GTX Semiconductor | 2023-07-28 | X Ringing(Rpv) | 0.508
| Shanghai GTX Semiconductor | 2023-07-28 | X Non-linearity error | 1.8
| Shanghai GTX Semiconductor | 2023-07-28 | X Detector noise (XY Servo on) | 0.369
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Repeatability @1st position | 0.07
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Repeatability @ 4th position | 0.07
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Repeatability @ 3rd position | 0.06
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Repeatability @ 2nd position | 0.06
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Accuracy @1st position | 0.82
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Accuracy @ 4th position | 1.23
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Accuracy @ 3rd position | 0.6
| Shanghai GTX Semiconductor | 2023-07-28 | X axis, Accuracy @ 2nd position | 0.43
| Shanghai GTX Semiconductor | 2023-07-28 | X axis OPM | 0.527
| Shanghai GTX Semiconductor | 2023-07-28 | Under 50nm Step Height Sample (10 Repeats) | 18.068
| Shanghai GTX Semiconductor | 2023-07-28 | Under 50nm Step Height Sample (10 Repeats) | 18.033
| Shanghai GTX Semiconductor | 2023-07-28 | Torque Profile (50mm, 20um/s) | 3.416
| Shanghai GTX Semiconductor | 2023-07-28 | Torque Profile (50mm, 200um/s) | 3.193
| Shanghai GTX Semiconductor | 2023-07-28 | StrokeYMin | -153692
| Shanghai GTX Semiconductor | 2023-07-28 | StrokeYMax | 151801.4
| Shanghai GTX Semiconductor | 2023-07-28 | StrokeY | 305493.7
| Shanghai GTX Semiconductor | 2023-07-28 | StrokeXMin | -204060
| Shanghai GTX Semiconductor | 2023-07-28 | StrokeXMax | 203809.7
| Shanghai GTX Semiconductor | 2023-07-28 | StrokeX | 407869.2
| Shanghai GTX Semiconductor | 2023-07-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.639
| Shanghai GTX Semiconductor | 2023-07-28 | Roll motion to X+ full scan | 6.3
| Shanghai GTX Semiconductor | 2023-07-28 | RMS of ZeroScan | 0.061
| Shanghai GTX Semiconductor | 2023-07-28 | Profiler noise @200um | 0.157
| Shanghai GTX Semiconductor | 2023-07-28 | PosMin | -25.5
| Shanghai GTX Semiconductor | 2023-07-28 | PosMax | 25.4
| Shanghai GTX Semiconductor | 2023-07-28 | Position repeatability(10repeats, Position Based) | 0.8
| Shanghai GTX Semiconductor | 2023-07-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 8e-05
| Shanghai GTX Semiconductor | 2023-07-28 | OpticalAxisShiftPerFocusY | -0.0068
| Shanghai GTX Semiconductor | 2023-07-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00044
| Shanghai GTX Semiconductor | 2023-07-28 | OpticalAxisShiftPerFocusX | -0.0003
| Shanghai GTX Semiconductor | 2023-07-28 | OPM Repeatability @5mm | 3.219
| Shanghai GTX Semiconductor | 2023-07-28 | OPM Repeatability @25mm | 10.855
| Shanghai GTX Semiconductor | 2023-07-28 | OPM Repeatability @1mm | 1.33
| Shanghai GTX Semiconductor | 2023-07-28 | OPM Repeatability @10mm | 1.568
| Shanghai GTX Semiconductor | 2023-07-28 | NCM driving% with AC160TS cantilever | 35.89
| Shanghai GTX Semiconductor | 2023-07-28 | NCM driving% with AC160TS cantilever | 9.69
| Shanghai GTX Semiconductor | 2023-07-28 | Min. Vibration Transmissbility (~100Hz) | 0.018
| Shanghai GTX Semiconductor | 2023-07-28 | MaxLength | 51
| Shanghai GTX Semiconductor | 2023-07-28 | Max OPM @5mm | 24.465
| Shanghai GTX Semiconductor | 2023-07-28 | Max OPM @25mm | 87.431
| Shanghai GTX Semiconductor | 2023-07-28 | Max OPM @1mm | 7.519
| Shanghai GTX Semiconductor | 2023-07-28 | Max OPM @10mm | 59.318
| Shanghai GTX Semiconductor | 2023-07-28 | Lateral(C-D) sensitivity | 8.6
| Shanghai GTX Semiconductor | 2023-07-28 | Lateral(C-D) noise | 0.238
| Shanghai GTX Semiconductor | 2023-07-28 | Laser intensity with AC160 cantilever | 3.009
| Shanghai GTX Semiconductor | 2023-07-28 | Laser intensity with AC160 cantilever | 3.054
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from -Y to +Y (Y scan) | 41.75
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from -Y to +Y (X scan) | 13.92
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from -X to +X (Y scan) | 16.11
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from -X to +X (X scan) | 84.96
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from +Y to -Y (Y scan) | 35.89
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from +Y to -Y (X scan) | 18.31
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from +X to -X (Y scan) | 32.23
| Shanghai GTX Semiconductor | 2023-07-28 | Image offset after moving from +X to -X (X scan) | 51.27
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Z Height during Z Scan sweep | 17.2
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Z Height during Z Scan sweep | 40.4
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Z Detector during Z Scan sweep | 12.8
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Z Detector during Z Scan sweep | 13.4
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.21
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Y2 Detector during Y Scan sweep | 204.7587
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.75
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of Y1 Detector during Y Scan sweep | 215.2928
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of X2 Detector during X Scan sweep | 215.2701
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of X1 Detector Raw during X Scan sweep | 14.08
| Shanghai GTX Semiconductor | 2023-07-28 | Full range of X1 Detector during X Scan sweep | 208.9564
| Shanghai GTX Semiconductor | 2023-07-28 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| Shanghai GTX Semiconductor | 2023-07-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| Shanghai GTX Semiconductor | 2023-07-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.79
| Shanghai GTX Semiconductor | 2023-07-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.4
| Shanghai GTX Semiconductor | 2023-07-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Shanghai GTX Semiconductor | 2023-07-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Shanghai GTX Semiconductor | 2023-07-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.21
| Shanghai GTX Semiconductor | 2023-07-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Shanghai GTX Semiconductor | 2023-07-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Shanghai GTX Semiconductor | 2023-07-28 | A－B difference during 3 times run test | 0.55
| Shanghai GTX Semiconductor | 2023-07-28 | A＋B difference during 3 times run test | 0.08
| Shanghai GTX Semiconductor | 2023-07-28 | 50~200nm Step Height Sample (10 Repeats) | 183.538
| Shanghai GTX Semiconductor | 2023-07-28 | 50~200nm Step Height Sample (10 Repeats) | 183.499
| Shanghai GTX Semiconductor | 2023-07-28 | 200~500nm Step Height Sample (10 Repeats) | 458.424
| Shanghai GTX Semiconductor | 2023-07-28 | 200~500nm Step Height Sample (10 Repeats) | 459.189
| Shanghai GTX Semiconductor | 2023-07-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.3
| Shanghai GTX Semiconductor | 2023-07-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.3
| Shanghai GTX Semiconductor | 2023-07-28 | %GRR w/ roughness of non-patterened wafer | 5.7
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Z Scanner position at -3μm | 0.0022
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Z Scanner position at -3μm | 0.0151
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Z Scanner position at +3μm | 0.0009
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Z Scanner position at +3μm | 0.002
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Z Linearity by 3 different sample height | 1.01
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Z Linearity by 3 different sample height | 1.013
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Z Height noise | 0.017
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Z Height noise | 0.03
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y2 Detector noise (XY Servo on) | 0.28
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y2 Detector noise (XY Servo on) | 0.228
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y2 Detector noise (XY Servo off) | 0.18
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y2 Detector noise (XY Servo off) | 0.225
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y1 Detector noise (XY Servo on) | 0.28
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y1 Detector noise (XY Servo on) | 0.228
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y1 Detector noise (XY Servo off) | 0.17
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y1 Detector noise (XY Servo off) | 0.233
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y Scan Stroke | 173.7
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y Scan Stroke | 174.234
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y Detector noise (XY Servo on) | 0.22
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y Detector noise (XY Servo on) | 0.166
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y Detector noise (XY Servo off) | 0.13
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y Detector noise (XY Servo off) | 0.164
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Y axis OPM | 1.28
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Y axis OPM | 1.3434
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X2 Detector noise (XY Servo on) | 0.29
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X2 Detector noise (XY Servo on) | 0.235
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X2 Detector noise (XY Servo off) | 0.18
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X2 Detector noise (XY Servo off) | 0.238
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X1 Detector noise (XY Servo on) | 0.31
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X1 Detector noise (XY Servo on) | 0.231
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X1 Detector noise (XY Servo off) | 0.19
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X1 Detector noise (XY Servo off) | 0.234
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X Scan Stroke | 170.1
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X Scan Stroke | 173.448
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | -X Direction software Limit | -102403
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | -X Direction software Limit | -103009
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | -X Direction software Limit | -102393
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | -X Direction software Limit | -102826
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X Detector noise (XY Servo on) | 0.24
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X Detector noise (XY Servo on) | 0.173
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X Detector noise (XY Servo off) | 0.13
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X Detector noise (XY Servo off) | 0.174
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | X axis OPM | 0.77
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | X axis OPM | 0.9834
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.6
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Top) | 0.39
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Top) | 0.3
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Right) | 0.33
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Right) | 0.31
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Left) | 0.34
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Left) | 0.29
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Height of Wafer ZeroScan (Center) | 0.33
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of Wafer ZeroScan (Center) | 0.29
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.37
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.29
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.59
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | NCM driving% with AC160TS cantilever | 1.11
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | NCM driving% with AC160TS cantilever | 1.446
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | NCM Amplitude noise | 12.42
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | NCM Amplitude noise | 16.529
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.35
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.348
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.11
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.1
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 237.7
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.1
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 227
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.7
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 236.8
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.6
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 232.5
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238.2
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Lateral(C-D) sensitivity | 7.7
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Lateral(C-D) sensitivity | 16.5
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Lateral(C-D) noise | 0.137
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Lateral(C-D) noise | 0.097
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of Z Height during Z Scan sweep | 16.2
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of Z Height during Z Scan sweep | 16.3
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of Z Detector during Z Scan sweep | 11.5
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of Z Detector during Z Scan sweep | 12.5
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of Y1 Detector Raw during Y Scan sweep | 15
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of Y Detector during Y Scan sweep | 169.5
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of Y Detector during Y Scan sweep | 173.7
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of X2 Detector Raw during X Scan sweep | 14.8
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of X1 Detector Raw during X Scan sweep | 14.6
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | Full range of X Detector during X Scan sweep | 171.8
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | Full range of X Detector during X Scan sweep | 173.8
| Chip Hub for Integrated Photonics Xplore (CHIPX) | 2023-07-28 | A-B noise | 17.45
| Hebei Synlight Crystal Co.Ltd #6 | 2023-07-28 | A-B noise | 10.517
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Linearity by Z scanner position | 0.59
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Linearity by 4 different sample height | 0.999
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Height noise | 0.03
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.1
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y2 Detector noise (XY Servo on) | 0.249
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 242.9
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y1 Detector noise (XY Servo on) | 0.256
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 244
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 167.3
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y Ringing(Rpv) | 0.324
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y Non-linearity error | 2.15
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y Detector noise (XY Servo on) | 0.147
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Repeatability @1st position | 0.06
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Repeatability @ 4th position | 0.08
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Repeatability @ 3rd position | 0.05
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Repeatability @ 2nd position | 0.08
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Accuracy @1st position | 1.15
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Accuracy @ 4th position | 0.7
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Accuracy @ 3rd position | 0.73
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis, Accuracy @ 2nd position | 0.9
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Y axis OPM | 1.203
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X2 Detector noise (XY Servo on) | 0.317
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 251.1
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X1 Detector noise (XY Servo on) | 0.31
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 242.8
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 169.2
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X Ringing(Rpv) | 0.489
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X Non-linearity error | 2.36
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X Detector noise (XY Servo on) | 0.23
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Repeatability @1st position | 0.05
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Repeatability @ 4th position | 0.05
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Repeatability @ 3rd position | 0.04
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Repeatability @ 2nd position | 0.05
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Accuracy @1st position | 0.32
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Accuracy @ 4th position | 1.18
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Accuracy @ 3rd position | 0.31
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis, Accuracy @ 2nd position | 0.26
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | X axis OPM | 0.862
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | StrokeYMin | -108830
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | StrokeYMax | 105382
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | StrokeY | 214212
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | StrokeXMin | -140658
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | StrokeXMax | 137942
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | StrokeX | 278600
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.8
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | RMS of ZeroScan | 0.038
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | OpticalAxisShiftPerFocusY | 0.008
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 8e-05
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | OpticalAxisShiftPerFocusX | 0.0035
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | NCM driving% with AC160TS cantilever | 8.36
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Min. Vibration Transmissbility (~100Hz) | 0.001
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Lateral(C-D) sensitivity | 5.8
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Lateral(C-D) noise | 0.13
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Laser intensity with AC160 cantilever | 3.205
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from -Y to +Y (Y scan) | 169.19
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from -Y to +Y (X scan) | 109.86
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from -X to +X (Y scan) | 88.62
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from -X to +X (X scan) | 19.04
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from +Y to -Y (Y scan) | 84.23
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from +Y to -Y (X scan) | 39.55
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from +X to -X (Y scan) | 157.47
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Image offset after moving from +X to -X (X scan) | 16.11
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of Z Height during Z Scan sweep | 16.8
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of Z Detector during Z Scan sweep | 13.28
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of Y2 Detector Raw during Y Scan sweep | 14.26
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of Y2 Detector during Y Scan sweep | 186.3166
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of Y1 Detector Raw during Y Scan sweep | 14.28
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of Y1 Detector during Y Scan sweep | 180.1679
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of X2 Detector Raw during X Scan sweep | 13.98
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of X2 Detector during X Scan sweep | 181.1193
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of X1 Detector Raw during X Scan sweep | 14.04
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | Full range of X1 Detector during X Scan sweep | 188.8821
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.63
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | A－B difference during 3 times run test | 0.67
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | A＋B difference during 3 times run test | 0.15
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | 50~200nm Step Height Sample (10 Repeats) | 183.315
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | 200~500nm Step Height Sample (10 Repeats) | 460.1947
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 9
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.7
| Huzhou Tony Semiconductor Technology Co., Ltd | 2023-07-27 | %GRR w/ roughness of non-patterened wafer | 10.1
| 부경대학교 화학과 | 2023-07-26 | Z Scanner position at -3μm | 0.0029
| 부경대학교 화학과 | 2023-07-26 | Z Scanner position at +3μm | 0.0016
| 부경대학교 화학과 | 2023-07-26 | Z Linearity by 3 different sample height | 0.994
| 부경대학교 화학과 | 2023-07-26 | Z Height noise | 0.02
| 부경대학교 화학과 | 2023-07-26 | Y Scan Stroke | 55
| 부경대학교 화학과 | 2023-07-26 | Y Detector noise (XY Servo on) | 0.28
| 부경대학교 화학과 | 2023-07-26 | Y Detector noise (XY Servo off) | 0.28
| 부경대학교 화학과 | 2023-07-26 | Y axis OPM | 0.58
| 부경대학교 화학과 | 2023-07-26 | X Scan Stroke | 55
| 부경대학교 화학과 | 2023-07-26 | X Detector noise (XY Servo on) | 0.28
| 부경대학교 화학과 | 2023-07-26 | X Detector noise (XY Servo off) | 0.28
| 부경대학교 화학과 | 2023-07-26 | X axis OPM | 0.91
| 부경대학교 화학과 | 2023-07-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 부경대학교 화학과 | 2023-07-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.38
| 부경대학교 화학과 | 2023-07-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| 부경대학교 화학과 | 2023-07-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| 부경대학교 화학과 | 2023-07-26 | NCM driving% with AC160TS cantilever | 0.83
| 부경대학교 화학과 | 2023-07-26 | NCM Amplitude noise | 12.51
| 부경대학교 화학과 | 2023-07-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.12
| 부경대학교 화학과 | 2023-07-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.755
| 부경대학교 화학과 | 2023-07-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.6
| 부경대학교 화학과 | 2023-07-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99.2
| 부경대학교 화학과 | 2023-07-26 | Lateral(C-D) sensitivity | 10.4
| 부경대학교 화학과 | 2023-07-26 | Lateral(C-D) noise | 0.075
| 부경대학교 화학과 | 2023-07-26 | Full range of Z Height during Z Scan sweep | 15.9
| 부경대학교 화학과 | 2023-07-26 | Full range of Z Detector during Z Scan sweep | 11
| 부경대학교 화학과 | 2023-07-26 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| 부경대학교 화학과 | 2023-07-26 | Full range of Y Detector during Y Scan sweep | 72.3
| 부경대학교 화학과 | 2023-07-26 | Full range of X Detector Raw during X Scan sweep | 14.5
| 부경대학교 화학과 | 2023-07-26 | Full range of X Detector during X Scan sweep | 71.6
| 부경대학교 화학과 | 2023-07-26 | A-B noise | 8.15
| CXMT Hefei #7 | 2023-07-25 | Z Linearity by Z scanner position | 0.02
| CXMT Hefei #7 | 2023-07-25 | Z Linearity by 5 different sample height | 1.006
| CXMT Hefei #7 | 2023-07-25 | Z Height noise | 0.036
| CXMT Hefei #7 | 2023-07-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.84
| CXMT Hefei #7 | 2023-07-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.82
| CXMT Hefei #7 | 2023-07-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.64
| CXMT Hefei #7 | 2023-07-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 9.12
| CXMT Hefei #7 | 2023-07-25 | Y2 Detector noise (XY Servo on) | 0.289
| CXMT Hefei #7 | 2023-07-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 279.972
| CXMT Hefei #7 | 2023-07-25 | Y1 Detector noise (XY Servo on) | 0.22
| CXMT Hefei #7 | 2023-07-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 272.635
| CXMT Hefei #7 | 2023-07-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 197.737
| CXMT Hefei #7 | 2023-07-25 | Y Non-linearity error | 3.53
| CXMT Hefei #7 | 2023-07-25 | Y Detector noise (XY Servo on) | 0.306
| CXMT Hefei #7 | 2023-07-25 | Y axis, Repeatability @1st position | 0.03
| CXMT Hefei #7 | 2023-07-25 | Y axis, Repeatability @ 4th position | 0.05
| CXMT Hefei #7 | 2023-07-25 | Y axis, Repeatability @ 3rd position | 0.03
| CXMT Hefei #7 | 2023-07-25 | Y axis, Repeatability @ 2nd position | 0.03
| CXMT Hefei #7 | 2023-07-25 | Y axis, Accuracy @1st position | 0.26
| CXMT Hefei #7 | 2023-07-25 | Y axis, Accuracy @ 4th position | 0.47
| CXMT Hefei #7 | 2023-07-25 | Y axis, Accuracy @ 3rd position | 0.48
| CXMT Hefei #7 | 2023-07-25 | Y axis, Accuracy @ 2nd position | 0.41
| CXMT Hefei #7 | 2023-07-25 | Y axis OPM | 0.577
| CXMT Hefei #7 | 2023-07-25 | X2 Detector noise (XY Servo on) | 0.287
| CXMT Hefei #7 | 2023-07-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 272.02
| CXMT Hefei #7 | 2023-07-25 | X1 Detector noise (XY Servo on) | 0.313
| CXMT Hefei #7 | 2023-07-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 271.525
| CXMT Hefei #7 | 2023-07-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.311
| CXMT Hefei #7 | 2023-07-25 | X Non-linearity error | 3.27
| CXMT Hefei #7 | 2023-07-25 | X Detector noise (XY Servo on) | 0.244
| CXMT Hefei #7 | 2023-07-25 | X axis, Repeatability @1st position | 0.05
| CXMT Hefei #7 | 2023-07-25 | X axis, Repeatability @ 4th position | 0.04
| CXMT Hefei #7 | 2023-07-25 | X axis, Repeatability @ 3rd position | 0.04
| CXMT Hefei #7 | 2023-07-25 | X axis, Repeatability @ 2nd position | 0.06
| CXMT Hefei #7 | 2023-07-25 | X axis, Accuracy @1st position | 0.24
| CXMT Hefei #7 | 2023-07-25 | X axis, Accuracy @ 4th position | 0.3
| CXMT Hefei #7 | 2023-07-25 | X axis, Accuracy @ 3rd position | 0.25
| CXMT Hefei #7 | 2023-07-25 | X axis, Accuracy @ 2nd position | 0.43
| CXMT Hefei #7 | 2023-07-25 | X axis OPM | 0.419
| CXMT Hefei #7 | 2023-07-25 | Under 50nm Step Height Sample (10 Repeats) | 18.267
| CXMT Hefei #7 | 2023-07-25 | StrokeYMin | -153847
| CXMT Hefei #7 | 2023-07-25 | StrokeYMax | 150353
| CXMT Hefei #7 | 2023-07-25 | StrokeY | 304200
| CXMT Hefei #7 | 2023-07-25 | StrokeXMin | -204042
| CXMT Hefei #7 | 2023-07-25 | StrokeXMax | 203862
| CXMT Hefei #7 | 2023-07-25 | StrokeX | 407904
| CXMT Hefei #7 | 2023-07-25 | Roll motion to X+ full scan | 18.5
| CXMT Hefei #7 | 2023-07-25 | RMS of ZeroScan | 0.047
| CXMT Hefei #7 | 2023-07-25 | PosMin | -26.01
| CXMT Hefei #7 | 2023-07-25 | PosMax | 26.104
| CXMT Hefei #7 | 2023-07-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00018
| CXMT Hefei #7 | 2023-07-25 | OpticalAxisShiftPerFocusY | 0.0035
| CXMT Hefei #7 | 2023-07-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00021
| CXMT Hefei #7 | 2023-07-25 | OpticalAxisShiftPerFocusX | -0.0071
| CXMT Hefei #7 | 2023-07-25 | NCM driving% with AC160TS cantilever | 16
| CXMT Hefei #7 | 2023-07-25 | Min. Vibration Transmissbility (~100Hz) | 0.072
| CXMT Hefei #7 | 2023-07-25 | MaxLength | 50
| CXMT Hefei #7 | 2023-07-25 | Max OPM @5mm | 6.647
| CXMT Hefei #7 | 2023-07-25 | Max OPM @25mm | 90.136
| CXMT Hefei #7 | 2023-07-25 | Max OPM @1mm | 3.075
| CXMT Hefei #7 | 2023-07-25 | Max OPM @10mm | 19.483
| CXMT Hefei #7 | 2023-07-25 | Lateral(C-D) sensitivity | 3.6
| CXMT Hefei #7 | 2023-07-25 | Lateral(C-D) noise | 0.248
| CXMT Hefei #7 | 2023-07-25 | Laser intensity with AC160 cantilever | 3.2
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from -Y to +Y (Y scan) | 24.16
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from -Y to +Y (X scan) | 9.52
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from -X to +X (Y scan) | 145.75
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from -X to +X (X scan) | 90.82
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from +Y to -Y (Y scan) | 65.18
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from +Y to -Y (X scan) | 5.12
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from +X to -X (Y scan) | 111.32
| CXMT Hefei #7 | 2023-07-25 | Image offset after moving from +X to -X (X scan) | 161.87
| CXMT Hefei #7 | 2023-07-25 | Full range of Z Height during Z Scan sweep | 16.5
| CXMT Hefei #7 | 2023-07-25 | Full range of Z Detector during Z Scan sweep | 11.5
| CXMT Hefei #7 | 2023-07-25 | Full range of Y2 Detector Raw during Y Scan sweep | 13.9
| CXMT Hefei #7 | 2023-07-25 | Full range of Y2 Detector during Y Scan sweep | 213.1725
| CXMT Hefei #7 | 2023-07-25 | Full range of Y1 Detector Raw during Y Scan sweep | 13.9
| CXMT Hefei #7 | 2023-07-25 | Full range of Y1 Detector during Y Scan sweep | 208.8653
| CXMT Hefei #7 | 2023-07-25 | Full range of X2 Detector Raw during X Scan sweep | 14
| CXMT Hefei #7 | 2023-07-25 | Full range of X2 Detector during X Scan sweep | 206.4977
| CXMT Hefei #7 | 2023-07-25 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| CXMT Hefei #7 | 2023-07-25 | Full range of X1 Detector during X Scan sweep | 204.9345
| CXMT Hefei #7 | 2023-07-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.18
| CXMT Hefei #7 | 2023-07-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.97
| CXMT Hefei #7 | 2023-07-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.68
| CXMT Hefei #7 | 2023-07-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| CXMT Hefei #7 | 2023-07-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.1
| CXMT Hefei #7 | 2023-07-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.55
| CXMT Hefei #7 | 2023-07-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| CXMT Hefei #7 | 2023-07-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| CXMT Hefei #7 | 2023-07-25 | A－B difference during 3 times run test | 0.54
| CXMT Hefei #7 | 2023-07-25 | A＋B difference during 3 times run test | 0.09
| CXMT Hefei #7 | 2023-07-25 | 50~200nm Step Height Sample (10 Repeats) | 178.716
| CXMT Hefei #7 | 2023-07-25 | 200~500nm Step Height Sample (10 Repeats) | 459.978
| CXMT Hefei #7 | 2023-07-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.3
| CXMT Hefei #7 | 2023-07-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 8.3
| CXMT Hefei #7 | 2023-07-25 | %GRR w/ roughness of non-patterened wafer | 9.3
| Tokuyama Corporation | 2023-07-21 | Z Scanner position at -3μm | 0.0036
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Z Scanner position at -3μm | 0.0002
| Tokuyama Corporation | 2023-07-21 | Z Scanner position at +3μm | 0.0078
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Z Scanner position at +3μm | 0.0059
| Tokuyama Corporation | 2023-07-21 | Z Linearity by 3 different sample height | 1.006
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Z Linearity by 3 different sample height | 1.007
| Tokuyama Corporation | 2023-07-21 | Z Height noise | 0.018
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Z Height noise | 0.019
| Tokuyama Corporation | 2023-07-21 | Y Scan Stroke | 55
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Y Scan Stroke | 55
| Tokuyama Corporation | 2023-07-21 | Y Detector noise (XY Servo on) | 0.26
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Y Detector noise (XY Servo on) | 0.27
| Tokuyama Corporation | 2023-07-21 | Y Detector noise (XY Servo off) | 0.26
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Y Detector noise (XY Servo off) | 0.26
| Tokuyama Corporation | 2023-07-21 | Y axis OPM | 0.93
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Y axis OPM | 0.78
| Tokuyama Corporation | 2023-07-21 | X Scan Stroke | 55
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | X Scan Stroke | 55
| Tokuyama Corporation | 2023-07-21 | X Detector noise (XY Servo on) | 0.27
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | X Detector noise (XY Servo on) | 0.27
| Tokuyama Corporation | 2023-07-21 | X Detector noise (XY Servo off) | 0.26
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | X Detector noise (XY Servo off) | 0.26
| Tokuyama Corporation | 2023-07-21 | X axis OPM | 0.62
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | X axis OPM | 0.94
| Tokuyama Corporation | 2023-07-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Tokuyama Corporation | 2023-07-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Tokuyama Corporation | 2023-07-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Tokuyama Corporation | 2023-07-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| Tokuyama Corporation | 2023-07-21 | NCM driving% with AC160TS cantilever | 2.54
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | NCM driving% with AC160TS cantilever | 0.52
| Tokuyama Corporation | 2023-07-21 | NCM Amplitude noise | 13.72
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | NCM Amplitude noise | 14.13
| Tokuyama Corporation | 2023-07-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.7
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| Tokuyama Corporation | 2023-07-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.1
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.9
| Tokuyama Corporation | 2023-07-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 107.9
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.4
| Tokuyama Corporation | 2023-07-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.1
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.7
| Tokuyama Corporation | 2023-07-21 | Lateral(C-D) sensitivity | 9.8
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Lateral(C-D) sensitivity | 12.1
| Tokuyama Corporation | 2023-07-21 | Lateral(C-D) noise | 0.067
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Lateral(C-D) noise | 0.081
| Tokuyama Corporation | 2023-07-21 | Full range of Z Height during Z Scan sweep | 16.9
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Full range of Z Height during Z Scan sweep | 16.3
| Tokuyama Corporation | 2023-07-21 | Full range of Z Detector during Z Scan sweep | 12.5
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Full range of Z Detector during Z Scan sweep | 12.1
| Tokuyama Corporation | 2023-07-21 | Full range of Y Detector Raw during Y Scan sweep | 13.8
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Tokuyama Corporation | 2023-07-21 | Full range of Y Detector during Y Scan sweep | 74.4
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Full range of Y Detector during Y Scan sweep | 74.7
| Tokuyama Corporation | 2023-07-21 | Full range of X Detector Raw during X Scan sweep | 14
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Full range of X Detector Raw during X Scan sweep | 13.9
| Tokuyama Corporation | 2023-07-21 | Full range of X Detector during X Scan sweep | 74.4
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | Full range of X Detector during X Scan sweep | 74
| Tokuyama Corporation | 2023-07-21 | A-B noise | 8.78
| 제주대학교 화학코스메틱스학과 | 2023-07-21 | A-B noise | 10.24
| TYSiC #2 | 2023-07-20 | Z Scanner position at -3μm | 0.0016
| TYSiC #2 | 2023-07-20 | Z Scanner position at +3μm | 0.0082
| TYSiC #2 | 2023-07-20 | Z Linearity by 3 different sample height | 1.019
| TYSiC #2 | 2023-07-20 | Z Height noise | 0.028
| TYSiC #2 | 2023-07-20 | Y Scan Stroke | 133
| TYSiC #2 | 2023-07-20 | Y Detector noise (XY Servo on) | 0.29
| TYSiC #2 | 2023-07-20 | Y Detector noise (XY Servo off) | 0.29
| TYSiC #2 | 2023-07-20 | Y axis OPM | 1.41
| TYSiC #2 | 2023-07-20 | X Scan Stroke | 130
| TYSiC #2 | 2023-07-20 | -X Direction software Limit | -100771
| TYSiC #2 | 2023-07-20 | -X Direction software Limit | -101039
| TYSiC #2 | 2023-07-20 | X Detector noise (XY Servo on) | 0.3
| TYSiC #2 | 2023-07-20 | X Detector noise (XY Servo off) | 0.28
| TYSiC #2 | 2023-07-20 | X axis OPM | 0.94
| TYSiC #2 | 2023-07-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| TYSiC #2 | 2023-07-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| TYSiC #2 | 2023-07-20 | RMS of Z Height of Wafer Zeroscan (Top) | 0.28
| TYSiC #2 | 2023-07-20 | RMS of Z Height of Wafer Zeroscan (Right) | 0.27
| TYSiC #2 | 2023-07-20 | RMS of Z Height of Wafer Zeroscan (Left) | 0.27
| TYSiC #2 | 2023-07-20 | RMS of Z Height of Wafer ZeroScan (Center) | 0.27
| TYSiC #2 | 2023-07-20 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.28
| TYSiC #2 | 2023-07-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| TYSiC #2 | 2023-07-20 | NCM driving% with AC160TS cantilever | 1.21
| TYSiC #2 | 2023-07-20 | NCM Amplitude noise | 14.49
| TYSiC #2 | 2023-07-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.245
| TYSiC #2 | 2023-07-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.085
| TYSiC #2 | 2023-07-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 189
| TYSiC #2 | 2023-07-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 184
| TYSiC #2 | 2023-07-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| TYSiC #2 | 2023-07-20 | Lateral(C-D) sensitivity | 11.3
| TYSiC #2 | 2023-07-20 | Lateral(C-D) noise | 0.109
| TYSiC #2 | 2023-07-20 | Full range of Z Height during Z Scan sweep | 15.7
| TYSiC #2 | 2023-07-20 | Full range of Z Detector during Z Scan sweep | 10.8
| TYSiC #2 | 2023-07-20 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| TYSiC #2 | 2023-07-20 | Full range of Y Detector during Y Scan sweep | 133
| TYSiC #2 | 2023-07-20 | Full range of X Detector Raw during X Scan sweep | 14.1
| TYSiC #2 | 2023-07-20 | Full range of X Detector during X Scan sweep | 130
| TYSiC #2 | 2023-07-20 | A-B noise | 9.17
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Z Linearity by 3 different sample height | 0.988
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Z Height noise | 0.022
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Y Detector noise (XY Servo on) | 0.29
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Y Detector noise (XY Servo off) | 0.28
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Y axis OPM | 0.86
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | X Detector noise (XY Servo on) | 0.29
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | X Detector noise (XY Servo off) | 0.29
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | X axis OPM | 1.66
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | NCM driving% with AC160TS cantilever | 2.16
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | NCM Amplitude noise | 7.66
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.412
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.838
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 153.4
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 151.3
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Full range of Z Height during Z Scan sweep | 16.7
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Lateral(C-D) sensitivity | 8.3
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Lateral(C-D) noise | 0.095
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Full range of Z Detector during Z Scan sweep | 14
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Full range of Y Detector during Y Scan sweep | 115.7
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Full range of X Detector Raw during X Scan sweep | 15.2
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Full range of X Detector during X Scan sweep | 115
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Approach position at -3μm | 0.0052
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | Approach position at +3μm | 0.0003
| Zhejiang University_Center of Electron Microscopy | 2023-07-18 | A-B noise | 5.95
| OLED AOMORI | 2023-07-17 | Z Scanner position at -3μm | 0.0002
| OLED AOMORI | 2023-07-17 | Z Scanner position at +3μm | 0.0014
| OLED AOMORI | 2023-07-17 | Z Linearity by 3 different sample height | 1.005
| OLED AOMORI | 2023-07-17 | Z Height noise | 0.036
| OLED AOMORI | 2023-07-17 | Y Scan Stroke | 55
| OLED AOMORI | 2023-07-17 | Y Detector noise (XY Servo on) | 0.35
| OLED AOMORI | 2023-07-17 | Y Detector noise (XY Servo off) | 0.35
| OLED AOMORI | 2023-07-17 | Y axis OPM | 0.42
| OLED AOMORI | 2023-07-17 | X Scan Stroke | 55
| OLED AOMORI | 2023-07-17 | X Detector noise (XY Servo on) | 0.36
| OLED AOMORI | 2023-07-17 | X Detector noise (XY Servo off) | 0.36
| OLED AOMORI | 2023-07-17 | X axis OPM | 0.48
| OLED AOMORI | 2023-07-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| OLED AOMORI | 2023-07-17 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| OLED AOMORI | 2023-07-17 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| OLED AOMORI | 2023-07-17 | NCM driving% with AC160TS cantilever | 2.99
| OLED AOMORI | 2023-07-17 | NCM Amplitude noise | 15.08
| OLED AOMORI | 2023-07-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.6
| OLED AOMORI | 2023-07-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.5
| OLED AOMORI | 2023-07-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 91.4
| OLED AOMORI | 2023-07-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 92.6
| OLED AOMORI | 2023-07-17 | Lateral(C-D) sensitivity | 13.9
| OLED AOMORI | 2023-07-17 | Lateral(C-D) noise | 0.06
| OLED AOMORI | 2023-07-17 | Full range of Z Height during Z Scan sweep | 16.6
| OLED AOMORI | 2023-07-17 | Full range of Z Detector during Z Scan sweep | 11.3
| OLED AOMORI | 2023-07-17 | Full range of Y Detector Raw during Y Scan sweep | 15.8
| OLED AOMORI | 2023-07-17 | Full range of Y Detector during Y Scan sweep | 71.9
| OLED AOMORI | 2023-07-17 | Full range of X Detector Raw during X Scan sweep | 15.6
| OLED AOMORI | 2023-07-17 | Full range of X Detector during X Scan sweep | 72.2
| OLED AOMORI | 2023-07-17 | A-B noise | 9.75
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Z Scanner position at -3μm | 0.0053
| TYSiC #1 | 2023-07-14 | Z Scanner position at -3μm | 0.0006
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Z Scanner position at +3μm | 0.0084
| TYSiC #1 | 2023-07-14 | Z Scanner position at +3μm | 0.0168
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Z Linearity by 3 different sample height | 0.988
| TYSiC #1 | 2023-07-14 | Z Linearity by 3 different sample height | 0.987
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Z Height noise | 0.02
| TYSiC #1 | 2023-07-14 | Z Height noise | 0.017
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y2 Detector noise (XY Servo on) | 0.42
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y2 Detector noise (XY Servo off) | 0.29
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y1 Detector noise (XY Servo on) | 0.37
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y1 Detector noise (XY Servo off) | 0.28
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y Scan Stroke | 172.7
| TYSiC #1 | 2023-07-14 | Y Scan Stroke | 136.4
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y Detector noise (XY Servo on) | 0.3
| TYSiC #1 | 2023-07-14 | Y Detector noise (XY Servo on) | 0.27
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y Detector noise (XY Servo off) | 0.2
| TYSiC #1 | 2023-07-14 | Y Detector noise (XY Servo off) | 0.27
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Y axis OPM | 0.55
| TYSiC #1 | 2023-07-14 | Y axis OPM | 0.51
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X2 Detector noise (XY Servo on) | 0.41
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X2 Detector noise (XY Servo off) | 0.27
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X1 Detector noise (XY Servo on) | 0.42
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X1 Detector noise (XY Servo off) | 0.32
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X Scan Stroke | 175.1
| TYSiC #1 | 2023-07-14 | X Scan Stroke | 131.4
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | -X Direction software Limit | -102352
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | -X Direction software Limit | -102938
| TYSiC #1 | 2023-07-14 | -X Direction software Limit | -101189
| TYSiC #1 | 2023-07-14 | -X Direction software Limit | -100881
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X Detector noise (XY Servo on) | 0.32
| TYSiC #1 | 2023-07-14 | X Detector noise (XY Servo on) | 0.27
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X Detector noise (XY Servo off) | 0.21
| TYSiC #1 | 2023-07-14 | X Detector noise (XY Servo off) | 0.27
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | X axis OPM | 0.78
| TYSiC #1 | 2023-07-14 | X axis OPM | 0.7
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| TYSiC #1 | 2023-07-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.6
| TYSiC #1 | 2023-07-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of Wafer Zeroscan (Top) | 0.36
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of Wafer Zeroscan (Right) | 0.3
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of Wafer Zeroscan (Left) | 0.31
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of Wafer ZeroScan (Center) | 0.31
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.32
| TYSiC #1 | 2023-07-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| TYSiC #1 | 2023-07-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | NCM driving% with AC160TS cantilever | 1.76
| TYSiC #1 | 2023-07-14 | NCM driving% with AC160TS cantilever | 1.41
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | NCM Amplitude noise | 0.61
| TYSiC #1 | 2023-07-14 | NCM Amplitude noise | 13.2
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.5
| TYSiC #1 | 2023-07-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.1
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.4
| TYSiC #1 | 2023-07-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 29
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 220.3
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 222.4
| TYSiC #1 | 2023-07-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 189.4
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 224.7
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 224.5
| TYSiC #1 | 2023-07-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 183.8
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Lateral(C-D) sensitivity | 0.117
| TYSiC #1 | 2023-07-14 | Lateral(C-D) sensitivity | 11.6
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Lateral(C-D) noise | 8.82
| TYSiC #1 | 2023-07-14 | Lateral(C-D) noise | 0.135
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of Z Height during Z Scan sweep | 15.5
| TYSiC #1 | 2023-07-14 | Full range of Z Height during Z Scan sweep | 15.6
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of Z Detector during Z Scan sweep | 13.2
| TYSiC #1 | 2023-07-14 | Full range of Z Detector during Z Scan sweep | 10.7
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of Y2 Detector Raw during Y Scan sweep | 15.7
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of Y1 Detector Raw during Y Scan sweep | 15.6
| TYSiC #1 | 2023-07-14 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of Y Detector during Y Scan sweep | 172.9
| TYSiC #1 | 2023-07-14 | Full range of Y Detector during Y Scan sweep | 134.7
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of X2 Detector Raw during X Scan sweep | 15.7
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of X1 Detector Raw during X Scan sweep | 15.5
| TYSiC #1 | 2023-07-14 | Full range of X Detector Raw during X Scan sweep | 14.1
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | Full range of X Detector during X Scan sweep | 175.3
| TYSiC #1 | 2023-07-14 | Full range of X Detector during X Scan sweep | 130
| Nanjing University_Electronic Science & Engineering | 2023-07-14 | A-B noise | 12.93
| TYSiC #1 | 2023-07-14 | A-B noise | 8.6
| PS UK Demo | 2023-07-13 | Z Linearity by 3 different sample height | 1.011
| PS UK Demo | 2023-07-13 | Z Height noise | 0.019
| PS UK Demo | 2023-07-13 | Y Detector noise (XY Servo on) | 0.27
| PS UK Demo | 2023-07-13 | Y Detector noise (XY Servo off) | 0.27
| PS UK Demo | 2023-07-13 | Y axis OPM | 1.52
| PS UK Demo | 2023-07-13 | X Detector noise (XY Servo on) | 0.27
| PS UK Demo | 2023-07-13 | X Detector noise (XY Servo off) | 0.27
| PS UK Demo | 2023-07-13 | X axis OPM | 1.43
| PS UK Demo | 2023-07-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| PS UK Demo | 2023-07-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| PS UK Demo | 2023-07-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| PS UK Demo | 2023-07-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.25
| PS UK Demo | 2023-07-13 | NCM driving% with AC160TS cantilever | 2.08
| PS UK Demo | 2023-07-13 | NCM Amplitude noise | 7.42
| PS UK Demo | 2023-07-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.077
| PS UK Demo | 2023-07-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.601
| PS UK Demo | 2023-07-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 146.8
| PS UK Demo | 2023-07-13 | Full range of Z Height during Z Scan sweep | 15.9
| PS UK Demo | 2023-07-13 | Lateral(C-D) sensitivity | 6.6
| PS UK Demo | 2023-07-13 | Lateral(C-D) noise | 0.115
| PS UK Demo | 2023-07-13 | Full range of Z Detector during Z Scan sweep | 14.7
| PS UK Demo | 2023-07-13 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| PS UK Demo | 2023-07-13 | Full range of Y Detector during Y Scan sweep | 111
| PS UK Demo | 2023-07-13 | Full range of X Detector Raw during X Scan sweep | 15.4
| PS UK Demo | 2023-07-13 | Full range of X Detector during X Scan sweep | 112.8
| PS UK Demo | 2023-07-13 | Approach position at -3μm | 0.019
| PS UK Demo | 2023-07-13 | Approach position at +3μm | 0.0185
| PS UK Demo | 2023-07-13 | A-B noise | 12.77
| SMOC | 2023-07-10 | Z Linearity by Z scanner position | 0.62
| SMOC | 2023-07-10 | Z Linearity by 5 different sample height | 1.003
| SMOC | 2023-07-10 | Z Height noise | 0.041
| SMOC | 2023-07-10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.7
| SMOC | 2023-07-10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.62
| SMOC | 2023-07-10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.11
| SMOC | 2023-07-10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.09
| SMOC | 2023-07-10 | Y2 Detector noise (XY Servo on) | 0.39
| SMOC | 2023-07-10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 275.4
| SMOC | 2023-07-10 | Y1 Detector noise (XY Servo on) | 0.39
| SMOC | 2023-07-10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 278.6
| SMOC | 2023-07-10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 194.7
| SMOC | 2023-07-10 | Y Ringing(Rpv) | 0.09
| SMOC | 2023-07-10 | Y Non-linearity error | 2.5
| SMOC | 2023-07-10 | Y Detector noise (XY Servo on) | 0.31
| SMOC | 2023-07-10 | Y axis, Repeatability @1st position | 0.06
| SMOC | 2023-07-10 | Y axis, Repeatability @ 4th position | 0.05
| SMOC | 2023-07-10 | Y axis, Repeatability @ 3rd position | 0.07
| SMOC | 2023-07-10 | Y axis, Repeatability @ 2nd position | 0.06
| SMOC | 2023-07-10 | Y axis, Accuracy @1st position | 0.61
| SMOC | 2023-07-10 | Y axis, Accuracy @ 4th position | 0.85
| SMOC | 2023-07-10 | Y axis, Accuracy @ 3rd position | 0.94
| SMOC | 2023-07-10 | Y axis, Accuracy @ 2nd position | 0.74
| SMOC | 2023-07-10 | Y axis OPM | 0.37
| SMOC | 2023-07-10 | X2 Detector noise (XY Servo on) | 0.32
| SMOC | 2023-07-10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 288.5
| SMOC | 2023-07-10 | X1 Detector noise (XY Servo on) | 0.37
| SMOC | 2023-07-10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 283.8
| SMOC | 2023-07-10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 196.7
| SMOC | 2023-07-10 | X Ringing(Rpv) | 0.197
| SMOC | 2023-07-10 | X Non-linearity error | 3.13
| SMOC | 2023-07-10 | X Detector noise (XY Servo on) | 0.31
| SMOC | 2023-07-10 | X axis, Repeatability @1st position | 0.04
| SMOC | 2023-07-10 | X axis, Repeatability @ 4th position | 0.05
| SMOC | 2023-07-10 | X axis, Repeatability @ 3rd position | 0.05
| SMOC | 2023-07-10 | X axis, Repeatability @ 2nd position | 0.05
| SMOC | 2023-07-10 | X axis, Accuracy @1st position | 0.2
| SMOC | 2023-07-10 | X axis, Accuracy @ 4th position | 0.34
| SMOC | 2023-07-10 | X axis, Accuracy @ 3rd position | 0.27
| SMOC | 2023-07-10 | X axis, Accuracy @ 2nd position | 0.47
| SMOC | 2023-07-10 | X axis OPM | 0.24
| SMOC | 2023-07-10 | Under 50nm Step Height Sample (10 Repeats) | 18.3
| SMOC | 2023-07-10 | Torque Profile (50mm, 20um/s) | 5.13
| SMOC | 2023-07-10 | Torque Profile (50mm, 200um/s) | 4.178
| SMOC | 2023-07-10 | StrokeYMin | -153110
| SMOC | 2023-07-10 | StrokeYMax | 150170
| SMOC | 2023-07-10 | StrokeY | 303280
| SMOC | 2023-07-10 | StrokeXMin | -203460
| SMOC | 2023-07-10 | StrokeXMax | 203490
| SMOC | 2023-07-10 | StrokeX | 406950
| SMOC | 2023-07-10 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.57
| SMOC | 2023-07-10 | Roll motion to X+ full scan | 19.4
| SMOC | 2023-07-10 | RMS of ZeroScan | 0.046
| SMOC | 2023-07-10 | Profiler noise @200um | 0.128
| SMOC | 2023-07-10 | PosMin | -26
| SMOC | 2023-07-10 | PosMax | 26.03
| SMOC | 2023-07-10 | Position repeatability(10repeats, Position Based) | 0.18
| SMOC | 2023-07-10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| SMOC | 2023-07-10 | OpticalAxisShiftPerFocusY | 0.0029
| SMOC | 2023-07-10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00037
| SMOC | 2023-07-10 | OpticalAxisShiftPerFocusX | -0.0007
| SMOC | 2023-07-10 | OPM Repeatability @5mm | 0.94
| SMOC | 2023-07-10 | OPM Repeatability @25mm | 5.67
| SMOC | 2023-07-10 | OPM Repeatability @1mm | 0.39
| SMOC | 2023-07-10 | OPM Repeatability @10mm | 1.2
| SMOC | 2023-07-10 | NCM driving% with AC160TS cantilever | 14.78
| SMOC | 2023-07-10 | Min. Vibration Transmissbility (~100Hz) | 0.07
| SMOC | 2023-07-10 | MaxLength | 50
| SMOC | 2023-07-10 | Max OPM @5mm | 15.49
| SMOC | 2023-07-10 | Max OPM @25mm | 221.59
| SMOC | 2023-07-10 | Max OPM @1mm | 4.19
| SMOC | 2023-07-10 | Max OPM @10mm | 22.88
| SMOC | 2023-07-10 | Lateral(C-D) sensitivity | 3.8
| SMOC | 2023-07-10 | Lateral(C-D) noise | 0.242
| SMOC | 2023-07-10 | Laser intensity with AC160 cantilever | 2.77
| SMOC | 2023-07-10 | Image offset after moving from -Y to +Y (Y scan) | 40.28
| SMOC | 2023-07-10 | Image offset after moving from -Y to +Y (X scan) | 31.49
| SMOC | 2023-07-10 | Image offset after moving from -X to +X (Y scan) | 17.58
| SMOC | 2023-07-10 | Image offset after moving from -X to +X (X scan) | 14.65
| SMOC | 2023-07-10 | Image offset after moving from +Y to -Y (Y scan) | 120.12
| SMOC | 2023-07-10 | Image offset after moving from +Y to -Y (X scan) | 6.59
| SMOC | 2023-07-10 | Image offset after moving from +X to -X (Y scan) | 14.65
| SMOC | 2023-07-10 | Image offset after moving from +X to -X (X scan) | 60.79
| SMOC | 2023-07-10 | Full range of Z Height during Z Scan sweep | 17.1
| SMOC | 2023-07-10 | Full range of Z Detector during Z Scan sweep | 12.4
| SMOC | 2023-07-10 | Full range of Y2 Detector Raw during Y Scan sweep | 13.44
| SMOC | 2023-07-10 | Full range of Y2 Detector during Y Scan sweep | 201.5066
| SMOC | 2023-07-10 | Full range of Y1 Detector Raw during Y Scan sweep | 13.74
| SMOC | 2023-07-10 | Full range of Y1 Detector during Y Scan sweep | 203.5423
| SMOC | 2023-07-10 | Full range of X2 Detector Raw during X Scan sweep | 13.84
| SMOC | 2023-07-10 | Full range of X2 Detector during X Scan sweep | 215.7954
| SMOC | 2023-07-10 | Full range of X1 Detector Raw during X Scan sweep | 13.82
| SMOC | 2023-07-10 | Full range of X1 Detector during X Scan sweep | 209.1507
| SMOC | 2023-07-10 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| SMOC | 2023-07-10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.09
| SMOC | 2023-07-10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.47
| SMOC | 2023-07-10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.67
| SMOC | 2023-07-10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| SMOC | 2023-07-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| SMOC | 2023-07-10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.09
| SMOC | 2023-07-10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| SMOC | 2023-07-10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| SMOC | 2023-07-10 | A－B difference during 3 times run test | 0.34
| SMOC | 2023-07-10 | A＋B difference during 3 times run test | 0.04
| SMOC | 2023-07-10 | 50~200nm Step Height Sample (10 Repeats) | 183.32
| SMOC | 2023-07-10 | 200~500nm Step Height Sample (10 Repeats) | 464.07
| SMOC | 2023-07-10 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.3
| SMOC | 2023-07-10 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3
| SMOC | 2023-07-10 | %GRR w/ roughness of non-patterened wafer | 8.8
| Samsung P3 | 2023-06-30 | Z Linearity by Z scanner position | 0.57
| Samsung P3 | 2023-06-30 | Z Linearity by 5 different sample height | 0.999
| Samsung P3 | 2023-06-30 | Z Height noise | 0.039
| Samsung P3 | 2023-06-30 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.54
| Samsung P3 | 2023-06-30 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.58
| Samsung P3 | 2023-06-30 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.65
| Samsung P3 | 2023-06-30 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.634
| Samsung P3 | 2023-06-30 | Y2 Detector noise (XY Servo on) | 0.7
| Samsung P3 | 2023-06-30 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 259.5
| Samsung P3 | 2023-06-30 | Y1 Detector noise (XY Servo on) | 0.64
| Samsung P3 | 2023-06-30 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 258.9
| Samsung P3 | 2023-06-30 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 184.4
| Samsung P3 | 2023-06-30 | Y Ringing(Rpv) | 0.12
| Samsung P3 | 2023-06-30 | Y Non-linearity error | 2.94
| Samsung P3 | 2023-06-30 | Y Detector noise (XY Servo on) | 0.6
| Samsung P3 | 2023-06-30 | Y axis, Repeatability @1st position | 0.03
| Samsung P3 | 2023-06-30 | Y axis, Repeatability @ 4th position | 0.06
| Samsung P3 | 2023-06-30 | Y axis, Repeatability @ 3rd position | 0.05
| Samsung P3 | 2023-06-30 | Y axis, Repeatability @ 2nd position | 0.02
| Samsung P3 | 2023-06-30 | Y axis, Accuracy @1st position | 0.42
| Samsung P3 | 2023-06-30 | Y axis, Accuracy @ 4th position | 0.59
| Samsung P3 | 2023-06-30 | Y axis, Accuracy @ 3rd position | 0.64
| Samsung P3 | 2023-06-30 | Y axis, Accuracy @ 2nd position | 0.53
| Samsung P3 | 2023-06-30 | Y axis OPM | 1.49
| Samsung P3 | 2023-06-30 | X2 Detector noise (XY Servo on) | 0.6
| Samsung P3 | 2023-06-30 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 262.4
| Samsung P3 | 2023-06-30 | X1 Detector noise (XY Servo on) | 0.54
| Samsung P3 | 2023-06-30 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 258.9
| Samsung P3 | 2023-06-30 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 189.8
| Samsung P3 | 2023-06-30 | X Ringing(Rpv) | 0.17
| Samsung P3 | 2023-06-30 | X Non-linearity error | 2.2
| Samsung P3 | 2023-06-30 | X Detector noise (XY Servo on) | 0.49
| Samsung P3 | 2023-06-30 | X axis, Repeatability @1st position | 0.03
| Samsung P3 | 2023-06-30 | X axis, Repeatability @ 4th position | 0.03
| Samsung P3 | 2023-06-30 | X axis, Repeatability @ 3rd position | 0.03
| Samsung P3 | 2023-06-30 | X axis, Repeatability @ 2nd position | 0.03
| Samsung P3 | 2023-06-30 | X axis, Accuracy @1st position | 0.61
| Samsung P3 | 2023-06-30 | X axis, Accuracy @ 4th position | 0.95
| Samsung P3 | 2023-06-30 | X axis, Accuracy @ 3rd position | 0.32
| Samsung P3 | 2023-06-30 | X axis, Accuracy @ 2nd position | 0.46
| Samsung P3 | 2023-06-30 | X axis OPM | 2.85
| Samsung P3 | 2023-06-30 | Under 50nm Step Height Sample (10 Repeats) | 18.136
| Samsung P3 | 2023-06-30 | Torque Profile (50mm, 20um/s) | 2.952
| Samsung P3 | 2023-06-30 | Torque Profile (50mm, 200um/s) | 2.492
| Samsung P3 | 2023-06-30 | StrokeYMin | -152000
| Samsung P3 | 2023-06-30 | StrokeYMax | 150000
| Samsung P3 | 2023-06-30 | StrokeY | 302000
| Samsung P3 | 2023-06-30 | StrokeXMin | -201000
| Samsung P3 | 2023-06-30 | StrokeXMax | 252000
| Samsung P3 | 2023-06-30 | StrokeX | 453000
| Samsung P3 | 2023-06-30 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.32
| Samsung P3 | 2023-06-30 | Roll motion to X+ full scan | 2.1
| Samsung P3 | 2023-06-30 | RMS of ZeroScan | 0.04
| Samsung P3 | 2023-06-30 | Profiler noise @200um | 0.143
| Samsung P3 | 2023-06-30 | PosMin | -26.057
| Samsung P3 | 2023-06-30 | PosMax | 26.155
| Samsung P3 | 2023-06-30 | Position repeatability(10repeats, Position Based) | 0.2
| Samsung P3 | 2023-06-30 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | -0.003
| Samsung P3 | 2023-06-30 | OpticalAxisShiftPerFocusY | -0.0027
| Samsung P3 | 2023-06-30 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | -0.0011
| Samsung P3 | 2023-06-30 | OpticalAxisShiftPerFocusX | 0.0009
| Samsung P3 | 2023-06-30 | OPM Repeatability @5mm | 2.018
| Samsung P3 | 2023-06-30 | OPM Repeatability @25mm | 3.963
| Samsung P3 | 2023-06-30 | OPM Repeatability @1mm | 1.255
| Samsung P3 | 2023-06-30 | OPM Repeatability @10mm | 2.26
| Samsung P3 | 2023-06-30 | NCM driving% with AC160TS cantilever | 28.9
| Samsung P3 | 2023-06-30 | Min. Vibration Transmissbility (~100Hz) | 0.002
| Samsung P3 | 2023-06-30 | MaxLength | 51
| Samsung P3 | 2023-06-30 | Max OPM @5mm | 29.41
| Samsung P3 | 2023-06-30 | Max OPM @25mm | 146.717
| Samsung P3 | 2023-06-30 | Max OPM @1mm | 8.031
| Samsung P3 | 2023-06-30 | Max OPM @10mm | 50.531
| Samsung P3 | 2023-06-30 | Lateral(C-D) sensitivity | 8.2
| Samsung P3 | 2023-06-30 | Lateral(C-D) noise | 0.245
| Samsung P3 | 2023-06-30 | Laser intensity with AC160 cantilever | 3
| Samsung P3 | 2023-06-30 | Image offset after moving from -Y to +Y (Y scan) | 73.97
| Samsung P3 | 2023-06-30 | Image offset after moving from -Y to +Y (X scan) | 22.71
| Samsung P3 | 2023-06-30 | Image offset after moving from -X to +X (Y scan) | 51.27
| Samsung P3 | 2023-06-30 | Image offset after moving from -X to +X (X scan) | 82.03
| Samsung P3 | 2023-06-30 | Image offset after moving from +Y to -Y (Y scan) | 65.19
| Samsung P3 | 2023-06-30 | Image offset after moving from +Y to -Y (X scan) | 13.92
| Samsung P3 | 2023-06-30 | Image offset after moving from +X to -X (Y scan) | 73.24
| Samsung P3 | 2023-06-30 | Image offset after moving from +X to -X (X scan) | 12.45
| Samsung P3 | 2023-06-30 | Full range of Z Height during Z Scan sweep | 17
| Samsung P3 | 2023-06-30 | Full range of Z Detector during Z Scan sweep | 12.13
| Samsung P3 | 2023-06-30 | Full range of Y2 Detector Raw during Y Scan sweep | 13.71
| Samsung P3 | 2023-06-30 | Full range of Y2 Detector during Y Scan sweep | 197.6127
| Samsung P3 | 2023-06-30 | Full range of Y1 Detector Raw during Y Scan sweep | 13.63
| Samsung P3 | 2023-06-30 | Full range of Y1 Detector during Y Scan sweep | 188.1608
| Samsung P3 | 2023-06-30 | Full range of X2 Detector Raw during X Scan sweep | 14.27
| Samsung P3 | 2023-06-30 | Full range of X2 Detector during X Scan sweep | 205.2912
| Samsung P3 | 2023-06-30 | Full range of X1 Detector Raw during X Scan sweep | 14.13
| Samsung P3 | 2023-06-30 | Full range of X1 Detector during X Scan sweep | 195.5264
| Samsung P3 | 2023-06-30 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| Samsung P3 | 2023-06-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Samsung P3 | 2023-06-30 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.89
| Samsung P3 | 2023-06-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.66
| Samsung P3 | 2023-06-30 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Samsung P3 | 2023-06-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| Samsung P3 | 2023-06-30 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.23
| Samsung P3 | 2023-06-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| Samsung P3 | 2023-06-30 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Samsung P3 | 2023-06-30 | A－B difference during 3 times run test | 0.24
| Samsung P3 | 2023-06-30 | A＋B difference during 3 times run test | 0.04
| Samsung P3 | 2023-06-30 | 50~200nm Step Height Sample (10 Repeats) | 183.178
| Samsung P3 | 2023-06-30 | 200~500nm Step Height Sample (10 Repeats) | 462.296
| Samsung P3 | 2023-06-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 2.5
| Samsung P3 | 2023-06-30 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.4
| Samsung P3 | 2023-06-30 | %GRR w/ roughness of non-patterened wafer | 4.3
| Sicentury Semiconductor Technology | 2023-06-30 | Z Scanner position at -3μm | 0.009205
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Z Scanner position at -3μm | 0.0007
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Z Scanner position at -3μm | 0.0009
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Z Scanner position at -3μm | 0.001882
| Sicentury Semiconductor Technology | 2023-06-30 | Z Scanner position at +3μm | 0.000229
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Z Scanner position at +3μm | 0.0059
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Z Scanner position at +3μm | 0.0068
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Z Scanner position at +3μm | 0.009631
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Z Linearity by Z Scanner position
| Sicentury Semiconductor Technology | 2023-06-30 | Z Linearity by 3 different sample height | 1.0099
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Z Linearity by 3 different sample height | 1.016
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Z Linearity by 3 different sample height | 1
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Z Linearity by 3 different sample height | 0.99791
| Sicentury Semiconductor Technology | 2023-06-30 | Z Height noise | 0.021
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Z Height noise | 0.018
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Z Height noise | 0.019
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Z Height noise | 0.036
| Sicentury Semiconductor Technology | 2023-06-30 | Y2 Detector noise (XY Servo on) | 0.274
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y2 Detector noise (XY Servo on) | 0.25
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y2 Detector noise (XY Servo on) | 0.36
| Sicentury Semiconductor Technology | 2023-06-30 | Y2 Detector noise (XY Servo off) | 0.196
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y2 Detector noise (XY Servo off) | 0.25
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y2 Detector noise (XY Servo off) | 0.29
| Sicentury Semiconductor Technology | 2023-06-30 | Y1 Detector noise (XY Servo on) | 0.283
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y1 Detector noise (XY Servo on) | 0.27
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y1 Detector noise (XY Servo on) | 0.32
| Sicentury Semiconductor Technology | 2023-06-30 | Y1 Detector noise (XY Servo off) | 0.192
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y1 Detector noise (XY Servo off) | 0.28
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y1 Detector noise (XY Servo off) | 0.22
| Sicentury Semiconductor Technology | 2023-06-30 | Y Scan Stroke | 167.6
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y Scan Stroke | 162.7
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Y Scan Stroke | 55
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y Scan Stroke | 174.1
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | -Y Direction software Limit | -75508.8
| Sicentury Semiconductor Technology | 2023-06-30 | Y Detector noise (XY Servo on) | 0.219
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y Detector noise (XY Servo on) | 0.19
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Y Detector noise (XY Servo on) | 0.36
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y Detector noise (XY Servo on) | 0.27
| Sicentury Semiconductor Technology | 2023-06-30 | Y Detector noise (XY Servo off) | 0.153
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y Detector noise (XY Servo off) | 0.19
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Y Detector noise (XY Servo off) | 0.38
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y Detector noise (XY Servo off) | 0.19
| Sicentury Semiconductor Technology | 2023-06-30 | Y axis OPM | 0.8952
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Y axis OPM | 1.07
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Y axis OPM | 0.48
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Y axis OPM | 0.9896
| Sicentury Semiconductor Technology | 2023-06-30 | X2 Detector noise (XY Servo on) | 0.289
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X2 Detector noise (XY Servo on) | 0.24
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X2 Detector noise (XY Servo on) | 0.38
| Sicentury Semiconductor Technology | 2023-06-30 | X2 Detector noise (XY Servo off) | 0.213
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X2 Detector noise (XY Servo off) | 0.24
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X2 Detector noise (XY Servo off) | 0.28
| Sicentury Semiconductor Technology | 2023-06-30 | X1 Detector noise (XY Servo on) | 0.306
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X1 Detector noise (XY Servo on) | 0.26
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X1 Detector noise (XY Servo on) | 0.39
| Sicentury Semiconductor Technology | 2023-06-30 | X1 Detector noise (XY Servo off) | 0.216
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X1 Detector noise (XY Servo off) | 0.3
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X1 Detector noise (XY Servo off) | 0.29
| Sicentury Semiconductor Technology | 2023-06-30 | X Scan Stroke | 169.2
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X Scan Stroke | 167.4
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | X Scan Stroke | 55
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X Scan Stroke | 173.1
| Sicentury Semiconductor Technology | 2023-06-30 | -X Direction software Limit | -102094
| Sicentury Semiconductor Technology | 2023-06-30 | -X Direction software Limit | -103296
| Zhejiang Jinray Electronic Technology | 2023-06-30 | -X Direction software Limit | -102069
| Zhejiang Jinray Electronic Technology | 2023-06-30 | -X Direction software Limit | -103240
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | -X Direction software Limit | -75470
| Sicentury Semiconductor Technology | 2023-06-30 | X Detector noise (XY Servo on) | 0.242
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X Detector noise (XY Servo on) | 0.19
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | X Detector noise (XY Servo on) | 0.36
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X Detector noise (XY Servo on) | 0.31
| Sicentury Semiconductor Technology | 2023-06-30 | X Detector noise (XY Servo off) | 0.178
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X Detector noise (XY Servo off) | 0.2
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | X Detector noise (XY Servo off) | 0.37
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X Detector noise (XY Servo off) | 0.21
| Sicentury Semiconductor Technology | 2023-06-30 | X axis OPM | 1.1832
| Zhejiang Jinray Electronic Technology | 2023-06-30 | X axis OPM | 0.88
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | X axis OPM | 0.38
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | X axis OPM | 0.9158
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Top) | 0.32
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Top) | 0.36
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Right) | 0.31
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Right) | 0.32
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Left) | 0.31
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Left) | 0.39
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of Wafer ZeroScan (Center) | 0.3
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of Wafer ZeroScan (Center) | 0.37
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.31
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.37
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Sicentury Semiconductor Technology | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Zhejiang Jinray Electronic Technology | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.35
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of Wafer Zeroscan(-60000,0) | 0.38
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of Wafer Zeroscan(60000,0) | 0.35
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of Wafer Zeroscan(0,-60000) | 0.37
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of Wafer Zeroscan(0,60000) | 0.38
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | RMS of Z Drive of Wafer ZeroScan (0,0) | 0.37
| Sicentury Semiconductor Technology | 2023-06-30 | NCM driving% with AC160TS cantilever | 1.132
| Zhejiang Jinray Electronic Technology | 2023-06-30 | NCM driving% with AC160TS cantilever | 0.58
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | NCM driving% with AC160TS cantilever | 1.5
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | NCM driving% with AC160TS cantilever | 1.068
| Sicentury Semiconductor Technology | 2023-06-30 | NCM Amplitude noise | 12.192
| Zhejiang Jinray Electronic Technology | 2023-06-30 | NCM Amplitude noise | 14.33
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | NCM Amplitude noise | 16.67
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | NCM Amplitude noise | 14.463
| Sicentury Semiconductor Technology | 2023-06-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.255
| Zhejiang Jinray Electronic Technology | 2023-06-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.5
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.24
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.412
| Sicentury Semiconductor Technology | 2023-06-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.157
| Zhejiang Jinray Electronic Technology | 2023-06-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.4
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.297
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.679
| Sicentury Semiconductor Technology | 2023-06-30 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 231.1
| Zhejiang Jinray Electronic Technology | 2023-06-30 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 229.2
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.6
| Sicentury Semiconductor Technology | 2023-06-30 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 228
| Zhejiang Jinray Electronic Technology | 2023-06-30 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 215.7
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 240.3
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 97
| Sicentury Semiconductor Technology | 2023-06-30 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 232.1
| Zhejiang Jinray Electronic Technology | 2023-06-30 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 241.1
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 238.7
| Sicentury Semiconductor Technology | 2023-06-30 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 228.4
| Zhejiang Jinray Electronic Technology | 2023-06-30 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 220.8
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 245.6
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101
| Sicentury Semiconductor Technology | 2023-06-30 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Sicentury Semiconductor Technology | 2023-06-30 | Lateral(C-D) sensitivity | 9.9
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Lateral(C-D) sensitivity | 13.1
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Lateral(C-D) sensitivity | 12.5
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Lateral(C-D) sensitivity | 12.5
| Sicentury Semiconductor Technology | 2023-06-30 | Lateral(C-D) noise | 0.12
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Lateral(C-D) noise | 0.117
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Lateral(C-D) noise | 0.057
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Lateral(C-D) noise | 0.085
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of Z Height during Z Scan sweep | 16.1
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of Z Height during Z Scan sweep | 16.4
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Full range of Z Height during Z Scan sweep | 16
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of Z Height during Z Scan sweep | 16.5
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of Z Detector during Z Scan sweep | 11.4
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of Z Detector during Z Scan sweep | 11.5
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Full range of Z Detector during Z Scan sweep | 11.3
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of Z Detector during Z Scan sweep | 12.4
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of Y Detector during Y Scan sweep | 167.6
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of Y Detector during Y Scan sweep | 160.7
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Full range of Y Detector during Y Scan sweep | 71.6
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of Y Detector during Y Scan sweep | 176.3
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of X2 Detector Raw during X Scan sweep | 14.7
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of X1 Detector Raw during X Scan sweep | 14.7
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of X1 Detector Raw during X Scan sweep | 14.8
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Full range of X Detector Raw during X Scan sweep | 14.1
| Sicentury Semiconductor Technology | 2023-06-30 | Full range of X Detector during X Scan sweep | 169.1
| Zhejiang Jinray Electronic Technology | 2023-06-30 | Full range of X Detector during X Scan sweep | 166.5
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | Full range of X Detector during X Scan sweep | 75.1
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | Full range of X Detector during X Scan sweep | 175.1
| Sicentury Semiconductor Technology | 2023-06-30 | A-B noise | 7.724
| Zhejiang Jinray Electronic Technology | 2023-06-30 | A-B noise | 8.39
| 전남대학교_석유화학신소재공학과 | 2023-06-30 | A-B noise | 8.89
| Sun Yat-sen University of ShenZhen Lab | 2023-06-30 | A-B noise | 12.448
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Z Linearity by Z scanner position | 0.34
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Z Linearity by 5 different sample height | 1
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Z Height noise | 0.026
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 20.04
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 8.81
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 65.1
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y Ringing(Rpv) | 0.852
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y Detector noise (XY Servo on) | 0.19
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 83.7781
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Repeatability @1st position | 0.04
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Repeatability @ 4th position | 0.04
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Repeatability @ 3rd position | 0.05
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Repeatability @ 2nd position | 0.04
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Accuracy @1st position | 0.56
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Accuracy @ 4th position | 0.63
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Accuracy @ 3rd position | 0.72
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis, Accuracy @ 2nd position | 0.54
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Y axis OPM | 0.37
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 55
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X Ringing(Rpv) | 0.732
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X Detector noise (XY Servo on) | 0.18
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 90.9637
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Repeatability @1st position | 0.04
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Repeatability @ 4th position | 0.05
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Repeatability @ 3rd position | 0.03
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Repeatability @ 2nd position | 0.03
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Accuracy @1st position | 0.19
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Accuracy @ 4th position | 0.26
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Accuracy @ 3rd position | 0.25
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis, Accuracy @ 2nd position | 0.23
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | X axis OPM | 1.53
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Under 50nm Step Height Sample (10 Repeats) | 18.9
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | StrokeYMin | -153961
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | StrokeYMax | 150081
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | StrokeY | 304041.8
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | StrokeXMin | -204385
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | StrokeXMax | 203520.4
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | StrokeX | 407905.3
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | RMS of ZeroScan | 0.05
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | OpticalAxisShiftPerFocusY | 0.0016
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | OpticalAxisShiftPerFocusX | 0.0064
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | NCM driving% with AC160TS cantilever | 2.54
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Min. Vibration Transmissbility (~100Hz) | 0.01
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Mask Roughness Repeatability (10Repeats, σ ) | 0.002
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Mask Roughness Repeatability (10Repeats, RSD%) | 2.71
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Mask exchange time | 3.519
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Laser intensity with AC160 cantilever | 2.89
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from -Y to +Y (Y scan) | 34.18
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from -Y to +Y (X scan) | 39.06
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from -X to +X (Y scan) | 53.71
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from -X to +X (X scan) | 102.54
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from +Y to -Y (Y scan) | 78.13
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from +Y to -Y (X scan) | 39.06
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from +X to -X (Y scan) | 92.77
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Image offset after moving from +X to -X (X scan) | 161.13
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Full range of Z Height during Z Scan sweep | 14
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Full range of Z Height during Z Scan sweep | 65.0877
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Full range of Z Detector during Z Scan sweep | 13.21
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Full range of Y Detector Raw during Y Scan sweep | 15.2
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Full range of Y Detector during Y Scan sweep | 67.5942
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | Full range of X Detector Raw during X Scan sweep | 14.81
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | A－B difference during 3 times run test | 0.24
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | A＋B difference during 3 times run test | 0.15
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | 50~200nm Step Height Sample (10 Repeats) | 183.4
| Si En(Qingdao) Intergrated Circuits Co.,Ltd | 2023-06-29 | 200~500nm Step Height Sample (10 Repeats) | 461
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Linearity by Z scanner position | 0.69
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Linearity by Z scanner position | 0.6
| CXMT Beijing #3 | 2023-06-28 | Z Linearity by Z scanner position | 0.02
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Linearity by 5 different sample height | 1.003
| OmniVision Semiconductor Shanghai #2_1 | 2023-06-28 | Z Linearity by 5 different sample height | 1.0016
| CXMT Beijing #3 | 2023-06-28 | Z Linearity by 5 different sample height | 1.005
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Height noise | 0.037
| CXMT Beijing #3 | 2023-06-28 | Z Height noise | 0.038
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7842
| CXMT Beijing #3 | 2023-06-28 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.95
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.433
| CXMT Beijing #3 | 2023-06-28 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.96
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8056
| CXMT Beijing #3 | 2023-06-28 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.38
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.8181
| CXMT Beijing #3 | 2023-06-28 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.7
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y2 Detector noise (XY Servo on) | 0.331
| CXMT Beijing #3 | 2023-06-28 | Y2 Detector noise (XY Servo on) | 0.305
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 283.64
| CXMT Beijing #3 | 2023-06-28 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 276.631
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y1 Detector noise (XY Servo on) | 0.354
| CXMT Beijing #3 | 2023-06-28 | Y1 Detector noise (XY Servo on) | 0.381
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 269.952
| CXMT Beijing #3 | 2023-06-28 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 288.1858
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 189.56
| CXMT Beijing #3 | 2023-06-28 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 190.121
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y Ringing(Rpv) | 0.267
| CXMT Beijing #3 | 2023-06-28 | Y Ringing(Rpv) | 0.213
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y Non-linearity error | 2.83
| CXMT Beijing #3 | 2023-06-28 | Y Non-linearity error | 2.88
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y Detector noise (XY Servo on) | 0.261
| CXMT Beijing #3 | 2023-06-28 | Y Detector noise (XY Servo on) | 0.258
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Repeatability @1st position | 0.03
| CXMT Beijing #3 | 2023-06-28 | Y axis, Repeatability @1st position | 0.03
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Repeatability @ 4th position | 0.03
| CXMT Beijing #3 | 2023-06-28 | Y axis, Repeatability @ 4th position | 0.04
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Repeatability @ 3rd position | 0.04
| CXMT Beijing #3 | 2023-06-28 | Y axis, Repeatability @ 3rd position | 0.04
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Repeatability @ 2nd position | 0.04
| CXMT Beijing #3 | 2023-06-28 | Y axis, Repeatability @ 2nd position | 0.04
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Accuracy @1st position | 0.67
| CXMT Beijing #3 | 2023-06-28 | Y axis, Accuracy @1st position | 1.09
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Accuracy @ 4th position | 0.54
| CXMT Beijing #3 | 2023-06-28 | Y axis, Accuracy @ 4th position | 1.05
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Accuracy @ 3rd position | 0.51
| CXMT Beijing #3 | 2023-06-28 | Y axis, Accuracy @ 3rd position | 0.98
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis, Accuracy @ 2nd position | 0.69
| CXMT Beijing #3 | 2023-06-28 | Y axis, Accuracy @ 2nd position | 0.92
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Y axis OPM | 0.268
| CXMT Beijing #3 | 2023-06-28 | Y axis OPM | 1.427
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X2 Detector noise (XY Servo on) | 0.37
| CXMT Beijing #3 | 2023-06-28 | X2 Detector noise (XY Servo on) | 0.34
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 291.39
| CXMT Beijing #3 | 2023-06-28 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 276.8298
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X1 Detector noise (XY Servo on) | 0.423
| CXMT Beijing #3 | 2023-06-28 | X1 Detector noise (XY Servo on) | 0.334
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 272.918
| CXMT Beijing #3 | 2023-06-28 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 273.874
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 182.94
| CXMT Beijing #3 | 2023-06-28 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 190.384
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X Ringing(Rpv) | 0.335
| CXMT Beijing #3 | 2023-06-28 | X Ringing(Rpv) | 0.159
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X Non-linearity error | 3.31
| CXMT Beijing #3 | 2023-06-28 | X Non-linearity error | 1.48
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X Detector noise (XY Servo on) | 0.315
| CXMT Beijing #3 | 2023-06-28 | X Detector noise (XY Servo on) | 0.264
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Repeatability @1st position | 0.03
| CXMT Beijing #3 | 2023-06-28 | X axis, Repeatability @1st position | 0.02
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Repeatability @ 4th position | 0.02
| CXMT Beijing #3 | 2023-06-28 | X axis, Repeatability @ 4th position | 0.02
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Repeatability @ 3rd position | 0.03
| CXMT Beijing #3 | 2023-06-28 | X axis, Repeatability @ 3rd position | 0.03
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Repeatability @ 2nd position | 0.02
| CXMT Beijing #3 | 2023-06-28 | X axis, Repeatability @ 2nd position | 0.03
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Accuracy @1st position | 0.18
| CXMT Beijing #3 | 2023-06-28 | X axis, Accuracy @1st position | 0.23
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Accuracy @ 4th position | 0.22
| CXMT Beijing #3 | 2023-06-28 | X axis, Accuracy @ 4th position | 0.44
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Accuracy @ 3rd position | 0.15
| CXMT Beijing #3 | 2023-06-28 | X axis, Accuracy @ 3rd position | 0.29
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis, Accuracy @ 2nd position | 0.16
| CXMT Beijing #3 | 2023-06-28 | X axis, Accuracy @ 2nd position | 0.21
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | X axis OPM | 1.935
| CXMT Beijing #3 | 2023-06-28 | X axis OPM | 0.625
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Under 50nm Step Height Sample (10 Repeats) | 18.16
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Under 50nm Step Height Sample (10 Repeats) | 18.269
| CXMT Beijing #3 | 2023-06-28 | Under 50nm Step Height Sample (10 Repeats) | 18.385
| CXMT Beijing #3 | 2023-06-28 | Torque Profile (50mm, 20um/s) | 3.3
| CXMT Beijing #3 | 2023-06-28 | Torque Profile (50mm, 200um/s) | 2.193
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | StrokeYMin | -153787
| CXMT Beijing #3 | 2023-06-28 | StrokeYMin | -153948
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | StrokeYMax | 150497.1
| CXMT Beijing #3 | 2023-06-28 | StrokeYMax | 150607
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | StrokeY | 304284.3
| CXMT Beijing #3 | 2023-06-28 | StrokeY | 304555
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | StrokeXMin | -204532
| CXMT Beijing #3 | 2023-06-28 | StrokeXMin | -203808
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | StrokeXMax | 203568.6
| CXMT Beijing #3 | 2023-06-28 | StrokeXMax | 204107
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | StrokeX | 408100.3
| CXMT Beijing #3 | 2023-06-28 | StrokeX | 407915
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.79
| CXMT Beijing #3 | 2023-06-28 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.538
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | RMS of ZeroScan | 0.047
| CXMT Beijing #3 | 2023-06-28 | RMS of ZeroScan | 0.048
| CXMT Beijing #3 | 2023-06-28 | PosMin | -26.042
| CXMT Beijing #3 | 2023-06-28 | PosMax | 26.035
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 7e-05
| CXMT Beijing #3 | 2023-06-28 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00026
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | OpticalAxisShiftPerFocusY | -0.0054
| CXMT Beijing #3 | 2023-06-28 | OpticalAxisShiftPerFocusY | -0.0025
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00016
| CXMT Beijing #3 | 2023-06-28 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00046
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | OpticalAxisShiftPerFocusX | -0.0017
| CXMT Beijing #3 | 2023-06-28 | OpticalAxisShiftPerFocusX | -0.0005
| CXMT Beijing #3 | 2023-06-28 | OPM Repeatability @5mm | 0.701
| CXMT Beijing #3 | 2023-06-28 | OPM Repeatability @25mm | 5.426
| CXMT Beijing #3 | 2023-06-28 | OPM Repeatability @1mm | 0.551
| CXMT Beijing #3 | 2023-06-28 | OPM Repeatability @10mm | 1.213
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | NCM driving% with AC160TS cantilever | 19.2
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | NCM driving% with AC160TS cantilever | 2.86
| CXMT Beijing #3 | 2023-06-28 | NCM driving% with AC160TS cantilever | 11.17
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Min. Vibration Transmissbility (~100Hz) | 0.003
| CXMT Beijing #3 | 2023-06-28 | Min. Vibration Transmissbility (~100Hz) | 0.001
| CXMT Beijing #3 | 2023-06-28 | MaxLength | 51
| CXMT Beijing #3 | 2023-06-28 | Max OPM @5mm | 21.785
| CXMT Beijing #3 | 2023-06-28 | Max OPM @25mm | 172.862
| CXMT Beijing #3 | 2023-06-28 | Max OPM @1mm | 5.738
| CXMT Beijing #3 | 2023-06-28 | Max OPM @10mm | 42.709
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Lateral(C-D) sensitivity | 5.9
| CXMT Beijing #3 | 2023-06-28 | Lateral(C-D) sensitivity | 3.7
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Lateral(C-D) noise | 0.19
| CXMT Beijing #3 | 2023-06-28 | Lateral(C-D) noise | 0.153
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Laser intensity with AC160 cantilever | 3
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Laser intensity with AC160 cantilever | 3
| CXMT Beijing #3 | 2023-06-28 | Laser intensity with AC160 cantilever | 3.18
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from -Y to +Y (Y scan) | 76.9
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from -Y to +Y (Y scan) | 82.76
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from -Y to +Y (X scan) | 3.66
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from -Y to +Y (X scan) | 21.97
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from -X to +X (Y scan) | 73.24
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from -X to +X (Y scan) | 61.52
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from -X to +X (X scan) | 100.34
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from -X to +X (X scan) | 76.9
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from +Y to -Y (Y scan) | 60.79
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from +Y to -Y (Y scan) | 34.42
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from +Y to -Y (X scan) | 27.1
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from +Y to -Y (X scan) | 44.68
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from +X to -X (Y scan) | 40.28
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from +X to -X (Y scan) | 61.52
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Image offset after moving from +X to -X (X scan) | 82.03
| CXMT Beijing #3 | 2023-06-28 | Image offset after moving from +X to -X (X scan) | 13.92
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Z Height during Z Scan sweep | 16.3
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Z Height during Z Scan sweep | 42.2
| CXMT Beijing #3 | 2023-06-28 | Full range of Z Height during Z Scan sweep | 16.6
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Z Detector during Z Scan sweep | 12.2
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Z Detector during Z Scan sweep | 13.61
| CXMT Beijing #3 | 2023-06-28 | Full range of Z Detector during Z Scan sweep | 12.8
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Y2 Detector Raw during Y Scan sweep | 13.01
| CXMT Beijing #3 | 2023-06-28 | Full range of Y2 Detector Raw during Y Scan sweep | 13.82
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Y2 Detector during Y Scan sweep | 210.2865
| CXMT Beijing #3 | 2023-06-28 | Full range of Y2 Detector during Y Scan sweep | 204.7083
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Y1 Detector Raw during Y Scan sweep | 13.22
| CXMT Beijing #3 | 2023-06-28 | Full range of Y1 Detector Raw during Y Scan sweep | 13.99
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of Y1 Detector during Y Scan sweep | 211.353
| CXMT Beijing #3 | 2023-06-28 | Full range of Y1 Detector during Y Scan sweep | 218.236
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of X2 Detector Raw during X Scan sweep | 12.79
| CXMT Beijing #3 | 2023-06-28 | Full range of X2 Detector Raw during X Scan sweep | 13.8
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of X2 Detector during X Scan sweep | 205.772
| CXMT Beijing #3 | 2023-06-28 | Full range of X2 Detector during X Scan sweep | 210.8437
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of X1 Detector Raw during X Scan sweep | 12.58
| CXMT Beijing #3 | 2023-06-28 | Full range of X1 Detector Raw during X Scan sweep | 13.56
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | Full range of X1 Detector during X Scan sweep | 215.4347
| CXMT Beijing #3 | 2023-06-28 | Full range of X1 Detector during X Scan sweep | 206.545
| CXMT Beijing #3 | 2023-06-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| CXMT Beijing #3 | 2023-06-28 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.38
| CXMT Beijing #3 | 2023-06-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.38
| CXMT Beijing #3 | 2023-06-28 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| CXMT Beijing #3 | 2023-06-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| CXMT Beijing #3 | 2023-06-28 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.39
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.15
| CXMT Beijing #3 | 2023-06-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.33
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| CXMT Beijing #3 | 2023-06-28 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | A－B difference during 3 times run test | 0.66
| CXMT Beijing #3 | 2023-06-28 | A－B difference during 3 times run test | 0.42
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | A＋B difference during 3 times run test | 0.07
| CXMT Beijing #3 | 2023-06-28 | A＋B difference during 3 times run test | 0.25
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | 50~200nm Step Height Sample (10 Repeats) | 183.07
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | 50~200nm Step Height Sample (10 Repeats) | 182.684
| CXMT Beijing #3 | 2023-06-28 | 50~200nm Step Height Sample (10 Repeats) | 183.724
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | 200~500nm Step Height Sample (10 Repeats) | 462.68
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | 200~500nm Step Height Sample (10 Repeats) | 462.187
| CXMT Beijing #3 | 2023-06-28 | 200~500nm Step Height Sample (10 Repeats) | 462.501
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 3.4
| CXMT Beijing #3 | 2023-06-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.8
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3.9
| CXMT Beijing #3 | 2023-06-28 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 3
| OmniVision Semiconductor Shanghai #2 | 2023-06-28 | %GRR w/ roughness of non-patterened wafer | 5.1
| CXMT Beijing #3 | 2023-06-28 | %GRR w/ roughness of non-patterened wafer | 5.7
| Polytechnic University of Milan #2 | 2023-06-28 | Z Scanner position at -3μm | 0.0033
| 목포대학교_첨단재료공학과 | 2023-06-28 | Z Scanner position at -3μm | 0.0042
| Polytechnic University of Milan #2 | 2023-06-28 | Z Scanner position at +3μm | 0.0095
| 목포대학교_첨단재료공학과 | 2023-06-28 | Z Scanner position at +3μm | 0.0129
| Polytechnic University of Milan #2 | 2023-06-28 | Z Linearity by 3 different sample height | 0.996
| 목포대학교_첨단재료공학과 | 2023-06-28 | Z Linearity by 3 different sample height | 1.003
| Polytechnic University of Milan #2 | 2023-06-28 | Z Height noise | 0.017
| 목포대학교_첨단재료공학과 | 2023-06-28 | Z Height noise | 0.029
| 목포대학교_첨단재료공학과 | 2023-06-28 | Y Scan Stroke | 55
| Polytechnic University of Milan #2 | 2023-06-28 | Y Detector noise (XY Servo on) | 0.28
| 목포대학교_첨단재료공학과 | 2023-06-28 | Y Detector noise (XY Servo on) | 0.37
| Polytechnic University of Milan #2 | 2023-06-28 | Y Detector noise (XY Servo off) | 0.26
| 목포대학교_첨단재료공학과 | 2023-06-28 | Y Detector noise (XY Servo off) | 0.4
| Polytechnic University of Milan #2 | 2023-06-28 | Y axis OPM | 0.74
| 목포대학교_첨단재료공학과 | 2023-06-28 | Y axis OPM | 0.12
| 목포대학교_첨단재료공학과 | 2023-06-28 | X Scan Stroke | 55
| Polytechnic University of Milan #2 | 2023-06-28 | X Detector noise (XY Servo on) | 0.28
| 목포대학교_첨단재료공학과 | 2023-06-28 | X Detector noise (XY Servo on) | 0.4
| Polytechnic University of Milan #2 | 2023-06-28 | X Detector noise (XY Servo off) | 0.25
| 목포대학교_첨단재료공학과 | 2023-06-28 | X Detector noise (XY Servo off) | 0.43
| Polytechnic University of Milan #2 | 2023-06-28 | X axis OPM | 0.76
| 목포대학교_첨단재료공학과 | 2023-06-28 | X axis OPM | 0.18
| Polytechnic University of Milan #2 | 2023-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Polytechnic University of Milan #2 | 2023-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| 목포대학교_첨단재료공학과 | 2023-06-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 목포대학교_첨단재료공학과 | 2023-06-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Polytechnic University of Milan #2 | 2023-06-28 | NCM driving% with AC160TS cantilever | 0.83
| 목포대학교_첨단재료공학과 | 2023-06-28 | NCM driving% with AC160TS cantilever | 1.57
| Polytechnic University of Milan #2 | 2023-06-28 | NCM Amplitude noise | 14.11
| 목포대학교_첨단재료공학과 | 2023-06-28 | NCM Amplitude noise | 17.91
| Polytechnic University of Milan #2 | 2023-06-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.5
| 목포대학교_첨단재료공학과 | 2023-06-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| Polytechnic University of Milan #2 | 2023-06-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.9
| 목포대학교_첨단재료공학과 | 2023-06-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.9
| Polytechnic University of Milan #2 | 2023-06-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102.5
| 목포대학교_첨단재료공학과 | 2023-06-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 95.9
| Polytechnic University of Milan #2 | 2023-06-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.7
| 목포대학교_첨단재료공학과 | 2023-06-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 99.6
| Polytechnic University of Milan #2 | 2023-06-28 | Lateral(C-D) sensitivity | 6.9
| 목포대학교_첨단재료공학과 | 2023-06-28 | Lateral(C-D) sensitivity | 16.6
| Polytechnic University of Milan #2 | 2023-06-28 | Lateral(C-D) noise | 0.071
| Polytechnic University of Milan #2 | 2023-06-28 | Full range of Z Height during Z Scan sweep | 16.6
| 목포대학교_첨단재료공학과 | 2023-06-28 | Full range of Z Height during Z Scan sweep | 15.9
| Polytechnic University of Milan #2 | 2023-06-28 | Full range of Z Detector during Z Scan sweep | 11.4
| 목포대학교_첨단재료공학과 | 2023-06-28 | Full range of Z Detector during Z Scan sweep | 11
| Polytechnic University of Milan #2 | 2023-06-28 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| 목포대학교_첨단재료공학과 | 2023-06-28 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Polytechnic University of Milan #2 | 2023-06-28 | Full range of Y Detector during Y Scan sweep | 73.7
| 목포대학교_첨단재료공학과 | 2023-06-28 | Full range of Y Detector during Y Scan sweep | 67.1
| Polytechnic University of Milan #2 | 2023-06-28 | Full range of X Detector Raw during X Scan sweep | 14
| 목포대학교_첨단재료공학과 | 2023-06-28 | Full range of X Detector Raw during X Scan sweep | 13.3
| 목포대학교_첨단재료공학과 | 2023-06-28 | Full range of X Detector during X Scan sweep | 66.1
| Polytechnic University of Milan #2 | 2023-06-28 | A-B noise | 7.11
| 목포대학교_첨단재료공학과 | 2023-06-28 | A-B noise | 11.32
| 삼성전자 | 2023-06-27 | Z Linearity by 3 different sample height | 0.998
| 삼성전자 | 2023-06-27 | Z Height noise | 0.022
| 삼성전자 | 2023-06-27 | Y Detector noise (XY Servo on) | 0.29
| 삼성전자 | 2023-06-27 | Y Detector noise (XY Servo off) | 0.29
| 삼성전자 | 2023-06-27 | Y axis OPM | 0.6
| 삼성전자 | 2023-06-27 | X Detector noise (XY Servo on) | 0.28
| 삼성전자 | 2023-06-27 | X Detector noise (XY Servo off) | 0.28
| 삼성전자 | 2023-06-27 | X axis OPM | 1.31
| 삼성전자 | 2023-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| 삼성전자 | 2023-06-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 삼성전자 | 2023-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| 삼성전자 | 2023-06-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.2
| 삼성전자 | 2023-06-27 | NCM driving% with AC160TS cantilever | 11.543
| 삼성전자 | 2023-06-27 | NCM Amplitude noise | 10.23
| 삼성전자 | 2023-06-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.329
| 삼성전자 | 2023-06-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.851
| 삼성전자 | 2023-06-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 158.5
| 삼성전자 | 2023-06-27 | Full range of Z Height during Z Scan sweep | 16.1
| 삼성전자 | 2023-06-27 | Lateral(C-D) sensitivity | 6.7
| 삼성전자 | 2023-06-27 | Lateral(C-D) noise | 0.121
| 삼성전자 | 2023-06-27 | Full range of Z Detector during Z Scan sweep | 13
| 삼성전자 | 2023-06-27 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| 삼성전자 | 2023-06-27 | Full range of Y Detector during Y Scan sweep | 115
| 삼성전자 | 2023-06-27 | Full range of X Detector Raw during X Scan sweep | 14.5
| 삼성전자 | 2023-06-27 | Full range of X Detector during X Scan sweep | 112.5
| 삼성전자 | 2023-06-27 | Approach position at -3μm | 0.0059
| 삼성전자 | 2023-06-27 | Approach position at +3μm | 0.0136
| 삼성전자 | 2023-06-27 | A-B noise | 15.25
| 안동대학교_첨단재료공학과 | 2023-06-26 | Z Scanner position at -3μm | 0.0023
| 안동대학교_첨단재료공학과 | 2023-06-26 | Z Scanner position at +3μm | 0.0055
| 안동대학교_첨단재료공학과 | 2023-06-26 | Z Linearity by 3 different sample height | 0.991
| 안동대학교_첨단재료공학과 | 2023-06-26 | Z Height noise | 0.027
| 안동대학교_첨단재료공학과 | 2023-06-26 | Y Detector noise (XY Servo on) | 0.27
| 안동대학교_첨단재료공학과 | 2023-06-26 | Y Detector noise (XY Servo off) | 0.26
| 안동대학교_첨단재료공학과 | 2023-06-26 | Y axis OPM | 0.78
| 안동대학교_첨단재료공학과 | 2023-06-26 | X Detector noise (XY Servo on) | 0.27
| 안동대학교_첨단재료공학과 | 2023-06-26 | X Detector noise (XY Servo off) | 0.27
| 안동대학교_첨단재료공학과 | 2023-06-26 | X axis OPM | 0.94
| 안동대학교_첨단재료공학과 | 2023-06-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 안동대학교_첨단재료공학과 | 2023-06-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 안동대학교_첨단재료공학과 | 2023-06-26 | NCM driving% with AC160TS cantilever | 1.6
| 안동대학교_첨단재료공학과 | 2023-06-26 | NCM Amplitude noise | 12.21
| 안동대학교_첨단재료공학과 | 2023-06-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.9
| 안동대학교_첨단재료공학과 | 2023-06-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.6
| 안동대학교_첨단재료공학과 | 2023-06-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.4
| 안동대학교_첨단재료공학과 | 2023-06-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.7
| 안동대학교_첨단재료공학과 | 2023-06-26 | Lateral(C-D) sensitivity | 15.7
| 안동대학교_첨단재료공학과 | 2023-06-26 | Lateral(C-D) noise | 0.06
| 안동대학교_첨단재료공학과 | 2023-06-26 | Full range of Z Height during Z Scan sweep | 15.8
| 안동대학교_첨단재료공학과 | 2023-06-26 | Full range of Z Detector during Z Scan sweep | 11.4
| 안동대학교_첨단재료공학과 | 2023-06-26 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| 안동대학교_첨단재료공학과 | 2023-06-26 | Full range of Y Detector during Y Scan sweep | 72.6
| 안동대학교_첨단재료공학과 | 2023-06-26 | Full range of X Detector Raw during X Scan sweep | 14.2
| 안동대학교_첨단재료공학과 | 2023-06-26 | Full range of X Detector during X Scan sweep | 71.1
| 안동대학교_첨단재료공학과 | 2023-06-26 | A-B noise | 13.41
| Eindhoven University of Technology | 2023-06-22 | Z Scanner position at -3μm | 0.0053
| Eindhoven University of Technology | 2023-06-22 | Z Scanner position at +3μm | 0.0143
| Eindhoven University of Technology | 2023-06-22 | Z Linearity by 3 different sample height | 1.002
| Eindhoven University of Technology | 2023-06-22 | Z Height noise | 0.019
| Eindhoven University of Technology | 2023-06-22 | Y Detector noise (XY Servo on) | 0.25
| Eindhoven University of Technology | 2023-06-22 | Y Detector noise (XY Servo off) | 0.24
| Eindhoven University of Technology | 2023-06-22 | Y axis OPM | 0.76
| Eindhoven University of Technology | 2023-06-22 | X Detector noise (XY Servo on) | 0.25
| Eindhoven University of Technology | 2023-06-22 | X Detector noise (XY Servo off) | 0.25
| Eindhoven University of Technology | 2023-06-22 | X axis OPM | 0.66
| Eindhoven University of Technology | 2023-06-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Eindhoven University of Technology | 2023-06-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| Eindhoven University of Technology | 2023-06-22 | NCM driving% with AC160TS cantilever | 1.52
| Eindhoven University of Technology | 2023-06-22 | NCM Amplitude noise | 12.55
| Eindhoven University of Technology | 2023-06-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.389
| Eindhoven University of Technology | 2023-06-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.032
| Eindhoven University of Technology | 2023-06-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105.8
| Eindhoven University of Technology | 2023-06-22 | Lateral(C-D) sensitivity | 8.9
| Eindhoven University of Technology | 2023-06-22 | Lateral(C-D) noise | 0.098
| Eindhoven University of Technology | 2023-06-22 | Full range of Z Height during Z Scan sweep | 16.1
| Eindhoven University of Technology | 2023-06-22 | Full range of Z Detector during Z Scan sweep | 11.6
| Eindhoven University of Technology | 2023-06-22 | Full range of Y Detector Raw during Y Scan sweep | 14
| Eindhoven University of Technology | 2023-06-22 | Full range of Y Detector during Y Scan sweep | 79.1
| Eindhoven University of Technology | 2023-06-22 | Full range of X Detector Raw during X Scan sweep | 14.3
| Eindhoven University of Technology | 2023-06-22 | Full range of X Detector during X Scan sweep | 77.7
| Eindhoven University of Technology | 2023-06-22 | A-B noise | 7.6
| 에이프로세미콘 | 2023-06-20 | Z Scanner position at -3μm | 0.008
| 에이프로세미콘 | 2023-06-20 | Z Scanner position at +3μm | 0.0072
| 에이프로세미콘 | 2023-06-20 | Z Linearity by 3 different sample height | 0.992
| 에이프로세미콘 | 2023-06-20 | Z Height noise | 0.018
| 에이프로세미콘 | 2023-06-20 | Y2 Detector noise (XY Servo on) | 0.18
| 에이프로세미콘 | 2023-06-20 | Y2 Detector noise (XY Servo off) | 0.18
| 에이프로세미콘 | 2023-06-20 | Y1 Detector noise (XY Servo on) | 0.18
| 에이프로세미콘 | 2023-06-20 | Y1 Detector noise (XY Servo off) | 0.17
| 에이프로세미콘 | 2023-06-20 | Y Scan Stroke | 173.6
| 에이프로세미콘 | 2023-06-20 | Y Detector noise (XY Servo on) | 0.13
| 에이프로세미콘 | 2023-06-20 | Y Detector noise (XY Servo off) | 0.13
| 에이프로세미콘 | 2023-06-20 | Y axis OPM | 0.7
| 에이프로세미콘 | 2023-06-20 | X2 Detector noise (XY Servo on) | 0.2
| 에이프로세미콘 | 2023-06-20 | X2 Detector noise (XY Servo off) | 0.18
| 에이프로세미콘 | 2023-06-20 | X1 Detector noise (XY Servo on) | 0.2
| 에이프로세미콘 | 2023-06-20 | X1 Detector noise (XY Servo off) | 0.18
| 에이프로세미콘 | 2023-06-20 | X Scan Stroke | 165.6
| 에이프로세미콘 | 2023-06-20 | -X Direction software Limit | -102644
| 에이프로세미콘 | 2023-06-20 | -X Direction software Limit | -102754
| 에이프로세미콘 | 2023-06-20 | X Detector noise (XY Servo on) | 0.15
| 에이프로세미콘 | 2023-06-20 | X Detector noise (XY Servo off) | 0.13
| 에이프로세미콘 | 2023-06-20 | X axis OPM | 1.42
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of Wafer Zeroscan (Top) | 0.34
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of Wafer Zeroscan (Right) | 0.3
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of Wafer Zeroscan (Left) | 0.32
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of Wafer ZeroScan (Center) | 0.3
| 에이프로세미콘 | 2023-06-20 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.32
| 에이프로세미콘 | 2023-06-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| 에이프로세미콘 | 2023-06-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| 에이프로세미콘 | 2023-06-20 | NCM driving% with AC160TS cantilever | 1.17
| 에이프로세미콘 | 2023-06-20 | NCM Amplitude noise | 12.37
| 에이프로세미콘 | 2023-06-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.6
| 에이프로세미콘 | 2023-06-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.8
| 에이프로세미콘 | 2023-06-20 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.8
| 에이프로세미콘 | 2023-06-20 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 244.7
| 에이프로세미콘 | 2023-06-20 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 233.7
| 에이프로세미콘 | 2023-06-20 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 234.2
| 에이프로세미콘 | 2023-06-20 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.04
| 에이프로세미콘 | 2023-06-20 | Lateral(C-D) sensitivity | 9.5
| 에이프로세미콘 | 2023-06-20 | Lateral(C-D) noise | 0.121
| 에이프로세미콘 | 2023-06-20 | Full range of Z Height during Z Scan sweep | 16.7
| 에이프로세미콘 | 2023-06-20 | Full range of Z Detector during Z Scan sweep | 12.5
| 에이프로세미콘 | 2023-06-20 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| 에이프로세미콘 | 2023-06-20 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| 에이프로세미콘 | 2023-06-20 | Full range of Y Detector during Y Scan sweep | 176.5
| 에이프로세미콘 | 2023-06-20 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| 에이프로세미콘 | 2023-06-20 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| 에이프로세미콘 | 2023-06-20 | Full range of X Detector during X Scan sweep | 168.4
| 에이프로세미콘 | 2023-06-20 | A-B noise | 6.27
| University of Bergamo | 2023-06-15 | Z Scanner position at -3μm | 0.0008
| AIST | 2023-06-15 | Z Scanner position at -3μm | -0.0103
| University of Bergamo | 2023-06-15 | Z Scanner position at +3μm | 0.0012
| AIST | 2023-06-15 | Z Scanner position at +3μm | 0.0027
| University of Bergamo | 2023-06-15 | Z Linearity by 3 different sample height | 0.986
| AIST | 2023-06-15 | Z Linearity by 3 different sample height | 0.993
| University of Bergamo | 2023-06-15 | Z Height noise | 0.02
| AIST | 2023-06-15 | Z Height noise | 0.022
| AIST | 2023-06-15 | Y Scan Stroke | 55
| University of Bergamo | 2023-06-15 | Y Detector noise (XY Servo on) | 0.34
| AIST | 2023-06-15 | Y Detector noise (XY Servo on) | 0.33
| University of Bergamo | 2023-06-15 | Y Detector noise (XY Servo off) | 0.34
| AIST | 2023-06-15 | Y Detector noise (XY Servo off) | 0.33
| University of Bergamo | 2023-06-15 | Y axis OPM | 0.55
| AIST | 2023-06-15 | Y axis OPM | 0.41
| AIST | 2023-06-15 | X Scan Stroke | 55
| University of Bergamo | 2023-06-15 | X Detector noise (XY Servo on) | 0.33
| AIST | 2023-06-15 | X Detector noise (XY Servo on) | 0.32
| University of Bergamo | 2023-06-15 | X Detector noise (XY Servo off) | 0.36
| AIST | 2023-06-15 | X Detector noise (XY Servo off) | 0.32
| University of Bergamo | 2023-06-15 | X axis OPM | 0.55
| AIST | 2023-06-15 | X axis OPM | 0.23
| University of Bergamo | 2023-06-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| AIST | 2023-06-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| University of Bergamo | 2023-06-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| AIST | 2023-06-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| AIST | 2023-06-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| University of Bergamo | 2023-06-15 | NCM driving% with AC160TS cantilever | 5.15
| AIST | 2023-06-15 | NCM driving% with AC160TS cantilever | 1.72
| University of Bergamo | 2023-06-15 | NCM Amplitude noise | 14.47
| AIST | 2023-06-15 | NCM Amplitude noise | 15.2
| University of Bergamo | 2023-06-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.7
| AIST | 2023-06-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| University of Bergamo | 2023-06-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.7
| AIST | 2023-06-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.3
| University of Bergamo | 2023-06-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.8
| AIST | 2023-06-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 97.2
| University of Bergamo | 2023-06-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 167.2
| AIST | 2023-06-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 97.6
| University of Bergamo | 2023-06-15 | Lateral(C-D) sensitivity | 13
| AIST | 2023-06-15 | Lateral(C-D) sensitivity | 8.5
| University of Bergamo | 2023-06-15 | Lateral(C-D) noise | 0.097
| AIST | 2023-06-15 | Lateral(C-D) noise | 0.106
| University of Bergamo | 2023-06-15 | Full range of Z Height during Z Scan sweep | 17.1
| AIST | 2023-06-15 | Full range of Z Height during Z Scan sweep | 16.1
| University of Bergamo | 2023-06-15 | Full range of Z Detector during Z Scan sweep | 12.3
| AIST | 2023-06-15 | Full range of Z Detector during Z Scan sweep | 12.5
| University of Bergamo | 2023-06-15 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| AIST | 2023-06-15 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| University of Bergamo | 2023-06-15 | Full range of Y Detector during Y Scan sweep | 116.8
| AIST | 2023-06-15 | Full range of Y Detector during Y Scan sweep | 71.3
| University of Bergamo | 2023-06-15 | Full range of X Detector Raw during X Scan sweep | 14.1
| AIST | 2023-06-15 | Full range of X Detector Raw during X Scan sweep | 14.5
| University of Bergamo | 2023-06-15 | Full range of X Detector during X Scan sweep | 117.6
| AIST | 2023-06-15 | Full range of X Detector during X Scan sweep | 70.7
| University of Bergamo | 2023-06-15 | A-B noise | 11.16
| AIST | 2023-06-15 | A-B noise | 10.74
| Ruhr-University of Bochum | 2023-06-14 | Z Scanner position at -3μm | 0.0017
| PSI DEMO Replacement | 2023-06-14 | Z Scanner position at -3μm | 0.0019
| Ruhr-University of Bochum | 2023-06-14 | Z Scanner position at +3μm | 0.0048
| PSI DEMO Replacement | 2023-06-14 | Z Scanner position at +3μm | 0.002
| Ruhr-University of Bochum | 2023-06-14 | Z Linearity by 3 different sample height | 1.007
| PSI DEMO Replacement | 2023-06-14 | Z Linearity by 3 different sample height | 1.002
| Ruhr-University of Bochum | 2023-06-14 | Z Height noise | 0.033
| PSI DEMO Replacement | 2023-06-14 | Z Height noise | 0.03
| Ruhr-University of Bochum | 2023-06-14 | Y2 Detector noise (XY Servo on) | 0.23
| Ruhr-University of Bochum | 2023-06-14 | Y2 Detector noise (XY Servo off) | 0.23
| Ruhr-University of Bochum | 2023-06-14 | Y1 Detector noise (XY Servo on) | 0.23
| Ruhr-University of Bochum | 2023-06-14 | Y1 Detector noise (XY Servo off) | 0.22
| Ruhr-University of Bochum | 2023-06-14 | Y Scan Stroke | 177.5
| Ruhr-University of Bochum | 2023-06-14 | Y Detector noise (XY Servo on) | 0.16
| PSI DEMO Replacement | 2023-06-14 | Y Detector noise (XY Servo on) | 0.32
| Ruhr-University of Bochum | 2023-06-14 | Y Detector noise (XY Servo off) | 0.17
| PSI DEMO Replacement | 2023-06-14 | Y Detector noise (XY Servo off) | 0.29
| Ruhr-University of Bochum | 2023-06-14 | Y axis OPM | 0.5
| PSI DEMO Replacement | 2023-06-14 | Y axis OPM | 0.45
| Ruhr-University of Bochum | 2023-06-14 | X2 Detector noise (XY Servo on) | 0.22
| Ruhr-University of Bochum | 2023-06-14 | X2 Detector noise (XY Servo off) | 0.22
| Ruhr-University of Bochum | 2023-06-14 | X1 Detector noise (XY Servo on) | 0.23
| Ruhr-University of Bochum | 2023-06-14 | X1 Detector noise (XY Servo off) | 0.23
| Ruhr-University of Bochum | 2023-06-14 | X Scan Stroke | 167.1
| Ruhr-University of Bochum | 2023-06-14 | -X Direction software Limit | -101530
| Ruhr-University of Bochum | 2023-06-14 | -X Direction software Limit | -102779
| Ruhr-University of Bochum | 2023-06-14 | X Detector noise (XY Servo on) | 0.17
| PSI DEMO Replacement | 2023-06-14 | X Detector noise (XY Servo on) | 0.31
| Ruhr-University of Bochum | 2023-06-14 | X Detector noise (XY Servo off) | 0.17
| PSI DEMO Replacement | 2023-06-14 | X Detector noise (XY Servo off) | 0.29
| Ruhr-University of Bochum | 2023-06-14 | X axis OPM | 1.2
| PSI DEMO Replacement | 2023-06-14 | X axis OPM | 0.67
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| PSI DEMO Replacement | 2023-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| PSI DEMO Replacement | 2023-06-14 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.49
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of Wafer Zeroscan (Top) | 0.35
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of Wafer Zeroscan (Right) | 0.27
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of Wafer Zeroscan (Left) | 0.29
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of Wafer ZeroScan (Center) | 0.27
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.32
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Ruhr-University of Bochum | 2023-06-14 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| Ruhr-University of Bochum | 2023-06-14 | NCM driving% with AC160TS cantilever | 0.753
| PSI DEMO Replacement | 2023-06-14 | NCM driving% with AC160TS cantilever | 3.3
| Ruhr-University of Bochum | 2023-06-14 | NCM Amplitude noise | 16.82
| PSI DEMO Replacement | 2023-06-14 | NCM Amplitude noise | 12.56
| Ruhr-University of Bochum | 2023-06-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.259
| PSI DEMO Replacement | 2023-06-14 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.3
| Ruhr-University of Bochum | 2023-06-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.914
| PSI DEMO Replacement | 2023-06-14 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.8
| Ruhr-University of Bochum | 2023-06-14 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 249.6
| Ruhr-University of Bochum | 2023-06-14 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 245.7
| PSI DEMO Replacement | 2023-06-14 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 176.7
| Ruhr-University of Bochum | 2023-06-14 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 228.7
| Ruhr-University of Bochum | 2023-06-14 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 233.6
| PSI DEMO Replacement | 2023-06-14 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 172.7
| Ruhr-University of Bochum | 2023-06-14 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Ruhr-University of Bochum | 2023-06-14 | Lateral(C-D) sensitivity | 16.159
| PSI DEMO Replacement | 2023-06-14 | Lateral(C-D) sensitivity | 5.4
| Ruhr-University of Bochum | 2023-06-14 | Lateral(C-D) noise | 0.109
| PSI DEMO Replacement | 2023-06-14 | Lateral(C-D) noise | 0.083
| Ruhr-University of Bochum | 2023-06-14 | Full range of Z Height during Z Scan sweep | 15.8
| PSI DEMO Replacement | 2023-06-14 | Full range of Z Height during Z Scan sweep | 16.4
| Ruhr-University of Bochum | 2023-06-14 | Full range of Z Detector during Z Scan sweep | 11.7
| PSI DEMO Replacement | 2023-06-14 | Full range of Z Detector during Z Scan sweep | 11.3
| Ruhr-University of Bochum | 2023-06-14 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Ruhr-University of Bochum | 2023-06-14 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| PSI DEMO Replacement | 2023-06-14 | Full range of Y Detector Raw during Y Scan sweep | 13.2
| Ruhr-University of Bochum | 2023-06-14 | Full range of Y Detector during Y Scan sweep | 177.1
| PSI DEMO Replacement | 2023-06-14 | Full range of Y Detector during Y Scan sweep | 116.2
| Ruhr-University of Bochum | 2023-06-14 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| Ruhr-University of Bochum | 2023-06-14 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| PSI DEMO Replacement | 2023-06-14 | Full range of X Detector Raw during X Scan sweep | 13.2
| Ruhr-University of Bochum | 2023-06-14 | Full range of X Detector during X Scan sweep | 166.4
| PSI DEMO Replacement | 2023-06-14 | Full range of X Detector during X Scan sweep | 114.2
| Ruhr-University of Bochum | 2023-06-14 | A-B noise | 8.99
| PSI DEMO Replacement | 2023-06-14 | A-B noise | 11.85
| Teknotip Cankiri University | 2023-06-13 | Z Scanner position at -3μm | 0.0148
| Teknotip Cankiri University | 2023-06-13 | Z Scanner position at +3μm | 0.0068
| Teknotip Cankiri University | 2023-06-13 | Z Linearity by 3 different sample height | 1.011
| Universitas Andalas | 2023-06-13 | Z Linearity by 3 different sample height | 1.016
| Teknotip Cankiri University | 2023-06-13 | Z Height noise | 0.032
| Universitas Andalas | 2023-06-13 | Z Height noise | 0.024
| Teknotip Cankiri University | 2023-06-13 | Y Detector noise (XY Servo on) | 0.25
| Universitas Andalas | 2023-06-13 | Y Detector noise (XY Servo on) | 0.3
| Teknotip Cankiri University | 2023-06-13 | Y Detector noise (XY Servo off) | 0.25
| Universitas Andalas | 2023-06-13 | Y Detector noise (XY Servo off) | 0.3
| Teknotip Cankiri University | 2023-06-13 | Y axis OPM | 0.43
| Universitas Andalas | 2023-06-13 | Y axis OPM | 0.91
| Teknotip Cankiri University | 2023-06-13 | X Detector noise (XY Servo on) | 0.26
| Universitas Andalas | 2023-06-13 | X Detector noise (XY Servo on) | 0.31
| Teknotip Cankiri University | 2023-06-13 | X Detector noise (XY Servo off) | 0.25
| Universitas Andalas | 2023-06-13 | X Detector noise (XY Servo off) | 0.3
| Teknotip Cankiri University | 2023-06-13 | X axis OPM | 0.56
| Universitas Andalas | 2023-06-13 | X axis OPM | 0.65
| Teknotip Cankiri University | 2023-06-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Universitas Andalas | 2023-06-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| Teknotip Cankiri University | 2023-06-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.66
| Universitas Andalas | 2023-06-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| Universitas Andalas | 2023-06-13 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Universitas Andalas | 2023-06-13 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.32
| Teknotip Cankiri University | 2023-06-13 | NCM driving% with AC160TS cantilever | 1.16
| Universitas Andalas | 2023-06-13 | NCM driving% with AC160TS cantilever | 3.175
| Teknotip Cankiri University | 2023-06-13 | NCM Amplitude noise | 14.78
| Universitas Andalas | 2023-06-13 | NCM Amplitude noise | 5.8
| Teknotip Cankiri University | 2023-06-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.115
| Universitas Andalas | 2023-06-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.391
| Teknotip Cankiri University | 2023-06-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 30.267
| Universitas Andalas | 2023-06-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.816
| Universitas Andalas | 2023-06-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 158.6
| Teknotip Cankiri University | 2023-06-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.6
| Universitas Andalas | 2023-06-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 158
| Universitas Andalas | 2023-06-13 | Full range of Z Height during Z Scan sweep | 16.4
| Teknotip Cankiri University | 2023-06-13 | Lateral(C-D) sensitivity | 33.693
| Universitas Andalas | 2023-06-13 | Lateral(C-D) sensitivity | 8.2
| Teknotip Cankiri University | 2023-06-13 | Lateral(C-D) noise | 0.085
| Universitas Andalas | 2023-06-13 | Lateral(C-D) noise | 0.118
| Teknotip Cankiri University | 2023-06-13 | Full range of Z Height during Z Scan sweep | 16.4
| Teknotip Cankiri University | 2023-06-13 | Full range of Z Detector during Z Scan sweep | 10.8
| Universitas Andalas | 2023-06-13 | Full range of Z Detector during Z Scan sweep | 13.2
| Teknotip Cankiri University | 2023-06-13 | Full range of Y Detector Raw during Y Scan sweep | 15.5
| Universitas Andalas | 2023-06-13 | Full range of Y Detector Raw during Y Scan sweep | 14.5
| Teknotip Cankiri University | 2023-06-13 | Full range of Y Detector during Y Scan sweep | 76.1
| Universitas Andalas | 2023-06-13 | Full range of Y Detector during Y Scan sweep | 114.8
| Teknotip Cankiri University | 2023-06-13 | Full range of X Detector Raw during X Scan sweep | 14.8
| Universitas Andalas | 2023-06-13 | Full range of X Detector Raw during X Scan sweep | 14.5
| Teknotip Cankiri University | 2023-06-13 | Full range of X Detector during X Scan sweep | 75.9
| Universitas Andalas | 2023-06-13 | Full range of X Detector during X Scan sweep | 114.6
| Universitas Andalas | 2023-06-13 | Approach position at -3μm | 0.0057
| Universitas Andalas | 2023-06-13 | Approach position at +3μm | 0.0016
| Teknotip Cankiri University | 2023-06-13 | A-B noise | 11.29
| Universitas Andalas | 2023-06-13 | A-B noise | 11.36
| SPTS | 2023-06-09 | Z Scanner position at -3μm | 0.0009
| SPTS | 2023-06-09 | Z Scanner position at +3μm | 0.0062
| SPTS | 2023-06-09 | Z Linearity by 3 different sample height | 0.99
| SPTS | 2023-06-09 | Z Height noise | 0.018
| SPTS | 2023-06-09 | Y Scan Stroke | 55
| SPTS | 2023-06-09 | Y Detector noise (XY Servo on) | 0.33
| SPTS | 2023-06-09 | Y Detector noise (XY Servo off) | 0.35
| SPTS | 2023-06-09 | Y axis OPM | 0.28
| SPTS | 2023-06-09 | X Scan Stroke | 55
| SPTS | 2023-06-09 | X Detector noise (XY Servo on) | 0.34
| SPTS | 2023-06-09 | X Detector noise (XY Servo off) | 0.35
| SPTS | 2023-06-09 | X axis OPM | 0.37
| SPTS | 2023-06-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| SPTS | 2023-06-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| SPTS | 2023-06-09 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| SPTS | 2023-06-09 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| SPTS | 2023-06-09 | NCM driving% with AC160TS cantilever | 0.51
| SPTS | 2023-06-09 | NCM Amplitude noise | 10.46
| SPTS | 2023-06-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.5
| SPTS | 2023-06-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 28
| SPTS | 2023-06-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105.7
| SPTS | 2023-06-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 106.4
| SPTS | 2023-06-09 | Lateral(C-D) sensitivity | 7.8
| SPTS | 2023-06-09 | Lateral(C-D) noise | 0.037
| SPTS | 2023-06-09 | Full range of Z Height during Z Scan sweep | 16.6
| SPTS | 2023-06-09 | Full range of Z Detector during Z Scan sweep | 11.8
| SPTS | 2023-06-09 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| SPTS | 2023-06-09 | Full range of Y Detector during Y Scan sweep | 70.9
| SPTS | 2023-06-09 | Full range of X Detector Raw during X Scan sweep | 13.4
| SPTS | 2023-06-09 | Full range of X Detector during X Scan sweep | 71.4
| TDK | 2023-06-02 | Z Linearity by Z scanner position | 0.835
| TDK | 2023-06-02 | Z Linearity by 4 different sample height | 1.005
| TDK | 2023-06-02 | Z Height noise | 0.035
| TDK | 2023-06-02 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.8691
| TDK | 2023-06-02 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 29.0384
| TDK | 2023-06-02 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.053
| TDK | 2023-06-02 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.9833
| TDK | 2023-06-02 | Y2 Detector noise (XY Servo on) | 0.319
| TDK | 2023-06-02 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 248.9
| TDK | 2023-06-02 | Y1 Detector noise (XY Servo on) | 0.32
| TDK | 2023-06-02 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 259.1
| TDK | 2023-06-02 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.9
| TDK | 2023-06-02 | Y Ringing(Rpv) | 0.136
| TDK | 2023-06-02 | Y Detector noise (XY Servo on) | 0.246
| TDK | 2023-06-02 | Y axis OPM | 0.499
| TDK | 2023-06-02 | X2 Detector noise (XY Servo on) | 0.298
| TDK | 2023-06-02 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 250
| TDK | 2023-06-02 | X1 Detector noise (XY Servo on) | 0.31
| TDK | 2023-06-02 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 252.4
| TDK | 2023-06-02 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 172.3
| TDK | 2023-06-02 | X Ringing(Rpv) | 0.185
| TDK | 2023-06-02 | X Detector noise (XY Servo on) | 0.236
| TDK | 2023-06-02 | X axis OPM | 0.705
| TDK | 2023-06-02 | StrokeYMin | -102726
| TDK | 2023-06-02 | StrokeYMax | 102721.9
| TDK | 2023-06-02 | StrokeY | 205447.5
| TDK | 2023-06-02 | StrokeXMin | -116470
| TDK | 2023-06-02 | StrokeXMax | 116436.3
| TDK | 2023-06-02 | StrokeX | 232906.3
| TDK | 2023-06-02 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.7
| TDK | 2023-06-02 | RMS of ZeroScan | 0.046
| TDK | 2023-06-02 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| TDK | 2023-06-02 | OpticalAxisShiftPerFocusY | 0.0004
| TDK | 2023-06-02 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 5e-05
| TDK | 2023-06-02 | OpticalAxisShiftPerFocusX | 0.0003
| TDK | 2023-06-02 | NCM driving% with AC160TS cantilever | 8.87
| TDK | 2023-06-02 | Lateral(C-D) sensitivity | 5.9
| TDK | 2023-06-02 | Lateral(C-D) noise | 0.104
| TDK | 2023-06-02 | Laser intensity with AC160 cantilever | 3.108
| TDK | 2023-06-02 | Full range of Z Height during Z Scan sweep | 16.8642
| TDK | 2023-06-02 | Full range of Z Detector during Z Scan sweep | 11.63
| TDK | 2023-06-02 | Full range of Y2 Detector Raw during Y Scan sweep | 13.26
| TDK | 2023-06-02 | Full range of Y2 Detector during Y Scan sweep | 179.6681
| TDK | 2023-06-02 | Full range of Y1 Detector Raw during Y Scan sweep | 13.2
| TDK | 2023-06-02 | Full range of Y1 Detector during Y Scan sweep | 180.8138
| TDK | 2023-06-02 | Full range of X2 Detector Raw during X Scan sweep | 13.52
| TDK | 2023-06-02 | Full range of X2 Detector during X Scan sweep | 184.89
| TDK | 2023-06-02 | Full range of X1 Detector Raw during X Scan sweep | 13.51
| TDK | 2023-06-02 | Full range of X1 Detector during X Scan sweep | 181.909
| TDK | 2023-06-02 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| TDK | 2023-06-02 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.76
| TDK | 2023-06-02 | A－B difference during 3 times run test | 0.32
| TDK | 2023-06-02 | A＋B difference during 3 times run test | 0.15
| TDK | 2023-06-02 | 50~200nm Step Height Sample (10 Repeats) | 183.421
| TDK | 2023-06-02 | 200~500nm Step Height Sample (10 Repeats) | 462.926
| TDK | 2023-06-02 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 3.4
| TDK | 2023-06-02 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.2
| TDK | 2023-06-02 | %GRR w/ roughness of non-patterened wafer | 4
| Westlake University_School of Science | 2023-06-01 | Z Scanner position at -3μm | 0.0016
| Westlake University_School of Science | 2023-06-01 | Z Scanner position at +3μm | 0.002
| Westlake University_School of Science | 2023-06-01 | Z Linearity by 3 different sample height | 0.993
| Westlake University_School of Science | 2023-06-01 | Z Height noise | 0.022
| Westlake University_School of Science | 2023-06-01 | Y Detector noise (XY Servo on) | 0.17
| Westlake University_School of Science | 2023-06-01 | Y Detector noise (XY Servo off) | 0.17
| Westlake University_School of Science | 2023-06-01 | Y axis OPM | 0.79
| Westlake University_School of Science | 2023-06-01 | X Detector noise (XY Servo on) | 0.16
| Westlake University_School of Science | 2023-06-01 | X Detector noise (XY Servo off) | 0.16
| Westlake University_School of Science | 2023-06-01 | X axis OPM | 1.04
| Westlake University_School of Science | 2023-06-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Westlake University_School of Science | 2023-06-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Westlake University_School of Science | 2023-06-01 | NCM driving% with AC160TS cantilever | 1.74
| Westlake University_School of Science | 2023-06-01 | NCM Amplitude noise | 11.42
| Westlake University_School of Science | 2023-06-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| Westlake University_School of Science | 2023-06-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.7
| Westlake University_School of Science | 2023-06-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 172.9
| Westlake University_School of Science | 2023-06-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 176
| Westlake University_School of Science | 2023-06-01 | Lateral(C-D) sensitivity | 8.1
| Westlake University_School of Science | 2023-06-01 | Lateral(C-D) noise | 0.08
| Westlake University_School of Science | 2023-06-01 | Full range of Z Height during Z Scan sweep | 15.9
| Westlake University_School of Science | 2023-06-01 | Full range of Z Detector during Z Scan sweep | 11.4
| Westlake University_School of Science | 2023-06-01 | Full range of Y Detector Raw during Y Scan sweep | 13.4
| Westlake University_School of Science | 2023-06-01 | Full range of Y Detector during Y Scan sweep | 116
| Westlake University_School of Science | 2023-06-01 | Full range of X Detector Raw during X Scan sweep | 13.2
| Westlake University_School of Science | 2023-06-01 | Full range of X Detector during X Scan sweep | 115.7
| Westlake University_School of Science | 2023-06-01 | A-B noise | 7.43
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Z Scanner position at -3μm | 0.0028
| HC SemiTek | 2023-05-30 | Z Scanner position at -3μm | 0.0063
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Z Scanner position at +3μm | 0.0032
| HC SemiTek | 2023-05-30 | Z Scanner position at +3μm | 0.0152
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Z Linearity by 3 different sample height | 0.991
| HC SemiTek | 2023-05-30 | Z Linearity by 3 different sample height | 1
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Z Height noise | 0.032
| HC SemiTek | 2023-05-30 | Z Height noise | 0.035
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y2 Detector noise (XY Servo on) | 0.22
| HC SemiTek | 2023-05-30 | Y2 Detector noise (XY Servo on) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y2 Detector noise (XY Servo off) | 0.22
| HC SemiTek | 2023-05-30 | Y2 Detector noise (XY Servo off) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y1 Detector noise (XY Servo on) | 0.22
| HC SemiTek | 2023-05-30 | Y1 Detector noise (XY Servo on) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y1 Detector noise (XY Servo off) | 0.21
| HC SemiTek | 2023-05-30 | Y1 Detector noise (XY Servo off) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y Scan Stroke | 176.6
| HC SemiTek | 2023-05-30 | Y Scan Stroke | 172.9
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y Detector noise (XY Servo on) | 0.16
| HC SemiTek | 2023-05-30 | Y Detector noise (XY Servo on) | 0.15
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y Detector noise (XY Servo off) | 0.15
| HC SemiTek | 2023-05-30 | Y Detector noise (XY Servo off) | 0.14
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Y axis OPM | 1.29
| HC SemiTek | 2023-05-30 | Y axis OPM | 1.74
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X2 Detector noise (XY Servo on) | 0.22
| HC SemiTek | 2023-05-30 | X2 Detector noise (XY Servo on) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X2 Detector noise (XY Servo off) | 0.21
| HC SemiTek | 2023-05-30 | X2 Detector noise (XY Servo off) | 0.18
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X1 Detector noise (XY Servo on) | 0.23
| HC SemiTek | 2023-05-30 | X1 Detector noise (XY Servo on) | 0.2
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X1 Detector noise (XY Servo off) | 0.22
| HC SemiTek | 2023-05-30 | X1 Detector noise (XY Servo off) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X Scan Stroke | 177.3
| HC SemiTek | 2023-05-30 | X Scan Stroke | 170.9
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | -X Direction software Limit | -102363
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | -X Direction software Limit | -103304
| HC SemiTek | 2023-05-30 | -X Direction software Limit | -102093
| HC SemiTek | 2023-05-30 | -X Direction software Limit | -103356
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X Detector noise (XY Servo on) | 0.17
| HC SemiTek | 2023-05-30 | X Detector noise (XY Servo on) | 0.14
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X Detector noise (XY Servo off) | 0.16
| HC SemiTek | 2023-05-30 | X Detector noise (XY Servo off) | 0.14
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | X axis OPM | 1.52
| HC SemiTek | 2023-05-30 | X axis OPM | 0.61
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| HC SemiTek | 2023-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.34
| HC SemiTek | 2023-05-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.54
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Top) | 0.3
| HC SemiTek | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Top) | 0.46
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Right) | 0.27
| HC SemiTek | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Right) | 0.37
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Left) | 0.28
| HC SemiTek | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Left) | 0.4
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of Wafer ZeroScan (Center) | 0.28
| HC SemiTek | 2023-05-30 | RMS of Z Height of Wafer ZeroScan (Center) | 0.37
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.3
| HC SemiTek | 2023-05-30 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.42
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| HC SemiTek | 2023-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| HC SemiTek | 2023-05-30 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | NCM driving% with AC160TS cantilever | 0.91
| HC SemiTek | 2023-05-30 | NCM driving% with AC160TS cantilever | 1.93
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | NCM Amplitude noise | 13.09
| HC SemiTek | 2023-05-30 | NCM Amplitude noise | 13.73
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| HC SemiTek | 2023-05-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.277
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.2
| HC SemiTek | 2023-05-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.83
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 252
| HC SemiTek | 2023-05-30 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 244.9
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 240.5
| HC SemiTek | 2023-05-30 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 238.8
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.5
| HC SemiTek | 2023-05-30 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 239.8
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249
| HC SemiTek | 2023-05-30 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 241.7
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| HC SemiTek | 2023-05-30 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.09
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Lateral(C-D) sensitivity | 6.5
| HC SemiTek | 2023-05-30 | Lateral(C-D) sensitivity | 4.7
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Lateral(C-D) noise | 0.099
| HC SemiTek | 2023-05-30 | Lateral(C-D) noise | 0.122
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of Z Height during Z Scan sweep | 16.5
| HC SemiTek | 2023-05-30 | Full range of Z Height during Z Scan sweep | 15.3
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of Z Detector during Z Scan sweep | 12.1
| HC SemiTek | 2023-05-30 | Full range of Z Detector during Z Scan sweep | 10.8
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| HC SemiTek | 2023-05-30 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| HC SemiTek | 2023-05-30 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of Y Detector during Y Scan sweep | 176.9
| HC SemiTek | 2023-05-30 | Full range of Y Detector during Y Scan sweep | 173.3
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| HC SemiTek | 2023-05-30 | Full range of X2 Detector Raw during X Scan sweep | 14.3
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| HC SemiTek | 2023-05-30 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | Full range of X Detector during X Scan sweep | 177.8
| HC SemiTek | 2023-05-30 | Full range of X Detector during X Scan sweep | 171.9
| Shenzhen Founder Microelectronics International(FMIC) #2 | 2023-05-30 | A-B noise | 9.11
| HC SemiTek | 2023-05-30 | A-B noise | 12.77
| PSE Demo | 2023-05-26 | Z Scanner position at -3μm | 0.0047
| PSE Demo | 2023-05-26 | Z Scanner position at +3μm | 0.0154
| PSE Demo | 2023-05-26 | Z Linearity by 3 different sample height | 1
| PSE Demo | 2023-05-26 | Z Height noise | 0.036
| PSE Demo | 2023-05-26 | Y2 Detector noise (XY Servo on) | 0.19
| PSE Demo | 2023-05-26 | Y2 Detector noise (XY Servo off) | 0.17
| PSE Demo | 2023-05-26 | Y1 Detector noise (XY Servo on) | 0.19
| PSE Demo | 2023-05-26 | Y1 Detector noise (XY Servo off) | 0.17
| PSE Demo | 2023-05-26 | Y Scan Stroke | 171.3
| PSE Demo | 2023-05-26 | Y Detector noise (XY Servo on) | 0.15
| PSE Demo | 2023-05-26 | Y Detector noise (XY Servo off) | 0.13
| PSE Demo | 2023-05-26 | Y axis OPM | 1.15
| PSE Demo | 2023-05-26 | X2 Detector noise (XY Servo on) | 0.2
| PSE Demo | 2023-05-26 | X2 Detector noise (XY Servo off) | 0.24
| PSE Demo | 2023-05-26 | X1 Detector noise (XY Servo on) | 0.19
| PSE Demo | 2023-05-26 | X1 Detector noise (XY Servo off) | 0.19
| PSE Demo | 2023-05-26 | X Scan Stroke | 177.01
| PSE Demo | 2023-05-26 | -X Direction software Limit | -102384
| PSE Demo | 2023-05-26 | -X Direction software Limit | -103603
| PSE Demo | 2023-05-26 | X Detector noise (XY Servo on) | 0.15
| PSE Demo | 2023-05-26 | X Detector noise (XY Servo off) | 0.16
| PSE Demo | 2023-05-26 | X axis OPM | 0.82
| PSE Demo | 2023-05-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| PSE Demo | 2023-05-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| PSE Demo | 2023-05-26 | RMS of Z Height of Wafer Zeroscan (Top) | 0.35
| PSE Demo | 2023-05-26 | RMS of Z Height of Wafer Zeroscan (Right) | 0.35
| PSE Demo | 2023-05-26 | RMS of Z Height of Wafer Zeroscan (Left) | 0.31
| PSE Demo | 2023-05-26 | RMS of Z Height of Wafer ZeroScan (Center) | 0.38
| PSE Demo | 2023-05-26 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.35
| PSE Demo | 2023-05-26 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| PSE Demo | 2023-05-26 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| PSE Demo | 2023-05-26 | NCM driving% with AC160TS cantilever | 0.72
| PSE Demo | 2023-05-26 | NCM Amplitude noise | 13.23
| PSE Demo | 2023-05-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.486
| PSE Demo | 2023-05-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.725
| PSE Demo | 2023-05-26 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 242.6
| PSE Demo | 2023-05-26 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 234.9
| PSE Demo | 2023-05-26 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 241.7
| PSE Demo | 2023-05-26 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 246.6
| PSE Demo | 2023-05-26 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| PSE Demo | 2023-05-26 | Lateral(C-D) sensitivity | 7.3
| PSE Demo | 2023-05-26 | Lateral(C-D) noise | 0.141
| PSE Demo | 2023-05-26 | Full range of Z Height during Z Scan sweep | 16
| PSE Demo | 2023-05-26 | Full range of Z Detector during Z Scan sweep | 11.1
| PSE Demo | 2023-05-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14.4
| PSE Demo | 2023-05-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.3
| PSE Demo | 2023-05-26 | Full range of Y Detector during Y Scan sweep | 171.9
| PSE Demo | 2023-05-26 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| PSE Demo | 2023-05-26 | Full range of X1 Detector Raw during X Scan sweep | 14.5
| PSE Demo | 2023-05-26 | Full range of X Detector during X Scan sweep | 177.1
| PSE Demo | 2023-05-26 | A-B noise | 9.37
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Linearity by Z scanner position | 0.46
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Linearity by Z scanner position | 0.464
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Linearity by 5 different sample height | 1.006
| PXW SEMICONDUCTOR MANUFACTURING_1 | 2023-05-25 | Z Linearity by 5 different sample height | 1.0039
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Height noise | 0.036
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.9909
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.8152
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.2088
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2077
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y2 Detector noise (XY Servo on) | 0.52
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 281.6511
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y1 Detector noise (XY Servo on) | 0.57
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 273.7581
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 196.197
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y Ringing(Rpv) | 0.763
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y Non-linearity error | 2.29
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y Detector noise (XY Servo on) | 0.45
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Repeatability @1st position | 0.06
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Repeatability @ 4th position | 0.08
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Repeatability @ 3rd position | 0.05
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Repeatability @ 2nd position | 0.08
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Accuracy @1st position | 0.69
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Accuracy @ 4th position | 0.7
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Accuracy @ 3rd position | 0.54
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis, Accuracy @ 2nd position | 0.61
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Y axis OPM | 0.77
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X2 Detector noise (XY Servo on) | 0.59
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 281.5453
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X1 Detector noise (XY Servo on) | 0.36
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 256.7892
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 193.326
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X Ringing(Rpv) | 0.777
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X Non-linearity error | 3.19
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X Detector noise (XY Servo on) | 0.39
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Repeatability @1st position | 0.05
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Repeatability @ 4th position | 0.05
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Repeatability @ 3rd position | 0.04
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Repeatability @ 2nd position | 0.05
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Accuracy @1st position | 0.25
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Accuracy @ 4th position | 0.31
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Accuracy @ 3rd position | 0.27
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis, Accuracy @ 2nd position | 0.27
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | X axis OPM | 1.719
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Under 50nm Step Height Sample (10 Repeats) | 18.203
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Under 50nm Step Height Sample (10 Repeats) | 18.126
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Torque Profile (50mm, 20um/s) | 2.763
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Torque Profile (50mm, 200um/s) | 1.875
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | StrokeYMin | -153221
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | StrokeYMax | 150155
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | StrokeY | 303376
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | StrokeXMin | -204314
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | StrokeXMax | 203568
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | StrokeX | 407882
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.625
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | RMS of ZeroScan | 0.059
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | PosMin | -26
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | PosMax | 26
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00013
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OpticalAxisShiftPerFocusY | 0.0049
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00038
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OpticalAxisShiftPerFocusX | 0.0085
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OPM Repeatability @5mm | 1.288
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OPM Repeatability @25mm | 4.56
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OPM Repeatability @1mm | 0.7152
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | OPM Repeatability @10mm | 1.717
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | NCM driving% with AC160TS cantilever | 5.38
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | NCM driving% with AC160TS cantilever | 3.97
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Min. Vibration Transmissbility (~100Hz) | 0.076
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | MaxLength | 52
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Max OPM @5mm | 40.072
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Max OPM @25mm | 232.224
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Max OPM @1mm | 8.718
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Max OPM @10mm | 66.085
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Lateral(C-D) sensitivity | 45.53
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Lateral(C-D) noise | 0.188
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Laser intensity with AC160 cantilever | 3.053
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Laser intensity with AC160 cantilever | 2.993
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from -Y to +Y (Y scan) | 32.23
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from -Y to +Y (X scan) | 16.11
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from -X to +X (Y scan) | 63.72
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from -X to +X (X scan) | 125.98
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from +Y to -Y (Y scan) | 30.03
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from +Y to -Y (X scan) | 11.72
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from +X to -X (Y scan) | 93.75
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Image offset after moving from +X to -X (X scan) | 115.72
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Z Height during Z Scan sweep | 17.1
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Z Height during Z Scan sweep | 38.9
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Z Detector during Z Scan sweep | 11.5
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Z Detector during Z Scan sweep | 12.57
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.24
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Y2 Detector during Y Scan sweep | 208.2039
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14.75
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of Y1 Detector during Y Scan sweep | 211.9769
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of X2 Detector Raw during X Scan sweep | 14.34
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of X2 Detector during X Scan sweep | 210.2943
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of X1 Detector Raw during X Scan sweep | 14.62
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Full range of X1 Detector during X Scan sweep | 207.9595
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.72
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.72
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.86
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.62
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.11
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | A－B difference during 3 times run test | 0.77
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | A＋B difference during 3 times run test | 0.19
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | 50~200nm Step Height Sample (10 Repeats) | 183.959
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | 50~200nm Step Height Sample (10 Repeats) | 183.712
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | 200~500nm Step Height Sample (10 Repeats) | 463.996
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | 200~500nm Step Height Sample (10 Repeats) | 464.016
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.7
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 4.7
| PXW SEMICONDUCTOR MANUFACTURING | 2023-05-25 | %GRR w/ roughness of non-patterened wafer | 6.7
| Mahidol University | 2023-05-24 | Z Scanner position at -3μm | 0.0111
| 성균관대학교 정보통신대학 | 2023-05-24 | Z Scanner position at -3μm | 0.0046
| KAIST(물리학과) | 2023-05-24 | Z Scanner position at -3μm | -0.0019
| Mahidol University | 2023-05-24 | Z Scanner position at +3μm | 0.0024
| 성균관대학교 정보통신대학 | 2023-05-24 | Z Scanner position at +3μm | 0.0073
| KAIST(물리학과) | 2023-05-24 | Z Scanner position at +3μm | -0.0039
| Mahidol University | 2023-05-24 | Z Linearity by 3 different sample height | 1.014
| 성균관대학교 정보통신대학 | 2023-05-24 | Z Linearity by 3 different sample height | 0.997
| KAIST(물리학과) | 2023-05-24 | Z Linearity by 3 different sample height | 1.01
| Mahidol University | 2023-05-24 | Z Height noise | 0.027
| 성균관대학교 정보통신대학 | 2023-05-24 | Z Height noise | 0.022
| KAIST(물리학과) | 2023-05-24 | Z Height noise | 0.022
| Mahidol University | 2023-05-24 | Y2 Detector noise (XY Servo on) | 0.3
| Mahidol University | 2023-05-24 | Y2 Detector noise (XY Servo off) | 0.18
| Mahidol University | 2023-05-24 | Y1 Detector noise (XY Servo on) | 0.3
| Mahidol University | 2023-05-24 | Y1 Detector noise (XY Servo off) | 0.18
| Mahidol University | 2023-05-24 | Y Scan Stroke | 172.6
| KAIST(물리학과) | 2023-05-24 | Y Scan Stroke | 55
| Mahidol University | 2023-05-24 | Y Detector noise (XY Servo on) | 0.23
| 성균관대학교 정보통신대학 | 2023-05-24 | Y Detector noise (XY Servo on) | 0.26
| KAIST(물리학과) | 2023-05-24 | Y Detector noise (XY Servo on) | 0.35
| Mahidol University | 2023-05-24 | Y Detector noise (XY Servo off) | 0.13
| 성균관대학교 정보통신대학 | 2023-05-24 | Y Detector noise (XY Servo off) | 0.25
| KAIST(물리학과) | 2023-05-24 | Y Detector noise (XY Servo off) | 0.37
| Mahidol University | 2023-05-24 | Y axis OPM | 1.09
| 성균관대학교 정보통신대학 | 2023-05-24 | Y axis OPM | 0.46
| KAIST(물리학과) | 2023-05-24 | Y axis OPM | 0.41
| Mahidol University | 2023-05-24 | X2 Detector noise (XY Servo on) | 0.31
| Mahidol University | 2023-05-24 | X2 Detector noise (XY Servo off) | 0.19
| Mahidol University | 2023-05-24 | X1 Detector noise (XY Servo on) | 0.32
| Mahidol University | 2023-05-24 | X1 Detector noise (XY Servo off) | 0.18
| Mahidol University | 2023-05-24 | X Scan Stroke | 172.9
| KAIST(물리학과) | 2023-05-24 | X Scan Stroke | 55
| Mahidol University | 2023-05-24 | -X Direction software Limit | -102019
| Mahidol University | 2023-05-24 | -X Direction software Limit | -104951
| Mahidol University | 2023-05-24 | X Detector noise (XY Servo on) | 0.25
| 성균관대학교 정보통신대학 | 2023-05-24 | X Detector noise (XY Servo on) | 0.26
| KAIST(물리학과) | 2023-05-24 | X Detector noise (XY Servo on) | 0.32
| Mahidol University | 2023-05-24 | X Detector noise (XY Servo off) | 0.14
| 성균관대학교 정보통신대학 | 2023-05-24 | X Detector noise (XY Servo off) | 0.24
| KAIST(물리학과) | 2023-05-24 | X Detector noise (XY Servo off) | 0.32
| Mahidol University | 2023-05-24 | X axis OPM | 1.08
| 성균관대학교 정보통신대학 | 2023-05-24 | X axis OPM | 0.71
| KAIST(물리학과) | 2023-05-24 | X axis OPM | 0.5
| Mahidol University | 2023-05-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 성균관대학교 정보통신대학 | 2023-05-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| KAIST(물리학과) | 2023-05-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.16
| Mahidol University | 2023-05-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| 성균관대학교 정보통신대학 | 2023-05-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| KAIST(물리학과) | 2023-05-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| Mahidol University | 2023-05-24 | RMS of Z Height of Wafer Zeroscan (Top) | 0.29
| Mahidol University | 2023-05-24 | RMS of Z Height of Wafer Zeroscan (Right) | 0.29
| Mahidol University | 2023-05-24 | RMS of Z Height of Wafer Zeroscan (Left) | 0.29
| Mahidol University | 2023-05-24 | RMS of Z Height of Wafer ZeroScan (Center) | 0.36
| Mahidol University | 2023-05-24 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.31
| Mahidol University | 2023-05-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| KAIST(물리학과) | 2023-05-24 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Mahidol University | 2023-05-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.53
| KAIST(물리학과) | 2023-05-24 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.59
| Mahidol University | 2023-05-24 | NCM driving% with AC160TS cantilever | 0.25
| 성균관대학교 정보통신대학 | 2023-05-24 | NCM driving% with AC160TS cantilever | 5.25
| KAIST(물리학과) | 2023-05-24 | NCM driving% with AC160TS cantilever | 0.58
| Mahidol University | 2023-05-24 | NCM Amplitude noise | 15.28
| 성균관대학교 정보통신대학 | 2023-05-24 | NCM Amplitude noise | 8.95
| KAIST(물리학과) | 2023-05-24 | NCM Amplitude noise | 15.25
| Mahidol University | 2023-05-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| 성균관대학교 정보통신대학 | 2023-05-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| KAIST(물리학과) | 2023-05-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.6
| Mahidol University | 2023-05-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.3
| 성균관대학교 정보통신대학 | 2023-05-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.3
| KAIST(물리학과) | 2023-05-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.1
| Mahidol University | 2023-05-24 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 245.7
| Mahidol University | 2023-05-24 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 240.4
| 성균관대학교 정보통신대학 | 2023-05-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 109.3
| KAIST(물리학과) | 2023-05-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 96.6
| Mahidol University | 2023-05-24 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 245.5
| Mahidol University | 2023-05-24 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 245.2
| 성균관대학교 정보통신대학 | 2023-05-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 108.5
| KAIST(물리학과) | 2023-05-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 96.5
| Mahidol University | 2023-05-24 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Mahidol University | 2023-05-24 | Lateral(C-D) sensitivity | 39.2
| 성균관대학교 정보통신대학 | 2023-05-24 | Lateral(C-D) sensitivity | 7.8
| KAIST(물리학과) | 2023-05-24 | Lateral(C-D) sensitivity | 6.1
| Mahidol University | 2023-05-24 | Lateral(C-D) noise | 0.115
| 성균관대학교 정보통신대학 | 2023-05-24 | Lateral(C-D) noise | 0.075
| KAIST(물리학과) | 2023-05-24 | Lateral(C-D) noise | 0.123
| Mahidol University | 2023-05-24 | Full range of Z Height during Z Scan sweep | 15.4
| 성균관대학교 정보통신대학 | 2023-05-24 | Full range of Z Height during Z Scan sweep | 15.9
| KAIST(물리학과) | 2023-05-24 | Full range of Z Height during Z Scan sweep | 16.9
| Mahidol University | 2023-05-24 | Full range of Z Detector during Z Scan sweep | 11.1
| 성균관대학교 정보통신대학 | 2023-05-24 | Full range of Z Detector during Z Scan sweep | 11.2
| KAIST(물리학과) | 2023-05-24 | Full range of Z Detector during Z Scan sweep | 12
| Mahidol University | 2023-05-24 | Full range of Y2 Detector Raw during Y Scan sweep | 13.9
| Mahidol University | 2023-05-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| 성균관대학교 정보통신대학 | 2023-05-24 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| KAIST(물리학과) | 2023-05-24 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Mahidol University | 2023-05-24 | Full range of Y Detector during Y Scan sweep | 171.7
| 성균관대학교 정보통신대학 | 2023-05-24 | Full range of Y Detector during Y Scan sweep | 75.4
| KAIST(물리학과) | 2023-05-24 | Full range of Y Detector during Y Scan sweep | 64.8
| Mahidol University | 2023-05-24 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| Mahidol University | 2023-05-24 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| 성균관대학교 정보통신대학 | 2023-05-24 | Full range of X Detector Raw during X Scan sweep | 13.7
| KAIST(물리학과) | 2023-05-24 | Full range of X Detector Raw during X Scan sweep | 14.8
| Mahidol University | 2023-05-24 | Full range of X Detector during X Scan sweep | 173.6
| 성균관대학교 정보통신대학 | 2023-05-24 | Full range of X Detector during X Scan sweep | 73.9
| KAIST(물리학과) | 2023-05-24 | Full range of X Detector during X Scan sweep | 66.5
| Mahidol University | 2023-05-24 | A-B noise | 11.41
| 성균관대학교 정보통신대학 | 2023-05-24 | A-B noise | 4.65
| KAIST(물리학과) | 2023-05-24 | A-B noise | 13.83
| Hangzhou HFC Semiconductor | 2023-05-23 | Z Scanner position at -3μm | 0.0022
| UNC | 2023-05-23 | Z Scanner position at -3μm | 0.0007
| Hangzhou HFC Semiconductor | 2023-05-23 | Z Scanner position at +3μm | 0.0003
| UNC | 2023-05-23 | Z Scanner position at +3μm | 0.0011
| Hangzhou HFC Semiconductor | 2023-05-23 | Z Linearity by 3 different sample height | 1.003
| UNC | 2023-05-23 | Z Linearity by 3 different sample height | 1.018
| Hangzhou HFC Semiconductor | 2023-05-23 | Z Height noise | 0.037
| UNC | 2023-05-23 | Z Height noise | 0.032
| Hangzhou HFC Semiconductor | 2023-05-23 | Y2 Detector noise (XY Servo on) | 0.28
| Hangzhou HFC Semiconductor | 2023-05-23 | Y2 Detector noise (XY Servo off) | 0.27
| Hangzhou HFC Semiconductor | 2023-05-23 | Y1 Detector noise (XY Servo on) | 0.24
| Hangzhou HFC Semiconductor | 2023-05-23 | Y1 Detector noise (XY Servo off) | 0.29
| Hangzhou HFC Semiconductor | 2023-05-23 | Y Scan Stroke | 197.4
| UNC | 2023-05-23 | Y Scan Stroke | 185.1
| Hangzhou HFC Semiconductor | 2023-05-23 | Y Detector noise (XY Servo on) | 0.21
| UNC | 2023-05-23 | Y Detector noise (XY Servo on) | 0.35
| Hangzhou HFC Semiconductor | 2023-05-23 | Y Detector noise (XY Servo off) | 0.21
| UNC | 2023-05-23 | Y Detector noise (XY Servo off) | 0.28
| Hangzhou HFC Semiconductor | 2023-05-23 | Y axis OPM | 1.36
| UNC | 2023-05-23 | Y axis OPM | 0.98
| Hangzhou HFC Semiconductor | 2023-05-23 | X2 Detector noise (XY Servo on) | 0.3
| Hangzhou HFC Semiconductor | 2023-05-23 | X2 Detector noise (XY Servo off) | 0.3
| Hangzhou HFC Semiconductor | 2023-05-23 | X1 Detector noise (XY Servo on) | 0.3
| Hangzhou HFC Semiconductor | 2023-05-23 | X1 Detector noise (XY Servo off) | 0.29
| Hangzhou HFC Semiconductor | 2023-05-23 | X Scan Stroke | 197
| UNC | 2023-05-23 | X Scan Stroke | 182.5
| Hangzhou HFC Semiconductor | 2023-05-23 | -X Direction software Limit | -150299
| Hangzhou HFC Semiconductor | 2023-05-23 | -X Direction software Limit | -150001
| Hangzhou HFC Semiconductor | 2023-05-23 | X Detector noise (XY Servo on) | 0.23
| UNC | 2023-05-23 | X Detector noise (XY Servo on) | 0.34
| Hangzhou HFC Semiconductor | 2023-05-23 | X Detector noise (XY Servo off) | 0.23
| Hangzhou HFC Semiconductor | 2023-05-23 | X axis OPM | 2.59
| UNC | 2023-05-23 | X axis OPM | 0.61
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| UNC | 2023-05-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.4
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.42
| UNC | 2023-05-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.63
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of Wafer Zeroscan (Top) | 0.35
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of Wafer Zeroscan (Right) | 0.27
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of Wafer Zeroscan (Left) | 0.32
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of Wafer ZeroScan (Center) | 0.31
| Hangzhou HFC Semiconductor | 2023-05-23 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.33
| UNC | 2023-05-23 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.37
| UNC | 2023-05-23 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.62
| Hangzhou HFC Semiconductor | 2023-05-23 | NCM driving% with AC160TS cantilever | 1.683
| UNC | 2023-05-23 | NCM driving% with AC160TS cantilever | 1.58
| Hangzhou HFC Semiconductor | 2023-05-23 | NCM Amplitude noise | 14.84
| UNC | 2023-05-23 | NCM Amplitude noise | 13.36
| Hangzhou HFC Semiconductor | 2023-05-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.382
| UNC | 2023-05-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.516
| Hangzhou HFC Semiconductor | 2023-05-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.715
| UNC | 2023-05-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.955
| Hangzhou HFC Semiconductor | 2023-05-23 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 278.6
| Hangzhou HFC Semiconductor | 2023-05-23 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 281.3
| UNC | 2023-05-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 262.5
| Hangzhou HFC Semiconductor | 2023-05-23 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 278.6
| Hangzhou HFC Semiconductor | 2023-05-23 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 280.2
| UNC | 2023-05-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 257.1
| Hangzhou HFC Semiconductor | 2023-05-23 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.08
| Hangzhou HFC Semiconductor | 2023-05-23 | Lateral(C-D) sensitivity | 16.4
| UNC | 2023-05-23 | Lateral(C-D) sensitivity | 6
| Hangzhou HFC Semiconductor | 2023-05-23 | Lateral(C-D) noise | 0.101
| UNC | 2023-05-23 | Lateral(C-D) noise | 0.073
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of Z Height during Z Scan sweep | 16.3
| UNC | 2023-05-23 | Full range of Z Height during Z Scan sweep | 16.5
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of Z Detector during Z Scan sweep | 10.9
| UNC | 2023-05-23 | Full range of Z Detector during Z Scan sweep | 12.2
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| UNC | 2023-05-23 | Full range of Y Detector Raw during Y Scan sweep | 14
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of Y Detector during Y Scan sweep | 196.6
| UNC | 2023-05-23 | Full range of Y Detector during Y Scan sweep | 184.4
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| UNC | 2023-05-23 | Full range of X Detector Raw during X Scan sweep | 14.1
| Hangzhou HFC Semiconductor | 2023-05-23 | Full range of X Detector during X Scan sweep | 197
| UNC | 2023-05-23 | Full range of X Detector during X Scan sweep | 182
| Hangzhou HFC Semiconductor | 2023-05-23 | A-B noise | 10.16
| UNC | 2023-05-23 | A-B noise | 8.77
| Swansea University | 2023-05-19 | Z Scanner position at -3μm | 0.0029
| Swansea University | 2023-05-19 | Z Scanner position at +3μm | 0.008
| Swansea University | 2023-05-19 | Z Linearity by 3 different sample height | 1.003
| Swansea University | 2023-05-19 | Z Height noise | 0.022
| Swansea University | 2023-05-19 | Y Detector noise (XY Servo on) | 0.28
| Swansea University | 2023-05-19 | Y Detector noise (XY Servo off) | 0.27
| Swansea University | 2023-05-19 | Y axis OPM | 0.44
| Swansea University | 2023-05-19 | X Detector noise (XY Servo on) | 0.29
| Swansea University | 2023-05-19 | X Detector noise (XY Servo off) | 0.27
| Swansea University | 2023-05-19 | X axis OPM | 0.49
| Swansea University | 2023-05-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.24
| Swansea University | 2023-05-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Swansea University | 2023-05-19 | NCM driving% with AC160TS cantilever | 1.9
| Swansea University | 2023-05-19 | NCM Amplitude noise | 11.5
| Swansea University | 2023-05-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.9
| Swansea University | 2023-05-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.5
| Swansea University | 2023-05-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 102.2
| Swansea University | 2023-05-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.3
| Swansea University | 2023-05-19 | Lateral(C-D) sensitivity | 8.2
| Swansea University | 2023-05-19 | Lateral(C-D) noise | 0.049
| Swansea University | 2023-05-19 | Full range of Z Height during Z Scan sweep | 17.3
| Swansea University | 2023-05-19 | Full range of Z Detector during Z Scan sweep | 13
| Swansea University | 2023-05-19 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| Swansea University | 2023-05-19 | Full range of Y Detector during Y Scan sweep | 73.3
| Swansea University | 2023-05-19 | Full range of X Detector Raw during X Scan sweep | 14.1
| Swansea University | 2023-05-19 | Full range of X Detector during X Scan sweep | 71.9
| Swansea University | 2023-05-19 | A-B noise | 7.38
| Quantlase Lab | 2023-05-18 | Z Scanner position at -3μm | 0.0011
| Quantlase Lab | 2023-05-18 | Z Scanner position at +3μm | 0.0056
| Quantlase Lab | 2023-05-18 | Z Linearity by 3 different sample height | 0.991
| Quantlase Lab | 2023-05-18 | Z Height noise | 0.024
| Quantlase Lab | 2023-05-18 | Y Scan Stroke | 190.5
| Quantlase Lab | 2023-05-18 | Y Detector noise (XY Servo on) | 0.35
| Quantlase Lab | 2023-05-18 | Y Detector noise (XY Servo off) | 0.28
| Quantlase Lab | 2023-05-18 | Y axis OPM | 1.25
| Quantlase Lab | 2023-05-18 | X Scan Stroke | 189.8
| Quantlase Lab | 2023-05-18 | X Detector noise (XY Servo on) | 0.33
| Quantlase Lab | 2023-05-18 | X Detector noise (XY Servo off) | 0.27
| Quantlase Lab | 2023-05-18 | X axis OPM | 0.82
| Quantlase Lab | 2023-05-18 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.48
| Quantlase Lab | 2023-05-18 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.55
| Quantlase Lab | 2023-05-18 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.46
| Quantlase Lab | 2023-05-18 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.52
| Quantlase Lab | 2023-05-18 | NCM driving% with AC160TS cantilever | 0.5
| Quantlase Lab | 2023-05-18 | NCM Amplitude noise | 11.79
| Quantlase Lab | 2023-05-18 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| Quantlase Lab | 2023-05-18 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.1
| Quantlase Lab | 2023-05-18 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 266.4
| Quantlase Lab | 2023-05-18 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 261.8
| Quantlase Lab | 2023-05-18 | Lateral(C-D) sensitivity | 61.198
| Quantlase Lab | 2023-05-18 | Lateral(C-D) noise | 0.059
| Quantlase Lab | 2023-05-18 | Full range of Z Height during Z Scan sweep | 16.8
| Quantlase Lab | 2023-05-18 | Full range of Z Detector during Z Scan sweep | 12.8
| Quantlase Lab | 2023-05-18 | Full range of Y Detector during Y Scan sweep | 190.3
| Quantlase Lab | 2023-05-18 | Full range of X Detector Raw during X Scan sweep | 14.5
| Quantlase Lab | 2023-05-18 | Full range of X Detector during X Scan sweep | 189.9
| Quantlase Lab | 2023-05-18 | A-B noise | 4.43
| A Inc._Aurora | 2023-05-16 | Z Linearity by Z scanner position | 1.2
| A Inc._Aurora | 2023-05-16 | Z Linearity by 4 different sample height | 1.001
| A Inc._Aurora | 2023-05-16 | Z Height noise | 0.04
| A Inc._Aurora | 2023-05-16 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.93
| A Inc._Aurora | 2023-05-16 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.84
| A Inc._Aurora | 2023-05-16 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.63
| A Inc._Aurora | 2023-05-16 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.85
| A Inc._Aurora | 2023-05-16 | Z Detector offset | 80
| A Inc._Aurora | 2023-05-16 | Y2 Detector noise (XY Servo on) | 0.29
| A Inc._Aurora | 2023-05-16 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 250.9
| A Inc._Aurora | 2023-05-16 | Y1 Detector noise (XY Servo on) | 0.3
| A Inc._Aurora | 2023-05-16 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 241.39
| A Inc._Aurora | 2023-05-16 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 168.14
| A Inc._Aurora | 2023-05-16 | Y Ringing(Rpv) | 0.44
| A Inc._Aurora | 2023-05-16 | Y Non-linearity error | 2.39
| A Inc._Aurora | 2023-05-16 | Y Detector noise (XY Servo on) | 0.21
| A Inc._Aurora | 2023-05-16 | Y axis, Repeatability @1st position | 0.06
| A Inc._Aurora | 2023-05-16 | Y axis, Repeatability @ 4th position | 0.04
| A Inc._Aurora | 2023-05-16 | Y axis, Repeatability @ 3rd position | 0.05
| A Inc._Aurora | 2023-05-16 | Y axis, Repeatability @ 2nd position | 0.03
| A Inc._Aurora | 2023-05-16 | Y axis, Accuracy @1st position | 0.83
| A Inc._Aurora | 2023-05-16 | Y axis, Accuracy @ 4th position | 0.76
| A Inc._Aurora | 2023-05-16 | Y axis, Accuracy @ 3rd position | 1.1
| A Inc._Aurora | 2023-05-16 | Y axis, Accuracy @ 2nd position | 1.07
| A Inc._Aurora | 2023-05-16 | Y axis OPM | 0.26
| A Inc._Aurora | 2023-05-16 | XY Orthogonality | 89.9
| A Inc._Aurora | 2023-05-16 | X2 Detector noise (XY Servo on) | 0.28
| A Inc._Aurora | 2023-05-16 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 240.03
| A Inc._Aurora | 2023-05-16 | X1 Detector noise (XY Servo on) | 0.29
| A Inc._Aurora | 2023-05-16 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 247.16
| A Inc._Aurora | 2023-05-16 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 166.54
| A Inc._Aurora | 2023-05-16 | X Ringing(Rpv) | 0.81
| A Inc._Aurora | 2023-05-16 | X Non-linearity error | 2.59
| A Inc._Aurora | 2023-05-16 | X Detector noise (XY Servo on) | 0.21
| A Inc._Aurora | 2023-05-16 | X axis, Repeatability @1st position | 0.07
| A Inc._Aurora | 2023-05-16 | X axis, Repeatability @ 4th position | 0.05
| A Inc._Aurora | 2023-05-16 | X axis, Repeatability @ 3rd position | 0.08
| A Inc._Aurora | 2023-05-16 | X axis, Repeatability @ 2nd position | 0.07
| A Inc._Aurora | 2023-05-16 | X axis, Accuracy @1st position | 1.49
| A Inc._Aurora | 2023-05-16 | X axis, Accuracy @ 4th position | 1.24
| A Inc._Aurora | 2023-05-16 | X axis, Accuracy @ 3rd position | 1.29
| A Inc._Aurora | 2023-05-16 | X axis, Accuracy @ 2nd position | 1.3
| A Inc._Aurora | 2023-05-16 | X axis OPM | 0.3
| A Inc._Aurora | 2023-05-16 | Torque Profile (10mm, 20um/s) | 2.27
| A Inc._Aurora | 2023-05-16 | Torque Profile (10mm, 200um/s) | 2.49
| A Inc._Aurora | 2023-05-16 | StrokeYMin | -108093
| A Inc._Aurora | 2023-05-16 | StrokeYMax | 104378
| A Inc._Aurora | 2023-05-16 | StrokeY | 212471
| A Inc._Aurora | 2023-05-16 | StrokeXMin | -141117
| A Inc._Aurora | 2023-05-16 | StrokeXMax | 137695
| A Inc._Aurora | 2023-05-16 | StrokeX | 278812
| A Inc._Aurora | 2023-05-16 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.61
| A Inc._Aurora | 2023-05-16 | Roll motion to X+ full scan | 21.3
| A Inc._Aurora | 2023-05-16 | RMS of ZeroScan | 0.042
| A Inc._Aurora | 2023-05-16 | Profiler noise @200um | 0.19
| A Inc._Aurora | 2023-05-16 | PosMin | -6.56
| A Inc._Aurora | 2023-05-16 | PosMax | 6.5
| A Inc._Aurora | 2023-05-16 | Position repeatability(10repeats, Position Based) | 0.18
| A Inc._Aurora | 2023-05-16 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.0001
| A Inc._Aurora | 2023-05-16 | OpticalAxisShiftPerFocusY | 0.0017
| A Inc._Aurora | 2023-05-16 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 6e-05
| A Inc._Aurora | 2023-05-16 | OpticalAxisShiftPerFocusX | -0.0041
| A Inc._Aurora | 2023-05-16 | OPM Repeatability @5mm | 2.033
| A Inc._Aurora | 2023-05-16 | OPM Repeatability @1mm | 1.022
| A Inc._Aurora | 2023-05-16 | OPM Repeatability @10mm | 5.563
| A Inc._Aurora | 2023-05-16 | NCM driving% with AC160TS cantilever | 1.51
| A Inc._Aurora | 2023-05-16 | Min. Vibration Transmissbility (~100Hz) | 0.064
| A Inc._Aurora | 2023-05-16 | MaxLength | 50
| A Inc._Aurora | 2023-05-16 | Max OPM @5mm | 12.094
| A Inc._Aurora | 2023-05-16 | Max OPM @1mm | 7.621
| A Inc._Aurora | 2023-05-16 | Max OPM @10mm | 27.649
| A Inc._Aurora | 2023-05-16 | Lateral(C-D) sensitivity | 3.4
| A Inc._Aurora | 2023-05-16 | Lateral(C-D) noise | 0.125
| A Inc._Aurora | 2023-05-16 | Laser intensity with AC160 cantilever | 3.01
| A Inc._Aurora | 2023-05-16 | Image offset after moving from -Y to +Y (Y scan) | 57.86
| A Inc._Aurora | 2023-05-16 | Image offset after moving from -Y to +Y (X scan) | 36.62
| A Inc._Aurora | 2023-05-16 | Image offset after moving from -X to +X (Y scan) | 89.36
| A Inc._Aurora | 2023-05-16 | Image offset after moving from -X to +X (X scan) | 101.81
| A Inc._Aurora | 2023-05-16 | Image offset after moving from +Y to -Y (Y scan) | 57.86
| A Inc._Aurora | 2023-05-16 | Image offset after moving from +Y to -Y (X scan) | 49.07
| A Inc._Aurora | 2023-05-16 | Image offset after moving from +X to -X (Y scan) | 144.29
| A Inc._Aurora | 2023-05-16 | Image offset after moving from +X to -X (X scan) | 57.86
| A Inc._Aurora | 2023-05-16 | Full range of Z Height during Z Scan sweep | 16.8
| A Inc._Aurora | 2023-05-16 | Full range of Z Detector during Z Scan sweep | 13
| A Inc._Aurora | 2023-05-16 | Full range of Y2 Detector Raw during Y Scan sweep | 14.15
| A Inc._Aurora | 2023-05-16 | Full range of Y2 Detector during Y Scan sweep | 184.4646
| A Inc._Aurora | 2023-05-16 | Full range of Y1 Detector Raw during Y Scan sweep | 14.07
| A Inc._Aurora | 2023-05-16 | Full range of Y1 Detector during Y Scan sweep | 184.38
| A Inc._Aurora | 2023-05-16 | Full range of X2 Detector Raw during X Scan sweep | 14.26
| A Inc._Aurora | 2023-05-16 | Full range of X2 Detector during X Scan sweep | 187.3781
| A Inc._Aurora | 2023-05-16 | Full range of X1 Detector Raw during X Scan sweep | 14.21
| A Inc._Aurora | 2023-05-16 | Full range of X1 Detector during X Scan sweep | 187.2204
| A Inc._Aurora | 2023-05-16 | Check sample surface mapping | 12
| A Inc._Aurora | 2023-05-16 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| A Inc._Aurora | 2023-05-16 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| A Inc._Aurora | 2023-05-16 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.75
| A Inc._Aurora | 2023-05-16 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.14
| A Inc._Aurora | 2023-05-16 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| A Inc._Aurora | 2023-05-16 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.11
| A Inc._Aurora | 2023-05-16 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.15
| A Inc._Aurora | 2023-05-16 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| A Inc._Aurora | 2023-05-16 | A－B difference during 3 times run test | 0.83
| A Inc._Aurora | 2023-05-16 | A-B after Auto PSPD Alignment | -0.15
| A Inc._Aurora | 2023-05-16 | A＋B difference during 3 times run test | 0.12
| A Inc._Aurora | 2023-05-16 | 50~200nm Step Height Sample (10 Repeats) | 183.959
| A Inc._Aurora | 2023-05-16 | 200~500nm Step Height Sample (10 Repeats) | 461.437
| A Inc._Aurora | 2023-05-16 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.6
| A Inc._Aurora | 2023-05-16 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.6
| A Inc._Aurora | 2023-05-16 | %GRR w/ roughness of non-patterened wafer | 9
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Z Scanner position at -3μm | 0.0092
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Z Scanner position at +3μm | 0.0002
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Z Linearity by 3 different sample height | 1.01
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Z Height noise | 0.021
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y2 Detector noise (XY Servo on) | 0.27
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y2 Detector noise (XY Servo off) | 0.2
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y1 Detector noise (XY Servo on) | 0.28
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y1 Detector noise (XY Servo off) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y Scan Stroke | 167.6
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y Detector noise (XY Servo on) | 0.22
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y Detector noise (XY Servo off) | 0.15
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Y axis OPM | 0.9
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X2 Detector noise (XY Servo on) | 0.29
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X2 Detector noise (XY Servo off) | 0.21
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X1 Detector noise (XY Servo on) | 0.31
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X1 Detector noise (XY Servo off) | 0.22
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X Scan Stroke | 169.2
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | -X Direction software Limit | -102094
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | -X Direction software Limit | -103296
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X Detector noise (XY Servo on) | 0.24
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X Detector noise (XY Servo off) | 0.18
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | X axis OPM | 1.18
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.28
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.46
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of Wafer Zeroscan (Top) | 0.32
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of Wafer Zeroscan (Right) | 0.31
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of Wafer Zeroscan (Left) | 0.31
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of Wafer ZeroScan (Center) | 0.3
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.31
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.19
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | NCM driving% with AC160TS cantilever | 1.13
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | NCM Amplitude noise | 12.19
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.3
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.2
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 231.1
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 228
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 232.1
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 228.4
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Lateral(C-D) sensitivity | 9.9
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Lateral(C-D) noise | 0.12
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of Z Height during Z Scan sweep | 16.1
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of Z Detector during Z Scan sweep | 11.4
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of Y1 Detector Raw during Y Scan sweep | 14.6
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of Y Detector during Y Scan sweep | 167.6
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of X2 Detector Raw during X Scan sweep | 14.7
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of X1 Detector Raw during X Scan sweep | 14.7
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | Full range of X Detector during X Scan sweep | 169.1
| Shenzhen Founder Microelectronics International(FMIC) #1 | 2023-05-15 | A-B noise | 7.72
| CXT #2 | 2023-05-11 | Z Linearity by Z scanner position | 0.3
| CXT #3 | 2023-05-11 | Z Linearity by Z scanner position | 1.3
| CXT #2 | 2023-05-11 | Z Linearity by 5 different sample height | 1.002
| CXT #3 | 2023-05-11 | Z Linearity by 5 different sample height | 1.007
| CXT #2 | 2023-05-11 | Z Height noise | 0.04
| CXT #3 | 2023-05-11 | Z Height noise | 0.03
| CXT #2 | 2023-05-11 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 29.1
| CXT #3 | 2023-05-11 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 27.66
| CXT #2 | 2023-05-11 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.7
| CXT #3 | 2023-05-11 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.75
| CXT #2 | 2023-05-11 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.8
| CXT #3 | 2023-05-11 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.26
| CXT #2 | 2023-05-11 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.2
| CXT #3 | 2023-05-11 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.31
| CXT #2 | 2023-05-11 | Z Detector offset | 43.5
| CXT #3 | 2023-05-11 | Z Detector offset | 62.5
| CXT #2 | 2023-05-11 | Y2 Detector noise (XY Servo on) | 0.542
| CXT #3 | 2023-05-11 | Y2 Detector noise (XY Servo on) | 0.36
| CXT #2 | 2023-05-11 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 269.8
| CXT #3 | 2023-05-11 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 275
| CXT #2 | 2023-05-11 | Y1 Detector noise (XY Servo on) | 0.527
| CXT #3 | 2023-05-11 | Y1 Detector noise (XY Servo on) | 0.39
| CXT #2 | 2023-05-11 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 281.7
| CXT #3 | 2023-05-11 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 278.5
| CXT #2 | 2023-05-11 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 196.1
| CXT #3 | 2023-05-11 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 165.6
| CXT #2 | 2023-05-11 | Y Ringing(Rpv) | 0.777
| CXT #3 | 2023-05-11 | Y Ringing(Rpv) | 0.6
| CXT #2 | 2023-05-11 | Y Non-linearity error | 2.58
| CXT #3 | 2023-05-11 | Y Non-linearity error | 3.32
| CXT #2 | 2023-05-11 | Y Detector noise (XY Servo on) | 0.451
| CXT #3 | 2023-05-11 | Y Detector noise (XY Servo on) | 0.28
| CXT #2 | 2023-05-11 | Y axis, Repeatability @1st position | 0.04
| CXT #3 | 2023-05-11 | Y axis, Repeatability @1st position | 0.06
| CXT #2 | 2023-05-11 | Y axis, Repeatability @ 4th position | 0.03
| CXT #3 | 2023-05-11 | Y axis, Repeatability @ 4th position | 0.08
| CXT #2 | 2023-05-11 | Y axis, Repeatability @ 3rd position | 0.08
| CXT #3 | 2023-05-11 | Y axis, Repeatability @ 3rd position | 0.05
| CXT #2 | 2023-05-11 | Y axis, Repeatability @ 2nd position | 0.06
| CXT #3 | 2023-05-11 | Y axis, Repeatability @ 2nd position | 0.08
| CXT #2 | 2023-05-11 | Y axis, Accuracy @1st position | 1.11
| CXT #3 | 2023-05-11 | Y axis, Accuracy @1st position | 1.15
| CXT #2 | 2023-05-11 | Y axis, Accuracy @ 4th position | 0.28
| CXT #3 | 2023-05-11 | Y axis, Accuracy @ 4th position | 0.7
| CXT #2 | 2023-05-11 | Y axis, Accuracy @ 3rd position | 0.51
| CXT #3 | 2023-05-11 | Y axis, Accuracy @ 3rd position | 0.73
| CXT #2 | 2023-05-11 | Y axis, Accuracy @ 2nd position | 1.42
| CXT #3 | 2023-05-11 | Y axis, Accuracy @ 2nd position | 0.9
| CXT #2 | 2023-05-11 | Y axis OPM | 0.697
| CXT #3 | 2023-05-11 | Y axis OPM | 3.62
| CXT #2 | 2023-05-11 | XY Orthogonality | 89.7
| CXT #3 | 2023-05-11 | XY Orthogonality | 90.2
| CXT #2 | 2023-05-11 | X2 Detector noise (XY Servo on) | 0.491
| CXT #3 | 2023-05-11 | X2 Detector noise (XY Servo on) | 0.39
| CXT #2 | 2023-05-11 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 269.7
| CXT #3 | 2023-05-11 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 284.9
| CXT #2 | 2023-05-11 | X1 Detector noise (XY Servo on) | 0.471
| CXT #3 | 2023-05-11 | X1 Detector noise (XY Servo on) | 0.38
| CXT #2 | 2023-05-11 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 265.6
| CXT #3 | 2023-05-11 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 280.8
| CXT #2 | 2023-05-11 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 179.5
| CXT #3 | 2023-05-11 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 187.8
| CXT #2 | 2023-05-11 | X Ringing(Rpv) | 0.718
| CXT #3 | 2023-05-11 | X Ringing(Rpv) | 0.49
| CXT #2 | 2023-05-11 | X Non-linearity error | 2.27
| CXT #3 | 2023-05-11 | X Non-linearity error | 1.43
| CXT #2 | 2023-05-11 | X Detector noise (XY Servo on) | 0.39
| CXT #3 | 2023-05-11 | X Detector noise (XY Servo on) | 0.3
| CXT #2 | 2023-05-11 | X axis, Repeatability @1st position | 0.03
| CXT #3 | 2023-05-11 | X axis, Repeatability @1st position | 0.05
| CXT #2 | 2023-05-11 | X axis, Repeatability @ 4th position | 0.03
| CXT #3 | 2023-05-11 | X axis, Repeatability @ 4th position | 0.05
| CXT #2 | 2023-05-11 | X axis, Repeatability @ 3rd position | 0.04
| CXT #3 | 2023-05-11 | X axis, Repeatability @ 3rd position | 0.04
| CXT #2 | 2023-05-11 | X axis, Repeatability @ 2nd position | 0.03
| CXT #3 | 2023-05-11 | X axis, Repeatability @ 2nd position | 0.05
| CXT #2 | 2023-05-11 | X axis, Accuracy @1st position | 0.3
| CXT #3 | 2023-05-11 | X axis, Accuracy @1st position | 0.32
| CXT #2 | 2023-05-11 | X axis, Accuracy @ 4th position | 0.27
| CXT #3 | 2023-05-11 | X axis, Accuracy @ 4th position | 1.18
| CXT #2 | 2023-05-11 | X axis, Accuracy @ 3rd position | 1.16
| CXT #3 | 2023-05-11 | X axis, Accuracy @ 3rd position | 0.31
| CXT #2 | 2023-05-11 | X axis, Accuracy @ 2nd position | 0.34
| CXT #3 | 2023-05-11 | X axis, Accuracy @ 2nd position | 0.26
| CXT #2 | 2023-05-11 | X axis OPM | 0.595
| CXT #3 | 2023-05-11 | X axis OPM | 2.9
| CXT #2 | 2023-05-11 | Under 50nm Step Height Sample (10 Repeats) | 18.13
| CXT #3 | 2023-05-11 | Under 50nm Step Height Sample (10 Repeats) | 18.07
| CXT #2 | 2023-05-11 | Torque Profile (50mm, 20um/s) | 2.322
| CXT #3 | 2023-05-11 | Torque Profile (50mm, 20um/s) | 2.652
| CXT #2 | 2023-05-11 | Torque Profile (50mm, 200um/s) | 1.91
| CXT #3 | 2023-05-11 | Torque Profile (50mm, 200um/s) | 2.304
| CXT #2 | 2023-05-11 | StrokeYMin | -153973
| CXT #3 | 2023-05-11 | StrokeYMin | -154032
| CXT #2 | 2023-05-11 | StrokeYMax | 150452.2
| CXT #3 | 2023-05-11 | StrokeYMax | 150139
| CXT #2 | 2023-05-11 | StrokeY | 304425.2
| CXT #3 | 2023-05-11 | StrokeY | 304171
| CXT #2 | 2023-05-11 | StrokeXMin | -203900
| CXT #3 | 2023-05-11 | StrokeXMin | -204538
| CXT #2 | 2023-05-11 | StrokeXMax | 203997.1
| CXT #3 | 2023-05-11 | StrokeXMax | 203523
| CXT #2 | 2023-05-11 | StrokeX | 407896.9
| CXT #3 | 2023-05-11 | StrokeX | 408061
| CXT #2 | 2023-05-11 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.6
| CXT #3 | 2023-05-11 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.47
| CXT #2 | 2023-05-11 | Roll motion to X+ full scan | 2.1
| CXT #3 | 2023-05-11 | Roll motion to X+ full scan | 14.5
| CXT #2 | 2023-05-11 | RMS of ZeroScan | 0.055
| CXT #3 | 2023-05-11 | RMS of ZeroScan | 0.065
| CXT #2 | 2023-05-11 | Profiler noise @200um | 0.108
| CXT #3 | 2023-05-11 | Profiler noise @200um | 0.162
| CXT #2 | 2023-05-11 | PosMin | -26
| CXT #3 | 2023-05-11 | PosMin | -26
| CXT #2 | 2023-05-11 | PosMax | 25.9
| CXT #3 | 2023-05-11 | PosMax | 26.2
| CXT #2 | 2023-05-11 | Position repeatability(10repeats, Position Based) | 0.58
| CXT #3 | 2023-05-11 | Position repeatability(10repeats, Position Based) | 0.45
| CXT #2 | 2023-05-11 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 9e-05
| CXT #3 | 2023-05-11 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5.68e-05
| CXT #2 | 2023-05-11 | OpticalAxisShiftPerFocusY | -0.0016
| CXT #3 | 2023-05-11 | OpticalAxisShiftPerFocusY | -0.0024
| CXT #2 | 2023-05-11 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00014
| CXT #3 | 2023-05-11 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7.07e-05
| CXT #2 | 2023-05-11 | OpticalAxisShiftPerFocusX | -0.0006
| CXT #3 | 2023-05-11 | OpticalAxisShiftPerFocusX | 0.0018
| CXT #2 | 2023-05-11 | OPM Repeatability @5mm | 1.18
| CXT #3 | 2023-05-11 | OPM Repeatability @5mm | 1.268
| CXT #2 | 2023-05-11 | OPM Repeatability @25mm | 3.15
| CXT #3 | 2023-05-11 | OPM Repeatability @25mm | 3.131
| CXT #2 | 2023-05-11 | OPM Repeatability @1mm | 1.49
| CXT #3 | 2023-05-11 | OPM Repeatability @1mm | 0.397
| CXT #2 | 2023-05-11 | OPM Repeatability @10mm | 1.36
| CXT #3 | 2023-05-11 | OPM Repeatability @10mm | 1.103
| CXT #2 | 2023-05-11 | NCM driving% with AC160TS cantilever | 1.3
| CXT #3 | 2023-05-11 | NCM driving% with AC160TS cantilever | 26.33
| CXT #2 | 2023-05-11 | Min. Vibration Transmissbility (~100Hz) | 0.089
| CXT #3 | 2023-05-11 | Min. Vibration Transmissbility (~100Hz) | 0.095
| CXT #2 | 2023-05-11 | MaxLength | 52
| CXT #3 | 2023-05-11 | MaxLength | 52.2
| CXT #2 | 2023-05-11 | Max OPM @5mm | 39.87
| CXT #3 | 2023-05-11 | Max OPM @5mm | 12.349
| CXT #2 | 2023-05-11 | Max OPM @25mm | 228.38
| CXT #3 | 2023-05-11 | Max OPM @25mm | 216.732
| CXT #2 | 2023-05-11 | Max OPM @1mm | 14.43
| CXT #3 | 2023-05-11 | Max OPM @1mm | 4.881
| CXT #2 | 2023-05-11 | Max OPM @10mm | 65.41
| CXT #3 | 2023-05-11 | Max OPM @10mm | 58.611
| CXT #2 | 2023-05-11 | Lateral(C-D) sensitivity | 3.8
| CXT #3 | 2023-05-11 | Lateral(C-D) sensitivity | 3
| CXT #2 | 2023-05-11 | Lateral(C-D) noise | 0.222
| CXT #3 | 2023-05-11 | Lateral(C-D) noise | 0.169
| CXT #2 | 2023-05-11 | Laser intensity with AC160 cantilever | 2.91
| CXT #3 | 2023-05-11 | Laser intensity with AC160 cantilever | 3.021
| CXT #2 | 2023-05-11 | Image offset after moving from -Y to +Y (Y scan) | 93.02
| CXT #3 | 2023-05-11 | Image offset after moving from -Y to +Y (Y scan) | 66.65
| CXT #2 | 2023-05-11 | Image offset after moving from -Y to +Y (X scan) | 21.24
| CXT #3 | 2023-05-11 | Image offset after moving from -Y to +Y (X scan) | 10.25
| CXT #2 | 2023-05-11 | Image offset after moving from -X to +X (Y scan) | 146.48
| CXT #3 | 2023-05-11 | Image offset after moving from -X to +X (Y scan) | 21.97
| CXT #2 | 2023-05-11 | Image offset after moving from -X to +X (X scan) | 40.28
| CXT #3 | 2023-05-11 | Image offset after moving from -X to +X (X scan) | 134.03
| CXT #2 | 2023-05-11 | Image offset after moving from +Y to -Y (Y scan) | 108.4
| CXT #3 | 2023-05-11 | Image offset after moving from +Y to -Y (Y scan) | 21.97
| CXT #2 | 2023-05-11 | Image offset after moving from +Y to -Y (X scan) | 58.59
| CXT #3 | 2023-05-11 | Image offset after moving from +Y to -Y (X scan) | 33.69
| CXT #2 | 2023-05-11 | Image offset after moving from +X to -X (Y scan) | 87.16
| CXT #3 | 2023-05-11 | Image offset after moving from +X to -X (Y scan) | 183.11
| CXT #2 | 2023-05-11 | Image offset after moving from +X to -X (X scan) | 81.3
| CXT #3 | 2023-05-11 | Image offset after moving from +X to -X (X scan) | 33.69
| CXT #2 | 2023-05-11 | Full range of Z Height during Z Scan sweep | 16.4
| CXT #3 | 2023-05-11 | Full range of Z Height during Z Scan sweep | 16.53
| CXT #2 | 2023-05-11 | Full range of Z Detector during Z Scan sweep | 11.2
| CXT #3 | 2023-05-11 | Full range of Z Detector during Z Scan sweep | 11.9
| CXT #2 | 2023-05-11 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| CXT #3 | 2023-05-11 | Full range of Y2 Detector Raw during Y Scan sweep | 13.45
| CXT #2 | 2023-05-11 | Full range of Y2 Detector during Y Scan sweep | 203.8328
| CXT #3 | 2023-05-11 | Full range of Y2 Detector during Y Scan sweep | 195.6762
| CXT #2 | 2023-05-11 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| CXT #3 | 2023-05-11 | Full range of Y1 Detector Raw during Y Scan sweep | 13.18
| CXT #2 | 2023-05-11 | Full range of Y1 Detector during Y Scan sweep | 206.158
| CXT #3 | 2023-05-11 | Full range of Y1 Detector during Y Scan sweep | 200.9646
| CXT #2 | 2023-05-11 | Full range of X2 Detector Raw during X Scan sweep | 13.8
| CXT #3 | 2023-05-11 | Full range of X2 Detector Raw during X Scan sweep | 13.4
| CXT #2 | 2023-05-11 | Full range of X2 Detector during X Scan sweep | 195.411
| CXT #3 | 2023-05-11 | Full range of X2 Detector during X Scan sweep | 199.366
| CXT #2 | 2023-05-11 | Full range of X1 Detector Raw during X Scan sweep | 13.7
| CXT #3 | 2023-05-11 | Full range of X1 Detector Raw during X Scan sweep | 13.57
| CXT #2 | 2023-05-11 | Full range of X1 Detector during X Scan sweep | 195.9273
| CXT #3 | 2023-05-11 | Full range of X1 Detector during X Scan sweep | 198.1836
| CXT #3 | 2023-05-11 | Check sample surface mapping | 24
| CXT #2 | 2023-05-11 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| CXT #3 | 2023-05-11 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| CXT #2 | 2023-05-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| CXT #3 | 2023-05-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.06
| CXT #2 | 2023-05-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.3
| CXT #3 | 2023-05-11 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.34
| CXT #2 | 2023-05-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.42
| CXT #3 | 2023-05-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.46
| CXT #2 | 2023-05-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| CXT #3 | 2023-05-11 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| CXT #2 | 2023-05-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| CXT #3 | 2023-05-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.16
| CXT #2 | 2023-05-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| CXT #3 | 2023-05-11 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.89
| CXT #2 | 2023-05-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.07
| CXT #3 | 2023-05-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| CXT #2 | 2023-05-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.01
| CXT #3 | 2023-05-11 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| CXT #2 | 2023-05-11 | A－B difference during 3 times run test | 0.55
| CXT #3 | 2023-05-11 | A－B difference during 3 times run test | 0.41
| CXT #2 | 2023-05-11 | A-B after Auto PSPD Alignment | 0.02
| CXT #3 | 2023-05-11 | A-B after Auto PSPD Alignment | 0.12
| CXT #2 | 2023-05-11 | A＋B difference during 3 times run test | 0.04
| CXT #3 | 2023-05-11 | A＋B difference during 3 times run test | 0.16
| CXT #2 | 2023-05-11 | 50~200nm Step Height Sample (10 Repeats) | 182.97
| CXT #3 | 2023-05-11 | 50~200nm Step Height Sample (10 Repeats) | 183.48
| CXT #2 | 2023-05-11 | 200~500nm Step Height Sample (10 Repeats) | 462.41
| CXT #3 | 2023-05-11 | 200~500nm Step Height Sample (10 Repeats) | 464.38
| CXT #2 | 2023-05-11 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.1
| CXT #3 | 2023-05-11 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 4.5
| CXT #2 | 2023-05-11 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.8
| CXT #3 | 2023-05-11 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 2.6
| CXT #2 | 2023-05-11 | %GRR w/ roughness of non-patterened wafer | 5.8
| CXT #3 | 2023-05-11 | %GRR w/ roughness of non-patterened wafer | 5.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Z Scanner position at -3μm | 0.0065
| 동아대학교_신소재공학과 | 2023-05-10 | Z Scanner position at -3μm | 0.0084
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Z Scanner position at +3μm | -0.0032
| 동아대학교_신소재공학과 | 2023-05-10 | Z Scanner position at +3μm | -0.0087
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Z Linearity by 3 different sample height | 1.012
| 동아대학교_신소재공학과 | 2023-05-10 | Z Linearity by 3 different sample height | 1.009
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Z Height noise | 0.031
| 동아대학교_신소재공학과 | 2023-05-10 | Z Height noise | 0.021
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y2 Detector noise (XY Servo on) | 0.35
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y2 Detector noise (XY Servo off) | 0.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y1 Detector noise (XY Servo on) | 0.34
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y1 Detector noise (XY Servo off) | 0.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y Scan Stroke | 176.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y Detector noise (XY Servo on) | 0.27
| 동아대학교_신소재공학과 | 2023-05-10 | Y Detector noise (XY Servo on) | 0.29
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y Detector noise (XY Servo off) | 0.15
| 동아대학교_신소재공학과 | 2023-05-10 | Y Detector noise (XY Servo off) | 0.28
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Y axis OPM | 0.64
| 동아대학교_신소재공학과 | 2023-05-10 | Y axis OPM | 0.78
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X2 Detector noise (XY Servo on) | 0.37
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X2 Detector noise (XY Servo off) | 0.24
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X1 Detector noise (XY Servo on) | 0.36
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X1 Detector noise (XY Servo off) | 0.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X Scan Stroke | 178
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | -X Direction software Limit | -102476
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | -X Direction software Limit | -102681
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X Detector noise (XY Servo on) | 0.3
| 동아대학교_신소재공학과 | 2023-05-10 | X Detector noise (XY Servo on) | 0.28
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X Detector noise (XY Servo off) | 0.16
| 동아대학교_신소재공학과 | 2023-05-10 | X Detector noise (XY Servo off) | 0.28
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | X axis OPM | 0.86
| 동아대학교_신소재공학과 | 2023-05-10 | X axis OPM | 1.12
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| 동아대학교_신소재공학과 | 2023-05-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| 동아대학교_신소재공학과 | 2023-05-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of Wafer Zeroscan (Top) | 0.27
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of Wafer Zeroscan (Right) | 0.21
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of Wafer Zeroscan (Left) | 0.24
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of Wafer ZeroScan (Center) | 0.23
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.26
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.31
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | NCM driving% with AC160TS cantilever | 3.43
| 동아대학교_신소재공학과 | 2023-05-10 | NCM driving% with AC160TS cantilever | 2.42
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | NCM Amplitude noise | 14.97
| 동아대학교_신소재공학과 | 2023-05-10 | NCM Amplitude noise | 16.39
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.7
| 동아대학교_신소재공학과 | 2023-05-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.1
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.2
| 동아대학교_신소재공학과 | 2023-05-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 247.3
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 248.9
| 동아대학교_신소재공학과 | 2023-05-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 99.5
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 250.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249.4
| 동아대학교_신소재공학과 | 2023-05-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 103
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.06
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Lateral(C-D) sensitivity | 20.243
| 동아대학교_신소재공학과 | 2023-05-10 | Lateral(C-D) sensitivity | 21.805
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Lateral(C-D) noise | 0.1
| 동아대학교_신소재공학과 | 2023-05-10 | Lateral(C-D) noise | 0.027
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of Z Height during Z Scan sweep | 15.7
| 동아대학교_신소재공학과 | 2023-05-10 | Full range of Z Height during Z Scan sweep | 15.9
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of Z Detector during Z Scan sweep | 11.1
| 동아대학교_신소재공학과 | 2023-05-10 | Full range of Z Detector during Z Scan sweep | 11.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| 동아대학교_신소재공학과 | 2023-05-10 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of Y Detector during Y Scan sweep | 176.2
| 동아대학교_신소재공학과 | 2023-05-10 | Full range of Y Detector during Y Scan sweep | 72.9
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of X2 Detector Raw during X Scan sweep | 14.1
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| 동아대학교_신소재공학과 | 2023-05-10 | Full range of X Detector Raw during X Scan sweep | 14.4
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | Full range of X Detector during X Scan sweep | 178.1
| 동아대학교_신소재공학과 | 2023-05-10 | Full range of X Detector during X Scan sweep | 74.2
| Shenzhen Heavy Investment Tianke Semiconductor | 2023-05-10 | A-B noise | 11.56
| 동아대학교_신소재공학과 | 2023-05-10 | A-B noise | 13.67
| Intel Chandler | 2023-05-08 | Z Height noise | 0.028
| Intel Chandler | 2023-05-08 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.8924
| Intel Chandler | 2023-05-08 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.5209
| Intel Chandler | 2023-05-08 | Z Direction, Accuracy Test (Bias not significant) | 1.006
| Intel Chandler | 2023-05-08 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 210
| Intel Chandler | 2023-05-08 | Y Direction, Accuracy Test (Bias not significant) | 1.545
| Intel Chandler | 2023-05-08 | Y Detector noise (XY Servo on) | 0.379
| Intel Chandler | 2023-05-08 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 162.732
| Intel Chandler | 2023-05-08 | Y axis, Repeatability @1st position | 0.04
| Intel Chandler | 2023-05-08 | Y axis, Repeatability @ 4th position | 0.04
| Intel Chandler | 2023-05-08 | Y axis, Repeatability @ 3rd position | 0.04
| Intel Chandler | 2023-05-08 | Y axis, Repeatability @ 2nd position | 0.04
| Intel Chandler | 2023-05-08 | Y axis, Accuracy @1st position | 0.41
| Intel Chandler | 2023-05-08 | Y axis, Accuracy @ 4th position | 0.49
| Intel Chandler | 2023-05-08 | Y axis, Accuracy @ 3rd position | 0.45
| Intel Chandler | 2023-05-08 | Y axis, Accuracy @ 2nd position | 0.46
| Intel Chandler | 2023-05-08 | Y axis OPM | 3.3
| Intel Chandler | 2023-05-08 | XY Orthogonality | 90.2
| Intel Chandler | 2023-05-08 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 193.548
| Intel Chandler | 2023-05-08 | X Direction, Accuracy Test (Bias not significant) | 0.388
| Intel Chandler | 2023-05-08 | X Detector noise (XY Servo on) | 0.258
| Intel Chandler | 2023-05-08 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 162.906
| Intel Chandler | 2023-05-08 | X axis, Repeatability @1st position | 0.04
| Intel Chandler | 2023-05-08 | X axis, Repeatability @ 4th position | 0.03
| Intel Chandler | 2023-05-08 | X axis, Repeatability @ 3rd position | 0.02
| Intel Chandler | 2023-05-08 | X axis, Repeatability @ 2nd position | 0.03
| Intel Chandler | 2023-05-08 | X axis, Accuracy @1st position | 0.12
| Intel Chandler | 2023-05-08 | X axis, Accuracy @ 4th position | 0.11
| Intel Chandler | 2023-05-08 | X axis, Accuracy @ 3rd position | 0.08
| Intel Chandler | 2023-05-08 | X axis, Accuracy @ 2nd position | 0.09
| Intel Chandler | 2023-05-08 | X axis OPM | 4.086
| Intel Chandler | 2023-05-08 | StrokeYMin | -68130
| Intel Chandler | 2023-05-08 | StrokeYMax | 562363
| Intel Chandler | 2023-05-08 | StrokeY | 630493
| Intel Chandler | 2023-05-08 | StrokeXMin | -520665
| Intel Chandler | 2023-05-08 | StrokeXMax | 424359
| Intel Chandler | 2023-05-08 | StrokeX | 945024
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (SR Panel) | 15.5
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (Smooth Cu Panel, Trace) | 1.4
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (Smooth Cu Panel, Plane) | 6.2
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (SiNx Panel) | 16.8
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (Rough Cu Panel, Trace) | 14
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (Rough Cu Panel, Plane) | 4.7
| Intel Chandler | 2023-05-08 | Roughness Reproducibility Test (ABF Panel) | 22.3
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (SR Panel) | 8.9
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (Smooth Cu Panel, Trace) | 0.7
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (Smooth Cu Panel, Plane) | 2.3
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (SiNx Panel) | 19.99
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (Rough Cu Panel, Trace) | 4.3
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (Rough Cu Panel, Plane) | 9.6
| Intel Chandler | 2023-05-08 | Roughness Repeatability Test (ABF Panel) | 10.4
| Intel Chandler | 2023-05-08 | RMS of ZeroScan | 2.44
| Intel Chandler | 2023-05-08 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 6.75e-05
| Intel Chandler | 2023-05-08 | OpticalAxisShiftPerFocusY | -0.0014
| Intel Chandler | 2023-05-08 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 7.89e-05
| Intel Chandler | 2023-05-08 | OpticalAxisShiftPerFocusX | 0.006
| Intel Chandler | 2023-05-08 | NCM driving% with AC160TS cantilever | 18.21
| Intel Chandler | 2023-05-08 | Min. Vibration Transmissbility (~100Hz) | 0.08
| Intel Chandler | 2023-05-08 | Lift Pin Stroke Min Check in DB | -4450
| Intel Chandler | 2023-05-08 | Lift Pin Stroke Max Check in DB | 106000
| Intel Chandler | 2023-05-08 | Laser intensity with AC160 cantilever | 3.118
| Intel Chandler | 2023-05-08 | Full range of Z Height during Z Scan sweep | 17.39
| Intel Chandler | 2023-05-08 | Full range of Z Height during Z Scan sweep | 101.01
| Intel Chandler | 2023-05-08 | Full range of Y Detector Raw during Y Scan sweep | 13.84
| Intel Chandler | 2023-05-08 | Full range of Y Detector during Y Scan sweep | 101.01
| Intel Chandler | 2023-05-08 | Full range of X Detector Raw during X Scan sweep | 13.72
| Intel Chandler | 2023-05-08 | A－B difference during 3 times run test | 0
| Intel Chandler | 2023-05-08 | A-B after Auto PSPD Alignment | 0.17
| Intel Chandler | 2023-05-08 | A＋B difference during 3 times run test | 0.21
| Lanzhou University | 2023-05-03 | Z Scanner position at -3μm | 0.0046
| Tianjin Normal University | 2023-05-03 | Z Scanner position at -3μm | -0.0107
| Lanzhou University | 2023-05-03 | Z Scanner position at +3μm | 0.0129
| Tianjin Normal University | 2023-05-03 | Z Scanner position at +3μm | -0.0015
| Lanzhou University | 2023-05-03 | Z Linearity by 3 different sample height | 0.987
| Tianjin Normal University | 2023-05-03 | Z Linearity by 3 different sample height | 0.995
| Lanzhou University | 2023-05-03 | Z Height noise | 0.041
| Tianjin Normal University | 2023-05-03 | Z Height noise | 0.024
| Lanzhou University | 2023-05-03 | Y Detector noise (XY Servo on) | 0.31
| Tianjin Normal University | 2023-05-03 | Y Detector noise (XY Servo on) | 0.33
| Lanzhou University | 2023-05-03 | Y Detector noise (XY Servo off) | 0.28
| Tianjin Normal University | 2023-05-03 | Y Detector noise (XY Servo off) | 0.27
| Lanzhou University | 2023-05-03 | Y axis OPM | 0.59
| Tianjin Normal University | 2023-05-03 | Y axis OPM | 1.26
| Lanzhou University | 2023-05-03 | X Detector noise (XY Servo on) | 0.3
| Tianjin Normal University | 2023-05-03 | X Detector noise (XY Servo on) | 0.32
| Lanzhou University | 2023-05-03 | X Detector noise (XY Servo off) | 0.27
| Tianjin Normal University | 2023-05-03 | X Detector noise (XY Servo off) | 0.27
| Lanzhou University | 2023-05-03 | X axis OPM | 1.08
| Tianjin Normal University | 2023-05-03 | X axis OPM | 1.26
| Lanzhou University | 2023-05-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Tianjin Normal University | 2023-05-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Lanzhou University | 2023-05-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| Tianjin Normal University | 2023-05-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.48
| Lanzhou University | 2023-05-03 | NCM driving% with AC160TS cantilever | 7
| Tianjin Normal University | 2023-05-03 | NCM driving% with AC160TS cantilever | 0.53
| Lanzhou University | 2023-05-03 | NCM Amplitude noise | 14
| Tianjin Normal University | 2023-05-03 | NCM Amplitude noise | 13.61
| Lanzhou University | 2023-05-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.45
| Tianjin Normal University | 2023-05-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.7
| Lanzhou University | 2023-05-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.866
| Tianjin Normal University | 2023-05-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.8
| Lanzhou University | 2023-05-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 156.2
| Tianjin Normal University | 2023-05-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 165.1
| Lanzhou University | 2023-05-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 160
| Tianjin Normal University | 2023-05-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 164.2
| Lanzhou University | 2023-05-03 | Lateral(C-D) sensitivity | 6.9
| Tianjin Normal University | 2023-05-03 | Lateral(C-D) sensitivity | 4.2
| Lanzhou University | 2023-05-03 | Lateral(C-D) noise | 0.086
| Tianjin Normal University | 2023-05-03 | Lateral(C-D) noise | 0.091
| Lanzhou University | 2023-05-03 | Full range of Z Height during Z Scan sweep | 16.6
| Tianjin Normal University | 2023-05-03 | Full range of Z Height during Z Scan sweep | 17
| Lanzhou University | 2023-05-03 | Full range of Z Detector during Z Scan sweep | 11.5
| Tianjin Normal University | 2023-05-03 | Full range of Z Detector during Z Scan sweep | 12.5
| Lanzhou University | 2023-05-03 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Tianjin Normal University | 2023-05-03 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Lanzhou University | 2023-05-03 | Full range of Y Detector during Y Scan sweep | 112.2
| Tianjin Normal University | 2023-05-03 | Full range of Y Detector during Y Scan sweep | 116.2
| Lanzhou University | 2023-05-03 | Full range of X Detector Raw during X Scan sweep | 14.1
| Tianjin Normal University | 2023-05-03 | Full range of X Detector Raw during X Scan sweep | 14.2
| Lanzhou University | 2023-05-03 | Full range of X Detector during X Scan sweep | 112.9
| Tianjin Normal University | 2023-05-03 | Full range of X Detector during X Scan sweep | 116.8
| Lanzhou University | 2023-05-03 | A-B noise | 8
| Tianjin Normal University | 2023-05-03 | A-B noise | 10.73
| Jincheng industrial research institute | 2023-04-28 | Z Scanner position at -3μm | 0.012
| 한국화학연구원 | 2023-04-28 | Z Scanner position at -3μm | 0.0035
| Jincheng industrial research institute | 2023-04-28 | Z Scanner position at +3μm | 0.0047
| 한국화학연구원 | 2023-04-28 | Z Scanner position at +3μm | 0.0016
| Jincheng industrial research institute | 2023-04-28 | Z Linearity by 3 different sample height | 1.005
| 한국화학연구원 | 2023-04-28 | Z Linearity by 3 different sample height | 0.997
| Jincheng industrial research institute | 2023-04-28 | Z Height noise | 0.022
| 한국화학연구원 | 2023-04-28 | Z Height noise | 0.021
| Jincheng industrial research institute | 2023-04-28 | Y2 Detector noise (XY Servo on) | 0.27
| Jincheng industrial research institute | 2023-04-28 | Y2 Detector noise (XY Servo off) | 0.21
| Jincheng industrial research institute | 2023-04-28 | Y1 Detector noise (XY Servo on) | 0.27
| Jincheng industrial research institute | 2023-04-28 | Y1 Detector noise (XY Servo off) | 0.21
| Jincheng industrial research institute | 2023-04-28 | Y Scan Stroke | 170.2
| Jincheng industrial research institute | 2023-04-28 | Y Detector noise (XY Servo on) | 0.21
| 한국화학연구원 | 2023-04-28 | Y Detector noise (XY Servo on) | 0.27
| Jincheng industrial research institute | 2023-04-28 | Y Detector noise (XY Servo off) | 0.15
| 한국화학연구원 | 2023-04-28 | Y Detector noise (XY Servo off) | 0.27
| Jincheng industrial research institute | 2023-04-28 | Y axis OPM | 0.94
| 한국화학연구원 | 2023-04-28 | Y axis OPM | 1.39
| Jincheng industrial research institute | 2023-04-28 | X2 Detector noise (XY Servo on) | 0.28
| Jincheng industrial research institute | 2023-04-28 | X2 Detector noise (XY Servo off) | 0.21
| Jincheng industrial research institute | 2023-04-28 | X1 Detector noise (XY Servo on) | 0.27
| Jincheng industrial research institute | 2023-04-28 | X1 Detector noise (XY Servo off) | 0.2
| Jincheng industrial research institute | 2023-04-28 | X Scan Stroke | 170.6
| Jincheng industrial research institute | 2023-04-28 | -X Direction software Limit | -74948.8
| Jincheng industrial research institute | 2023-04-28 | -X Direction software Limit | -75009.4
| Jincheng industrial research institute | 2023-04-28 | X Detector noise (XY Servo on) | 0.22
| 한국화학연구원 | 2023-04-28 | X Detector noise (XY Servo on) | 0.28
| Jincheng industrial research institute | 2023-04-28 | X Detector noise (XY Servo off) | 0.15
| 한국화학연구원 | 2023-04-28 | X Detector noise (XY Servo off) | 0.27
| Jincheng industrial research institute | 2023-04-28 | X axis OPM | 0.7
| 한국화학연구원 | 2023-04-28 | X axis OPM | 0.96
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한국화학연구원 | 2023-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.33
| 한국화학연구원 | 2023-04-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of Wafer Zeroscan (Top) | 0.25
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of Wafer Zeroscan (Right) | 0.24
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of Wafer Zeroscan (Left) | 0.25
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of Wafer ZeroScan (Center) | 0.24
| Jincheng industrial research institute | 2023-04-28 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.24
| Jincheng industrial research institute | 2023-04-28 | NCM driving% with AC160TS cantilever | 1.5
| 한국화학연구원 | 2023-04-28 | NCM driving% with AC160TS cantilever | 0.46
| Jincheng industrial research institute | 2023-04-28 | NCM Amplitude noise | 15.31
| 한국화학연구원 | 2023-04-28 | NCM Amplitude noise | 11.86
| Jincheng industrial research institute | 2023-04-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| 한국화학연구원 | 2023-04-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.9
| Jincheng industrial research institute | 2023-04-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.4
| 한국화학연구원 | 2023-04-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.4
| Jincheng industrial research institute | 2023-04-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 241.4
| Jincheng industrial research institute | 2023-04-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 236.4
| 한국화학연구원 | 2023-04-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.5
| Jincheng industrial research institute | 2023-04-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 248.9
| Jincheng industrial research institute | 2023-04-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 238.6
| 한국화학연구원 | 2023-04-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.4
| Jincheng industrial research institute | 2023-04-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.01
| Jincheng industrial research institute | 2023-04-28 | Lateral(C-D) sensitivity | 10.934
| 한국화학연구원 | 2023-04-28 | Lateral(C-D) sensitivity | 5.6
| Jincheng industrial research institute | 2023-04-28 | Lateral(C-D) noise | 0.084
| 한국화학연구원 | 2023-04-28 | Lateral(C-D) noise | 0.08
| Jincheng industrial research institute | 2023-04-28 | Full range of Z Height during Z Scan sweep | 16.2
| 한국화학연구원 | 2023-04-28 | Full range of Z Height during Z Scan sweep | 17
| Jincheng industrial research institute | 2023-04-28 | Full range of Z Detector during Z Scan sweep | 11.8
| 한국화학연구원 | 2023-04-28 | Full range of Z Detector during Z Scan sweep | 11.9
| Jincheng industrial research institute | 2023-04-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Jincheng industrial research institute | 2023-04-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| 한국화학연구원 | 2023-04-28 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Jincheng industrial research institute | 2023-04-28 | Full range of Y Detector during Y Scan sweep | 169.8
| 한국화학연구원 | 2023-04-28 | Full range of Y Detector during Y Scan sweep | 74.3
| Jincheng industrial research institute | 2023-04-28 | Full range of X2 Detector Raw during X Scan sweep | 14
| Jincheng industrial research institute | 2023-04-28 | Full range of X1 Detector Raw during X Scan sweep | 14.2
| 한국화학연구원 | 2023-04-28 | Full range of X Detector Raw during X Scan sweep | 13.8
| Jincheng industrial research institute | 2023-04-28 | Full range of X Detector during X Scan sweep | 170.4
| 한국화학연구원 | 2023-04-28 | Full range of X Detector during X Scan sweep | 72.5
| Jincheng industrial research institute | 2023-04-28 | A-B noise | 8.9
| 한국화학연구원 | 2023-04-28 | A-B noise | 9.47
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Z Scanner position at -3μm | 0.0065
| Tel Aviv University | 2023-04-26 | Z Scanner position at -3μm | 0.0083
| Tel Aviv University_NX12 | 2023-04-26 | Z Scanner position at -3μm | 0.0117
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Z Scanner position at +3μm | 0.0098
| Tel Aviv University | 2023-04-26 | Z Scanner position at +3μm | 0.0088
| Tel Aviv University_NX12 | 2023-04-26 | Z Scanner position at +3μm | 0.0156
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Z Linearity by 3 different sample height | 0.998
| Tel Aviv University | 2023-04-26 | Z Linearity by 3 different sample height | 1.02
| Tel Aviv University_NX12 | 2023-04-26 | Z Linearity by 3 different sample height | 1.007
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Z Height noise | 0.03
| Tel Aviv University | 2023-04-26 | Z Height noise | 0.027
| Tel Aviv University_NX12 | 2023-04-26 | Z Height noise | 0.035
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y2 Detector noise (XY Servo on) | 0.34
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y2 Detector noise (XY Servo off) | 0.21
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y1 Detector noise (XY Servo on) | 0.37
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y1 Detector noise (XY Servo off) | 0.24
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y Scan Stroke | 166.3
| Tel Aviv University_NX12 | 2023-04-26 | Y Scan Stroke | 188.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y Detector noise (XY Servo on) | 0.28
| Tel Aviv University | 2023-04-26 | Y Detector noise (XY Servo on) | 0.37
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y Detector noise (XY Servo off) | 0.17
| Tel Aviv University | 2023-04-26 | Y Detector noise (XY Servo off) | 0.33
| Tel Aviv University_NX12 | 2023-04-26 | Y Detector noise (XY Servo off) | 0.29
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Y axis OPM | 0.94
| Tel Aviv University | 2023-04-26 | Y axis OPM | 1.21
| Tel Aviv University_NX12 | 2023-04-26 | Y axis OPM | 1.54
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X2 Detector noise (XY Servo on) | 0.37
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X2 Detector noise (XY Servo off) | 0.26
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X1 Detector noise (XY Servo on) | 0.38
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X1 Detector noise (XY Servo off) | 0.25
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X Scan Stroke | 169.3
| Tel Aviv University_NX12 | 2023-04-26 | X Scan Stroke | 189
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | -X Direction software Limit | -102378
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | -X Direction software Limit | -103265
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X Detector noise (XY Servo on) | 0.31
| Tel Aviv University | 2023-04-26 | X Detector noise (XY Servo on) | 0.4
| Tel Aviv University_NX12 | 2023-04-26 | X Detector noise (XY Servo on) | 0.35
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X Detector noise (XY Servo off) | 0.21
| Tel Aviv University | 2023-04-26 | X Detector noise (XY Servo off) | 0.34
| Tel Aviv University_NX12 | 2023-04-26 | X Detector noise (XY Servo off) | 0.29
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | X axis OPM | 1.31
| Tel Aviv University | 2023-04-26 | X axis OPM | 1.15
| Tel Aviv University_NX12 | 2023-04-26 | X axis OPM | 0.52
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Tel Aviv University | 2023-04-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.31
| Tel Aviv University_NX12 | 2023-04-26 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.38
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Tel Aviv University | 2023-04-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| Tel Aviv University_NX12 | 2023-04-26 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.47
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of Wafer Zeroscan (Top) | 0.27
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of Wafer Zeroscan (Right) | 0.24
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of Wafer Zeroscan (Left) | 0.25
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of Wafer ZeroScan (Center) | 0.24
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.27
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | NCM driving% with AC160TS cantilever | 0.186
| Tel Aviv University | 2023-04-26 | NCM driving% with AC160TS cantilever | 0.45
| Tel Aviv University_NX12 | 2023-04-26 | NCM driving% with AC160TS cantilever | 1.49
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | NCM Amplitude noise | 15.18
| Tel Aviv University | 2023-04-26 | NCM Amplitude noise | 15.07
| Tel Aviv University_NX12 | 2023-04-26 | NCM Amplitude noise | 14.38
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.411
| Tel Aviv University | 2023-04-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.1
| Tel Aviv University_NX12 | 2023-04-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.697
| Tel Aviv University | 2023-04-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.9
| Tel Aviv University_NX12 | 2023-04-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 240.5
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 233
| Tel Aviv University | 2023-04-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 160
| Tel Aviv University_NX12 | 2023-04-26 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 267
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 237.7
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 234.2
| Tel Aviv University | 2023-04-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.5
| Tel Aviv University_NX12 | 2023-04-26 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 266.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Lateral(C-D) sensitivity | 42.612
| Tel Aviv University | 2023-04-26 | Lateral(C-D) sensitivity | 6.9
| Tel Aviv University_NX12 | 2023-04-26 | Lateral(C-D) sensitivity | 46.016
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Lateral(C-D) noise | 0.107
| Tel Aviv University | 2023-04-26 | Lateral(C-D) noise | 0.1
| Tel Aviv University_NX12 | 2023-04-26 | Lateral(C-D) noise | 0.074
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of Z Height during Z Scan sweep | 16.6
| Tel Aviv University | 2023-04-26 | Full range of Z Height during Z Scan sweep | 15.8
| Tel Aviv University_NX12 | 2023-04-26 | Full range of Z Height during Z Scan sweep | 16.3
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of Z Detector during Z Scan sweep | 12
| Tel Aviv University | 2023-04-26 | Full range of Z Detector during Z Scan sweep | 12.7
| Tel Aviv University_NX12 | 2023-04-26 | Full range of Z Detector during Z Scan sweep | 11.1
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| Tel Aviv University | 2023-04-26 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| Tel Aviv University_NX12 | 2023-04-26 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of Y Detector during Y Scan sweep | 165.7
| Tel Aviv University_NX12 | 2023-04-26 | Full range of Y Detector during Y Scan sweep | 188.3
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of X1 Detector Raw during X Scan sweep | 14.5
| Tel Aviv University | 2023-04-26 | Full range of X Detector Raw during X Scan sweep | 14.1
| Tel Aviv University_NX12 | 2023-04-26 | Full range of X Detector Raw during X Scan sweep | 14.2
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | Full range of X Detector during X Scan sweep | 167.5
| Tel Aviv University | 2023-04-26 | Full range of X Detector during X Scan sweep | 112.9
| Tel Aviv University_NX12 | 2023-04-26 | Full range of X Detector during X Scan sweep | 188.6
| Hebei Synlight Crystal Co. Ltd #5 | 2023-04-26 | A-B noise | 10.19
| Tel Aviv University | 2023-04-26 | A-B noise | 9.4
| Tel Aviv University_NX12 | 2023-04-26 | A-B noise | 9.22
| Hangzhou Zeda Semiconductor | 2023-04-25 | Z Scanner position at -3μm | 0.0025
| Hangzhou Zeda Semiconductor | 2023-04-25 | Z Scanner position at +3μm | 0.0077
| Hangzhou Zeda Semiconductor | 2023-04-25 | Z Linearity by 3 different sample height | 0.983
| Hangzhou Zeda Semiconductor | 2023-04-25 | Z Height noise | 0.034
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y2 Detector noise (XY Servo on) | 0.23
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y2 Detector noise (XY Servo off) | 0.24
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y1 Detector noise (XY Servo on) | 0.22
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y1 Detector noise (XY Servo off) | 0.22
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y Scan Stroke | 190
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y Detector noise (XY Servo on) | 0.17
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y Detector noise (XY Servo off) | 0.18
| Hangzhou Zeda Semiconductor | 2023-04-25 | Y axis OPM | 0.7
| Hangzhou Zeda Semiconductor | 2023-04-25 | X2 Detector noise (XY Servo on) | 0.21
| Hangzhou Zeda Semiconductor | 2023-04-25 | X2 Detector noise (XY Servo off) | 0.22
| Hangzhou Zeda Semiconductor | 2023-04-25 | X1 Detector noise (XY Servo on) | 0.22
| Hangzhou Zeda Semiconductor | 2023-04-25 | X1 Detector noise (XY Servo off) | 0.23
| Hangzhou Zeda Semiconductor | 2023-04-25 | X Scan Stroke | 185.6
| Hangzhou Zeda Semiconductor | 2023-04-25 | -X Direction software Limit | -75265
| Hangzhou Zeda Semiconductor | 2023-04-25 | -X Direction software Limit | -75294
| Hangzhou Zeda Semiconductor | 2023-04-25 | X Detector noise (XY Servo on) | 0.16
| Hangzhou Zeda Semiconductor | 2023-04-25 | X Detector noise (XY Servo off) | 0.16
| Hangzhou Zeda Semiconductor | 2023-04-25 | X axis OPM | 0.62
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.4
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of Wafer Zeroscan (Top) | 0.32
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of Wafer Zeroscan (Right) | 0.32
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of Wafer Zeroscan (Left) | 0.31
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of Wafer ZeroScan (Center) | 0.3
| Hangzhou Zeda Semiconductor | 2023-04-25 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.3
| Hangzhou Zeda Semiconductor | 2023-04-25 | NCM driving% with AC160TS cantilever | 1.53
| Hangzhou Zeda Semiconductor | 2023-04-25 | NCM Amplitude noise | 15.98
| Hangzhou Zeda Semiconductor | 2023-04-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.5
| Hangzhou Zeda Semiconductor | 2023-04-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.4
| Hangzhou Zeda Semiconductor | 2023-04-25 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 256.1
| Hangzhou Zeda Semiconductor | 2023-04-25 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 248.8
| Hangzhou Zeda Semiconductor | 2023-04-25 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 243.7
| Hangzhou Zeda Semiconductor | 2023-04-25 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249.7
| Hangzhou Zeda Semiconductor | 2023-04-25 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.02
| Hangzhou Zeda Semiconductor | 2023-04-25 | Lateral(C-D) sensitivity | 9
| Hangzhou Zeda Semiconductor | 2023-04-25 | Lateral(C-D) noise | 0.111
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of Z Height during Z Scan sweep | 16.2
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of Z Detector during Z Scan sweep | 12.5
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.7
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of Y1 Detector Raw during Y Scan sweep | 15.2
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of Y Detector during Y Scan sweep | 188.5
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of X2 Detector Raw during X Scan sweep | 14.9
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of X1 Detector Raw during X Scan sweep | 15
| Hangzhou Zeda Semiconductor | 2023-04-25 | Full range of X Detector during X Scan sweep | 185.6
| Hangzhou Zeda Semiconductor | 2023-04-25 | A-B noise | 8.41
| TSMC AP6 | 2023-04-24 | Z Linearity by Z scanner position | 0.74
| TSMC AP6 | 2023-04-24 | Z Linearity by 5 different sample height | 1.006
| TSMC AP6 | 2023-04-24 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 60.51
| TSMC AP6 | 2023-04-24 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 60.6
| TSMC AP6 | 2023-04-24 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 27.3
| TSMC AP6 | 2023-04-24 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.19
| TSMC AP6 | 2023-04-24 | Z Detector Min | -0.26423
| TSMC AP6 | 2023-04-24 | Z Detector Max | 0.33389
| TSMC AP6 | 2023-04-24 | Y2 Detector noise (XY Servo on) | 0.774
| TSMC AP6 | 2023-04-24 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 267.5
| TSMC AP6 | 2023-04-24 | Y1 Detector noise (XY Servo on) | 0.761
| TSMC AP6 | 2023-04-24 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 276
| TSMC AP6 | 2023-04-24 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 187.3
| TSMC AP6 | 2023-04-24 | Y Ringing(Rpv) | 0.23
| TSMC AP6 | 2023-04-24 | Y Non-linearity error | 3.46
| TSMC AP6 | 2023-04-24 | Y Detector noise (XY Servo on) | 0.72
| TSMC AP6 | 2023-04-24 | Y axis, Repeatability @1st position | 0.05
| TSMC AP6 | 2023-04-24 | Y axis, Repeatability @ 4th position | 0.04
| TSMC AP6 | 2023-04-24 | Y axis, Repeatability @ 3rd position | 0.05
| TSMC AP6 | 2023-04-24 | Y axis, Repeatability @ 2nd position | 0.04
| TSMC AP6 | 2023-04-24 | Y axis, Accuracy @1st position | 0.36
| TSMC AP6 | 2023-04-24 | Y axis, Accuracy @ 4th position | 1.31
| TSMC AP6 | 2023-04-24 | Y axis, Accuracy @ 3rd position | 0.4
| TSMC AP6 | 2023-04-24 | Y axis, Accuracy @ 2nd position | 1.25
| TSMC AP6 | 2023-04-24 | Y axis OPM | 3.26
| TSMC AP6 | 2023-04-24 | X2 Detector noise (XY Servo on) | 0.48
| TSMC AP6 | 2023-04-24 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 268.3
| TSMC AP6 | 2023-04-24 | X1 Detector noise (XY Servo on) | 0.566
| TSMC AP6 | 2023-04-24 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 275.4
| TSMC AP6 | 2023-04-24 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 177.1
| TSMC AP6 | 2023-04-24 | X Ringing(Rpv) | 0.14
| TSMC AP6 | 2023-04-24 | X Non-linearity error | 3.52
| TSMC AP6 | 2023-04-24 | X Detector noise (XY Servo on) | 0.413
| TSMC AP6 | 2023-04-24 | X axis, Repeatability @1st position | 0.03
| TSMC AP6 | 2023-04-24 | X axis, Repeatability @ 4th position | 0.03
| TSMC AP6 | 2023-04-24 | X axis, Repeatability @ 3rd position | 0.03
| TSMC AP6 | 2023-04-24 | X axis, Repeatability @ 2nd position | 0.04
| TSMC AP6 | 2023-04-24 | X axis, Accuracy @1st position | 1.13
| TSMC AP6 | 2023-04-24 | X axis, Accuracy @ 4th position | 1.31
| TSMC AP6 | 2023-04-24 | X axis, Accuracy @ 3rd position | 1.07
| TSMC AP6 | 2023-04-24 | X axis, Accuracy @ 2nd position | 1.12
| TSMC AP6 | 2023-04-24 | X axis OPM | 1.28
| TSMC AP6 | 2023-04-24 | WLI_Scale(Z Calibration) | 31.6
| TSMC AP6 | 2023-04-24 | Under 50nm Step Height Sample (10 Repeats) | 18.33
| TSMC AP6 | 2023-04-24 | Torque Profile (50mm, 20um/s) | 2.272
| TSMC AP6 | 2023-04-24 | Torque Profile (50mm, 200um/s) | 2.099
| TSMC AP6 | 2023-04-24 | StrokeYMin | -152171
| TSMC AP6 | 2023-04-24 | StrokeYMax | 150943
| TSMC AP6 | 2023-04-24 | StrokeY | 303113.8
| TSMC AP6 | 2023-04-24 | StrokeXMin | -204033
| TSMC AP6 | 2023-04-24 | StrokeXMax | 202991.7
| TSMC AP6 | 2023-04-24 | StrokeX | 407024.3
| TSMC AP6 | 2023-04-24 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer) | 0.24
| TSMC AP6 | 2023-04-24 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer) | 1.35
| TSMC AP6 | 2023-04-24 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.63
| TSMC AP6 | 2023-04-24 | Roll motion to X+ full scan | 64.4
| TSMC AP6 | 2023-04-24 | RMS of ZeroScan | 0.067
| TSMC AP6 | 2023-04-24 | Profiler noise @200um | 0.122
| TSMC AP6 | 2023-04-24 | PosMin | -25.982
| TSMC AP6 | 2023-04-24 | PosMax | 26.008
| TSMC AP6 | 2023-04-24 | Position repeatability(10repeats, Position Based) | 0.68
| TSMC AP6 | 2023-04-24 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5e-05
| TSMC AP6 | 2023-04-24 | OpticalAxisShiftPerFocusY | -0.0002
| TSMC AP6 | 2023-04-24 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00033
| TSMC AP6 | 2023-04-24 | OpticalAxisShiftPerFocusX | 0.0016
| TSMC AP6 | 2023-04-24 | OPM Repeatability @5mm | 1.516
| TSMC AP6 | 2023-04-24 | OPM Repeatability @25mm | 10.337
| TSMC AP6 | 2023-04-24 | OPM Repeatability @1mm | 0.868
| TSMC AP6 | 2023-04-24 | OPM Repeatability @10mm | 2.406
| TSMC AP6 | 2023-04-24 | NCM driving% with AC160TS cantilever | 2.73
| TSMC AP6 | 2023-04-24 | Min. Vibration Transmissbility (~100Hz) | 0.001
| TSMC AP6 | 2023-04-24 | MaxLength | 50
| TSMC AP6 | 2023-04-24 | Max OPM @5mm | 18.505
| TSMC AP6 | 2023-04-24 | Max OPM @25mm | 156.602
| TSMC AP6 | 2023-04-24 | Max OPM @1mm | 9.533
| TSMC AP6 | 2023-04-24 | Max OPM @10mm | 37.969
| TSMC AP6 | 2023-04-24 | Limit range | 13568.25
| TSMC AP6 | 2023-04-24 | Limit range | -42032.1
| TSMC AP6 | 2023-04-24 | Lateral(C-D) sensitivity | 3
| TSMC AP6 | 2023-04-24 | Lateral(C-D) noise | 0.233
| TSMC AP6 | 2023-04-24 | Laser intensity with AC160 cantilever | 3
| TSMC AP6 | 2023-04-24 | Image offset after moving from -Y to +Y (Y scan) | 15.38
| TSMC AP6 | 2023-04-24 | Image offset after moving from -Y to +Y (X scan) | 14.65
| TSMC AP6 | 2023-04-24 | Image offset after moving from -X to +X (Y scan) | 126.71
| TSMC AP6 | 2023-04-24 | Image offset after moving from -X to +X (X scan) | 23.44
| TSMC AP6 | 2023-04-24 | Image offset after moving from +Y to -Y (Y scan) | 13.18
| TSMC AP6 | 2023-04-24 | Image offset after moving from +Y to -Y (X scan) | 46.14
| TSMC AP6 | 2023-04-24 | Image offset after moving from +X to -X (Y scan) | 108.4
| TSMC AP6 | 2023-04-24 | Image offset after moving from +X to -X (X scan) | 27.83
| TSMC AP6 | 2023-04-24 | Full range of Z Height during Z Scan sweep | 42.38
| TSMC AP6 | 2023-04-24 | Full range of Z Detector during Z Scan sweep | 13.97
| TSMC AP6 | 2023-04-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| TSMC AP6 | 2023-04-24 | Full range of Y1 Detector Raw during Y Scan sweep | 13.9
| TSMC AP6 | 2023-04-24 | Full range of X2 Detector Raw during X Scan sweep | 13.9
| TSMC AP6 | 2023-04-24 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| TSMC AP6 | 2023-04-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.132
| TSMC AP6 | 2023-04-24 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.736
| TSMC AP6 | 2023-04-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.837
| TSMC AP6 | 2023-04-24 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| TSMC AP6 | 2023-04-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.07
| TSMC AP6 | 2023-04-24 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.37
| TSMC AP6 | 2023-04-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.095
| TSMC AP6 | 2023-04-24 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.017
| TSMC AP6 | 2023-04-24 | A－B difference during 3 times run test | 0.5
| TSMC AP6 | 2023-04-24 | A＋B difference during 3 times run test | 0.02
| TSMC AP6 | 2023-04-24 | 50~200nm Step Height Sample (10 Repeats) | 184.13
| TSMC AP6 | 2023-04-24 | 200~500nm Step Height Sample (10 Repeats) | 460.16
| TSMC AP6 | 2023-04-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 8.2
| TSMC AP6 | 2023-04-24 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 11.9
| TSMC AP6 | 2023-04-24 | %GRR w/ roughness of non-patterened wafer | 14.4
| Hacettepe University | 2023-04-24 | Z Scanner position at -3μm | 0.013
| Hacettepe University | 2023-04-24 | Z Scanner position at +3μm | 0.0124
| Hacettepe University | 2023-04-24 | Z Linearity by 3 different sample height | 1.007
| Hacettepe University | 2023-04-24 | Z Height noise | 0.034
| Hacettepe University | 2023-04-24 | Y Scan Stroke | 188.2
| Hacettepe University | 2023-04-24 | Y Detector noise (XY Servo on) | 0.36
| Hacettepe University | 2023-04-24 | Y Detector noise (XY Servo off) | 0.29
| Hacettepe University | 2023-04-24 | Y axis OPM | 0.81
| Hacettepe University | 2023-04-24 | X Scan Stroke | 182
| Hacettepe University | 2023-04-24 | X Detector noise (XY Servo on) | 0.34
| Hacettepe University | 2023-04-24 | X Detector noise (XY Servo off) | 0.28
| Hacettepe University | 2023-04-24 | X axis OPM | 0.67
| Hacettepe University | 2023-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.48
| Hacettepe University | 2023-04-24 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.64
| Hacettepe University | 2023-04-24 | NCM driving% with AC160TS cantilever | 0.92
| Hacettepe University | 2023-04-24 | NCM Amplitude noise | 15.94
| Hacettepe University | 2023-04-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.151
| Hacettepe University | 2023-04-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.968
| Hacettepe University | 2023-04-24 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 256.1
| Hacettepe University | 2023-04-24 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 252.8
| Hacettepe University | 2023-04-24 | Lateral(C-D) sensitivity | 12.8
| Hacettepe University | 2023-04-24 | Lateral(C-D) noise | 0.087
| Hacettepe University | 2023-04-24 | Full range of Z Height during Z Scan sweep | 16.2
| Hacettepe University | 2023-04-24 | Full range of Z Detector during Z Scan sweep | 11.2
| Hacettepe University | 2023-04-24 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Hacettepe University | 2023-04-24 | Full range of Y Detector during Y Scan sweep | 187.7
| Hacettepe University | 2023-04-24 | Full range of X Detector Raw during X Scan sweep | 14.4
| Hacettepe University | 2023-04-24 | Full range of X Detector during X Scan sweep | 189.5
| Hacettepe University | 2023-04-24 | A-B noise | 8.63
| On-Semi | 2023-04-20 | Z Linearity by Z scanner position | 0.6
| On-Semi | 2023-04-20 | Z Linearity by Z scanner position | 0.793
| On-Semi | 2023-04-20 | Z Linearity by 5 different sample height | 1.003
| On-Semi_1 | 2023-04-20 | Z Linearity by 5 different sample height | 1.0002
| On-Semi | 2023-04-20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 64.8
| On-Semi | 2023-04-20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 64.2
| On-Semi | 2023-04-20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 26.5
| On-Semi | 2023-04-20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 27.2
| On-Semi | 2023-04-20 | Y2 Detector noise (XY Servo on) | 0.65
| On-Semi | 2023-04-20 | Y2 Detector noise (XY Servo on) | 0.809
| On-Semi | 2023-04-20 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 285.6
| On-Semi | 2023-04-20 | Y1 Detector noise (XY Servo on) | 0.62
| On-Semi | 2023-04-20 | Y1 Detector noise (XY Servo on) | 0.791
| On-Semi | 2023-04-20 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 279.6
| On-Semi | 2023-04-20 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 194.9
| On-Semi | 2023-04-20 | Y Ringing(Rpv) | 0.12
| On-Semi | 2023-04-20 | Y Ringing(Rpv) | 0.115
| On-Semi | 2023-04-20 | Y Non-linearity error | 2.63
| On-Semi | 2023-04-20 | Y Non-linearity error | 2.82
| On-Semi | 2023-04-20 | Y Detector noise (XY Servo on) | 0.56
| On-Semi | 2023-04-20 | Y Detector noise (XY Servo on) | 0.744
| On-Semi | 2023-04-20 | Y axis, Repeatability @1st position | 0.03
| On-Semi | 2023-04-20 | Y axis, Repeatability @ 4th position | 0.03
| On-Semi | 2023-04-20 | Y axis, Repeatability @ 3rd position | 0.03
| On-Semi | 2023-04-20 | Y axis, Repeatability @ 2nd position | 0.04
| On-Semi | 2023-04-20 | Y axis, Accuracy @1st position | 1.92
| On-Semi | 2023-04-20 | Y axis, Accuracy @ 4th position | 0.74
| On-Semi | 2023-04-20 | Y axis, Accuracy @ 3rd position | 0.65
| On-Semi | 2023-04-20 | Y axis, Accuracy @ 2nd position | 1.65
| On-Semi | 2023-04-20 | Y axis OPM | 0.25
| On-Semi | 2023-04-20 | Y axis OPM | 0.77
| On-Semi | 2023-04-20 | X2 Detector noise (XY Servo on) | 0.53
| On-Semi | 2023-04-20 | X2 Detector noise (XY Servo on) | 0.6
| On-Semi | 2023-04-20 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 287.6
| On-Semi | 2023-04-20 | X1 Detector noise (XY Servo on) | 0.5
| On-Semi | 2023-04-20 | X1 Detector noise (XY Servo on) | 0.571
| On-Semi | 2023-04-20 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 280.6
| On-Semi | 2023-04-20 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 194.2
| On-Semi | 2023-04-20 | X Ringing(Rpv) | 0.1
| On-Semi | 2023-04-20 | X Ringing(Rpv) | 0.228
| On-Semi | 2023-04-20 | X Non-linearity error | 1.84
| On-Semi | 2023-04-20 | X Non-linearity error | 2.66
| On-Semi | 2023-04-20 | X Detector noise (XY Servo on) | 0.41
| On-Semi | 2023-04-20 | X Detector noise (XY Servo on) | 0.499
| On-Semi | 2023-04-20 | X axis, Repeatability @1st position | 0.02
| On-Semi | 2023-04-20 | X axis, Repeatability @ 4th position | 0.06
| On-Semi | 2023-04-20 | X axis, Repeatability @ 3rd position | 0.03
| On-Semi | 2023-04-20 | X axis, Repeatability @ 2nd position | 0.04
| On-Semi | 2023-04-20 | X axis, Accuracy @1st position | 0.56
| On-Semi | 2023-04-20 | X axis, Accuracy @ 4th position | 0.39
| On-Semi | 2023-04-20 | X axis, Accuracy @ 3rd position | 0.42
| On-Semi | 2023-04-20 | X axis, Accuracy @ 2nd position | 0.91
| On-Semi | 2023-04-20 | X axis OPM | 0.39
| On-Semi | 2023-04-20 | X axis OPM | 1.834
| On-Semi | 2023-04-20 | Under 50nm Step Height Sample (10 Repeats) | 18.112
| On-Semi | 2023-04-20 | Under 50nm Step Height Sample (10 Repeats) | 18.372
| On-Semi | 2023-04-20 | Torque Profile (10mm, 20um/s) | 3.43
| On-Semi | 2023-04-20 | Torque Profile (10mm, 200um/s) | 2.45
| On-Semi | 2023-04-20 | StrokeYMin | -150100
| On-Semi | 2023-04-20 | StrokeYMax | 303800
| On-Semi | 2023-04-20 | StrokeY | 453900
| On-Semi | 2023-04-20 | StrokeXMin | -203970
| On-Semi | 2023-04-20 | StrokeXMax | 203860
| On-Semi | 2023-04-20 | StrokeX | 407830
| On-Semi | 2023-04-20 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.6
| On-Semi | 2023-04-20 | Roll motion to X+ full scan | 11.9
| On-Semi | 2023-04-20 | RMS of ZeroScan | 0.065
| On-Semi | 2023-04-20 | Profiler noise @200um | 0.153
| On-Semi | 2023-04-20 | PosMin | -5.5
| On-Semi | 2023-04-20 | PosMax | 5.5
| On-Semi | 2023-04-20 | Position repeatability(10repeats, Position Based) | 0.4
| On-Semi | 2023-04-20 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 0.00012
| On-Semi | 2023-04-20 | OpticalAxisShiftPerFocusY | 0.0001
| On-Semi | 2023-04-20 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00025
| On-Semi | 2023-04-20 | OpticalAxisShiftPerFocusX | -0.0011
| On-Semi | 2023-04-20 | OPM Repeatability @5mm | 2.714
| On-Semi | 2023-04-20 | OPM Repeatability @1mm | 0.934
| On-Semi | 2023-04-20 | OPM Repeatability @10mm | 5.41
| On-Semi | 2023-04-20 | NCM driving% with AC160TS cantilever | 44
| On-Semi | 2023-04-20 | NCM driving% with AC160TS cantilever | 22
| On-Semi | 2023-04-20 | Min. Vibration Transmissbility (~100Hz) | 0.076
| On-Semi | 2023-04-20 | MaxLength | 11
| On-Semi | 2023-04-20 | Max OPM @5mm | 11.44
| On-Semi | 2023-04-20 | Max OPM @1mm | 2.753
| On-Semi | 2023-04-20 | Max OPM @10mm | 17.836
| On-Semi | 2023-04-20 | Lateral(C-D) sensitivity | 12.3
| On-Semi | 2023-04-20 | Lateral(C-D) noise | 0.203
| On-Semi | 2023-04-20 | Laser intensity with AC160 cantilever | 3.05
| On-Semi | 2023-04-20 | Laser intensity with AC160 cantilever | 3.2
| On-Semi | 2023-04-20 | Image offset after moving from -Y to +Y (Y scan) | 71
| On-Semi | 2023-04-20 | Image offset after moving from -Y to +Y (X scan) | 25
| On-Semi | 2023-04-20 | Image offset after moving from -X to +X (Y scan) | 104
| On-Semi | 2023-04-20 | Image offset after moving from -X to +X (X scan) | 75
| On-Semi | 2023-04-20 | Image offset after moving from +Y to -Y (Y scan) | 109
| On-Semi | 2023-04-20 | Image offset after moving from +Y to -Y (X scan) | 15
| On-Semi | 2023-04-20 | Image offset after moving from +X to -X (Y scan) | 62
| On-Semi | 2023-04-20 | Image offset after moving from +X to -X (X scan) | 125
| On-Semi | 2023-04-20 | Full range of Z Height during Z Scan sweep | 42.7
| On-Semi | 2023-04-20 | Full range of Z Height during Z Scan sweep | 42.7
| On-Semi | 2023-04-20 | Full range of Z Detector during Z Scan sweep | 13.2
| On-Semi | 2023-04-20 | Full range of Z Detector during Z Scan sweep | 13
| On-Semi | 2023-04-20 | Full range of Y2 Detector Raw during Y Scan sweep | 14
| On-Semi | 2023-04-20 | Full range of Y2 Detector Raw during Y Scan sweep | 13.4
| On-Semi | 2023-04-20 | Full range of Y2 Detector during Y Scan sweep | 214.3914
| On-Semi | 2023-04-20 | Full range of Y1 Detector Raw during Y Scan sweep | 14.1
| On-Semi | 2023-04-20 | Full range of Y1 Detector Raw during Y Scan sweep | 13.7
| On-Semi | 2023-04-20 | Full range of Y1 Detector during Y Scan sweep | 208.1827
| On-Semi | 2023-04-20 | Full range of X2 Detector Raw during X Scan sweep | 14
| On-Semi | 2023-04-20 | Full range of X2 Detector Raw during X Scan sweep | 13.1
| On-Semi | 2023-04-20 | Full range of X2 Detector during X Scan sweep | 218.9253
| On-Semi | 2023-04-20 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| On-Semi | 2023-04-20 | Full range of X1 Detector Raw during X Scan sweep | 13.9
| On-Semi | 2023-04-20 | Full range of X1 Detector during X Scan sweep | 208.7852
| On-Semi | 2023-04-20 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 570
| On-Semi | 2023-04-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| On-Semi | 2023-04-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.61
| On-Semi | 2023-04-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.7
| On-Semi | 2023-04-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| On-Semi | 2023-04-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| On-Semi | 2023-04-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.24
| On-Semi | 2023-04-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.31
| On-Semi | 2023-04-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.06
| On-Semi | 2023-04-20 | A－B difference during 3 times run test | 0.71
| On-Semi | 2023-04-20 | A＋B difference during 3 times run test | 0.11
| On-Semi | 2023-04-20 | 50~200nm Step Height Sample (10 Repeats) | 184.785
| On-Semi | 2023-04-20 | 50~200nm Step Height Sample (10 Repeats) | 184.064
| On-Semi | 2023-04-20 | 200~500nm Step Height Sample (10 Repeats) | 463.413
| On-Semi | 2023-04-20 | 200~500nm Step Height Sample (10 Repeats) | 462.618
| On-Semi | 2023-04-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 6.3
| On-Semi | 2023-04-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 18.3
| On-Semi | 2023-04-20 | %GRR w/ roughness of non-patterened wafer | 19.5
| Sincere International Trading Company Limited | 2023-04-19 | Z Scanner position at -3μm | 0.0052
| University of Arizona | 2023-04-19 | Z Scanner position at -3μm | -0.002
| Sincere International Trading Company Limited | 2023-04-19 | Z Scanner position at +3μm | 0.0073
| University of Arizona | 2023-04-19 | Z Scanner position at +3μm | -0.0027
| Sincere International Trading Company Limited | 2023-04-19 | Z Linearity by 3 different sample height | 1.018
| University of Arizona | 2023-04-19 | Z Linearity by 3 different sample height | 1.002
| Sincere International Trading Company Limited | 2023-04-19 | Z Height noise | 0.018
| University of Arizona | 2023-04-19 | Z Height noise | 0.035
| Sincere International Trading Company Limited | 2023-04-19 | Y2 Detector noise (XY Servo on) | 0.43
| Sincere International Trading Company Limited | 2023-04-19 | Y2 Detector noise (XY Servo off) | 0.33
| Sincere International Trading Company Limited | 2023-04-19 | Y1 Detector noise (XY Servo on) | 0.38
| Sincere International Trading Company Limited | 2023-04-19 | Y1 Detector noise (XY Servo off) | 0.29
| Sincere International Trading Company Limited | 2023-04-19 | Y Scan Stroke | 170.3
| University of Arizona | 2023-04-19 | Y Scan Stroke | 55
| Sincere International Trading Company Limited | 2023-04-19 | Y Detector noise (XY Servo on) | 0.31
| University of Arizona | 2023-04-19 | Y Detector noise (XY Servo on) | 0.37
| Sincere International Trading Company Limited | 2023-04-19 | Y Detector noise (XY Servo off) | 0.23
| University of Arizona | 2023-04-19 | Y Detector noise (XY Servo off) | 0.38
| Sincere International Trading Company Limited | 2023-04-19 | Y axis OPM | 1.86
| University of Arizona | 2023-04-19 | Y axis OPM | 0.47
| Sincere International Trading Company Limited | 2023-04-19 | X2 Detector noise (XY Servo on) | 0.29
| Sincere International Trading Company Limited | 2023-04-19 | X2 Detector noise (XY Servo off) | 0.19
| Sincere International Trading Company Limited | 2023-04-19 | X1 Detector noise (XY Servo on) | 0.31
| Sincere International Trading Company Limited | 2023-04-19 | X1 Detector noise (XY Servo off) | 0.18
| Sincere International Trading Company Limited | 2023-04-19 | X Scan Stroke | 177.9
| University of Arizona | 2023-04-19 | X Scan Stroke | 55
| Sincere International Trading Company Limited | 2023-04-19 | -X Direction software Limit | -102000
| Sincere International Trading Company Limited | 2023-04-19 | -X Direction software Limit | -104500
| Sincere International Trading Company Limited | 2023-04-19 | X Detector noise (XY Servo on) | 0.25
| University of Arizona | 2023-04-19 | X Detector noise (XY Servo on) | 0.36
| Sincere International Trading Company Limited | 2023-04-19 | X Detector noise (XY Servo off) | 0.14
| University of Arizona | 2023-04-19 | X Detector noise (XY Servo off) | 0.39
| Sincere International Trading Company Limited | 2023-04-19 | X axis OPM | 0.68
| University of Arizona | 2023-04-19 | X axis OPM | 0.8
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.33
| University of Arizona | 2023-04-19 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.6
| University of Arizona | 2023-04-19 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.69
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of Wafer Zeroscan (Top) | 0.3
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of Wafer Zeroscan (Right) | 0.32
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of Wafer Zeroscan (Left) | 0.31
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of Wafer ZeroScan (Center) | 0.3
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.31
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Sincere International Trading Company Limited | 2023-04-19 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.58
| Sincere International Trading Company Limited | 2023-04-19 | NCM driving% with AC160TS cantilever | 0.67
| University of Arizona | 2023-04-19 | NCM driving% with AC160TS cantilever | 1.37
| Sincere International Trading Company Limited | 2023-04-19 | NCM Amplitude noise | 15.7
| University of Arizona | 2023-04-19 | NCM Amplitude noise | 16.412
| Sincere International Trading Company Limited | 2023-04-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| University of Arizona | 2023-04-19 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| Sincere International Trading Company Limited | 2023-04-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.2
| University of Arizona | 2023-04-19 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.2
| Sincere International Trading Company Limited | 2023-04-19 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 234.8
| Sincere International Trading Company Limited | 2023-04-19 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 240
| University of Arizona | 2023-04-19 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 92.1
| Sincere International Trading Company Limited | 2023-04-19 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 249.7
| Sincere International Trading Company Limited | 2023-04-19 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 241.9
| University of Arizona | 2023-04-19 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.6
| Sincere International Trading Company Limited | 2023-04-19 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.09
| Sincere International Trading Company Limited | 2023-04-19 | Lateral(C-D) sensitivity | 8.5
| University of Arizona | 2023-04-19 | Lateral(C-D) sensitivity | 10.6
| Sincere International Trading Company Limited | 2023-04-19 | Lateral(C-D) noise | 0.126
| University of Arizona | 2023-04-19 | Lateral(C-D) noise | 0.12
| Sincere International Trading Company Limited | 2023-04-19 | Full range of Z Height during Z Scan sweep | 16.4
| University of Arizona | 2023-04-19 | Full range of Z Height during Z Scan sweep | 16.4
| Sincere International Trading Company Limited | 2023-04-19 | Full range of Z Detector during Z Scan sweep | 12.5
| University of Arizona | 2023-04-19 | Full range of Z Detector during Z Scan sweep | 11.3
| Sincere International Trading Company Limited | 2023-04-19 | Full range of Y2 Detector Raw during Y Scan sweep | 14.2
| Sincere International Trading Company Limited | 2023-04-19 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| University of Arizona | 2023-04-19 | Full range of Y Detector Raw during Y Scan sweep | 15.4
| Sincere International Trading Company Limited | 2023-04-19 | Full range of Y Detector during Y Scan sweep | 170.3
| University of Arizona | 2023-04-19 | Full range of Y Detector during Y Scan sweep | 70.8
| Sincere International Trading Company Limited | 2023-04-19 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Sincere International Trading Company Limited | 2023-04-19 | Full range of X1 Detector Raw during X Scan sweep | 14.7
| University of Arizona | 2023-04-19 | Full range of X Detector Raw during X Scan sweep | 14.7
| Sincere International Trading Company Limited | 2023-04-19 | Full range of X Detector during X Scan sweep | 177.9
| University of Arizona | 2023-04-19 | Full range of X Detector during X Scan sweep | 74.9
| Sincere International Trading Company Limited | 2023-04-19 | A-B noise | 11.25
| University of Arizona | 2023-04-19 | A-B noise | 9.73
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Z Linearity by 5 different sample height | 1.006
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Z Height noise | 0.04
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 28.57
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.62
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.68
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.7
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y2 Detector noise (XY Servo on) | 0.28
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 264.472
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y1 Detector noise (XY Servo on) | 0.233
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 267.32
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 144.642
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y Ringing(Rpv) | 0.817
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y Non-linearity error | -3.35
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y Detector noise (XY Servo on) | 0.223
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Repeatability @1st position | 0.06
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Repeatability @ 4th position | 0.04
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Repeatability @ 3rd position | 0.04
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Repeatability @ 2nd position | 0.05
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Accuracy @1st position | 0.59
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Accuracy @ 4th position | 0.57
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Accuracy @ 3rd position | 0.53
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis, Accuracy @ 2nd position | 0.56
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Y axis OPM | 0.999
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X2 Detector noise (XY Servo on) | 0.277
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 266.24
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X1 Detector noise (XY Servo on) | 0.271
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 256.449
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 189.777
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X Ringing(Rpv) | 0.678
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X Non-linearity error | 1.54
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X Detector noise (XY Servo on) | 0.215
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Repeatability @1st position | 0.06
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Repeatability @ 4th position | 0.04
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Repeatability @ 3rd position | 0.04
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Repeatability @ 2nd position | 0.04
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Accuracy @1st position | 0.55
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Accuracy @ 4th position | 0.23
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Accuracy @ 3rd position | 0.22
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis, Accuracy @ 2nd position | 0.2
| Wuhan YMTC Fab2 #2 | 2023-04-18 | X axis OPM | 1.47
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Under 50nm Step Height Sample (10 Repeats) | 18.384
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Torque Profile (50mm, 20um/s) | 2.947
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Torque Profile (50mm, 200um/s) | 2.083
| Wuhan YMTC Fab2 #2 | 2023-04-18 | StrokeYMin | -153850
| Wuhan YMTC Fab2 #2 | 2023-04-18 | StrokeYMax | 150370
| Wuhan YMTC Fab2 #2 | 2023-04-18 | StrokeY | 304220
| Wuhan YMTC Fab2 #2 | 2023-04-18 | StrokeXMin | -204530
| Wuhan YMTC Fab2 #2 | 2023-04-18 | StrokeXMax | 203590
| Wuhan YMTC Fab2 #2 | 2023-04-18 | StrokeX | 408120
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.19
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Roll motion to X+ full scan | 26.5
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Profiler noise @200um | 0.155
| Wuhan YMTC Fab2 #2 | 2023-04-18 | PosMin | -25.861
| Wuhan YMTC Fab2 #2 | 2023-04-18 | PosMax | 25.654
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Position repeatability(10repeats, Position Based) | 0.24
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5.2e-05
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OpticalAxisShiftPerFocusY | -0.002
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 4.2e-05
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OpticalAxisShiftPerFocusX | -0.0018
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OPM Repeatability @5mm | 1.8
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OPM Repeatability @25mm | 10.35
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OPM Repeatability @1mm | 0.744
| Wuhan YMTC Fab2 #2 | 2023-04-18 | OPM Repeatability @10mm | 1.478
| Wuhan YMTC Fab2 #2 | 2023-04-18 | NCM driving% with AC160TS cantilever | 12.59
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Min. Vibration Transmissbility (~100Hz) | 0.098
| Wuhan YMTC Fab2 #2 | 2023-04-18 | MaxLength | 51.515
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Max OPM @5mm | 14.676
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Max OPM @25mm | 193.116
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Max OPM @1mm | 6.376
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Max OPM @10mm | 33.071
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Lateral(C-D) sensitivity | 14.4
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Lateral(C-D) noise | 0.2
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Laser intensity with AC160 cantilever | 3.013
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from -Y to +Y (Y scan) | 23.44
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from -Y to +Y (X scan) | 23.44
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from -X to +X (Y scan) | 72.51
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from -X to +X (X scan) | 54.93
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from +Y to -Y (Y scan) | 46.14
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from +Y to -Y (X scan) | 17.58
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from +X to -X (Y scan) | 65.92
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Image offset after moving from +X to -X (X scan) | 103.27
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of Z Height during Z Scan sweep | 18.7
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of Z Detector during Z Scan sweep | 13.37
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of Y2 Detector Raw during Y Scan sweep | 13.82
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of Y2 Detector during Y Scan sweep | 202.4277
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of Y1 Detector during Y Scan sweep | 200.7743
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of X2 Detector Raw during X Scan sweep | 13.54
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of X2 Detector during X Scan sweep | 205.0262
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of X1 Detector Raw during X Scan sweep | 13.34
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Full range of X1 Detector during X Scan sweep | 200.7377
| Wuhan YMTC Fab2 #2 | 2023-04-18 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type) | 800
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.43
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.72
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.13
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.7
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.17
| Wuhan YMTC Fab2 #2 | 2023-04-18 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Wuhan YMTC Fab2 #2 | 2023-04-18 | A－B difference during 3 times run test | 0.81
| Wuhan YMTC Fab2 #2 | 2023-04-18 | A＋B difference during 3 times run test | 0.1
| Wuhan YMTC Fab2 #2 | 2023-04-18 | 50~200nm Step Height Sample (10 Repeats) | 178.381
| Wuhan YMTC Fab2 #2 | 2023-04-18 | 200~500nm Step Height Sample (10 Repeats) | 460.262
| Wuhan YMTC Fab2 #2 | 2023-04-18 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 6.6
| Wuhan YMTC Fab2 #2 | 2023-04-18 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 14.3
| Wuhan YMTC Fab2 #2 | 2023-04-18 | %GRR w/ roughness of non-patterened wafer | 15.7
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Z Linearity by 3 different sample height | 0.999
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Z Height noise | 0.021
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Y Detector noise (XY Servo on) | 0.3
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Y Detector noise (XY Servo off) | 0.25
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Y axis OPM | 0.71
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | X Detector noise (XY Servo on) | 0.3
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | X Detector noise (XY Servo off) | 0.25
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | X axis OPM | 0.58
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | NCM driving% with AC160TS cantilever | 6.617
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | NCM Amplitude noise | 8.9
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.54
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 157.7
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 157.3
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Full range of Z Height during Z Scan sweep | 16
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Lateral(C-D) sensitivity | 6.2
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Lateral(C-D) noise | 0.135
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Full range of Z Detector during Z Scan sweep | 14.1
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Full range of Y Detector Raw during Y Scan sweep | 14.7
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Full range of Y Detector during Y Scan sweep | 115.3
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Full range of X Detector Raw during X Scan sweep | 14.5
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Full range of X Detector during X Scan sweep | 114.85
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Approach position at -3μm | 0.0056
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | Approach position at +3μm | 0.012
| Ecole centrale de Lyon (LTDS) | 2023-04-17 | A-B noise | 6.9
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Z Linearity by 5 different sample height | 1.006
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Z Height noise | 0.03
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 23.73
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.71
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.26
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.27
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y2 Detector noise (XY Servo on) | 0.38
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 282.76
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y1 Detector noise (XY Servo on) | 0.34
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 288.21
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 188.77
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y Ringing(Rpv) | 0.5
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y Non-linearity error | 2.58
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y Detector noise (XY Servo on) | 0.27
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Repeatability @1st position | 0.05
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Repeatability @ 4th position | 0.03
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Repeatability @ 3rd position | 0.04
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Repeatability @ 2nd position | 0.04
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Accuracy @1st position | 1.75
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Accuracy @ 4th position | 0.6
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Accuracy @ 3rd position | 0.47
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis, Accuracy @ 2nd position | 1.77
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Y axis OPM | 0.41
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X2 Detector noise (XY Servo on) | 0.41
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 279.73
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X1 Detector noise (XY Servo on) | 0.33
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 279.51
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 188
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X Ringing(Rpv) | 0.69
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X Non-linearity error | 2.84
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X Detector noise (XY Servo on) | 0.28
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Repeatability @1st position | 0.03
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Repeatability @ 4th position | 0.05
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Repeatability @ 3rd position | 0.03
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Repeatability @ 2nd position | 0.05
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Accuracy @1st position | 1.09
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Accuracy @ 4th position | 0.47
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Accuracy @ 3rd position | 0.35
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis, Accuracy @ 2nd position | 0.43
| Wuhan YMTC Fab2 #3 | 2023-04-13 | X axis OPM | 0.97
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Under 50nm Step Height Sample (10 Repeats) | 18.23
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Torque Profile (50mm, 20um/s) | 3.184
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Torque Profile (50mm, 200um/s) | 2.359
| Wuhan YMTC Fab2 #3 | 2023-04-13 | StrokeYMin | -153378
| Wuhan YMTC Fab2 #3 | 2023-04-13 | StrokeYMax | 150919.2
| Wuhan YMTC Fab2 #3 | 2023-04-13 | StrokeY | 304296.9
| Wuhan YMTC Fab2 #3 | 2023-04-13 | StrokeXMin | -203746
| Wuhan YMTC Fab2 #3 | 2023-04-13 | StrokeXMax | 204023
| Wuhan YMTC Fab2 #3 | 2023-04-13 | StrokeX | 407769
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Short-term stability w/ 450nm Step Height Sample (3σ) | 0.189
| Wuhan YMTC Fab2 #3 | 2023-04-13 | RMS of ZeroScan | 0.056
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Profiler noise @200um | 0.217
| Wuhan YMTC Fab2 #3 | 2023-04-13 | PosMin | -25.93
| Wuhan YMTC Fab2 #3 | 2023-04-13 | PosMax | 25.7
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) | 5.16e-05
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OpticalAxisShiftPerFocusY | 0.001
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) | 0.00012
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OpticalAxisShiftPerFocusX | -0.0013
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OPM Repeatability @5mm | 1.01
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OPM Repeatability @25mm | 9.823
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OPM Repeatability @1mm | 0.472
| Wuhan YMTC Fab2 #3 | 2023-04-13 | OPM Repeatability @10mm | 0.66
| Wuhan YMTC Fab2 #3 | 2023-04-13 | NCM driving% with AC160TS cantilever | 6.74
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Min. Vibration Transmissbility (~100Hz) | 0.095
| Wuhan YMTC Fab2 #3 | 2023-04-13 | MaxLength | 51
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Max OPM @5mm | 13.88
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Max OPM @25mm | 195.6
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Max OPM @1mm | 5.55
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Max OPM @10mm | 52.83
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Lateral(C-D) sensitivity | 4
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Lateral(C-D) noise | 0.185
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Laser intensity with AC160 cantilever | 3.009
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from -Y to +Y (Y scan) | 32.23
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from -Y to +Y (X scan) | 28.56
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from -X to +X (Y scan) | 81.3
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from -X to +X (X scan) | 43.21
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from +Y to -Y (Y scan) | 16.11
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from +Y to -Y (X scan) | 30.76
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from +X to -X (Y scan) | 38.09
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Image offset after moving from +X to -X (X scan) | 28.56
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of Z Height during Z Scan sweep | 15.86
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of Z Detector during Z Scan sweep | 13.4
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of Y2 Detector Raw during Y Scan sweep | 12.64
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of Y2 Detector during Y Scan sweep | 203.8451
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of Y1 Detector Raw during Y Scan sweep | 12.47
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of Y1 Detector during Y Scan sweep | 214.0561
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of X2 Detector Raw during X Scan sweep | 12.59
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of X2 Detector during X Scan sweep | 214.6342
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of X1 Detector Raw during X Scan sweep | 12.74
| Wuhan YMTC Fab2 #3 | 2023-04-13 | Full range of X1 Detector during X Scan sweep | 218.8309
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.67
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.29
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.05
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.02
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.12
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.1
| Wuhan YMTC Fab2 #3 | 2023-04-13 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| Wuhan YMTC Fab2 #3 | 2023-04-13 | 50~200nm Step Height Sample (10 Repeats) | 178.1
| Wuhan YMTC Fab2 #3 | 2023-04-13 | 200~500nm Step Height Sample (10 Repeats) | 459.7
| Wuhan YMTC Fab2 #3 | 2023-04-13 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 13.5
| Wuhan YMTC Fab2 #3 | 2023-04-13 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 5.5
| Wuhan YMTC Fab2 #3 | 2023-04-13 | %GRR w/ roughness of non-patterened wafer | 14.6
| Yachay Tech University | 2023-04-13 | Z Scanner position at -3μm | -0.61
| Yachay Tech University | 2023-04-13 | Z Scanner position at +3μm | -0.04
| Yachay Tech University | 2023-04-13 | Z Linearity by 3 different sample height | 0.9959
| Yachay Tech University | 2023-04-13 | Z Height noise | 0.041
| Yachay Tech University | 2023-04-13 | Y Scan Stroke | 55
| Yachay Tech University | 2023-04-13 | Y Detector noise (XY Servo on) | 0.326
| Yachay Tech University | 2023-04-13 | Y Detector noise (XY Servo off) | 0.324
| Yachay Tech University | 2023-04-13 | Y axis OPM | 0.192
| Yachay Tech University | 2023-04-13 | X Scan Stroke | 55
| Yachay Tech University | 2023-04-13 | X Detector noise (XY Servo on) | 0.329
| Yachay Tech University | 2023-04-13 | X Detector noise (XY Servo off) | 0.325
| Yachay Tech University | 2023-04-13 | X axis OPM | 0.436
| Yachay Tech University | 2023-04-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Yachay Tech University | 2023-04-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.67
| Yachay Tech University | 2023-04-13 | NCM driving% with AC160TS cantilever | 1.762
| Yachay Tech University | 2023-04-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.434
| Yachay Tech University | 2023-04-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.296
| Yachay Tech University | 2023-04-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 101.9
| Yachay Tech University | 2023-04-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.9
| Yachay Tech University | 2023-04-13 | Lateral(C-D) sensitivity | 13
| Yachay Tech University | 2023-04-13 | Lateral(C-D) noise | 0.033
| Yachay Tech University | 2023-04-13 | Full range of Z Height during Z Scan sweep | 16.7
| Yachay Tech University | 2023-04-13 | Full range of Z Detector during Z Scan sweep | 12.5
| Yachay Tech University | 2023-04-13 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Yachay Tech University | 2023-04-13 | Full range of Y Detector during Y Scan sweep | 71.1
| Yachay Tech University | 2023-04-13 | Full range of X Detector Raw during X Scan sweep | 14
| Yachay Tech University | 2023-04-13 | Full range of X Detector during X Scan sweep | 71.3
| Yachay Tech University | 2023-04-13 | A-B noise | 11.955
| University of Electronic Science and Technology of China | 2023-04-12 | Z Scanner position at -3μm | 0.0061
| University of Electronic Science and Technology of China | 2023-04-12 | Z Scanner position at +3μm | 0.0103
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Z Linearity by Z Scanner position | 515.1
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Z Linearity by 3 different sample height | 3
| University of Electronic Science and Technology of China | 2023-04-12 | Z Linearity by 3 different sample height | 0.987
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Z Height noise | 0.082
| University of Electronic Science and Technology of China | 2023-04-12 | Z Height noise | 0.021
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Y Detector noise (XY Servo on) | 0.307
| University of Electronic Science and Technology of China | 2023-04-12 | Y Detector noise (XY Servo on) | 0.37
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Y Detector noise (XY Servo off) | 0.259
| University of Electronic Science and Technology of China | 2023-04-12 | Y Detector noise (XY Servo off) | 0.32
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Y axis OPM | 0.262
| University of Electronic Science and Technology of China | 2023-04-12 | Y axis OPM | 1.24
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | X Detector noise (XY Servo on) | 0.291
| University of Electronic Science and Technology of China | 2023-04-12 | X Detector noise (XY Servo on) | 0.39
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | X Detector noise (XY Servo off) | 0.263
| University of Electronic Science and Technology of China | 2023-04-12 | X Detector noise (XY Servo off) | 0.33
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | X axis OPM | 0.867
| University of Electronic Science and Technology of China | 2023-04-12 | X axis OPM | 1.05
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.49
| University of Electronic Science and Technology of China | 2023-04-12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| University of Electronic Science and Technology of China | 2023-04-12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| University of Electronic Science and Technology of China | 2023-04-12 | NCM driving% with AC160TS cantilever | 0.35
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | NCM Amplitude noise | 12
| University of Electronic Science and Technology of China | 2023-04-12 | NCM Amplitude noise | 16.4
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 65.928
| University of Electronic Science and Technology of China | 2023-04-12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.304
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 42.7
| University of Electronic Science and Technology of China | 2023-04-12 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.864
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 159
| University of Electronic Science and Technology of China | 2023-04-12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 168.4
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 159.4
| University of Electronic Science and Technology of China | 2023-04-12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 164.9
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Lateral(C-D) sensitivity | 7.7
| University of Electronic Science and Technology of China | 2023-04-12 | Lateral(C-D) sensitivity | 16.1
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Lateral(C-D) noise | 0.102
| University of Electronic Science and Technology of China | 2023-04-12 | Lateral(C-D) noise | 0.087
| University of Electronic Science and Technology of China | 2023-04-12 | Full range of Z Height during Z Scan sweep | 16.2
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Full range of Z Detector during Z Scan sweep | 25.855
| University of Electronic Science and Technology of China | 2023-04-12 | Full range of Z Detector during Z Scan sweep | 11.6
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Full range of Y Detector Raw during Y Scan sweep | 14.3
| University of Electronic Science and Technology of China | 2023-04-12 | Full range of Y Detector Raw during Y Scan sweep | 14
| University of Electronic Science and Technology of China | 2023-04-12 | Full range of Y Detector during Y Scan sweep | 117.7
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Full range of X Detector Raw during X Scan sweep | 14.1
| University of Electronic Science and Technology of China | 2023-04-12 | Full range of X Detector Raw during X Scan sweep | 14.1
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Full range of X Detector during X Scan sweep | 114
| University of Electronic Science and Technology of China | 2023-04-12 | Full range of X Detector during X Scan sweep | 115.9
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Approach position at -5μm | 0.0148
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | Approach position at +5μm | 0.0107
| Dongguan Huasheng Vacuum Coating Technology | 2023-04-12 | A-B noise | 9
| University of Electronic Science and Technology of China | 2023-04-12 | A-B noise | 10.24
| 전남대학교 신소재공학부 | 2023-04-10 | Z Scanner position at -3μm | 0.0032
| 전남대학교 신소재공학부 | 2023-04-10 | Z Scanner position at +3μm | 0.0029
| 전남대학교 신소재공학부 | 2023-04-10 | Z Linearity by 3 different sample height | 1.007
| 전남대학교 신소재공학부 | 2023-04-10 | Z Height noise | 0.034
| 전남대학교 신소재공학부 | 2023-04-10 | Y Detector noise (XY Servo on) | 0.26
| 전남대학교 신소재공학부 | 2023-04-10 | Y Detector noise (XY Servo off) | 0.244
| 전남대학교 신소재공학부 | 2023-04-10 | Y axis OPM | 0.426
| 전남대학교 신소재공학부 | 2023-04-10 | X Detector noise (XY Servo on) | 0.263
| 전남대학교 신소재공학부 | 2023-04-10 | X Detector noise (XY Servo off) | 0.244
| 전남대학교 신소재공학부 | 2023-04-10 | X axis OPM | 0.268
| 전남대학교 신소재공학부 | 2023-04-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| 전남대학교 신소재공학부 | 2023-04-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 전남대학교 신소재공학부 | 2023-04-10 | NCM driving% with AC160TS cantilever | 1.297
| 전남대학교 신소재공학부 | 2023-04-10 | NCM Amplitude noise | 17.35
| 전남대학교 신소재공학부 | 2023-04-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.323
| 전남대학교 신소재공학부 | 2023-04-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.498
| 전남대학교 신소재공학부 | 2023-04-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.8
| 전남대학교 신소재공학부 | 2023-04-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.5
| 전남대학교 신소재공학부 | 2023-04-10 | Lateral(C-D) sensitivity | 20.125
| 전남대학교 신소재공학부 | 2023-04-10 | Lateral(C-D) noise | 0.06
| 전남대학교 신소재공학부 | 2023-04-10 | Full range of Z Height during Z Scan sweep | 16.4
| 전남대학교 신소재공학부 | 2023-04-10 | Full range of Z Detector during Z Scan sweep | 11.5
| 전남대학교 신소재공학부 | 2023-04-10 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| 전남대학교 신소재공학부 | 2023-04-10 | Full range of Y Detector during Y Scan sweep | 75.4
| 전남대학교 신소재공학부 | 2023-04-10 | Full range of X Detector Raw during X Scan sweep | 14.2
| 전남대학교 신소재공학부 | 2023-04-10 | Full range of X Detector during X Scan sweep | 75.2
| 전남대학교 신소재공학부 | 2023-04-10 | A-B noise | 9.84
| MSSCORPS | 2023-04-07 | Z Linearity by 3 different sample height | 0.992
| MSSCORPS | 2023-04-07 | Z Height noise | 0.02
| MSSCORPS | 2023-04-07 | Y Detector noise (XY Servo on) | 0.29
| MSSCORPS | 2023-04-07 | Y Detector noise (XY Servo off) | 0.28
| MSSCORPS | 2023-04-07 | Y axis OPM | 1.67
| MSSCORPS | 2023-04-07 | X Detector noise (XY Servo on) | 0.28
| MSSCORPS | 2023-04-07 | X Detector noise (XY Servo off) | 0.28
| MSSCORPS | 2023-04-07 | X axis OPM | 1.31
| MSSCORPS | 2023-04-07 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| MSSCORPS | 2023-04-07 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.21
| MSSCORPS | 2023-04-07 | NCM driving% with AC160TS cantilever | 4.721
| MSSCORPS | 2023-04-07 | NCM Amplitude noise | 8
| MSSCORPS | 2023-04-07 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.559
| MSSCORPS | 2023-04-07 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.835
| MSSCORPS | 2023-04-07 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 150.1
| MSSCORPS | 2023-04-07 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 150.6
| MSSCORPS | 2023-04-07 | Full range of Z Height during Z Scan sweep | 15.5
| MSSCORPS | 2023-04-07 | Lateral(C-D) sensitivity | 20.564
| MSSCORPS | 2023-04-07 | Lateral(C-D) noise | 0.137
| MSSCORPS | 2023-04-07 | Full range of Z Detector during Z Scan sweep | 14
| MSSCORPS | 2023-04-07 | Full range of Y Detector Raw during Y Scan sweep | 14
| MSSCORPS | 2023-04-07 | Full range of Y Detector during Y Scan sweep | 105.1
| MSSCORPS | 2023-04-07 | Full range of X Detector Raw during X Scan sweep | 13.8
| MSSCORPS | 2023-04-07 | Full range of X Detector during X Scan sweep | 104.55
| MSSCORPS | 2023-04-07 | Approach position at -3μm | -0.0057
| MSSCORPS | 2023-04-07 | Approach position at +3μm | -0.0118
| MSSCORPS | 2023-04-07 | A-B noise | 8.3
| PSE DEMO_NX10 | 2023-04-06 | Z Scanner position at -3μm | 0.0094
| Leica Biosystems | 2023-04-06 | Z Scanner position at -3μm | 0.0042
| PSE DEMO_NX10 | 2023-04-06 | Z Scanner position at +3μm | 0.0077
| Leica Biosystems | 2023-04-06 | Z Scanner position at +3μm | 0.0021
| PSE DEMO_NX10 | 2023-04-06 | Z Linearity by 3 different sample height | 0.997
| Leica Biosystems | 2023-04-06 | Z Linearity by 3 different sample height | 0.996
| PSE DEMO_NX10 | 2023-04-06 | Z Height noise | 0.022
| Leica Biosystems | 2023-04-06 | Z Height noise | 0.033
| PSE DEMO_NX10 | 2023-04-06 | Y Detector noise (XY Servo on) | 0.24
| Leica Biosystems | 2023-04-06 | Y Detector noise (XY Servo on) | 0.272
| PSE DEMO_NX10 | 2023-04-06 | Y Detector noise (XY Servo off) | 0.22
| Leica Biosystems | 2023-04-06 | Y Detector noise (XY Servo off) | 0.253
| PSE DEMO_NX10 | 2023-04-06 | Y axis OPM | 0.42
| Leica Biosystems | 2023-04-06 | Y axis OPM | 0.393
| PSE DEMO_NX10 | 2023-04-06 | X Detector noise (XY Servo on) | 0.24
| Leica Biosystems | 2023-04-06 | X Detector noise (XY Servo on) | 0.276
| PSE DEMO_NX10 | 2023-04-06 | X Detector noise (XY Servo off) | 0.22
| Leica Biosystems | 2023-04-06 | X Detector noise (XY Servo off) | 0.256
| PSE DEMO_NX10 | 2023-04-06 | X axis OPM | 0.4
| Leica Biosystems | 2023-04-06 | X axis OPM | 0.516
| PSE DEMO_NX10 | 2023-04-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| Leica Biosystems | 2023-04-06 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| PSE DEMO_NX10 | 2023-04-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| Leica Biosystems | 2023-04-06 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| PSE DEMO_NX10 | 2023-04-06 | NCM driving% with AC160TS cantilever | 0.73
| Leica Biosystems | 2023-04-06 | NCM driving% with AC160TS cantilever | 2.921
| PSE DEMO_NX10 | 2023-04-06 | NCM Amplitude noise | 15.06
| Leica Biosystems | 2023-04-06 | NCM Amplitude noise | 15.636
| PSE DEMO_NX10 | 2023-04-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.5
| Leica Biosystems | 2023-04-06 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.645
| PSE DEMO_NX10 | 2023-04-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.5
| Leica Biosystems | 2023-04-06 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.999
| PSE DEMO_NX10 | 2023-04-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105.3
| Leica Biosystems | 2023-04-06 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 103.8
| PSE DEMO_NX10 | 2023-04-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 102.3
| Leica Biosystems | 2023-04-06 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.8
| PSE DEMO_NX10 | 2023-04-06 | Lateral(C-D) sensitivity | 14.1
| Leica Biosystems | 2023-04-06 | Lateral(C-D) sensitivity | 33.259
| PSE DEMO_NX10 | 2023-04-06 | Lateral(C-D) noise | 0.064
| Leica Biosystems | 2023-04-06 | Lateral(C-D) noise | 0.055
| PSE DEMO_NX10 | 2023-04-06 | Full range of Z Height during Z Scan sweep | 16.7
| Leica Biosystems | 2023-04-06 | Full range of Z Height during Z Scan sweep | 16.8
| PSE DEMO_NX10 | 2023-04-06 | Full range of Z Detector during Z Scan sweep | 12.3
| Leica Biosystems | 2023-04-06 | Full range of Z Detector during Z Scan sweep | 12.4
| PSE DEMO_NX10 | 2023-04-06 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| Leica Biosystems | 2023-04-06 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| PSE DEMO_NX10 | 2023-04-06 | Full range of Y Detector during Y Scan sweep | 74.8
| Leica Biosystems | 2023-04-06 | Full range of Y Detector during Y Scan sweep | 76.5
| PSE DEMO_NX10 | 2023-04-06 | Full range of X Detector Raw during X Scan sweep | 14.7
| Leica Biosystems | 2023-04-06 | Full range of X Detector Raw during X Scan sweep | 14.6
| PSE DEMO_NX10 | 2023-04-06 | Full range of X Detector during X Scan sweep | 75.8
| Leica Biosystems | 2023-04-06 | Full range of X Detector during X Scan sweep | 77
| PSE DEMO_NX10 | 2023-04-06 | A-B noise | 10.87
| Leica Biosystems | 2023-04-06 | A-B noise | 7.682
| National Institutes of Technology (NITs) | 2023-04-05 | Z Scanner position at -3μm | 0.0036
| National Institutes of Technology (NITs) | 2023-04-05 | Z Scanner position at +3μm | 0.0119
| National Institutes of Technology (NITs) | 2023-04-05 | Z Linearity by 3 different sample height | 1
| National Institutes of Technology (NITs) | 2023-04-05 | Z Height noise | 0.034
| National Institutes of Technology (NITs) | 2023-04-05 | Y Scan Stroke | 189
| National Institutes of Technology (NITs) | 2023-04-05 | Y Detector noise (XY Servo on) | 0.36
| National Institutes of Technology (NITs) | 2023-04-05 | Y axis OPM | 1.72
| National Institutes of Technology (NITs) | 2023-04-05 | X Scan Stroke | 180.3
| National Institutes of Technology (NITs) | 2023-04-05 | X Detector noise (XY Servo on) | 0.34
| National Institutes of Technology (NITs) | 2023-04-05 | X Detector noise (XY Servo off) | 0.32
| National Institutes of Technology (NITs) | 2023-04-05 | X axis OPM | 1.12
| National Institutes of Technology (NITs) | 2023-04-05 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.36
| National Institutes of Technology (NITs) | 2023-04-05 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.78
| National Institutes of Technology (NITs) | 2023-04-05 | NCM driving% with AC160TS cantilever | 2.446
| National Institutes of Technology (NITs) | 2023-04-05 | NCM Amplitude noise | 13.63
| National Institutes of Technology (NITs) | 2023-04-05 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.992
| National Institutes of Technology (NITs) | 2023-04-05 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.93
| National Institutes of Technology (NITs) | 2023-04-05 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 254.6
| National Institutes of Technology (NITs) | 2023-04-05 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 252.8
| National Institutes of Technology (NITs) | 2023-04-05 | Lateral(C-D) sensitivity | 14.3
| National Institutes of Technology (NITs) | 2023-04-05 | Lateral(C-D) noise | 0.068
| National Institutes of Technology (NITs) | 2023-04-05 | Full range of Z Height during Z Scan sweep | 16.1
| National Institutes of Technology (NITs) | 2023-04-05 | Full range of Z Detector during Z Scan sweep | 12.4
| National Institutes of Technology (NITs) | 2023-04-05 | Full range of Y Detector Raw during Y Scan sweep | 14.8
| National Institutes of Technology (NITs) | 2023-04-05 | Full range of Y Detector during Y Scan sweep | 187.2
| National Institutes of Technology (NITs) | 2023-04-05 | Full range of X Detector Raw during X Scan sweep | 14.2
| National Institutes of Technology (NITs) | 2023-04-05 | Full range of X Detector during X Scan sweep | 179.5
| National Institutes of Technology (NITs) | 2023-04-05 | A-B noise | 9.18
| Fudan University(Department of Physics) | 2023-04-03 | Z Scanner position at -3μm | 0.003
| Fudan University(Department of Physics) | 2023-04-03 | Z Scanner position at +3μm | -0.012
| Fudan University(Department of Physics) | 2023-04-03 | Z Linearity by 3 different sample height | 1.001
| Fudan University(Department of Physics) | 2023-04-03 | Z Height noise | 0.04
| Fudan University(Department of Physics) | 2023-04-03 | Y Detector noise (XY Servo on) | 0.3
| Fudan University(Department of Physics) | 2023-04-03 | Y Detector noise (XY Servo off) | 0.26
| Fudan University(Department of Physics) | 2023-04-03 | Y axis OPM | 0.4
| Fudan University(Department of Physics) | 2023-04-03 | X Detector noise (XY Servo on) | 0.31
| Fudan University(Department of Physics) | 2023-04-03 | X Detector noise (XY Servo off) | 0.26
| Fudan University(Department of Physics) | 2023-04-03 | X axis OPM | 0.9
| Fudan University(Department of Physics) | 2023-04-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.32
| Fudan University(Department of Physics) | 2023-04-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| Fudan University(Department of Physics) | 2023-04-03 | NCM driving% with AC160TS cantilever | 0.65
| Fudan University(Department of Physics) | 2023-04-03 | NCM Amplitude noise | 14.6
| Fudan University(Department of Physics) | 2023-04-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.6
| Fudan University(Department of Physics) | 2023-04-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.4
| Fudan University(Department of Physics) | 2023-04-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.9
| Fudan University(Department of Physics) | 2023-04-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.9
| Fudan University(Department of Physics) | 2023-04-03 | Lateral(C-D) sensitivity | 5.6
| Fudan University(Department of Physics) | 2023-04-03 | Lateral(C-D) noise | 0.102
| Fudan University(Department of Physics) | 2023-04-03 | Full range of Z Height during Z Scan sweep | 16.8
| Fudan University(Department of Physics) | 2023-04-03 | Full range of Z Detector during Z Scan sweep | 11.4
| Fudan University(Department of Physics) | 2023-04-03 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Fudan University(Department of Physics) | 2023-04-03 | Full range of Y Detector during Y Scan sweep | 115.3
| Fudan University(Department of Physics) | 2023-04-03 | Full range of X Detector Raw during X Scan sweep | 14.1
| Fudan University(Department of Physics) | 2023-04-03 | Full range of X Detector during X Scan sweep | 114.3
| Fudan University(Department of Physics) | 2023-04-03 | A-B noise | 11.9
| 경북대 하이브리드부품연구원 | 2023-03-31 | Z Scanner position at -3μm | 0.006
| 경북대 하이브리드부품연구원 | 2023-03-31 | Z Scanner position at +3μm | 0.012
| 경북대 하이브리드부품연구원 | 2023-03-31 | Z Linearity by 3 different sample height | 0.986
| 경북대 하이브리드부품연구원 | 2023-03-31 | Z Height noise | 0.034
| 경북대 하이브리드부품연구원 | 2023-03-31 | Y Scan Stroke | 182.8
| 경북대 하이브리드부품연구원 | 2023-03-31 | Y Detector noise (XY Servo on) | 0.39
| 경북대 하이브리드부품연구원 | 2023-03-31 | Y Detector noise (XY Servo off) | 0.31
| 경북대 하이브리드부품연구원 | 2023-03-31 | Y axis OPM | 1.2
| 경북대 하이브리드부품연구원 | 2023-03-31 | X Scan Stroke | 181
| 경북대 하이브리드부품연구원 | 2023-03-31 | X Detector noise (XY Servo off) | 0.31
| 경북대 하이브리드부품연구원 | 2023-03-31 | X axis OPM | 1
| 경북대 하이브리드부품연구원 | 2023-03-31 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.43
| 경북대 하이브리드부품연구원 | 2023-03-31 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.43
| 경북대 하이브리드부품연구원 | 2023-03-31 | NCM driving% with AC160TS cantilever | 0.35
| 경북대 하이브리드부품연구원 | 2023-03-31 | NCM Amplitude noise | 17
| 경북대 하이브리드부품연구원 | 2023-03-31 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.397
| 경북대 하이브리드부품연구원 | 2023-03-31 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.736
| 경북대 하이브리드부품연구원 | 2023-03-31 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 250.4
| 경북대 하이브리드부품연구원 | 2023-03-31 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 249.7
| 경북대 하이브리드부품연구원 | 2023-03-31 | Lateral(C-D) sensitivity | 16.373
| 경북대 하이브리드부품연구원 | 2023-03-31 | Lateral(C-D) noise | 0.086
| 경북대 하이브리드부품연구원 | 2023-03-31 | Full range of Z Height during Z Scan sweep | 16.4
| 경북대 하이브리드부품연구원 | 2023-03-31 | Full range of Z Detector during Z Scan sweep | 11.4
| 경북대 하이브리드부품연구원 | 2023-03-31 | Full range of Y Detector Raw during Y Scan sweep | 14.6
| 경북대 하이브리드부품연구원 | 2023-03-31 | Full range of Y Detector during Y Scan sweep | 183.1
| 경북대 하이브리드부품연구원 | 2023-03-31 | Full range of X Detector Raw during X Scan sweep | 14.5
| 경북대 하이브리드부품연구원 | 2023-03-31 | Full range of X Detector during X Scan sweep | 181.3
| 경북대 하이브리드부품연구원 | 2023-03-31 | A-B noise | 14
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Z Scanner position at -3μm | 0.0028
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Z Scanner position at +3μm | 0.0058
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Z Linearity by 3 different sample height | 0.985
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Z Height noise | 0.021
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Y Scan Stroke | 55
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Y Detector noise (XY Servo on) | 0.36
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Y Detector noise (XY Servo off) | 0.35
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Y axis OPM | 1.1
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | X Scan Stroke | 55
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | X Detector noise (XY Servo on) | 0.36
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | X Detector noise (XY Servo off) | 0.35
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | X axis OPM | 0.7
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.44
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | NCM driving% with AC160TS cantilever | 0.61
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | NCM Amplitude noise | 14.298
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.6
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.9
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 98.3
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.1
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Lateral(C-D) sensitivity | 5.4
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Lateral(C-D) noise | 0.058
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Full range of Z Height during Z Scan sweep | 16.3
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Full range of Z Detector during Z Scan sweep | 11.3
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Full range of Y Detector during Y Scan sweep | 70.3
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Full range of X Detector Raw during X Scan sweep | 14.2
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | Full range of X Detector during X Scan sweep | 70.1
| 선문대학교(디스플레이반도체공학과) | 2023-03-30 | A-B noise | 9.754
| Zhonghuan semiconductor | 2023-03-29 | Z Scanner position at -3μm | 0.0016
| Zhonghuan semiconductor | 2023-03-29 | Z Scanner position at +3μm | 0.0194
| Zhonghuan semiconductor | 2023-03-29 | Z Linearity by 3 different sample height | 0.999
| Zhonghuan semiconductor | 2023-03-29 | Z Height noise | 0.017
| Zhonghuan semiconductor | 2023-03-29 | Y2 Detector noise (XY Servo on) | 0.23
| Zhonghuan semiconductor | 2023-03-29 | Y2 Detector noise (XY Servo off) | 0.14
| Zhonghuan semiconductor | 2023-03-29 | Y1 Detector noise (XY Servo on) | 0.21
| Zhonghuan semiconductor | 2023-03-29 | Y1 Detector noise (XY Servo off) | 0.13
| Zhonghuan semiconductor | 2023-03-29 | Y Scan Stroke | 181.4
| Zhonghuan semiconductor | 2023-03-29 | Y Detector noise (XY Servo on) | 0.18
| Zhonghuan semiconductor | 2023-03-29 | Y Detector noise (XY Servo off) | 0.1
| Zhonghuan semiconductor | 2023-03-29 | Y axis OPM | 0.37
| Zhonghuan semiconductor | 2023-03-29 | X2 Detector noise (XY Servo on) | 0.23
| Zhonghuan semiconductor | 2023-03-29 | X2 Detector noise (XY Servo off) | 0.14
| Zhonghuan semiconductor | 2023-03-29 | X1 Detector noise (XY Servo on) | 0.23
| Zhonghuan semiconductor | 2023-03-29 | X1 Detector noise (XY Servo off) | 0.14
| Zhonghuan semiconductor | 2023-03-29 | X Scan Stroke | 184.7
| Zhonghuan semiconductor | 2023-03-29 | -X Direction software Limit | -102160
| Zhonghuan semiconductor | 2023-03-29 | -X Direction software Limit | -102832
| Zhonghuan semiconductor | 2023-03-29 | X Detector noise (XY Servo on) | 0.18
| Zhonghuan semiconductor | 2023-03-29 | X Detector noise (XY Servo off) | 0.1
| Zhonghuan semiconductor | 2023-03-29 | X axis OPM | 0.9
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.34
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.73
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of Wafer Zeroscan (Top) | 0.4
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of Wafer Zeroscan (Right) | 0.45
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of Wafer Zeroscan (Left) | 0.41
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of Wafer ZeroScan (Center) | 0.4
| Zhonghuan semiconductor | 2023-03-29 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.38
| Zhonghuan semiconductor | 2023-03-29 | NCM driving% with AC160TS cantilever | 3.39
| Zhonghuan semiconductor | 2023-03-29 | NCM Amplitude noise | 15.13
| Zhonghuan semiconductor | 2023-03-29 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.2
| Zhonghuan semiconductor | 2023-03-29 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.3
| Zhonghuan semiconductor | 2023-03-29 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 249.7
| Zhonghuan semiconductor | 2023-03-29 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 254.2
| Zhonghuan semiconductor | 2023-03-29 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 258.4
| Zhonghuan semiconductor | 2023-03-29 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 254.6
| Zhonghuan semiconductor | 2023-03-29 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.07
| Zhonghuan semiconductor | 2023-03-29 | Lateral(C-D) sensitivity | 4.8
| Zhonghuan semiconductor | 2023-03-29 | Lateral(C-D) noise | 0.142
| Zhonghuan semiconductor | 2023-03-29 | Full range of Z Height during Z Scan sweep | 16.2
| Zhonghuan semiconductor | 2023-03-29 | Full range of Z Detector during Z Scan sweep | 11.5
| Zhonghuan semiconductor | 2023-03-29 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| Zhonghuan semiconductor | 2023-03-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.4
| Zhonghuan semiconductor | 2023-03-29 | Full range of Y Detector during Y Scan sweep | 182.8
| Zhonghuan semiconductor | 2023-03-29 | Full range of X2 Detector Raw during X Scan sweep | 14.4
| Zhonghuan semiconductor | 2023-03-29 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| Zhonghuan semiconductor | 2023-03-29 | Full range of X Detector during X Scan sweep | 184.7
| Zhonghuan semiconductor | 2023-03-29 | A-B noise | 9.72
| OCI 주식회사 | 2023-03-28 | Z Scanner position at -3μm | 0.0019
| OCI 주식회사 | 2023-03-28 | Z Scanner position at +3μm | 0.0007
| OCI 주식회사 | 2023-03-28 | Z Linearity by 3 different sample height | 1.007
| OCI 주식회사 | 2023-03-28 | Z Height noise | 0.034
| OCI 주식회사 | 2023-03-28 | Y2 Detector noise (XY Servo on) | 0.34
| OCI 주식회사 | 2023-03-28 | Y2 Detector noise (XY Servo off) | 0.21
| OCI 주식회사 | 2023-03-28 | Y1 Detector noise (XY Servo on) | 0.34
| OCI 주식회사 | 2023-03-28 | Y1 Detector noise (XY Servo off) | 0.21
| OCI 주식회사 | 2023-03-28 | Y Scan Stroke | 173.2
| OCI 주식회사 | 2023-03-28 | Y Detector noise (XY Servo on) | 0.27
| OCI 주식회사 | 2023-03-28 | Y Detector noise (XY Servo off) | 0.15
| OCI 주식회사 | 2023-03-28 | Y axis OPM | 0.93
| OCI 주식회사 | 2023-03-28 | X2 Detector noise (XY Servo on) | 0.35
| OCI 주식회사 | 2023-03-28 | X2 Detector noise (XY Servo off) | 0.21
| OCI 주식회사 | 2023-03-28 | X1 Detector noise (XY Servo on) | 0.34
| OCI 주식회사 | 2023-03-28 | X1 Detector noise (XY Servo off) | 0.21
| OCI 주식회사 | 2023-03-28 | X Scan Stroke | 178.2
| OCI 주식회사 | 2023-03-28 | -X Direction software Limit | -102129
| OCI 주식회사 | 2023-03-28 | -X Direction software Limit | -103162
| OCI 주식회사 | 2023-03-28 | X Detector noise (XY Servo on) | 0.27
| OCI 주식회사 | 2023-03-28 | X Detector noise (XY Servo off) | 0.15
| OCI 주식회사 | 2023-03-28 | X axis OPM | 1.25
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.3
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.75
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of Wafer Zeroscan (Top) | 0.31
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of Wafer Zeroscan (Right) | 0.3
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of Wafer Zeroscan (Left) | 0.3
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of Wafer ZeroScan (Center) | 0.29
| OCI 주식회사 | 2023-03-28 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.28
| OCI 주식회사 | 2023-03-28 | NCM driving% with AC160TS cantilever | 3.44
| OCI 주식회사 | 2023-03-28 | NCM Amplitude noise | 14
| OCI 주식회사 | 2023-03-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.112
| OCI 주식회사 | 2023-03-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.641
| OCI 주식회사 | 2023-03-28 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 237.7
| OCI 주식회사 | 2023-03-28 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 243.1
| OCI 주식회사 | 2023-03-28 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 242.6
| OCI 주식회사 | 2023-03-28 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 249.6
| OCI 주식회사 | 2023-03-28 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| OCI 주식회사 | 2023-03-28 | Lateral(C-D) sensitivity | 4.3
| OCI 주식회사 | 2023-03-28 | Lateral(C-D) noise | 0.12
| OCI 주식회사 | 2023-03-28 | Full range of Z Height during Z Scan sweep | 16
| OCI 주식회사 | 2023-03-28 | Full range of Z Detector during Z Scan sweep | 11.5
| OCI 주식회사 | 2023-03-28 | Full range of Y2 Detector Raw during Y Scan sweep | 14.5
| OCI 주식회사 | 2023-03-28 | Full range of Y1 Detector Raw during Y Scan sweep | 14.5
| OCI 주식회사 | 2023-03-28 | Full range of Y Detector during Y Scan sweep | 173
| OCI 주식회사 | 2023-03-28 | Full range of X2 Detector Raw during X Scan sweep | 14.5
| OCI 주식회사 | 2023-03-28 | Full range of X1 Detector Raw during X Scan sweep | 14.4
| OCI 주식회사 | 2023-03-28 | Full range of X Detector during X Scan sweep | 176.6
| OCI 주식회사 | 2023-03-28 | A-B noise | 9
| QMOT | 2023-03-27 | Z Height noise | 0.0341
| QMOT | 2023-03-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.4527
| QMOT | 2023-03-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.4408
| QMOT | 2023-03-27 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 62.9316
| QMOT | 2023-03-27 | Y Detector noise (XY Servo on) | 0.1641
| QMOT | 2023-03-27 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 88.1475
| QMOT | 2023-03-27 | Y axis, Repeatability @1st position | 0.05
| QMOT | 2023-03-27 | Y axis, Repeatability @ 4th position | 0.05
| QMOT | 2023-03-27 | Y axis, Repeatability @ 3rd position | 0.05
| QMOT | 2023-03-27 | Y axis, Repeatability @ 2nd position | 0.04
| QMOT | 2023-03-27 | Y axis, Accuracy @1st position | 0.54
| QMOT | 2023-03-27 | Y axis, Accuracy @ 4th position | 0.6
| QMOT | 2023-03-27 | Y axis, Accuracy @ 3rd position | 0.56
| QMOT | 2023-03-27 | Y axis, Accuracy @ 2nd position | 0.51
| QMOT | 2023-03-27 | Y axis OPM | 3.6115
| QMOT | 2023-03-27 | XY Orthogonality | 90.2
| QMOT | 2023-03-27 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 61.5488
| QMOT | 2023-03-27 | X Detector noise (XY Servo on) | 0.1455
| QMOT | 2023-03-27 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 82.84
| QMOT | 2023-03-27 | X axis, Repeatability @1st position | 0.03
| QMOT | 2023-03-27 | X axis, Repeatability @ 4th position | 0.05
| QMOT | 2023-03-27 | X axis, Repeatability @ 3rd position | 0.03
| QMOT | 2023-03-27 | X axis, Repeatability @ 2nd position | 0.03
| QMOT | 2023-03-27 | X axis, Accuracy @1st position | 0.19
| QMOT | 2023-03-27 | X axis, Accuracy @ 4th position | 0.38
| QMOT | 2023-03-27 | X axis, Accuracy @ 3rd position | 0.27
| QMOT | 2023-03-27 | X axis, Accuracy @ 2nd position | 0.22
| QMOT | 2023-03-27 | X axis OPM | 0.7554
| QMOT | 2023-03-27 | RMS of ZeroScan | 0.41
| QMOT | 2023-03-27 | Image offset after moving from -Y to +Y (Y scan) | 29.2969
| QMOT | 2023-03-27 | Image offset after moving from -Y to +Y (X scan) | 19.5312
| QMOT | 2023-03-27 | Image offset after moving from -X to +X (Y scan) | 102.5391
| QMOT | 2023-03-27 | Image offset after moving from -X to +X (X scan) | 43.9453
| QMOT | 2023-03-27 | Image offset after moving from +Y to -Y (Y scan) | 43.9453
| QMOT | 2023-03-27 | Image offset after moving from +Y to -Y (X scan) | 63.4766
| QMOT | 2023-03-27 | Image offset after moving from +X to -X (Y scan) | 112.3047
| QMOT | 2023-03-27 | Image offset after moving from +X to -X (X scan) | 63.4766
| QMOT | 2023-03-27 | Full range of Z Height during Z Scan sweep | 17.1643
| QMOT | 2023-03-27 | Full range of Z Height during Z Scan sweep | 75.6632
| QMOT | 2023-03-27 | Full range of Z Detector during Z Scan sweep | 13.87
| QMOT | 2023-03-27 | Full range of Y Detector Raw during Y Scan sweep | 14.45
| QMOT | 2023-03-27 | Full range of Y Detector during Y Scan sweep | 64.4823
| QMOT | 2023-03-27 | Full range of X Detector Raw during X Scan sweep | 14.42
| QMOT | 2023-03-27 | A－B difference during 3 times run test | 0.2
| QMOT | 2023-03-27 | A＋B difference during 3 times run test | 0.11
| 한국기계연구원 | 2023-03-27 | Z Scanner position at -3μm | 0.008
| 한국기계연구원 | 2023-03-27 | Z Scanner position at +3μm | 0.0247
| 한국기계연구원 | 2023-03-27 | Z Linearity by 3 different sample height | 1.005
| 한국기계연구원 | 2023-03-27 | Z Height noise | 0.021
| 한국기계연구원 | 2023-03-27 | Y Detector noise (XY Servo on) | 0.26
| 한국기계연구원 | 2023-03-27 | Y Detector noise (XY Servo off) | 0.25
| 한국기계연구원 | 2023-03-27 | Y axis OPM | 0.18
| 한국기계연구원 | 2023-03-27 | X Detector noise (XY Servo off) | 0.26
| 한국기계연구원 | 2023-03-27 | X axis OPM | 1
| 한국기계연구원 | 2023-03-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.25
| 한국기계연구원 | 2023-03-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.3
| 한국기계연구원 | 2023-03-27 | NCM driving% with AC160TS cantilever | 0.56
| 한국기계연구원 | 2023-03-27 | NCM Amplitude noise | 16.1
| 한국기계연구원 | 2023-03-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.4
| 한국기계연구원 | 2023-03-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.6
| 한국기계연구원 | 2023-03-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 94.9
| 한국기계연구원 | 2023-03-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 95.2
| 한국기계연구원 | 2023-03-27 | Lateral(C-D) sensitivity | 19.723
| 한국기계연구원 | 2023-03-27 | Lateral(C-D) noise | 0.07
| 한국기계연구원 | 2023-03-27 | Full range of Z Height during Z Scan sweep | 16.8
| 한국기계연구원 | 2023-03-27 | Full range of Z Detector during Z Scan sweep | 12.6
| 한국기계연구원 | 2023-03-27 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| 한국기계연구원 | 2023-03-27 | Full range of Y Detector during Y Scan sweep | 65.6
| 한국기계연구원 | 2023-03-27 | Full range of X Detector Raw during X Scan sweep | 14.4
| 한국기계연구원 | 2023-03-27 | Full range of X Detector during X Scan sweep | 67.6
| 한국기계연구원 | 2023-03-27 | A-B noise | 8.7
| Park Systems Europe GmbH | 2023-03-23 | Z Scanner position at -3μm | 0.0031
| Park Systems Europe GmbH | 2023-03-23 | Z Scanner position at +3μm | 0.0082
| Park Systems Europe GmbH | 2023-03-23 | Z Linearity by 3 different sample height | 1.001
| PSI | 2023-03-23 | Z Linearity by 3 different sample height | 0.983
| Park Systems Europe GmbH | 2023-03-23 | Z Height noise | 0.022
| PSI | 2023-03-23 | Z Height noise | 0.028
| Park Systems Europe GmbH | 2023-03-23 | Y Detector noise (XY Servo on) | 0.248
| PSI | 2023-03-23 | Y Detector noise (XY Servo on) | 0.33
| Park Systems Europe GmbH | 2023-03-23 | Y Detector noise (XY Servo off) | 0.246
| PSI | 2023-03-23 | Y Detector noise (XY Servo off) | 0.27
| Park Systems Europe GmbH | 2023-03-23 | Y axis OPM | 0.57
| PSI | 2023-03-23 | Y axis OPM | 0.76
| Park Systems Europe GmbH | 2023-03-23 | X Detector noise (XY Servo on) | 0.268
| PSI | 2023-03-23 | X Detector noise (XY Servo on) | 0.34
| Park Systems Europe GmbH | 2023-03-23 | X Detector noise (XY Servo off) | 0.249
| PSI | 2023-03-23 | X Detector noise (XY Servo off) | 0.28
| Park Systems Europe GmbH | 2023-03-23 | X axis OPM | 0.449
| PSI | 2023-03-23 | X axis OPM | 0.65
| Park Systems Europe GmbH | 2023-03-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| PSI | 2023-03-23 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| Park Systems Europe GmbH | 2023-03-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.36
| PSI | 2023-03-23 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.28
| Park Systems Europe GmbH | 2023-03-23 | NCM driving% with AC160TS cantilever | 1.702
| PSI | 2023-03-23 | NCM driving% with AC160TS cantilever | 11.78
| Park Systems Europe GmbH | 2023-03-23 | NCM Amplitude noise | 10.664
| PSI | 2023-03-23 | NCM Amplitude noise | 10.4
| Park Systems Europe GmbH | 2023-03-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.559
| PSI | 2023-03-23 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.9
| Park Systems Europe GmbH | 2023-03-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.146
| PSI | 2023-03-23 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.86
| Park Systems Europe GmbH | 2023-03-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.7
| PSI | 2023-03-23 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 155.9
| Park Systems Europe GmbH | 2023-03-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 105.8
| PSI | 2023-03-23 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 153.4
| PSI | 2023-03-23 | Full range of Z Height during Z Scan sweep | 15.8
| Park Systems Europe GmbH | 2023-03-23 | Lateral(C-D) sensitivity | 7.6
| PSI | 2023-03-23 | Lateral(C-D) sensitivity | 13.406
| Park Systems Europe GmbH | 2023-03-23 | Lateral(C-D) noise | 0.097
| PSI | 2023-03-23 | Lateral(C-D) noise | 0.122
| Park Systems Europe GmbH | 2023-03-23 | Full range of Z Height during Z Scan sweep | 16.3
| Park Systems Europe GmbH | 2023-03-23 | Full range of Z Detector during Z Scan sweep | 11.3
| PSI | 2023-03-23 | Full range of Z Detector during Z Scan sweep | 13.2
| Park Systems Europe GmbH | 2023-03-23 | Full range of Y Detector Raw during Y Scan sweep | 14.1
| PSI | 2023-03-23 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| Park Systems Europe GmbH | 2023-03-23 | Full range of Y Detector during Y Scan sweep | 74.9
| PSI | 2023-03-23 | Full range of Y Detector during Y Scan sweep | 115.4
| Park Systems Europe GmbH | 2023-03-23 | Full range of X Detector Raw during X Scan sweep | 14.2
| PSI | 2023-03-23 | Full range of X Detector Raw during X Scan sweep | 14.8
| Park Systems Europe GmbH | 2023-03-23 | Full range of X Detector during X Scan sweep | 75.2
| PSI | 2023-03-23 | Full range of X Detector during X Scan sweep | 114.8
| PSI | 2023-03-23 | Approach position at -3μm | 0.0011
| PSI | 2023-03-23 | Approach position at +3μm | 0.0225
| Park Systems Europe GmbH | 2023-03-23 | A-B noise | 10.034
| PSI | 2023-03-23 | A-B noise | 10.5
| Park Systems UK | 2023-03-22 | Z Scanner position at -3μm | 0.0128
| Park Systems UK | 2023-03-22 | Z Scanner position at +3μm | 0.0027
| Park Systems UK | 2023-03-22 | Z Linearity by 3 different sample height | 0.98
| Park Systems UK | 2023-03-22 | Z Height noise | 0.019
| Park Systems UK | 2023-03-22 | Y2 Detector noise (XY Servo on) | 0.24
| Park Systems UK | 2023-03-22 | Y2 Detector noise (XY Servo off) | 0.13
| Park Systems UK | 2023-03-22 | Y1 Detector noise (XY Servo on) | 0.25
| Park Systems UK | 2023-03-22 | Y1 Detector noise (XY Servo off) | 0.14
| Park Systems UK | 2023-03-22 | Y Scan Stroke | 179.6
| Park Systems UK | 2023-03-22 | Y Detector noise (XY Servo on) | 0.18
| Park Systems UK | 2023-03-22 | Y Detector noise (XY Servo off) | 0.1
| Park Systems UK | 2023-03-22 | Y axis OPM | 0.8
| Park Systems UK | 2023-03-22 | X2 Detector noise (XY Servo on) | 0.32
| Park Systems UK | 2023-03-22 | X2 Detector noise (XY Servo off) | 0.16
| Park Systems UK | 2023-03-22 | X1 Detector noise (XY Servo on) | 0.28
| Park Systems UK | 2023-03-22 | X1 Detector noise (XY Servo off) | 0.14
| Park Systems UK | 2023-03-22 | X Scan Stroke | 174.7
| Park Systems UK | 2023-03-22 | -X Direction software Limit | -102238
| Park Systems UK | 2023-03-22 | -X Direction software Limit | -103365
| Park Systems UK | 2023-03-22 | X Detector noise (XY Servo on) | 0.25
| Park Systems UK | 2023-03-22 | X Detector noise (XY Servo off) | 0.11
| Park Systems UK | 2023-03-22 | X axis OPM | 0.56
| Park Systems UK | 2023-03-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.27
| Park Systems UK | 2023-03-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.45
| Park Systems UK | 2023-03-22 | RMS of Z Height of Wafer Zeroscan (Top) | 0.28
| Park Systems UK | 2023-03-22 | RMS of Z Height of Wafer Zeroscan (Right) | 0.25
| Park Systems UK | 2023-03-22 | RMS of Z Height of Wafer Zeroscan (Left) | 0.25
| Park Systems UK | 2023-03-22 | RMS of Z Height of Wafer ZeroScan (Center) | 0.27
| Park Systems UK | 2023-03-22 | RMS of Z Height of Wafer Zeroscan (Bottom) | 0.26
| Park Systems UK | 2023-03-22 | NCM driving% with AC160TS cantilever | 1.98
| Park Systems UK | 2023-03-22 | NCM Amplitude noise | 15.29
| Park Systems UK | 2023-03-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.7
| Park Systems UK | 2023-03-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.7
| Park Systems UK | 2023-03-22 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 251
| Park Systems UK | 2023-03-22 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 249.6
| Park Systems UK | 2023-03-22 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 244.3
| Park Systems UK | 2023-03-22 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 246.2
| Park Systems UK | 2023-03-22 | Max Wafer Zeroscan-Min Wafer Zeroscan | 0.03
| Park Systems UK | 2023-03-22 | Lateral(C-D) sensitivity | 15.25
| Park Systems UK | 2023-03-22 | Lateral(C-D) noise | 0.11
| Park Systems UK | 2023-03-22 | Full range of Z Height during Z Scan sweep | 15.9
| Park Systems UK | 2023-03-22 | Full range of Z Detector during Z Scan sweep | 13.2
| Park Systems UK | 2023-03-22 | Full range of Y2 Detector Raw during Y Scan sweep | 14.3
| Park Systems UK | 2023-03-22 | Full range of Y1 Detector Raw during Y Scan sweep | 14.2
| Park Systems UK | 2023-03-22 | Full range of Y Detector during Y Scan sweep | 178.2
| Park Systems UK | 2023-03-22 | Full range of X2 Detector Raw during X Scan sweep | 14.2
| Park Systems UK | 2023-03-22 | Full range of X1 Detector Raw during X Scan sweep | 14.1
| Park Systems UK | 2023-03-22 | Full range of X Detector during X Scan sweep | 174.7
| Park Systems UK | 2023-03-22 | A-B noise | 10.44
| Zhejiang University Department of Chemistry | 2023-03-20 | Z Scanner position at -3μm | 0.007797
| Zhejiang University Department of Chemistry | 2023-03-20 | Z Scanner position at +3μm | 0.010185
| Zhejiang University Department of Chemistry | 2023-03-20 | Z Linearity by 3 different sample height | 0.99426
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Z Linearity by 3 different sample height | 1.0056
| Zhejiang University Department of Chemistry | 2023-03-20 | Z Height noise | 0.032
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Z Height noise | 0.02
| Zhejiang University Department of Chemistry | 2023-03-20 | Y Detector noise (XY Servo on) | 0.301
| Zhejiang University Department of Chemistry | 2023-03-20 | Y Detector noise (XY Servo off) | 0.265
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Y Detector noise (XY Servo off) | 0.261
| Zhejiang University Department of Chemistry | 2023-03-20 | Y axis OPM | 1.08
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Y axis OPM | 1.245
| Zhejiang University Department of Chemistry | 2023-03-20 | X Detector noise (XY Servo on) | 0.295
| Zhejiang University Department of Chemistry | 2023-03-20 | X Detector noise (XY Servo off) | 0.258
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | X Detector noise (XY Servo off) | 0.252
| Zhejiang University Department of Chemistry | 2023-03-20 | X axis OPM | 1.419
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | X axis OPM | 1.005
| Zhejiang University Department of Chemistry | 2023-03-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.22
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.2
| Zhejiang University Department of Chemistry | 2023-03-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.53
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.15
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Zhejiang University Department of Chemistry | 2023-03-20 | NCM driving% with AC160TS cantilever | 1.06
| Zhejiang University Department of Chemistry | 2023-03-20 | NCM Amplitude noise | 18
| Zhejiang University Department of Chemistry | 2023-03-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.899
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.011
| Zhejiang University Department of Chemistry | 2023-03-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.4
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | MicrometerPerAdc of Z Height (Z Height calibration) | 23.192
| Zhejiang University Department of Chemistry | 2023-03-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 152.5
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 142.9
| Zhejiang University Department of Chemistry | 2023-03-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 147.8
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 145.7
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Full range of Z Height during Z Scan sweep | 15.9
| Zhejiang University Department of Chemistry | 2023-03-20 | Lateral(C-D) sensitivity | 5.3
| Zhejiang University Department of Chemistry | 2023-03-20 | Lateral(C-D) noise | 0.075
| Zhejiang University Department of Chemistry | 2023-03-20 | Full range of Z Height during Z Scan sweep | 15.8
| Zhejiang University Department of Chemistry | 2023-03-20 | Full range of Z Detector during Z Scan sweep | 11.5
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Full range of Z Detector during Z Scan sweep | 13.7
| Zhejiang University Department of Chemistry | 2023-03-20 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Full range of Y Detector Raw during Y Scan sweep | 15.5
| Zhejiang University Department of Chemistry | 2023-03-20 | Full range of Y Detector during Y Scan sweep | 105.1
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Full range of Y Detector during Y Scan sweep | 110
| Zhejiang University Department of Chemistry | 2023-03-20 | Full range of X Detector Raw during X Scan sweep | 14.8
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Full range of X Detector Raw during X Scan sweep | 15
| Zhejiang University Department of Chemistry | 2023-03-20 | Full range of X Detector during X Scan sweep | 107.2
| National Laboratory of Metrology and Testing (LNE) | 2023-03-20 | Full range of X Detector during X Scan sweep | 116.5
| Zhejiang University Department of Chemistry | 2023-03-20 | A-B noise | 15
| TSMC F12 EBO | 2023-03-16 | Z Height noise | 0.033
| TSMC F12 EBO | 2023-03-16 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.17
| TSMC F12 EBO | 2023-03-16 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3612
| TSMC F12 EBO | 2023-03-16 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 63.366
| TSMC F12 EBO | 2023-03-16 | Y Detector noise (XY Servo on) | 0.1717
| TSMC F12 EBO | 2023-03-16 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 88.2698
| TSMC F12 EBO | 2023-03-16 | Y axis, Repeatability @1st position | 0.03
| TSMC F12 EBO | 2023-03-16 | Y axis, Repeatability @ 4th position | 0.03
| TSMC F12 EBO | 2023-03-16 | Y axis, Repeatability @ 3rd position | 0.03
| TSMC F12 EBO | 2023-03-16 | Y axis, Repeatability @ 2nd position | 0.03
| TSMC F12 EBO | 2023-03-16 | Y axis, Accuracy @1st position | 0.19
| TSMC F12 EBO | 2023-03-16 | Y axis, Accuracy @ 4th position | 0.26
| TSMC F12 EBO | 2023-03-16 | Y axis, Accuracy @ 3rd position | 0.28
| TSMC F12 EBO | 2023-03-16 | Y axis, Accuracy @ 2nd position | 0.18
| TSMC F12 EBO | 2023-03-16 | Y axis OPM | 2.258
| TSMC F12 EBO | 2023-03-16 | XY Orthogonality | 89.7
| TSMC F12 EBO | 2023-03-16 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 64.2419
| TSMC F12 EBO | 2023-03-16 | X Detector noise (XY Servo on) | 0.1639
| TSMC F12 EBO | 2023-03-16 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 92.6855
| TSMC F12 EBO | 2023-03-16 | X axis, Repeatability @1st position | 0.02
| TSMC F12 EBO | 2023-03-16 | X axis, Repeatability @ 4th position | 0.07
| TSMC F12 EBO | 2023-03-16 | X axis, Repeatability @ 3rd position | 0.03
| TSMC F12 EBO | 2023-03-16 | X axis, Repeatability @ 2nd position | 0.04
| TSMC F12 EBO | 2023-03-16 | X axis, Accuracy @1st position | 0.1
| TSMC F12 EBO | 2023-03-16 | X axis, Accuracy @ 4th position | 0.25
| TSMC F12 EBO | 2023-03-16 | X axis, Accuracy @ 3rd position | 0.17
| TSMC F12 EBO | 2023-03-16 | X axis, Accuracy @ 2nd position | 0.2
| TSMC F12 EBO | 2023-03-16 | X axis OPM | 2.216
| TSMC F12 EBO | 2023-03-16 | Laser intensity with AC160 cantilever | 2.8
| TSMC F12 EBO | 2023-03-16 | Full range of Z Height during Z Scan sweep | 18.5309
| TSMC F12 EBO | 2023-03-16 | Full range of Z Height during Z Scan sweep | 62.1119
| TSMC F12 EBO | 2023-03-16 | Full range of Z Detector during Z Scan sweep | 14.74
| TSMC F12 EBO | 2023-03-16 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| TSMC F12 EBO | 2023-03-16 | Full range of Y Detector during Y Scan sweep | 62.1478
| TSMC F12 EBO | 2023-03-16 | Full range of X Detector Raw during X Scan sweep | 13.7
| TSMC F12 EBO | 2023-03-16 | A-B after Auto PSPD Alignment | 0.2
| INOMAR_NX10 | 2023-03-16 | Z Scanner position at -3μm | 0.0013
| INOMAR_NX10 | 2023-03-16 | Z Scanner position at +3μm | 0.0189
| INOMAR_NX10 | 2023-03-16 | Z Linearity by 3 different sample height | 1.012
| INOMAR_NX10 | 2023-03-16 | Z Height noise | 0.044
| INOMAR_NX10 | 2023-03-16 | Y Detector noise (XY Servo on) | 0.269
| INOMAR_NX10 | 2023-03-16 | Y Detector noise (XY Servo off) | 0.248
| INOMAR_NX10 | 2023-03-16 | Y axis OPM | 0.332
| INOMAR_NX10 | 2023-03-16 | X Detector noise (XY Servo on) | 0.277
| INOMAR_NX10 | 2023-03-16 | X Detector noise (XY Servo off) | 0.254
| INOMAR_NX10 | 2023-03-16 | X axis OPM | 0.302
| INOMAR_NX10 | 2023-03-16 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.29
| INOMAR_NX10 | 2023-03-16 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.39
| INOMAR_NX10 | 2023-03-16 | NCM driving% with AC160TS cantilever | 0.77
| INOMAR_NX10 | 2023-03-16 | NCM Amplitude noise | 14.749
| INOMAR_NX10 | 2023-03-16 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.367
| INOMAR_NX10 | 2023-03-16 | MicrometerPerAdc of Z Height (Z Height calibration) | 27.502
| INOMAR_NX10 | 2023-03-16 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 105
| INOMAR_NX10 | 2023-03-16 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 101.1
| INOMAR_NX10 | 2023-03-16 | Lateral(C-D) sensitivity | 8.1
| INOMAR_NX10 | 2023-03-16 | Lateral(C-D) noise | 0.078
| INOMAR_NX10 | 2023-03-16 | Full range of Z Height during Z Scan sweep | 17
| INOMAR_NX10 | 2023-03-16 | Full range of Z Detector during Z Scan sweep | 11.7
| INOMAR_NX10 | 2023-03-16 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| INOMAR_NX10 | 2023-03-16 | Full range of Y Detector during Y Scan sweep | 74.1
| INOMAR_NX10 | 2023-03-16 | Full range of X Detector during X Scan sweep | 74
| INOMAR_NX10 | 2023-03-16 | A-B noise | 9.33
| Tokyo institute of technology | 2023-03-13 | Z Scanner position at -3μm | 0.0022
| Tokyo institute of technology | 2023-03-13 | Z Scanner position at +3μm | -0.0162
| Tokyo institute of technology | 2023-03-13 | Z Linearity by 3 different sample height | 1.0018
| Tokyo institute of technology | 2023-03-13 | Z Height noise | 0.036
| Tokyo institute of technology | 2023-03-13 | Y Scan Stroke | 55
| Tokyo institute of technology | 2023-03-13 | Y Detector noise (XY Servo on) | 0.37
| Tokyo institute of technology | 2023-03-13 | Y Detector noise (XY Servo off) | 0.352
| Tokyo institute of technology | 2023-03-13 | Y axis OPM | 0.165
| Tokyo institute of technology | 2023-03-13 | X Scan Stroke | 55
| Tokyo institute of technology | 2023-03-13 | X Detector noise (XY Servo on) | 0.361
| Tokyo institute of technology | 2023-03-13 | X Detector noise (XY Servo off) | 0.35
| Tokyo institute of technology | 2023-03-13 | X axis OPM | 0.201
| Tokyo institute of technology | 2023-03-13 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.35
| Tokyo institute of technology | 2023-03-13 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.57
| Tokyo institute of technology | 2023-03-13 | NCM driving% with AC160TS cantilever | 1.083
| Tokyo institute of technology | 2023-03-13 | NCM Amplitude noise | 16.317
| Tokyo institute of technology | 2023-03-13 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.319
| Tokyo institute of technology | 2023-03-13 | MicrometerPerAdc of Z Height (Z Height calibration) | 28.618
| Tokyo institute of technology | 2023-03-13 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 100.1
| Tokyo institute of technology | 2023-03-13 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 98.8
| Tokyo institute of technology | 2023-03-13 | Full range of Z Height during Z Scan sweep | 16.4
| Tokyo institute of technology | 2023-03-13 | Full range of Z Detector during Z Scan sweep | 11.5
| Tokyo institute of technology | 2023-03-13 | Full range of Y Detector Raw during Y Scan sweep | 13.9
| Tokyo institute of technology | 2023-03-13 | Full range of Y Detector during Y Scan sweep | 69.2
| Tokyo institute of technology | 2023-03-13 | Full range of X Detector Raw during X Scan sweep | 14.1
| Tokyo institute of technology | 2023-03-13 | Full range of X Detector during X Scan sweep | 69.5
| Tokyo institute of technology | 2023-03-13 | A-B noise | 16.287
| University of Warwick | 2023-03-10 | Z Scanner position at -3μm | 0.003
| University of Warwick | 2023-03-10 | Z Scanner position at +3μm | 0.0091
| University of Warwick | 2023-03-10 | Z Linearity by 3 different sample height | 0.999
| University of Warwick | 2023-03-10 | Z Height noise | 0.036
| University of Warwick | 2023-03-10 | Y Detector noise (XY Servo on) | 0.242
| University of Warwick | 2023-03-10 | Y Detector noise (XY Servo off) | 0.228
| University of Warwick | 2023-03-10 | Y axis OPM | 0.511
| University of Warwick | 2023-03-10 | X Detector noise (XY Servo on) | 0.247
| University of Warwick | 2023-03-10 | X Detector noise (XY Servo off) | 0.227
| University of Warwick | 2023-03-10 | X axis OPM | 1.421
| University of Warwick | 2023-03-10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.26
| University of Warwick | 2023-03-10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.37
| University of Warwick | 2023-03-10 | NCM driving% with AC160TS cantilever | 3.354
| University of Warwick | 2023-03-10 | NCM Amplitude noise | 14.489
| University of Warwick | 2023-03-10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.299
| University of Warwick | 2023-03-10 | MicrometerPerAdc of Z Height (Z Height calibration) | 29.384
| University of Warwick | 2023-03-10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 108.4
| University of Warwick | 2023-03-10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 108.4
| University of Warwick | 2023-03-10 | Lateral(C-D) sensitivity | 5.2
| University of Warwick | 2023-03-10 | Lateral(C-D) noise | 0.081
| University of Warwick | 2023-03-10 | Full range of Z Height during Z Scan sweep | 16.2
| University of Warwick | 2023-03-10 | Full range of Z Detector during Z Scan sweep | 11
| University of Warwick | 2023-03-10 | Full range of Y Detector during Y Scan sweep | 75.8
| University of Warwick | 2023-03-10 | Full range of X Detector Raw during X Scan sweep | 14.1
| University of Warwick | 2023-03-10 | Full range of X Detector during X Scan sweep | 76.422
| University of Warwick | 2023-03-10 | A-B noise | 11.455
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Z Scanner position at -3μm | 0.0007
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Z Scanner position at +3μm | 0.0119
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Z Linearity by 3 different sample height | 0.993
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Z Height noise | 0.02
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Y Detector noise (XY Servo on) | 0.381
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Y Detector noise (XY Servo off) | 0.315
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Y axis OPM | 1.083
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | X Detector noise (XY Servo on) | 0.378
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | X Detector noise (XY Servo off) | 0.316
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | X axis OPM | 0.69
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.23
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.41
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | NCM driving% with AC160TS cantilever | 1.048
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | NCM Amplitude noise | 13.771
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.637
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | MicrometerPerAdc of Z Height (Z Height calibration) | 26.39
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 166.9
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 165.3
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Lateral(C-D) sensitivity | 2.1
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Lateral(C-D) noise | 0.093
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Full range of Z Height during Z Scan sweep | 16.9
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Full range of Z Detector during Z Scan sweep | 12.8
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Full range of Y Detector during Y Scan sweep | 119.6
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Full range of X Detector Raw during X Scan sweep | 14.5
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | Full range of X Detector during X Scan sweep | 119.9
| Jiangsu Xianke Semiconductor New Materials | 2023-03-09 | A-B noise | 13.544
| Nanoplexus Limited | 2023-02-27 | Z Linearity by 3 different sample height | 1.0155
| Northwest University | 2023-02-27 | Z Linearity by 3 different sample height | 0.985
| Nanoplexus Limited | 2023-02-27 | Z Height noise | 0.021
| Northwest University | 2023-02-27 | Z Height noise | 0.0827
| Northwest University | 2023-02-27 | Y Detector noise (XY Servo on) | 0.37
| Nanoplexus Limited | 2023-02-27 | Y Detector noise (XY Servo off) | 0.263
| Northwest University | 2023-02-27 | Y Detector noise (XY Servo off) | 0.35
| Nanoplexus Limited | 2023-02-27 | Y axis OPM | 0.745
| Northwest University | 2023-02-27 | Y axis OPM | 0.25
| Northwest University | 2023-02-27 | X Detector noise (XY Servo on) | 0.36
| Nanoplexus Limited | 2023-02-27 | X Detector noise (XY Servo off) | 0.254
| Northwest University | 2023-02-27 | X Detector noise (XY Servo off) | 0.35
| Nanoplexus Limited | 2023-02-27 | X axis OPM | 1.107
| Northwest University | 2023-02-27 | X axis OPM | 0.76
| Nanoplexus Limited | 2023-02-27 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.21
| Nanoplexus Limited | 2023-02-27 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.29
| Nanoplexus Limited | 2023-02-27 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.17
| Nanoplexus Limited | 2023-02-27 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| Northwest University | 2023-02-27 | NCM driving% with AC160TS cantilever | 3.31
| Nanoplexus Limited | 2023-02-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.095
| Northwest University | 2023-02-27 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 24.109
| Nanoplexus Limited | 2023-02-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.641
| Northwest University | 2023-02-27 | MicrometerPerAdc of Z Height (Z Height calibration) | 58.471
| Nanoplexus Limited | 2023-02-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 163.7
| Northwest University | 2023-02-27 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 161.2
| Nanoplexus Limited | 2023-02-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 154.2
| Northwest University | 2023-02-27 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 162.3
| Nanoplexus Limited | 2023-02-27 | Full range of Z Height during Z Scan sweep | 16.3
| Northwest University | 2023-02-27 | Full range of Z Height during Z Scan sweep | 40.5
| Nanoplexus Limited | 2023-02-27 | Full range of Z Detector during Z Scan sweep | 14.3
| Northwest University | 2023-02-27 | Full range of Z Detector during Z Scan sweep | 13.6
| Nanoplexus Limited | 2023-02-27 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Northwest University | 2023-02-27 | Full range of Y Detector Raw during Y Scan sweep | 14
| Nanoplexus Limited | 2023-02-27 | Full range of Y Detector during Y Scan sweep | 117.7
| Northwest University | 2023-02-27 | Full range of Y Detector during Y Scan sweep | 113
| Nanoplexus Limited | 2023-02-27 | Full range of X Detector Raw during X Scan sweep | 14.9
| Northwest University | 2023-02-27 | Full range of X Detector Raw during X Scan sweep | 14.1
| Nanoplexus Limited | 2023-02-27 | Full range of X Detector during X Scan sweep | 116.5
| Northwest University | 2023-02-27 | Full range of X Detector during X Scan sweep | 114.3
| Northwest University | 2023-02-27 | Approach position at -5μm | 0.0202
| Northwest University | 2023-02-27 | Approach position at +5μm | 0.0105
| AIST | 2023-02-20 | Z Linearity by 5 different sample height | 1.002
| AIST | 2023-02-20 | Z Height noise | 0.028
| AIST | 2023-02-20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.5231
| AIST | 2023-02-20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.4531
| AIST | 2023-02-20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.6636
| AIST | 2023-02-20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.6729
| AIST | 2023-02-20 | Y2 Detector noise (XY Servo on) | 0.639
| AIST | 2023-02-20 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 284.18
| AIST | 2023-02-20 | Y1 Detector noise (XY Servo on) | 0.62
| AIST | 2023-02-20 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 273.7
| AIST | 2023-02-20 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 196.4
| AIST | 2023-02-20 | Y Detector noise (XY Servo on) | 0.56
| AIST | 2023-02-20 | Y axis, Repeatability @1st position | 0.06
| AIST | 2023-02-20 | Y axis, Repeatability @ 4th position | 0.05
| AIST | 2023-02-20 | Y axis, Repeatability @ 3rd position | 0.06
| AIST | 2023-02-20 | Y axis, Repeatability @ 2nd position | 0.05
| AIST | 2023-02-20 | Y axis, Accuracy @1st position | 0.51
| AIST | 2023-02-20 | Y axis, Accuracy @ 4th position | 0.6
| AIST | 2023-02-20 | Y axis, Accuracy @ 3rd position | 0.62
| AIST | 2023-02-20 | Y axis, Accuracy @ 2nd position | 0.52
| AIST | 2023-02-20 | Y axis OPM | 0.543
| AIST | 2023-02-20 | XY Orthogonality | 90.1
| AIST | 2023-02-20 | X2 Detector noise (XY Servo on) | 0.358
| AIST | 2023-02-20 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 288
| AIST | 2023-02-20 | X1 Detector noise (XY Servo on) | 0.348
| AIST | 2023-02-20 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 276.25
| AIST | 2023-02-20 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.72
| AIST | 2023-02-20 | X Detector noise (XY Servo on) | 0.294
| AIST | 2023-02-20 | X axis, Repeatability @1st position | 0.04
| AIST | 2023-02-20 | X axis, Repeatability @ 4th position | 0.03
| AIST | 2023-02-20 | X axis, Repeatability @ 3rd position | 0.04
| AIST | 2023-02-20 | X axis, Repeatability @ 2nd position | 0.04
| AIST | 2023-02-20 | X axis, Accuracy @1st position | 0.21
| AIST | 2023-02-20 | X axis, Accuracy @ 4th position | 0.15
| AIST | 2023-02-20 | X axis, Accuracy @ 3rd position | 0.22
| AIST | 2023-02-20 | X axis, Accuracy @ 2nd position | 0.24
| AIST | 2023-02-20 | X axis OPM | 0.148
| AIST | 2023-02-20 | Under 50nm Step Height Sample (10 Repeats) | 91.857
| AIST | 2023-02-20 | OPM Repeatability @5mm | 0.72925
| AIST | 2023-02-20 | OPM Repeatability @25mm | 2.9364
| AIST | 2023-02-20 | OPM Repeatability @1mm | 0.68418
| AIST | 2023-02-20 | OPM Repeatability @10mm | 1.1957
| AIST | 2023-02-20 | NCM driving% with AC160TS cantilever | 17.1
| AIST | 2023-02-20 | Max OPM @5mm | 22.9871
| AIST | 2023-02-20 | Max OPM @25mm | 147.0129
| AIST | 2023-02-20 | Max OPM @1mm | 7.6705
| AIST | 2023-02-20 | Max OPM @10mm | 24.9336
| AIST | 2023-02-20 | Laser intensity with AC160 cantilever | 2.86
| AIST | 2023-02-20 | Image offset after moving from -Y to +Y (Y scan) | 55.6641
| AIST | 2023-02-20 | Image offset after moving from -Y to +Y (X scan) | 49.8047
| AIST | 2023-02-20 | Image offset after moving from -X to +X (Y scan) | 27.0996
| AIST | 2023-02-20 | Image offset after moving from -X to +X (X scan) | 80.5664
| AIST | 2023-02-20 | Image offset after moving from +Y to -Y (Y scan) | 73.9746
| AIST | 2023-02-20 | Image offset after moving from +Y to -Y (X scan) | 24.1699
| AIST | 2023-02-20 | Image offset after moving from +X to -X (Y scan) | 15.3809
| AIST | 2023-02-20 | Image offset after moving from +X to -X (X scan) | 35.1563
| AIST | 2023-02-20 | Full range of Z Height during Z Scan sweep | 16.8052
| AIST | 2023-02-20 | Full range of Z Detector during Z Scan sweep | 13.72
| AIST | 2023-02-20 | Full range of Y2 Detector Raw during Y Scan sweep | 13.72
| AIST | 2023-02-20 | Full range of Y2 Detector during Y Scan sweep | 196.0889
| AIST | 2023-02-20 | Full range of Y1 Detector Raw during Y Scan sweep | 14.13
| AIST | 2023-02-20 | Full range of Y1 Detector during Y Scan sweep | 196.0324
| AIST | 2023-02-20 | Full range of X2 Detector Raw during X Scan sweep | 13.71
| AIST | 2023-02-20 | Full range of X2 Detector during X Scan sweep | 199.4985
| AIST | 2023-02-20 | Full range of X1 Detector Raw during X Scan sweep | 14.17
| AIST | 2023-02-20 | Full range of X1 Detector during X Scan sweep | 199.1174
| AIST | 2023-02-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| AIST | 2023-02-20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.45
| AIST | 2023-02-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.41
| AIST | 2023-02-20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| AIST | 2023-02-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.03
| AIST | 2023-02-20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.14
| AIST | 2023-02-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.13
| AIST | 2023-02-20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.02
| AIST | 2023-02-20 | 50~200nm Step Height Sample (10 Repeats) | 179.565
| AIST | 2023-02-20 | 200~500nm Step Height Sample (10 Repeats) | 458.977
| AIST | 2023-02-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 16.4
| AIST | 2023-02-20 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 15.2
| AIST | 2023-02-20 | %GRR w/ roughness of non-patterened wafer | 22.3
| 울산대학교 | 2023-02-15 | Z Linearity by 3 different sample height | 1.002
| 울산대학교 | 2023-02-15 | Z Height noise | 0.02
| 울산대학교 | 2023-02-15 | Y Detector noise (XY Servo off) | 0.378
| 울산대학교 | 2023-02-15 | Y axis OPM | 0.899
| 울산대학교 | 2023-02-15 | X Detector noise (XY Servo off) | 0.384
| 울산대학교 | 2023-02-15 | X axis OPM | 1.521
| 울산대학교 | 2023-02-15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.51
| 울산대학교 | 2023-02-15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.4
| 울산대학교 | 2023-02-15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.5
| 울산대학교 | 2023-02-15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.087
| 울산대학교 | 2023-02-15 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.824
| 울산대학교 | 2023-02-15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 219.9
| 울산대학교 | 2023-02-15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 226.3
| 울산대학교 | 2023-02-15 | Full range of Z Height during Z Scan sweep | 16.2
| 울산대학교 | 2023-02-15 | Full range of Z Detector during Z Scan sweep | 14.2
| 울산대학교 | 2023-02-15 | Full range of Y Detector Raw during Y Scan sweep | 14.9
| 울산대학교 | 2023-02-15 | Full range of Y Detector during Y Scan sweep | 163.7
| 울산대학교 | 2023-02-15 | Full range of X Detector Raw during X Scan sweep | 14.2
| 울산대학교 | 2023-02-15 | Full range of X Detector during X Scan sweep | 161.1
| University of Reims | 2023-02-02 | Z Linearity by 3 different sample height | 1.0113
| University of Reims | 2023-02-02 | Z Height noise | 0.03
| University of Reims | 2023-02-02 | Y Detector noise (XY Servo off) | 0.284
| University of Reims | 2023-02-02 | Y axis OPM | 0.887
| University of Reims | 2023-02-02 | X Detector noise (XY Servo off) | 0.259
| University of Reims | 2023-02-02 | X axis OPM | 1.485
| University of Reims | 2023-02-02 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.91
| University of Reims | 2023-02-02 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.955
| University of Reims | 2023-02-02 | MicrometerPerAdc of Z Height (Z Height calibration) | 24.912
| University of Reims | 2023-02-02 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 221.9
| University of Reims | 2023-02-02 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 222.6
| University of Reims | 2023-02-02 | Full range of Z Height during Z Scan sweep | 15.9
| University of Reims | 2023-02-02 | Full range of Z Detector during Z Scan sweep | 12.7
| University of Reims | 2023-02-02 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| University of Reims | 2023-02-02 | Full range of Y Detector during Y Scan sweep | 157.9
| University of Reims | 2023-02-02 | Full range of X Detector Raw during X Scan sweep | 14.1
| University of Reims | 2023-02-02 | Full range of X Detector during X Scan sweep | 156.5
| 서울시립대 공동실습실험관 | 2023-02-01 | Z Linearity by 3 different sample height | 1.0002
| 서울시립대 공동실습실험관 | 2023-02-01 | Z Height noise | 0.02
| 서울시립대 공동실습실험관 | 2023-02-01 | Y Detector noise (XY Servo off) | 0.237
| 서울시립대 공동실습실험관 | 2023-02-01 | Y axis OPM | 1.576
| 서울시립대 공동실습실험관 | 2023-02-01 | X Detector noise (XY Servo off) | 0.249
| 서울시립대 공동실습실험관 | 2023-02-01 | X axis OPM | 1.428
| 서울시립대 공동실습실험관 | 2023-02-01 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 서울시립대 공동실습실험관 | 2023-02-01 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.26
| 서울시립대 공동실습실험관 | 2023-02-01 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.13
| 서울시립대 공동실습실험관 | 2023-02-01 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.23
| 서울시립대 공동실습실험관 | 2023-02-01 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.295
| 서울시립대 공동실습실험관 | 2023-02-01 | MicrometerPerAdc of Z Height (Z Height calibration) | 22.982
| 서울시립대 공동실습실험관 | 2023-02-01 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 144
| 서울시립대 공동실습실험관 | 2023-02-01 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 148.1
| 서울시립대 공동실습실험관 | 2023-02-01 | Full range of Z Height during Z Scan sweep | 16.6
| 서울시립대 공동실습실험관 | 2023-02-01 | Full range of Z Detector during Z Scan sweep | 14.4
| 서울시립대 공동실습실험관 | 2023-02-01 | Full range of Y Detector Raw during Y Scan sweep | 15.6
| 서울시립대 공동실습실험관 | 2023-02-01 | Full range of Y Detector during Y Scan sweep | 112.1
| 서울시립대 공동실습실험관 | 2023-02-01 | Full range of X Detector Raw during X Scan sweep | 14.6
| 서울시립대 공동실습실험관 | 2023-02-01 | Full range of X Detector during X Scan sweep | 110.4
| Samsung Mask | 2022-12-27 | Z Linearity by Z scanner position | 0.01
| Samsung Mask | 2022-12-27 | Z Linearity by 5 different sample height | 1.003
| Samsung Mask | 2022-12-27 | Z Height noise | 0.032
| Samsung Mask | 2022-12-27 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.5354
| Samsung Mask | 2022-12-27 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.1766
| Samsung Mask | 2022-12-27 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 122.771
| Samsung Mask | 2022-12-27 | Y Detector noise (XY Servo on) | 0.287
| Samsung Mask | 2022-12-27 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 192.492
| Samsung Mask | 2022-12-27 | Y axis, Repeatability @1st position | 0.13
| Samsung Mask | 2022-12-27 | Y axis, Repeatability @ 4th position | 0.14
| Samsung Mask | 2022-12-27 | Y axis, Repeatability @ 3rd position | 0.13
| Samsung Mask | 2022-12-27 | Y axis, Repeatability @ 2nd position | 0.12
| Samsung Mask | 2022-12-27 | Y axis, Accuracy @1st position | 0.52
| Samsung Mask | 2022-12-27 | Y axis, Accuracy @ 4th position | 0.69
| Samsung Mask | 2022-12-27 | Y axis, Accuracy @ 3rd position | 0.59
| Samsung Mask | 2022-12-27 | Y axis, Accuracy @ 2nd position | 0.54
| Samsung Mask | 2022-12-27 | Y axis OPM | 1.77
| Samsung Mask | 2022-12-27 | XY Orthogonality | 90.1
| Samsung Mask | 2022-12-27 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 122.373
| Samsung Mask | 2022-12-27 | X Detector noise (XY Servo on) | 0.282
| Samsung Mask | 2022-12-27 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 193.037
| Samsung Mask | 2022-12-27 | X axis, Repeatability @1st position | 0.02
| Samsung Mask | 2022-12-27 | X axis, Repeatability @ 4th position | 0.03
| Samsung Mask | 2022-12-27 | X axis, Repeatability @ 3rd position | 0.03
| Samsung Mask | 2022-12-27 | X axis, Repeatability @ 2nd position | 0.03
| Samsung Mask | 2022-12-27 | X axis, Accuracy @1st position | 0.1
| Samsung Mask | 2022-12-27 | X axis, Accuracy @ 4th position | 0.14
| Samsung Mask | 2022-12-27 | X axis, Accuracy @ 3rd position | 0.13
| Samsung Mask | 2022-12-27 | X axis, Accuracy @ 2nd position | 0.18
| Samsung Mask | 2022-12-27 | X axis OPM | 1.67
| Samsung Mask | 2022-12-27 | Under 50nm Step Height Sample (10 Repeats) | 16.186
| Samsung Mask | 2022-12-27 | Image offset after moving from -Y to +Y (Y scan) | 76.17
| Samsung Mask | 2022-12-27 | Image offset after moving from -Y to +Y (X scan) | 46.88
| Samsung Mask | 2022-12-27 | Image offset after moving from -X to +X (Y scan) | 29.3
| Samsung Mask | 2022-12-27 | Image offset after moving from -X to +X (X scan) | 62.26
| Samsung Mask | 2022-12-27 | Image offset after moving from +Y to -Y (Y scan) | 79.1
| Samsung Mask | 2022-12-27 | Image offset after moving from +Y to -Y (X scan) | 25.63
| Samsung Mask | 2022-12-27 | Image offset after moving from +X to -X (Y scan) | 33.69
| Samsung Mask | 2022-12-27 | Image offset after moving from +X to -X (X scan) | 87.89
| Samsung Mask | 2022-12-27 | Full range of Z Height during Z Scan sweep | 17.6948
| Samsung Mask | 2022-12-27 | Full range of Z Height during Z Scan sweep | 65.0058
| Samsung Mask | 2022-12-27 | Full range of Z Detector during Z Scan sweep | 13.85
| Samsung Mask | 2022-12-27 | Full range of Y Detector Raw during Y Scan sweep | 13.93
| Samsung Mask | 2022-12-27 | Full range of Y Detector during Y Scan sweep | 61.6962
| Samsung Mask | 2022-12-27 | Full range of X Detector Raw during X Scan sweep | 14.38
| Samsung Mask | 2022-12-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.257
| Samsung Mask | 2022-12-27 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 1.423
| Samsung Mask | 2022-12-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 1.126
| Samsung Mask | 2022-12-27 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.204
| Samsung Mask | 2022-12-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.042
| Samsung Mask | 2022-12-27 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.234
| Samsung Mask | 2022-12-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.345
| Samsung Mask | 2022-12-27 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.062
| Samsung Mask | 2022-12-27 | A-B after Auto PSPD Alignment | 0.086
| Samsung Mask | 2022-12-27 | 940nm Step Height Sample (10 Repeats) | 945.54
| Samsung Mask | 2022-12-27 | 200~500nm Step Height Sample (10 Repeats) | 455.552
| 주성엔지니어링 | 2022-12-22 | Z Linearity by 3 different sample height | 1.0088
| 주성엔지니어링 | 2022-12-22 | Z Height noise | 0.019
| 주성엔지니어링 | 2022-12-22 | Y Detector noise (XY Servo off) | 0.262
| 주성엔지니어링 | 2022-12-22 | Y axis OPM | 1.395
| 주성엔지니어링 | 2022-12-22 | X Detector noise (XY Servo off) | 0.253
| 주성엔지니어링 | 2022-12-22 | X axis OPM | 1.526
| 주성엔지니어링 | 2022-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| 주성엔지니어링 | 2022-12-22 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.27
| 주성엔지니어링 | 2022-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| 주성엔지니어링 | 2022-12-22 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.24
| 주성엔지니어링 | 2022-12-22 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.296
| 주성엔지니어링 | 2022-12-22 | MicrometerPerAdc of Z Height (Z Height calibration) | 20.792
| 주성엔지니어링 | 2022-12-22 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 136.1
| 주성엔지니어링 | 2022-12-22 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 151.4
| 주성엔지니어링 | 2022-12-22 | Full range of Z Height during Z Scan sweep | 16.7
| 주성엔지니어링 | 2022-12-22 | Full range of Z Detector during Z Scan sweep | 16
| 주성엔지니어링 | 2022-12-22 | Full range of Y Detector Raw during Y Scan sweep | 15.1
| 주성엔지니어링 | 2022-12-22 | Full range of Y Detector during Y Scan sweep | 107.8
| 주성엔지니어링 | 2022-12-22 | Full range of X Detector Raw during X Scan sweep | 13.8
| 주성엔지니어링 | 2022-12-22 | Full range of X Detector during X Scan sweep | 104.6
| Fujibo | 2022-12-01 | Z Linearity by 5 different sample height | 0.999
| Fujibo | 2022-12-01 | Z Height noise | 0.037
| Fujibo | 2022-12-01 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 24.1652
| Fujibo | 2022-12-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.1391
| Fujibo | 2022-12-01 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.9404
| Fujibo | 2022-12-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.9828
| Fujibo | 2022-12-01 | Y2 Detector noise (XY Servo on) | 0.697
| Fujibo | 2022-12-01 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 270.55
| Fujibo | 2022-12-01 | Y1 Detector noise (XY Servo on) | 0.668
| Fujibo | 2022-12-01 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 261.65
| Fujibo | 2022-12-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 186.59
| Fujibo | 2022-12-01 | Y Detector noise (XY Servo on) | 0.625
| Fujibo | 2022-12-01 | Y axis, Repeatability @1st position | 0.04
| Fujibo | 2022-12-01 | Y axis, Repeatability @ 4th position | 0.04
| Fujibo | 2022-12-01 | Y axis, Repeatability @ 3rd position | 0.05
| Fujibo | 2022-12-01 | Y axis, Repeatability @ 2nd position | 0.06
| Fujibo | 2022-12-01 | Y axis, Accuracy @1st position | 0.58
| Fujibo | 2022-12-01 | Y axis, Accuracy @ 4th position | 0.82
| Fujibo | 2022-12-01 | Y axis, Accuracy @ 3rd position | 0.81
| Fujibo | 2022-12-01 | Y axis, Accuracy @ 2nd position | 0.8
| Fujibo | 2022-12-01 | Y axis OPM | 0.68
| Fujibo | 2022-12-01 | XY Orthogonality | 89.9
| Fujibo | 2022-12-01 | X2 Detector noise (XY Servo on) | 0.482
| Fujibo | 2022-12-01 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 282.69
| Fujibo | 2022-12-01 | X1 Detector noise (XY Servo on) | 0.484
| Fujibo | 2022-12-01 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 278.69
| Fujibo | 2022-12-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 191.75
| Fujibo | 2022-12-01 | X Detector noise (XY Servo on) | 0.396
| Fujibo | 2022-12-01 | X axis, Repeatability @1st position | 0.04
| Fujibo | 2022-12-01 | X axis, Repeatability @ 4th position | 0.04
| Fujibo | 2022-12-01 | X axis, Repeatability @ 3rd position | 0.05
| Fujibo | 2022-12-01 | X axis, Repeatability @ 2nd position | 0.03
| Fujibo | 2022-12-01 | X axis, Accuracy @1st position | 0.35
| Fujibo | 2022-12-01 | X axis, Accuracy @ 4th position | 0.21
| Fujibo | 2022-12-01 | X axis, Accuracy @ 3rd position | 0.29
| Fujibo | 2022-12-01 | X axis, Accuracy @ 2nd position | 0.49
| Fujibo | 2022-12-01 | X axis OPM | 0.614
| Fujibo | 2022-12-01 | Under 50nm Step Height Sample (10 Repeats) | 18.284
| Fujibo | 2022-12-01 | RMS of ZeroScan | 0.05
| Fujibo | 2022-12-01 | OPM Repeatability @5mm | 0.91002
| Fujibo | 2022-12-01 | OPM Repeatability @25mm | 1.947
| Fujibo | 2022-12-01 | OPM Repeatability @1mm | 0.95079
| Fujibo | 2022-12-01 | OPM Repeatability @10mm | 0.23744
| Fujibo | 2022-12-01 | NCM driving% with AC160TS cantilever | 6.25
| Fujibo | 2022-12-01 | Max OPM @5mm | 27.7967
| Fujibo | 2022-12-01 | Max OPM @25mm | 56.3819
| Fujibo | 2022-12-01 | Max OPM @1mm | 10.2062
| Fujibo | 2022-12-01 | Max OPM @10mm | 14.1723
| Fujibo | 2022-12-01 | Laser intensity with AC160 cantilever | 3.05
| Fujibo | 2022-12-01 | Image offset after moving from -Y to +Y (Y scan) | 29.2969
| Fujibo | 2022-12-01 | Image offset after moving from -Y to +Y (X scan) | 15.3809
| Fujibo | 2022-12-01 | Image offset after moving from -X to +X (Y scan) | 32.959
| Fujibo | 2022-12-01 | Image offset after moving from -X to +X (X scan) | 28.5645
| Fujibo | 2022-12-01 | Image offset after moving from +Y to -Y (Y scan) | 21.2402
| Fujibo | 2022-12-01 | Image offset after moving from +Y to -Y (X scan) | 24.1699
| Fujibo | 2022-12-01 | Image offset after moving from +X to -X (Y scan) | 16.8457
| Fujibo | 2022-12-01 | Image offset after moving from +X to -X (X scan) | 25.6348
| Fujibo | 2022-12-01 | Full range of Z Height during Z Scan sweep | 16.8284
| Fujibo | 2022-12-01 | Full range of Z Detector during Z Scan sweep | 13.93
| Fujibo | 2022-12-01 | Full range of Y2 Detector Raw during Y Scan sweep | 14.16
| Fujibo | 2022-12-01 | Full range of Y2 Detector during Y Scan sweep | 207.1981
| Fujibo | 2022-12-01 | Full range of Y1 Detector Raw during Y Scan sweep | 14.85
| Fujibo | 2022-12-01 | Full range of Y1 Detector during Y Scan sweep | 207.8147
| Fujibo | 2022-12-01 | Full range of X2 Detector Raw during X Scan sweep | 14.33
| Fujibo | 2022-12-01 | Full range of X2 Detector during X Scan sweep | 207.6504
| Fujibo | 2022-12-01 | Full range of X1 Detector Raw during X Scan sweep | 14.23
| Fujibo | 2022-12-01 | Full range of X1 Detector during X Scan sweep | 206.4092
| Fujibo | 2022-12-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.18
| Fujibo | 2022-12-01 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.98
| Fujibo | 2022-12-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.8
| Fujibo | 2022-12-01 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.15
| Fujibo | 2022-12-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.04
| Fujibo | 2022-12-01 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.23
| Fujibo | 2022-12-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.28
| Fujibo | 2022-12-01 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.05
| Fujibo | 2022-12-01 | 50~200nm Step Height Sample (10 Repeats) | 184.099
| Fujibo | 2022-12-01 | 200~500nm Step Height Sample (10 Repeats) | 475.123
| Fujibo | 2022-12-01 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 12
| Fujibo | 2022-12-01 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 13.9
| Fujibo | 2022-12-01 | %GRR w/ roughness of non-patterened wafer | 18.4
| Canon(Hiratsuka #1) | 2022-11-29 | Z Linearity by 5 different sample height | 1
| Canon(Hiratsuka #2) | 2022-11-29 | Z Linearity by 5 different sample height | 1.002
| Canon(Hiratsuka #1) | 2022-11-29 | Z Height noise | 0.034
| Canon(Hiratsuka #2) | 2022-11-29 | Z Height noise | 0.033
| Canon(Hiratsuka #1) | 2022-11-29 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.5601
| Canon(Hiratsuka #1) | 2022-11-29 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.8877
| Canon(Hiratsuka #1) | 2022-11-29 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 11.1651
| Canon(Hiratsuka #1) | 2022-11-29 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.0649
| Canon(Hiratsuka #1) | 2022-11-29 | Y2 Detector noise (XY Servo on) | 0.764
| Canon(Hiratsuka #2) | 2022-11-29 | Y2 Detector noise (XY Servo on) | 0.612
| Canon(Hiratsuka #1) | 2022-11-29 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 259.595
| Canon(Hiratsuka #1) | 2022-11-29 | Y1 Detector noise (XY Servo on) | 0.632
| Canon(Hiratsuka #2) | 2022-11-29 | Y1 Detector noise (XY Servo on) | 0.567
| Canon(Hiratsuka #1) | 2022-11-29 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 274.206
| Canon(Hiratsuka #1) | 2022-11-29 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 184.353
| Canon(Hiratsuka #1) | 2022-11-29 | Y Detector noise (XY Servo on) | 0.674
| Canon(Hiratsuka #2) | 2022-11-29 | Y Detector noise (XY Servo on) | 0.539
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Repeatability @1st position | 0.08
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Repeatability @1st position | 0.07
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Repeatability @ 4th position | 0.05
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Repeatability @ 4th position | 0.05
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Repeatability @ 3rd position | 0.07
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Repeatability @ 3rd position | 0.04
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Repeatability @ 2nd position | 0.06
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Repeatability @ 2nd position | 0.04
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Accuracy @1st position | 1.69
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Accuracy @1st position | 1.08
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Accuracy @ 4th position | 1.25
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Accuracy @ 4th position | 0.76
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Accuracy @ 3rd position | 1.44
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Accuracy @ 3rd position | 1.26
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis, Accuracy @ 2nd position | 1.09
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis, Accuracy @ 2nd position | 1
| Canon(Hiratsuka #1) | 2022-11-29 | Y axis OPM | 0.934
| Canon(Hiratsuka #2) | 2022-11-29 | Y axis OPM | 1.666
| Canon(Hiratsuka #1) | 2022-11-29 | XY Orthogonality | 89.9
| Canon(Hiratsuka #2) | 2022-11-29 | XY Orthogonality | 90.3
| Canon(Hiratsuka #1) | 2022-11-29 | X2 Detector noise (XY Servo on) | 0.333
| Canon(Hiratsuka #2) | 2022-11-29 | X2 Detector noise (XY Servo on) | 0.441
| Canon(Hiratsuka #1) | 2022-11-29 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 275.923
| Canon(Hiratsuka #1) | 2022-11-29 | X1 Detector noise (XY Servo on) | 0.337
| Canon(Hiratsuka #2) | 2022-11-29 | X1 Detector noise (XY Servo on) | 0.428
| Canon(Hiratsuka #1) | 2022-11-29 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 274.983
| Canon(Hiratsuka #1) | 2022-11-29 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 185.765
| Canon(Hiratsuka #1) | 2022-11-29 | X Detector noise (XY Servo on) | 0.278
| Canon(Hiratsuka #2) | 2022-11-29 | X Detector noise (XY Servo on) | 0.366
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Repeatability @1st position | 0.04
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Repeatability @1st position | 0.09
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Repeatability @ 4th position | 0.04
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Repeatability @ 4th position | 0.06
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Repeatability @ 3rd position | 0.04
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Repeatability @ 3rd position | 0.06
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Repeatability @ 2nd position | 0.05
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Repeatability @ 2nd position | 0.06
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Accuracy @1st position | 0.22
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Accuracy @1st position | 0.78
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Accuracy @ 4th position | 0.38
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Accuracy @ 4th position | 0.51
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Accuracy @ 3rd position | 0.54
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Accuracy @ 3rd position | 0.46
| Canon(Hiratsuka #1) | 2022-11-29 | X axis, Accuracy @ 2nd position | 0.29
| Canon(Hiratsuka #2) | 2022-11-29 | X axis, Accuracy @ 2nd position | 0.38
| Canon(Hiratsuka #1) | 2022-11-29 | X axis OPM | 3.563
| Canon(Hiratsuka #2) | 2022-11-29 | X axis OPM | 0.9
| Canon(Hiratsuka #1) | 2022-11-29 | Under 50nm Step Height Sample (10 Repeats) | 18.35
| Canon(Hiratsuka #2) | 2022-11-29 | Under 50nm Step Height Sample (10 Repeats) | 18.292
| Canon(Hiratsuka #1) | 2022-11-29 | RMS of ZeroScan | 0.046
| Canon(Hiratsuka #2) | 2022-11-29 | RMS of ZeroScan | 0.061
| Canon(Hiratsuka #1) | 2022-11-29 | OPM Repeatability @5mm | 1.4561
| Canon(Hiratsuka #2) | 2022-11-29 | OPM Repeatability @5mm | 1.2336
| Canon(Hiratsuka #1) | 2022-11-29 | OPM Repeatability @25mm | 6.1931
| Canon(Hiratsuka #2) | 2022-11-29 | OPM Repeatability @25mm | 7.2344
| Canon(Hiratsuka #1) | 2022-11-29 | OPM Repeatability @1mm | 0.96912
| Canon(Hiratsuka #2) | 2022-11-29 | OPM Repeatability @1mm | 0.79
| Canon(Hiratsuka #1) | 2022-11-29 | OPM Repeatability @10mm | 1.6501
| Canon(Hiratsuka #2) | 2022-11-29 | OPM Repeatability @10mm | 1.695
| Canon(Hiratsuka #1) | 2022-11-29 | NCM driving% with AC160TS cantilever | 1.16
| Canon(Hiratsuka #2) | 2022-11-29 | NCM driving% with AC160TS cantilever | 1.14
| Canon(Hiratsuka #1) | 2022-11-29 | Max OPM @5mm | 24.2427
| Canon(Hiratsuka #2) | 2022-11-29 | Max OPM @5mm | 30.0886
| Canon(Hiratsuka #1) | 2022-11-29 | Max OPM @25mm | 115.0965
| Canon(Hiratsuka #2) | 2022-11-29 | Max OPM @25mm | 109.7431
| Canon(Hiratsuka #1) | 2022-11-29 | Max OPM @1mm | 11.2642
| Canon(Hiratsuka #2) | 2022-11-29 | Max OPM @1mm | 7.826
| Canon(Hiratsuka #1) | 2022-11-29 | Max OPM @10mm | 42.1987
| Canon(Hiratsuka #2) | 2022-11-29 | Max OPM @10mm | 40.595
| Canon(Hiratsuka #1) | 2022-11-29 | Laser intensity with AC160 cantilever | 3.1
| Canon(Hiratsuka #2) | 2022-11-29 | Laser intensity with AC160 cantilever | 3.1
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from -Y to +Y (Y scan) | 24.9023
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from -Y to +Y (Y scan) | 10.9863
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from -Y to +Y (X scan) | 14.6484
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from -Y to +Y (X scan) | 27.0996
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from -X to +X (Y scan) | 100.342
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from -X to +X (Y scan) | 20.5078
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from -X to +X (X scan) | 52.002
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from -X to +X (X scan) | 96.6797
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from +Y to -Y (Y scan) | 82.0313
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from +Y to -Y (Y scan) | 56.3965
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from +Y to -Y (X scan) | 3.6621
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from +Y to -Y (X scan) | 9.5215
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from +X to -X (Y scan) | 58.5938
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from +X to -X (Y scan) | 52.002
| Canon(Hiratsuka #1) | 2022-11-29 | Image offset after moving from +X to -X (X scan) | 60.791
| Canon(Hiratsuka #2) | 2022-11-29 | Image offset after moving from +X to -X (X scan) | 161.1328
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of Z Height during Z Scan sweep | 16.658
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of Z Height during Z Scan sweep | 16.8984
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of Z Detector during Z Scan sweep | 12.81
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of Z Detector during Z Scan sweep | 13.01
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of Y2 Detector Raw during Y Scan sweep | 14.67
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of Y2 Detector Raw during Y Scan sweep | 13.85
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of Y2 Detector during Y Scan sweep | 207.7432
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of Y2 Detector during Y Scan sweep | 207.3767
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.53
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of Y1 Detector Raw during Y Scan sweep | 14.04
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of Y1 Detector during Y Scan sweep | 199.0894
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of Y1 Detector during Y Scan sweep | 198.3016
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of X2 Detector Raw during X Scan sweep | 14.42
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of X2 Detector Raw during X Scan sweep | 14.39
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of X2 Detector during X Scan sweep | 190.8682
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of X2 Detector during X Scan sweep | 198.551
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of X1 Detector Raw during X Scan sweep | 14.25
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of X1 Detector Raw during X Scan sweep | 14.3
| Canon(Hiratsuka #1) | 2022-11-29 | Full range of X1 Detector during X Scan sweep | 204.9307
| Canon(Hiratsuka #2) | 2022-11-29 | Full range of X1 Detector during X Scan sweep | 203.061
| Canon(Hiratsuka #1) | 2022-11-29 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.11
| Canon(Hiratsuka #1) | 2022-11-29 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.63
| Canon(Hiratsuka #1) | 2022-11-29 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.58
| Canon(Hiratsuka #1) | 2022-11-29 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.08
| Canon(Hiratsuka #1) | 2022-11-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.12
| Canon(Hiratsuka #1) | 2022-11-29 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.66
| Canon(Hiratsuka #1) | 2022-11-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.23
| Canon(Hiratsuka #1) | 2022-11-29 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.03
| Canon(Hiratsuka #1) | 2022-11-29 | 50~200nm Step Height Sample (10 Repeats) | 184.18
| Canon(Hiratsuka #2) | 2022-11-29 | 50~200nm Step Height Sample (10 Repeats) | 184.122
| Canon(Hiratsuka #1) | 2022-11-29 | 200~500nm Step Height Sample (10 Repeats) | 460.377
| Canon(Hiratsuka #2) | 2022-11-29 | 200~500nm Step Height Sample (10 Repeats) | 462.716
| Canon(Hiratsuka #2) | 2022-11-29 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility | 5.1
| Canon(Hiratsuka #2) | 2022-11-29 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability | 24.2
| Canon(Hiratsuka #2) | 2022-11-29 | %GRR w/ roughness of non-patterened wafer | 24.7
| 경상국립대학교(화학과) | 2022-11-28 | Z Linearity by 3 different sample height | 0.9957
| 경상국립대학교(화학과) | 2022-11-28 | Z Height noise | 0.033
| 경상국립대학교(화학과) | 2022-11-28 | Y Detector noise (XY Servo off) | 0.418
| 경상국립대학교(화학과) | 2022-11-28 | Y axis OPM | 0.76
| 경상국립대학교(화학과) | 2022-11-28 | X Detector noise (XY Servo off) | 0.408
| 경상국립대학교(화학과) | 2022-11-28 | X axis OPM | 0.433
| 경상국립대학교(화학과) | 2022-11-28 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.44
| 경상국립대학교(화학과) | 2022-11-28 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.58
| 경상국립대학교(화학과) | 2022-11-28 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.41
| 경상국립대학교(화학과) | 2022-11-28 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.56
| 경상국립대학교(화학과) | 2022-11-28 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 10.398
| 경상국립대학교(화학과) | 2022-11-28 | MicrometerPerAdc of Z Height (Z Height calibration) | 25.653
| 경상국립대학교(화학과) | 2022-11-28 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 249
| 경상국립대학교(화학과) | 2022-11-28 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 245.7
| 경상국립대학교(화학과) | 2022-11-28 | Full range of Z Height during Z Scan sweep | 16.7
| 경상국립대학교(화학과) | 2022-11-28 | Full range of Z Detector during Z Scan sweep | 13
| 경상국립대학교(화학과) | 2022-11-28 | Full range of Y Detector Raw during Y Scan sweep | 13.5
| 경상국립대학교(화학과) | 2022-11-28 | Full range of Y Detector during Y Scan sweep | 167.6
| 경상국립대학교(화학과) | 2022-11-28 | Full range of X Detector Raw during X Scan sweep | 14
| 경상국립대학교(화학과) | 2022-11-28 | Full range of X Detector during X Scan sweep | 173.3
| Hoya | 2022-11-18 | Z Linearity by 5 different sample height | 1.0259
| Hoya | 2022-11-18 | Z Height noise | 0.032
| Hoya | 2022-11-18 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 23.6689
| Hoya | 2022-11-18 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 9.664
| Hoya | 2022-11-18 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 66.0169
| Hoya | 2022-11-18 | Y Detector noise (XY Servo on) | 0.216
| Hoya | 2022-11-18 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 94.2675
| Hoya | 2022-11-18 | Y axis, Repeatability @1st position | 0.05
| Hoya | 2022-11-18 | Y axis, Repeatability @ 4th position | 0.06
| Hoya | 2022-11-18 | Y axis, Repeatability @ 3rd position | 0.05
| Hoya | 2022-11-18 | Y axis, Repeatability @ 2nd position | 0.05
| Hoya | 2022-11-18 | Y axis, Accuracy @1st position | 0.32
| Hoya | 2022-11-18 | Y axis, Accuracy @ 4th position | 0.38
| Hoya | 2022-11-18 | Y axis, Accuracy @ 3rd position | 0.33
| Hoya | 2022-11-18 | Y axis, Accuracy @ 2nd position | 0.34
| Hoya | 2022-11-18 | Y axis OPM | 0.741
| Hoya | 2022-11-18 | XY Orthogonality | 90
| Hoya | 2022-11-18 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 65.0216
| Hoya | 2022-11-18 | X Detector noise (XY Servo on) | 0.184
| Hoya | 2022-11-18 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 88.5044
| Hoya | 2022-11-18 | X axis, Repeatability @1st position | 0.02
| Hoya | 2022-11-18 | X axis, Repeatability @ 4th position | 0.03
| Hoya | 2022-11-18 | X axis, Repeatability @ 3rd position | 0.03
| Hoya | 2022-11-18 | X axis, Repeatability @ 2nd position | 0.03
| Hoya | 2022-11-18 | X axis, Accuracy @1st position | 0.13
| Hoya | 2022-11-18 | X axis, Accuracy @ 4th position | 0.15
| Hoya | 2022-11-18 | X axis, Accuracy @ 3rd position | 0.14
| Hoya | 2022-11-18 | X axis, Accuracy @ 2nd position | 0.16
| Hoya | 2022-11-18 | X axis OPM | 0.551
| Hoya | 2022-11-18 | Under 50nm Step Height Sample (10 Repeats) | 19.658
| Hoya | 2022-11-18 | RMS of ZeroScan | 0.089
| Hoya | 2022-11-18 | NCM driving% with AC160TS cantilever | 4.31
| Hoya | 2022-11-18 | Laser intensity with AC160 cantilever | 2.994
| Hoya | 2022-11-18 | Full range of Z Height during Z Scan sweep | 15.2853
| Hoya | 2022-11-18 | Full range of Z Height during Z Scan sweep | 63.1523
| Hoya | 2022-11-18 | Full range of Z Detector during Z Scan sweep | 12.92
| Hoya | 2022-11-18 | Full range of Y Detector Raw during Y Scan sweep | 14.26
| Hoya | 2022-11-18 | Full range of Y Detector during Y Scan sweep | 64.3899
| Hoya | 2022-11-18 | Full range of X Detector Raw during X Scan sweep | 14.64
| Hoya | 2022-11-18 | 50~200nm Step Height Sample (10 Repeats) | 114.252
| Hoya | 2022-11-18 | 200~500nm Step Height Sample (10 Repeats) | 521.483
| Laboratory of Solid State Physics | 2022-11-03 | Z Linearity by 3 different sample height | 1.0151
| Laboratory of Solid State Physics | 2022-11-03 | Z Height noise | 0.025
| Laboratory of Solid State Physics | 2022-11-03 | Y Detector noise (XY Servo off) | 0.262
| Laboratory of Solid State Physics | 2022-11-03 | Y axis OPM | 0.49
| Laboratory of Solid State Physics | 2022-11-03 | X Detector noise (XY Servo off) | 0.258
| Laboratory of Solid State Physics | 2022-11-03 | X axis OPM | 0.442
| Laboratory of Solid State Physics | 2022-11-03 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.18
| Laboratory of Solid State Physics | 2022-11-03 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.22
| Laboratory of Solid State Physics | 2022-11-03 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off) | 0.14
| Laboratory of Solid State Physics | 2022-11-03 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on) | 0.18
| Laboratory of Solid State Physics | 2022-11-03 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 9.907
| Laboratory of Solid State Physics | 2022-11-03 | MicrometerPerAdc of Z Height (Z Height calibration) | 21.147
| Laboratory of Solid State Physics | 2022-11-03 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 162.4
| Laboratory of Solid State Physics | 2022-11-03 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 161.6
| Laboratory of Solid State Physics | 2022-11-03 | Full range of Z Height during Z Scan sweep | 15.6
| Laboratory of Solid State Physics | 2022-11-03 | Full range of Z Detector during Z Scan sweep | 14.8
| Laboratory of Solid State Physics | 2022-11-03 | Full range of Y Detector Raw during Y Scan sweep | 14.4
| Laboratory of Solid State Physics | 2022-11-03 | Full range of Y Detector during Y Scan sweep | 105.7
| Laboratory of Solid State Physics | 2022-11-03 | Full range of X Detector Raw during X Scan sweep | 14.7
| Laboratory of Solid State Physics | 2022-11-03 | Full range of X Detector during X Scan sweep | 117.85
| LG Display P9(H1) | 2022-09-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 27.082
| LG Display P9(H2) | 2022-09-01 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.7774
| LG Display P9(H1) | 2022-09-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 9.4117
| LG Display P9(H2) | 2022-09-01 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3203
| LG Display P9(H1) | 2022-09-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 110.33
| LG Display P9(H2) | 2022-09-01 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 100.368
| LG Display P9(H1) | 2022-09-01 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 163.61
| LG Display P9(H2) | 2022-09-01 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 153.416
| LG Display P9(H1) | 2022-09-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 106.54
| LG Display P9(H2) | 2022-09-01 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 358.367
| LG Display P9(H1) | 2022-09-01 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 158.21
| LG Display P9(H2) | 2022-09-01 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 159.013
| LG Display P9(H1) | 2022-09-01 | NCM driving% with AC160TS cantilever | 37.8
| LG Display P9(H2) | 2022-09-01 | NCM driving% with AC160TS cantilever | 11.5
| LG Display P9(H1) | 2022-09-01 | Laser intensity with AC160 cantilever | 2.572
| LG Display P9(H2) | 2022-09-01 | Laser intensity with AC160 cantilever | 2.895
| LG Display P9(H1) | 2022-09-01 | Full range of Z Height during Z Scan sweep | 15.9
| LG Display P9(H1) | 2022-09-01 | Full range of Z Height during Z Scan sweep | 114.5734
| LG Display P9(H2) | 2022-09-01 | Full range of Z Height during Z Scan sweep | 16.7478
| LG Display P9(H2) | 2022-09-01 | Full range of Z Height during Z Scan sweep | 120.2
| LG Display P9(H1) | 2022-09-01 | Full range of Y Detector Raw during Y Scan sweep | 14.23
| LG Display P9(H2) | 2022-09-01 | Full range of Y Detector Raw during Y Scan sweep | 14.27
| LG Display P9(H1) | 2022-09-01 | Full range of Y Detector during Y Scan sweep | 114.5201
| LG Display P9(H2) | 2022-09-01 | Full range of Y Detector during Y Scan sweep | 116.9
| LG Display P9(H1) | 2022-09-01 | Full range of X Detector Raw during X Scan sweep | 13.66
| LG Display P9(H2) | 2022-09-01 | Full range of X Detector Raw during X Scan sweep | 14.34
| LG Display P10(H1) | 2022-08-26 | Z Height noise | 0.033
| LG Display P10(H2) | 2022-08-26 | Z Height noise | 0.038
| LG Display P10(H1) | 2022-08-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.803
| LG Display P10(H2) | 2022-08-26 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 24.9784
| LG Display P10(H1) | 2022-08-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.068
| LG Display P10(H2) | 2022-08-26 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.3203
| LG Display P10(H1) | 2022-08-26 | Z Detector offset | 0.01
| LG Display P10(H2) | 2022-08-26 | Z Detector offset | 0.01
| LG Display P10(H1) | 2022-08-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 113.68
| LG Display P10(H2) | 2022-08-26 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 100.368
| LG Display P10(H1) | 2022-08-26 | Y Detector noise (XY Servo on) | 0.399
| LG Display P10(H2) | 2022-08-26 | Y Detector noise (XY Servo on) | 0.315
| LG Display P10(H1) | 2022-08-26 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 151.6
| LG Display P10(H1) | 2022-08-26 | Y axis OPM | 2.542
| LG Display P10(H2) | 2022-08-26 | Y axis OPM | 2.416
| LG Display P10(H1) | 2022-08-26 | XY Orthogonality | 90.2
| LG Display P10(H2) | 2022-08-26 | XY Orthogonality | 89.8
| LG Display P10(H1) | 2022-08-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 115.09
| LG Display P10(H2) | 2022-08-26 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 358.367
| LG Display P10(H1) | 2022-08-26 | X Detector noise (XY Servo on) | 0.233
| LG Display P10(H2) | 2022-08-26 | X Detector noise (XY Servo on) | 0.23
| LG Display P10(H1) | 2022-08-26 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 151.63
| LG Display P10(H2) | 2022-08-26 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 159.013
| LG Display P10(H1) | 2022-08-26 | X axis OPM | 3.808
| LG Display P10(H2) | 2022-08-26 | X axis OPM | 3.582
| LG Display P10(H1) | 2022-08-26 | NCM driving% with AC160TS cantilever | 12.96
| LG Display P10(H2) | 2022-08-26 | NCM driving% with AC160TS cantilever | 12.14
| LG Display P10(H1) | 2022-08-26 | Laser intensity with AC160 cantilever | 2.943
| LG Display P10(H2) | 2022-08-26 | Laser intensity with AC160 cantilever | 2.906
| LG Display P10(H1) | 2022-08-26 | Full range of Z Height during Z Scan sweep | 15.5062
| LG Display P10(H1) | 2022-08-26 | Full range of Z Height during Z Scan sweep | 122.2131
| LG Display P10(H2) | 2022-08-26 | Full range of Z Height during Z Scan sweep | 15.8633
| LG Display P10(H2) | 2022-08-26 | Full range of Z Height during Z Scan sweep | 112.8205
| LG Display P10(H1) | 2022-08-26 | Full range of Y Detector Raw during Y Scan sweep | 14.94
| LG Display P10(H2) | 2022-08-26 | Full range of Y Detector Raw during Y Scan sweep | 14.35
| LG Display P10(H1) | 2022-08-26 | Full range of Y Detector during Y Scan sweep | 114.618
| LG Display P10(H2) | 2022-08-26 | Full range of Y Detector during Y Scan sweep | 108.4151
| LG Display P10(H1) | 2022-08-26 | Full range of X Detector Raw during X Scan sweep | 15.04
| LG Display P10(H2) | 2022-08-26 | Full range of X Detector Raw during X Scan sweep | 14.02
| LG Display P10(H1) | 2022-08-26 | A-B after Auto PSPD Alignment | 0.048
| LG Display P10(H2) | 2022-08-26 | A-B after Auto PSPD Alignment | 0.075
| TSMC F12 P4 | 2022-08-25 | Z Linearity by 5 different sample height | 0.999
| TSMC F12 P4 | 2022-08-25 | Z Height noise | 0.034
| TSMC F12 P4 | 2022-08-25 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 25.9194
| TSMC F12 P4 | 2022-08-25 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 25.7884
| TSMC F12 P4 | 2022-08-25 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.6825
| TSMC F12 P4 | 2022-08-25 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.746
| TSMC F12 P4 | 2022-08-25 | Z Detector offset | -0.22
| TSMC F12 P4 | 2022-08-25 | Y2 Detector noise (XY Servo on) | 0.441
| TSMC F12 P4 | 2022-08-25 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 296.535
| TSMC F12 P4 | 2022-08-25 | Y1 Detector noise (XY Servo on) | 0.465
| TSMC F12 P4 | 2022-08-25 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 291.395
| TSMC F12 P4 | 2022-08-25 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 193.943
| TSMC F12 P4 | 2022-08-25 | Y Detector noise (XY Servo on) | 0.359
| TSMC F12 P4 | 2022-08-25 | Y axis, Repeatability @1st position | 0.05
| TSMC F12 P4 | 2022-08-25 | Y axis, Repeatability @ 4th position | 0.06
| TSMC F12 P4 | 2022-08-25 | Y axis, Repeatability @ 3rd position | 0.04
| TSMC F12 P4 | 2022-08-25 | Y axis, Repeatability @ 2nd position | 0.1
| TSMC F12 P4 | 2022-08-25 | Y axis, Accuracy @1st position | 0.52
| TSMC F12 P4 | 2022-08-25 | Y axis, Accuracy @ 4th position | 0.64
| TSMC F12 P4 | 2022-08-25 | Y axis, Accuracy @ 3rd position | 0.62
| TSMC F12 P4 | 2022-08-25 | Y axis, Accuracy @ 2nd position | 0.71
| TSMC F12 P4 | 2022-08-25 | Y axis OPM | 0.423
| TSMC F12 P4 | 2022-08-25 | X2 Detector noise (XY Servo on) | 0.378
| TSMC F12 P4 | 2022-08-25 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 296.619
| TSMC F12 P4 | 2022-08-25 | X1 Detector noise (XY Servo on) | 0.358
| TSMC F12 P4 | 2022-08-25 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 297.051
| TSMC F12 P4 | 2022-08-25 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 189.661
| TSMC F12 P4 | 2022-08-25 | X Detector noise (XY Servo on) | 0.292
| TSMC F12 P4 | 2022-08-25 | X axis, Repeatability @1st position | 0.05
| TSMC F12 P4 | 2022-08-25 | X axis, Repeatability @ 4th position | 0.03
| TSMC F12 P4 | 2022-08-25 | X axis, Repeatability @ 3rd position | 0.03
| TSMC F12 P4 | 2022-08-25 | X axis, Repeatability @ 2nd position | 0.09
| TSMC F12 P4 | 2022-08-25 | X axis, Accuracy @1st position | 0.24
| TSMC F12 P4 | 2022-08-25 | X axis, Accuracy @ 4th position | 0.18
| TSMC F12 P4 | 2022-08-25 | X axis, Accuracy @ 3rd position | 0.14
| TSMC F12 P4 | 2022-08-25 | X axis, Accuracy @ 2nd position | 0.38
| TSMC F12 P4 | 2022-08-25 | X axis OPM | 0.22
| TSMC F12 P4 | 2022-08-25 | OPM Repeatability @5mm | 1.0359
| TSMC F12 P4 | 2022-08-25 | OPM Repeatability @25mm | 4.1115
| TSMC F12 P4 | 2022-08-25 | OPM Repeatability @1mm | 0.75543
| TSMC F12 P4 | 2022-08-25 | OPM Repeatability @10mm | 1.237
| TSMC F12 P4 | 2022-08-25 | NCM driving% with AC160TS cantilever | 6.62
| TSMC F12 P4 | 2022-08-25 | Max OPM @5mm | 41.8995
| TSMC F12 P4 | 2022-08-25 | Max OPM @25mm | 704.4632
| TSMC F12 P4 | 2022-08-25 | Max OPM @1mm | 7.0323
| TSMC F12 P4 | 2022-08-25 | Max OPM @10mm | 341.427
| TSMC F12 P4 | 2022-08-25 | Laser intensity with AC160 cantilever | 3.1
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from -Y to +Y (Y scan) | 87.891
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from -Y to +Y (X scan) | 87.158
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from -X to +X (Y scan) | 26.367
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from -X to +X (X scan) | 62.988
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from +Y to -Y (Y scan) | 16.846
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from +Y to -Y (X scan) | 32.227
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from +X to -X (Y scan) | 17.578
| TSMC F12 P4 | 2022-08-25 | Image offset after moving from +X to -X (X scan) | 81.299
| TSMC F12 P4 | 2022-08-25 | Full range of Z Height during Z Scan sweep | 15.9738
| TSMC F12 P4 | 2022-08-25 | Full range of Z Detector during Z Scan sweep | 12.4
| TSMC F12 P4 | 2022-08-25 | Full range of Y2 Detector Raw during Y Scan sweep | 14.1
| TSMC F12 P4 | 2022-08-25 | Full range of Y2 Detector during Y Scan sweep | 200.1863
| TSMC F12 P4 | 2022-08-25 | Full range of Y1 Detector Raw during Y Scan sweep | 14
| TSMC F12 P4 | 2022-08-25 | Full range of Y1 Detector during Y Scan sweep | 203.9397
| TSMC F12 P4 | 2022-08-25 | Full range of X2 Detector Raw during X Scan sweep | 13.8
| TSMC F12 P4 | 2022-08-25 | Full range of X2 Detector during X Scan sweep | 195.9628
| TSMC F12 P4 | 2022-08-25 | Full range of X1 Detector Raw during X Scan sweep | 13.7
| TSMC F12 P4 | 2022-08-25 | Full range of X1 Detector during X Scan sweep | 198.3687
| TSMC F12 P4 | 2022-08-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.174
| TSMC F12 P4 | 2022-08-25 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.96
| TSMC F12 P4 | 2022-08-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.77
| TSMC F12 P4 | 2022-08-25 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.139
| TSMC F12 P4 | 2022-08-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) | 0.047
| TSMC F12 P4 | 2022-08-25 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.262
| TSMC F12 P4 | 2022-08-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.298
| TSMC F12 P4 | 2022-08-25 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) | 0.054
| TSMC F12 P4 | 2022-08-25 | A-B after Auto PSPD Alignment | 0.17
| TSMC F12 P4 | 2022-08-25 | A＋B difference during 3 times run test | 0.19
| SMIC Shanghai | 2022-08-19 | Z Height noise | 0.035
| SMIC Shanghai | 2022-08-19 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) | 26.7229
| SMIC Shanghai | 2022-08-19 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 26.5286
| SMIC Shanghai | 2022-08-19 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) | 10.7081
| SMIC Shanghai | 2022-08-19 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 10.7549
| SMIC Shanghai | 2022-08-19 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 278.486
| SMIC Shanghai | 2022-08-19 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 274.488
| SMIC Shanghai | 2022-08-19 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 193.647
| SMIC Shanghai | 2022-08-19 | Y axis, Repeatability @1st position | 0.09
| SMIC Shanghai | 2022-08-19 | Y axis, Repeatability @ 4th position | 0.07
| SMIC Shanghai | 2022-08-19 | Y axis, Repeatability @ 3rd position | 0.08
| SMIC Shanghai | 2022-08-19 | Y axis, Repeatability @ 2nd position | 0.06
| SMIC Shanghai | 2022-08-19 | Y axis, Accuracy @1st position | 0.45
| SMIC Shanghai | 2022-08-19 | Y axis, Accuracy @ 4th position | 0.38
| SMIC Shanghai | 2022-08-19 | Y axis, Accuracy @ 3rd position | 0.44
| SMIC Shanghai | 2022-08-19 | Y axis, Accuracy @ 2nd position | 0.26
| SMIC Shanghai | 2022-08-19 | Y axis OPM | 1.74
| SMIC Shanghai | 2022-08-19 | XY Orthogonality | 90
| SMIC Shanghai | 2022-08-19 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 273.853
| SMIC Shanghai | 2022-08-19 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 278.483
| SMIC Shanghai | 2022-08-19 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 193.235
| SMIC Shanghai | 2022-08-19 | X axis, Repeatability @1st position | 0.04
| SMIC Shanghai | 2022-08-19 | X axis, Repeatability @ 4th position | 0.05
| SMIC Shanghai | 2022-08-19 | X axis, Repeatability @ 3rd position | 0.06
| SMIC Shanghai | 2022-08-19 | X axis, Repeatability @ 2nd position | 0.06
| SMIC Shanghai | 2022-08-19 | X axis, Accuracy @1st position | 0.21
| SMIC Shanghai | 2022-08-19 | X axis, Accuracy @ 4th position | 0.24
| SMIC Shanghai | 2022-08-19 | X axis, Accuracy @ 3rd position | 0.24
| SMIC Shanghai | 2022-08-19 | X axis, Accuracy @ 2nd position | 0.27
| SMIC Shanghai | 2022-08-19 | X axis OPM | 1.23
| SMIC Shanghai | 2022-08-19 | OPM Repeatability @5mm | 1.1923
| SMIC Shanghai | 2022-08-19 | OPM Repeatability @25mm | 8.6728
| SMIC Shanghai | 2022-08-19 | OPM Repeatability @1mm | 0.76488
| SMIC Shanghai | 2022-08-19 | OPM Repeatability @10mm | 2.841
| SMIC Shanghai | 2022-08-19 | Max OPM @5mm | 28.969
| SMIC Shanghai | 2022-08-19 | Max OPM @25mm | 415.7567
| SMIC Shanghai | 2022-08-19 | Max OPM @1mm | 7.239
| SMIC Shanghai | 2022-08-19 | Max OPM @10mm | 89.9953
| SMIC Shanghai | 2022-08-19 | Laser intensity with AC160 cantilever | 2.98
| SMIC Shanghai | 2022-08-19 | Image offset after moving from -Y to +Y (Y scan) | 34.424
| SMIC Shanghai | 2022-08-19 | Image offset after moving from -Y to +Y (X scan) | 19.043
| SMIC Shanghai | 2022-08-19 | Image offset after moving from -X to +X (Y scan) | 70.312
| SMIC Shanghai | 2022-08-19 | Image offset after moving from -X to +X (X scan) | 9.521
| SMIC Shanghai | 2022-08-19 | Image offset after moving from +Y to -Y (Y scan) | 34.424
| SMIC Shanghai | 2022-08-19 | Image offset after moving from +Y to -Y (X scan) | 6.592
| SMIC Shanghai | 2022-08-19 | Image offset after moving from +X to -X (Y scan) | 72.51
| SMIC Shanghai | 2022-08-19 | Image offset after moving from +X to -X (X scan) | 55.664
| SMIC Shanghai | 2022-08-19 | Full range of Z Height during Z Scan sweep | 16.2222
| SMIC Shanghai | 2022-08-19 | Full range of Z Detector during Z Scan sweep | 12.24
| SMIC Shanghai | 2022-08-19 | Full range of Y2 Detector Raw during Y Scan sweep | 13.63
| SMIC Shanghai | 2022-08-19 | Full range of Y1 Detector Raw during Y Scan sweep | 13.98
| SMIC Shanghai | 2022-08-19 | Full range of Y1 Detector during Y Scan sweep | 212.4638
| SMIC Shanghai | 2022-08-19 | Full range of X2 Detector Raw during X Scan sweep | 13.68
| SMIC Shanghai | 2022-08-19 | Full range of X1 Detector Raw during X Scan sweep | 13.84
| SMIC Shanghai | 2022-08-19 | Full range of X1 Detector during X Scan sweep | 216.3269
| SMIC Shanghai | 2022-08-19 | A-B after Auto PSPD Alignment | -0.19
| NRMTC | 2022-08-17 | Z Linearity by Z scanner position | 0.02
| NRMTC | 2022-08-17 | Z Linearity by 5 different sample height | 1.006
| NRMTC | 2022-08-17 | Z Height noise | 0.037
| NRMTC | 2022-08-17 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.6576
| NRMTC | 2022-08-17 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 11.5112
| NRMTC | 2022-08-17 | Z Detector offset | 0.23
| NRMTC | 2022-08-17 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 59.879
| NRMTC | 2022-08-17 | Y Detector noise (XY Servo on) | 0.21
| NRMTC | 2022-08-17 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 90.5467
| NRMTC | 2022-08-17 | Y axis, Repeatability @1st position | 0.13
| NRMTC | 2022-08-17 | Y axis, Repeatability @ 4th position | 0.14
| NRMTC | 2022-08-17 | Y axis, Repeatability @ 3rd position | 0.13
| NRMTC | 2022-08-17 | Y axis, Repeatability @ 2nd position | 0.12
| NRMTC | 2022-08-17 | Y axis, Accuracy @1st position | 0.52
| NRMTC | 2022-08-17 | Y axis, Accuracy @ 4th position | 0.69
| NRMTC | 2022-08-17 | Y axis, Accuracy @ 3rd position | 0.59
| NRMTC | 2022-08-17 | Y axis, Accuracy @ 2nd position | 0.54
| NRMTC | 2022-08-17 | Y axis OPM | 1.942
| NRMTC | 2022-08-17 | XY Orthogonality | 90.2
| NRMTC | 2022-08-17 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 60.8646
| NRMTC | 2022-08-17 | X Detector noise (XY Servo on) | 0.176
| NRMTC | 2022-08-17 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 91.8677
| NRMTC | 2022-08-17 | X axis, Repeatability @1st position | 0.02
| NRMTC | 2022-08-17 | X axis, Repeatability @ 4th position | 0.03
| NRMTC | 2022-08-17 | X axis, Repeatability @ 3rd position | 0.03
| NRMTC | 2022-08-17 | X axis, Repeatability @ 2nd position | 0.03
| NRMTC | 2022-08-17 | X axis, Accuracy @1st position | 0.1
| NRMTC | 2022-08-17 | X axis, Accuracy @ 4th position | 0.14
| NRMTC | 2022-08-17 | X axis, Accuracy @ 3rd position | 0.13
| NRMTC | 2022-08-17 | X axis, Accuracy @ 2nd position | 0.18
| NRMTC | 2022-08-17 | X axis OPM | 1.411
| NRMTC | 2022-08-17 | Under 50nm Step Height Sample (10 Repeats) | 19.616
| NRMTC | 2022-08-17 | RMS of ZeroScan | 0.087
| NRMTC | 2022-08-17 | NCM driving% with AC160TS cantilever | 1.73
| NRMTC | 2022-08-17 | Laser intensity with AC160 cantilever | 2.966
| NRMTC | 2022-08-17 | Image offset after moving from -Y to +Y (Y scan) | 224.61
| NRMTC | 2022-08-17 | Image offset after moving from -Y to +Y (X scan) | 19.53
| NRMTC | 2022-08-17 | Image offset after moving from -X to +X (Y scan) | 131.84
| NRMTC | 2022-08-17 | Image offset after moving from -X to +X (X scan) | 83.01
| NRMTC | 2022-08-17 | Image offset after moving from +Y to -Y (Y scan) | 41.02
| NRMTC | 2022-08-17 | Image offset after moving from +Y to -Y (X scan) | 34.18
| NRMTC | 2022-08-17 | Image offset after moving from +X to -X (Y scan) | 126.95
| NRMTC | 2022-08-17 | Image offset after moving from +X to -X (X scan) | 161.13
| NRMTC | 2022-08-17 | Full range of Z Height during Z Scan sweep | 16.7355
| NRMTC | 2022-08-17 | Full range of Z Height during Z Scan sweep | 67.5342
| NRMTC | 2022-08-17 | Full range of Z Detector during Z Scan sweep | 12.94
| NRMTC | 2022-08-17 | Full range of Y Detector Raw during Y Scan sweep | 14.27
| NRMTC | 2022-08-17 | Full range of Y Detector during Y Scan sweep | 63.8992
| NRMTC | 2022-08-17 | Full range of X Detector Raw during X Scan sweep | 14.2
| NRMTC | 2022-08-17 | A-B after Auto PSPD Alignment | 0.024
| NRMTC | 2022-08-17 | 50~200nm Step Height Sample (10 Repeats) | 85.243
| Samsung SRD | 2022-04-15 | Z Linearity by Z scanner position | 2
| Samsung SRD | 2022-04-15 | Z Linearity by 5 different sample height | 1.0009
| Samsung SRD | 2022-04-15 | Z Height noise | 0.072
| Samsung SRD | 2022-04-15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 58.5942
| Samsung SRD | 2022-04-15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 25.2503
| Samsung SRD | 2022-04-15 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB) | 159.723
| Samsung SRD | 2022-04-15 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB) | 267.124
| Samsung SRD | 2022-04-15 | Y axis, Repeatability @1st position | 0.04
| Samsung SRD | 2022-04-15 | Y axis, Repeatability @ 4th position | 0.04
| Samsung SRD | 2022-04-15 | Y axis, Repeatability @ 3rd position | 0.03
| Samsung SRD | 2022-04-15 | Y axis, Repeatability @ 2nd position | 0.04
| Samsung SRD | 2022-04-15 | Y axis, Accuracy @1st position | 0.38
| Samsung SRD | 2022-04-15 | Y axis, Accuracy @ 4th position | 0.35
| Samsung SRD | 2022-04-15 | Y axis, Accuracy @ 3rd position | 0.28
| Samsung SRD | 2022-04-15 | Y axis, Accuracy @ 2nd position | 0.29
| Samsung SRD | 2022-04-15 | Y axis OPM | 0.373
| Samsung SRD | 2022-04-15 | XY Orthogonality | 90.1
| Samsung SRD | 2022-04-15 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB) | 142.311
| Samsung SRD | 2022-04-15 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB) | 279.541
| Samsung SRD | 2022-04-15 | X axis, Repeatability @1st position | 0.04
| Samsung SRD | 2022-04-15 | X axis, Repeatability @ 4th position | 0.03
| Samsung SRD | 2022-04-15 | X axis, Repeatability @ 3rd position | 0.03
| Samsung SRD | 2022-04-15 | X axis, Repeatability @ 2nd position | 0.05
| Samsung SRD | 2022-04-15 | X axis, Accuracy @1st position | 0.24
| Samsung SRD | 2022-04-15 | X axis, Accuracy @ 4th position | 0.16
| Samsung SRD | 2022-04-15 | X axis, Accuracy @ 3rd position | 0.15
| Samsung SRD | 2022-04-15 | X axis, Accuracy @ 2nd position | 0.29
| Samsung SRD | 2022-04-15 | X axis OPM | 0.196
| Samsung SRD | 2022-04-15 | OPM Repeatability @5mm | 3.1304
| Samsung SRD | 2022-04-15 | OPM Repeatability @25mm | 8.5318
| Samsung SRD | 2022-04-15 | OPM Repeatability @1mm | 1.1065
| Samsung SRD | 2022-04-15 | OPM Repeatability @10mm | 4.0371
| Samsung SRD | 2022-04-15 | Max OPM @5mm | 28.3297
| Samsung SRD | 2022-04-15 | Max OPM @25mm | 198.2548
| Samsung SRD | 2022-04-15 | Max OPM @1mm | 10.0719
| Samsung SRD | 2022-04-15 | Max OPM @10mm | 76.4336
| Samsung SRD | 2022-04-15 | Laser intensity with AC160 cantilever | 3.28
| Samsung SRD | 2022-04-15 | Image offset after moving from -Y to +Y (Y scan) | 42.48
| Samsung SRD | 2022-04-15 | Image offset after moving from -Y to +Y (X scan) | 19.78
| Samsung SRD | 2022-04-15 | Image offset after moving from -X to +X (Y scan) | 41.02
| Samsung SRD | 2022-04-15 | Image offset after moving from -X to +X (X scan) | 151.61
| Samsung SRD | 2022-04-15 | Image offset after moving from +Y to -Y (Y scan) | 75.44
| Samsung SRD | 2022-04-15 | Image offset after moving from +Y to -Y (X scan) | 8.79
| Samsung SRD | 2022-04-15 | Image offset after moving from +X to -X (Y scan) | 35.16
| Samsung SRD | 2022-04-15 | Image offset after moving from +X to -X (X scan) | 120.85
| Samsung SRD | 2022-04-15 | Full range of Z Height during Z Scan sweep | 39.852
| Samsung SRD | 2022-04-15 | Full range of Z Detector during Z Scan sweep | 13.59
| Samsung SRD | 2022-04-15 | Full range of Y2 Detector Raw during Y Scan sweep | 14.79
| Samsung SRD | 2022-04-15 | Full range of Y1 Detector Raw during Y Scan sweep | 15.01
| Samsung SRD | 2022-04-15 | Full range of Y1 Detector during Y Scan sweep | 241.64
| Samsung SRD | 2022-04-15 | Full range of X2 Detector Raw during X Scan sweep | 14.11
| Samsung SRD | 2022-04-15 | Full range of X1 Detector Raw during X Scan sweep | 14.47
| Samsung SRD | 2022-04-15 | Full range of X1 Detector during X Scan sweep | 234.58
| Samsung SRD | 2022-04-15 | 940nm Step Height Sample (10 Repeats) | 942.849
| Samsung SRD | 2022-04-15 | 50~200nm Step Height Sample (10 Repeats) | 15.898
| Samsung SRD | 2022-04-15 | 200~500nm Step Height Sample (10 Repeats) | 454.653
| Riken | 2022-03-14 | Z Linearity by Z scanner position | 0.02
| Riken | 2022-03-14 | Z Linearity by 5 different sample height | 0.999
| Riken | 2022-03-14 | Z Height noise | 0.045
| Riken | 2022-03-14 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 28.7189
| Riken | 2022-03-14 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 12.6678
| Riken | 2022-03-14 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB) | 61.9441
| Riken | 2022-03-14 | Y Detector noise (XY Servo on) | 0.158
| Riken | 2022-03-14 | Y Detector calibration value (MicrometerPerAdc in YDetector DB) | 81.991
| Riken | 2022-03-14 | Y axis, Repeatability @1st position | 0.03
| Riken | 2022-03-14 | Y axis, Repeatability @ 4th position | 0.03
| Riken | 2022-03-14 | Y axis, Repeatability @ 3rd position | 0.03
| Riken | 2022-03-14 | Y axis, Repeatability @ 2nd position | 0.03
| Riken | 2022-03-14 | Y axis, Accuracy @1st position | 0.56
| Riken | 2022-03-14 | Y axis, Accuracy @ 4th position | 0.64
| Riken | 2022-03-14 | Y axis, Accuracy @ 3rd position | 0.62
| Riken | 2022-03-14 | Y axis, Accuracy @ 2nd position | 0.58
| Riken | 2022-03-14 | Y axis OPM | 1.3486
| Riken | 2022-03-14 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB) | 60.571
| Riken | 2022-03-14 | X Detector noise (XY Servo on) | 0.164
| Riken | 2022-03-14 | X Detector calibration value (MicrometerPerAdc in XDetector DB) | 95.2491
| Riken | 2022-03-14 | X axis, Repeatability @1st position | 0.04
| Riken | 2022-03-14 | X axis, Repeatability @ 4th position | 0.03
| Riken | 2022-03-14 | X axis, Repeatability @ 3rd position | 0.04
| Riken | 2022-03-14 | X axis, Repeatability @ 2nd position | 0.04
| Riken | 2022-03-14 | X axis, Accuracy @1st position | 0.35
| Riken | 2022-03-14 | X axis, Accuracy @ 4th position | 0.16
| Riken | 2022-03-14 | X axis, Accuracy @ 3rd position | 0.26
| Riken | 2022-03-14 | X axis, Accuracy @ 2nd position | 0.35
| Riken | 2022-03-14 | X axis OPM | 1.0374
| Riken | 2022-03-14 | Under 50nm Step Height Sample (10 Repeats) | 16.085
| Riken | 2022-03-14 | Image offset after moving from -Y to +Y (Y scan) | 79.834
| Riken | 2022-03-14 | Image offset after moving from -Y to +Y (X scan) | 13.916
| Riken | 2022-03-14 | Image offset after moving from -X to +X (Y scan) | 46.875
| Riken | 2022-03-14 | Image offset after moving from -X to +X (X scan) | 10.9863
| Riken | 2022-03-14 | Image offset after moving from +Y to -Y (Y scan) | 2.9297
| Riken | 2022-03-14 | Image offset after moving from +Y to -Y (X scan) | 37.3535
| Riken | 2022-03-14 | Image offset after moving from +X to -X (Y scan) | 84.9609
| Riken | 2022-03-14 | Image offset after moving from +X to -X (X scan) | 72.5098
| Riken | 2022-03-14 | Full range of Z Height during Z Scan sweep | 19.7852
| Riken | 2022-03-14 | Full range of Z Height during Z Scan sweep | 60.7164
| Riken | 2022-03-14 | Full range of Z Detector during Z Scan sweep | 13.79
| Riken | 2022-03-14 | Full range of Y Detector Raw during Y Scan sweep | 12.62
| Riken | 2022-03-14 | Full range of Y Detector during Y Scan sweep | 58.4352
| Riken | 2022-03-14 | Full range of X Detector Raw during X Scan sweep | 13.16
| Riken | 2022-03-14 | 50~200nm Step Height Sample (10 Repeats) | 181.427
| Riken | 2022-03-14 | 200~500nm Step Height Sample (10 Repeats) | 455.154
| ASML (ST Instruments) | 2022-02-26 | Z Linearity by 3 different sample height | 1.018
| ASML (ST Instruments) | 2022-02-26 | Z Height noise | 0.081
| ASML (ST Instruments) | 2022-02-26 | Y Detector noise (XY Servo off) | 0.27
| ASML (ST Instruments) | 2022-02-26 | Y axis OPM | 0.672
| ASML (ST Instruments) | 2022-02-26 | X Detector noise (XY Servo off) | 0.265
| ASML (ST Instruments) | 2022-02-26 | X axis OPM | 1.773
| ASML (ST Instruments) | 2022-02-26 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 24.17
| ASML (ST Instruments) | 2022-02-26 | MicrometerPerAdc of Z Height (Z Height calibration) | 66.317
| ASML (ST Instruments) | 2022-02-26 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 171.8
| ASML (ST Instruments) | 2022-02-26 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 168.6
| ASML (ST Instruments) | 2022-02-26 | Full range of Z Height during Z Scan sweep | 41.4
| ASML (ST Instruments) | 2022-02-26 | Full range of Z Detector during Z Scan sweep | 13.3
| ASML (ST Instruments) | 2022-02-26 | Full range of Y1 Detector Raw during Y Scan sweep | 13.8
| ASML (ST Instruments) | 2022-02-26 | Full range of Y Detector during Y Scan sweep | 118
| ASML (ST Instruments) | 2022-02-26 | Full range of X1 Detector Raw during X Scan sweep | 14
| ASML (ST Instruments) | 2022-02-26 | Full range of X Detector during X Scan sweep | 117.8
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Z Linearity by 3 different sample height | 1.0217
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Z Height noise | 0.078
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Y Detector noise (XY Servo off) | 0.254
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Y axis OPM | 0.287
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | X Detector noise (XY Servo off) | 0.255
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | X axis OPM | 1.693
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 25.772
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | MicrometerPerAdc of Z Height (Z Height calibration) | 65.988
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration) | 106.9
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration) | 108.8
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Full range of Z Height during Z Scan sweep | 42.5
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Full range of Z Detector during Z Scan sweep | 12.8
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Full range of Y Detector Raw during Y Scan sweep | 14.2
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Full range of Y Detector during Y Scan sweep | 75.5
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Full range of X Detector Raw during X Scan sweep | 14
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Full range of X Detector during X Scan sweep | 75.8
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Approach position at -5μm | 0.0104
| Daikin Chemical Europe GmbH (PSE) | 2022-02-25 | Approach position at +5μm | 0.0096
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Z Linearity by 3 different sample height | 1.0045
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Z Height noise | 0.075
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y2 Detector noise (XY Servo on) | 0.272
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y2 Detector noise (XY Servo off) | 0.253
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y1 Detector noise (XY Servo on) | 0.265
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y1 Detector noise (XY Servo off) | 0.251
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y Detector noise (XY Servo on) | 0.202
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y Detector noise (XY Servo off) | 0.186
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Y axis OPM | 2.631
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X2 Detector noise (XY Servo on) | 0.26
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X2 Detector noise (XY Servo off) | 0.248
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X1 Detector noise (XY Servo on) | 0.262
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X1 Detector noise (XY Servo off) | 0.258
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X Detector noise (XY Servo on) | 0.194
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X Detector noise (XY Servo off) | 0.187
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | X axis OPM | 2.687
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration) | 25.764
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | MicrometerPerAdc of Z Height (Z Height calibration) | 63.983
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration) | 287.9
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration) | 284.6
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration) | 275.8
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration) | 275.7
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of Z Height during Z Scan sweep | 42.9
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of Z Detector during Z Scan sweep | 13.4
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of Y2 Detector Raw during Y Scan sweep | 14.6
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of Y1 Detector Raw during Y Scan sweep | 14.7
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of Y Detector during Y Scan sweep | 201.8
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of X2 Detector Raw during X Scan sweep | 14.7
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of X1 Detector Raw during X Scan sweep | 14.8
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Full range of X Detector during X Scan sweep | 201.7
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Approach position at -5μm | 0.0112
| Zhejiang Mdk Optical Semiconductor | 2022-02-24 | Approach position at +5μm | 0.0068

## Specs
Model | Check Item | LSL | USL | Target
NX-Wafer | Z Linearity by Z scanner position
NX-Wafer | Full range of Z Detector during Z Scan sweep
NX-Wafer | Z Height calibration error
NX-Wafer | Z Drive calibration error
NX-Wafer | Z Linearity by 4 different sample height
NX-Wafer | Full range of Z Height during Z Scan sweep
NX-Wafer | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer | Z Height noise
NX-Wafer | Wafer placement(Loading) error on the load port1 (X offset)
NX-Wafer | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Wafer | Wafer placement(Loading) error on the load port2 (X offset)
NX-Wafer | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Wafer | OpticalAxisShiftPerFocusX
NX-Wafer | OpticalAxisShiftPerFocusY
NX-Wafer | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Wafer | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Wafer | Lift pins synchronization offset check (X offset)
NX-Wafer | Lift pins synchronization offset check (Y offset)
NX-Wafer | Profiler noise @200um
NX-Wafer | Roll motion to X+ full scan
NX-Wafer | Position repeatability(10repeats, Position Based)
NX-Wafer | OPM Repeatability @5mm
NX-Wafer | OPM Repeatability @1mm
NX-Wafer | Max OPM @5mm
NX-Wafer | Max OPM @1mm
NX-Wafer | Torque Profile (10mm, 200um/s)
NX-Wafer | Torque Profile (10mm, 20um/s)
NX-Wafer | Lateral(C-D) sensitivity
NX-Wafer | Lateral(C-D) noise
NX-Wafer | Min. Vibration Transmissbility (~100Hz)
NX-Wafer | RMS of ZeroScan
NX-Wafer | LightStrengthGain
NX-Wafer | Full range of X Detector during X Scan sweep
NX-Wafer | Full range of X1 Detector during X Scan sweep
NX-Wafer | Full range of X2 Detector during X Scan sweep
NX-Wafer | Full range of Y Detector during Y Scan sweep
NX-Wafer | Full range of Y1 Detector during Y Scan sweep
NX-Wafer | Full range of Y2 Detector during Y Scan sweep
NX-Wafer | Full range of X1 Detector Raw during X Scan sweep
NX-Wafer | Full range of X2 Detector Raw during X Scan sweep
NX-Wafer | Full range of Y1 Detector Raw during Y Scan sweep
NX-Wafer | Full range of Y2 Detector Raw during Y Scan sweep
NX-Wafer | X Detector stroke calibration error (nominal scale error X)
NX-Wafer | Y Detector stroke calibration error (nominal scale error Y)
NX-Wafer | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Wafer | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Wafer | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Wafer | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Wafer | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Wafer | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Wafer | X Non-linearity error
NX-Wafer | Y Non-linearity error
NX-Wafer | X axis OPM
NX-Wafer | Y axis OPM
NX-Wafer | X Ringing(Rpv)
NX-Wafer | Y Ringing(Rpv)
NX-Wafer | X Detector noise (XY Servo on)
NX-Wafer | Y Detector noise (XY Servo on)
NX-Wafer | X1 Detector noise (XY Servo on)
NX-Wafer | Y1 Detector noise (XY Servo on)
NX-Wafer | X2 Detector noise (XY Servo on)
NX-Wafer | Y2 Detector noise (XY Servo on)
NX-Wafer | RMS error after Stage Calibration
NX-Wafer | Image offset after moving from -X to +X (X scan)
NX-Wafer | Image offset after moving from -Y to +Y (Y scan)
NX-Wafer | Image offset after moving from -X to +X (Y scan)
NX-Wafer | Image offset after moving from -Y to +Y (X scan)
NX-Wafer | Image offset after moving from +X to -X (X scan)
NX-Wafer | Image offset after moving from +Y to -Y (Y scan)
NX-Wafer | Image offset after moving from +X to -X (Y scan)
NX-Wafer | Image offset after moving from +Y to -Y (X scan)
NX-Wafer | XY Stage positioning X offset (Vision Pattern Matching)
NX-Wafer | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Wafer | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Wafer | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Wafer | XY Stage positioning X offset (In-Die Align)
NX-Wafer | XY Stage positioning Y offset (In-Die Align)
NX-Wafer | XY Stage positioning X 1σ (In-Die Align)
NX-Wafer | XY Stage positioning Y 1σ (In-Die Align)
NX-Wafer | XY Stage positioning X offset (LLC Translation)
NX-Wafer | XY Stage positioning Y offset (LLC Translation)
NX-Wafer | XY Stage positioning X 1σ (LLC Translation)
NX-Wafer | XY Stage positioning Y 1σ (LLC Translation)
NX-Wafer | XY Stage positioning X offset (Global Align)
NX-Wafer | XY Stage positioning Y offset (Global Align)
NX-Wafer | XY Stage positioning X 1σ (Global Align)
NX-Wafer | XY Stage positioning Y 1σ (Global Align)
NX-Wafer | Z Drift
NX-Wafer | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | Cycle time for 15-point measurement
NX-Wafer | Short-term stability using Step Height Sample (3σ)
NX-Wafer | Roughness Repeatability 1x1um(RSD%)
NX-Wafer | Roughness Repeatability 1x1um(1σ)
NX-Wafer | Roughness Repeatability 80x80um(RSD%)
NX-Wafer | Roughness Repeatability 80x80um(1σ)
NX-Wafer | Z Height calibrtaion error
NX-Wafer | Z Drive calibrtaion error
NX-Wafer | Z Linearity by 3 different sample height
NX-Wafer | Full range in voltage during Z Scan sweep
NX-Wafer | Full range in um during Z Scan sweep
NX-Wafer | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer | Lateral(C-D) sensitivity
NX-Wafer | Lateral(C-D) noise
NX-Wafer | A-B noise
NX-Wafer | NCM Amplitude noise
NX-Wafer | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Wafer | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Wafer | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Wafer | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Wafer | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Wafer | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Wafer | XY Orthogonality
NX-Wafer | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Wafer | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Wafer | Z linearity by sample height
NX-Wafer | 1) Z Scanner position at +3μm
NX-Wafer | 3) Z Scanner position at -3μm
NX-Wafer | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Wafer | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Wafer | X Scan Stroke
NX-Wafer | Y Scan Stroke
NX-Wafer | Full range of X Detector Raw during X Scan sweep
NX-Wafer | Full range of Y Detector Raw during Y Scan sweep
NX-Wafer | X Detector noise (XY Servo off)
NX-Wafer | Y Detector noise (XY Servo off)
NX-Wafer | X axis OPM
NX-Wafer | Y axis OPM
NX-Wafer | X Non-linearity
NX-Wafer | Y Non-linearity
NX-Wafer | OPM Repeatability @25mm
NX-Wafer | OPM Repeatability @10mm
NX-Wafer | Max OPM @25mm
NX-Wafer | Max OPM @10mm
NX-Wafer | Torque Profile (50mm, 200um/s)
NX-Wafer | Torque Profile (50mm, 20um/s)
NX-Wafer | SLD current limit check
NX-Wafer | SLD beam power at the end of the lens
NX-Wafer | 2) of 20nm
NX-Wafer | X Detector noise (XYServo off)
NX-Wafer | Y Detector noise (XYServo off)
NX-Wafer | X axis OPM raw Rpv
NX-Wafer | Y axis OPM raw Rpv
NX-Wafer | StageSafetyXMax
NX-Wafer | StageSafetyXMin
NX-Wafer | StageSafetyY
NX-Wafer | StageSafetyYMax
NX-Wafer | StageSafetyYMin
NX-Wafer | Check X axis upper limit sensor function
NX-Wafer | Check X axis lower limit sensor function
NX-Wafer | Check Y axis upper limit sensor function
NX-Wafer | Check Y axis lower limit sensor function
NX-Wafer | X axis Angle Range
NX-Wafer | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Wafer | X detector noise (XY Servo on)
NX-Wafer | Y detector noise (XY Servo on)
NX-Wafer | Full range in um during X Scan sweep (XY Servo off)
NX-Wafer | Full range in um during Y Scan sweep (XY Servo off)
NX-Wafer | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Wafer | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Wafer | X axis, Repeatability @1st position
NX-Wafer | X axis, Repeatability @ 2nd position
NX-Wafer | X axis, Repeatability @ 3rd position
NX-Wafer | X axis, Repeatability @ 4th position
NX-Wafer | X axis, Accuracy @1st position
NX-Wafer | X axis, Accuracy @ 2nd position
NX-Wafer | X axis, Accuracy @ 3rd position
NX-Wafer | X axis, Accuracy @ 4th position
NX-Wafer | Y axis, Repeatability @1st position
NX-Wafer | Y axis, Repeatability @ 2nd position
NX-Wafer | Y axis, Repeatability @ 3rd position
NX-Wafer | Y axis, Repeatability @ 4th position
NX-Wafer | Y axis, Accuracy @1st position
NX-Wafer | Y axis, Accuracy @ 2nd position
NX-Wafer | Y axis, Accuracy @ 3rd position
NX-Wafer | Y axis, Accuracy @ 4th position
NX-Wafer | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Wafer | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Wafer | Z-XY Orthogonality (X direction-Left)
NX-Wafer | Z-XY Orthogonality (X direction-Right)
NX-Wafer | Z-XY Orthogonality (Y direction-Left)
NX-Wafer | Z-XY Orthogonality (Y direction-Right)
NX-Wafer | X-Scan Repeatability (10 repeats, RSD%)
NX-Wafer | Y-Scan Repeatability (10 repeats, RSD%)
NX-Wafer | Mechanical Align Accuracy
NX-Wafer | Check the Multi Peattern Upload
NX-Wafer | X Detector noise (XY servo off)
NX-Wafer | Y Detector noise (XY servo off)
NX-Wafer | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer | Z Detector Max
NX-Wafer | Z Detector Min
NX-Wafer | WLI_Scale(Z Calibration)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Wafer | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Wafer | Laser intensity with AC160TS cantilever
NX-Wafer | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Wafer | NCM driving % at an amplitude of 1 nm
NX-Wafer | NCM driving % at an amplitude of 20 nm
NX-Wafer | NCM driving % at an amplitude of 100 nm
NX-Wafer | - 940nm Step Height Sample (10 Repeats)
NX-Wafer | - 200~500nm Step Height Sample (10 Repeats)
NX-Wafer | - 50~200nm Step Height Sample (10 Repeats)
NX-Wafer | - Under 50nm Step Height Sample (10 Repeats)
NX-Wafer | A-B after Auto PSPD Alignment
NX-Wafer | NCM driving% with AC160TS cantilever
NX-Wafer | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer | X Detector stroke calibration error (nominal scale error X)
NX-Wafer | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | Resonance Frequency in DB
NX-Wafer | Servo Cut-off Frequency in DB
NX-Wafer | Check chuck surface no damage
NX-Wafer | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Wafer | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Wafer | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Wafer | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Wafer | X1 Detector noise (XY servo off)
NX-Wafer | X2 Detector noise (XY servo off)
NX-Wafer | Y1 Detector noise (XY servo off)
NX-Wafer | Y2 Detector noise (XY servo off)
NX-Wafer | X Detector noise (XY servo on)
NX-Wafer | X1 Detector noise (XY servo on)
NX-Wafer | X2 Detector noise (XY servo on)
NX-Wafer | Y Detector noise (XY servo on)
NX-Wafer | Y1 Detector noise (XY servo on)
NX-Wafer | Y2 Detector noise (XY servo on)
NX-Wafer | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer | Z Detector offset
NX-Wafer | Check the full stroke image(servo on)
NX-Wafer | ResonanceFrequencyHz
NX-Wafer | ServoCutoffFrequencyHz
NX-Wafer | FullPixelHeight
NX-Wafer | Serial Port#
NX-Wafer | Off-axis Camera X Position(ΔX)
NX-Wafer | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Wafer | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Wafer | Z Scanner position at -3μm
NX-Wafer | Z Scanner position at +3μm
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer | Scan offset accuracy (Precision)
NX-Wafer | Aligned stage accuracy (X offset-Point Align1)
NX-Wafer | Aligned stage accuracy (Y offset-Point Align1)
NX-Wafer | Aligned stage accuracy (X offset-Point Align2)
NX-Wafer | Aligned stage accuracy (Y offset-Point Align2)
NX-Wafer | Aligned stage accuracy (X offset-Point Align3)
NX-Wafer | Aligned stage accuracy (Y offset-Point Align3)
NX-Wafer | Aligned stage accuracy (X offset)
NX-Wafer | Aligned stage accuracy (Y offset)
NX-Wafer | Imag offset after EFEM loading (X scan)
NX-Wafer | Imag offset after EFEM loading (Y scan)
NX-Wafer | Imag offset after Lock stage (X scan)
NX-Wafer | Imag offset after Lock stage (Y scan)
NX-Wafer | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Wafer | Mask Roughness Repeatability (10Repeats, σ )
NX-Wafer | Mask exchange time
NX-Wafer | WLI_Scale(C Calibration)
NX-Wafer | WLI_Scale(B Calibration)
NX-Wafer | WLI_Scale(A Calibration)
NX-Wafer | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Wafer | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Wafer | Port 1 lens(to OnAxis) Y Offset
NX-Wafer | Port 1 lens(to OnAxis) X Offset
NX-Wafer | P Gain
NX-Wafer | I Gain
NX-Wafer | D Gain
NX-Wafer | Check the Port 2 lens light strength
NX-Wafer | Check the Port 1 lens light strength
NX-Wafer | Check the focus deviation between the lenses(Δf)
NX-Wafer | Check the 25mm Moving – Y Axis Offset Check
NX-Wafer | C Scanner Detector Min
NX-Wafer | C Scanner Detector Max
NX-Wafer | B Scanner Detector Min
NX-Wafer | B Scanner Detector Max
NX-Wafer | Angle align between Port 2 lens and VLSI SHS: Y
NX-Wafer | Angle align between Port 2 lens and VLSI SHS: X
NX-Wafer | Angle align between Port 2 lens and Sample chuck: Y
NX-Wafer | Angle align between Port 2 lens and Sample chuck: X
NX-Wafer | Angle align between Port 1 lens and Sample chuck: Y
NX-Wafer | Angle align between Port 1 lens and Sample chuck: X
NX-Wafer | A Scanner Detector Min
NX-Wafer | A Scanner Detector Max
NX-Wafer | Y2 Detector noise (XY Servo off)
NX-Wafer | X2 Detector noise (XY Servo off)
NX-Wafer | Y axis OPM (-52000,-52000)
NX-Wafer | Y axis OPM (-52000,52000)
NX-Wafer | Y axis OPM (52000,-52000)
NX-Wafer | Y axis OPM (52000,52000)
NX-Wafer | Y axis OPM (0,0)
NX-Wafer | XY Orthogonality (-52000,-52000)
NX-Wafer | XY Orthogonality (-52000,52000)
NX-Wafer | XY Orthogonality (52000,-52000)
NX-Wafer | XY Orthogonality (52000,52000)
NX-Wafer | XY Orthogonality (0,0)
NX-Wafer | X axis OPM (-52000,-52000)
NX-Wafer | X axis OPM (-52000,52000)
NX-Wafer | X axis OPM (52000,-52000)
NX-Wafer | X axis OPM (52000,52000)
NX-Wafer | X axis OPM (0,0)
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Wafer | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Wafer | XY stage calibration
NX-Wafer | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Wafer | LD beam power at the end of the lens
NX-Wafer | Y Non-lineairty
NX-Wafer | X Non-lineairty
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Wafer | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Wafer | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Wafer | UpperLimit
NX-Wafer | Under 50nm Step Height Sample (10 Repeats)
NX-Wafer | Overlay Camera X Position(ΔX)
NX-Wafer | LowerLimit
NX-Wafer | Initial Z
NX-Wafer | 940nm Step Height Sample (10 Repeats)
NX-Wafer | 50~200nm Step Height Sample (10 Repeats)
NX-Wafer | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Wafer | StrokeYMin
NX-Wafer | StrokeYMax
NX-Wafer | StrokeY
NX-Wafer | StrokeXMin
NX-Wafer | StrokeXMax
NX-Wafer | StrokeX
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Wafer | Z Drift (Short Traveling)
NX-Wafer | 200~500nm Step Height Sample (10 Repeats)
NX-Wafer | XY orthogonality
NX-Wafer | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Wafer | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Wafer | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Wafer | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Wafer | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Wafer | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Wafer | Y1 Detector noise (XY Servo off)
NX-Wafer | X1 Detector noise (XY Servo off)
NX-Wafer | Maximum Y Scan range (Full Stroke)
NX-Wafer | Maximum X Scan range (Full Stroke)
NX-Wafer | Full range of Y2 Detector raw
NX-Wafer | Full range of Y1 Detector raw
NX-Wafer | Full range of X2 Detector raw
NX-Wafer | Full range of X1 Detector raw
NX-Wafer | Z Linearity by Z Scanner position
NX-Wafer | -Y Direction software Limit
NX-Wafer | -X Direction software Limit
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Wafer | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Wafer | %GRR w/ roughness of non-patterened wafer
NX-Wafer | Z Scanner position at -6μm
NX-Wafer | Z Scanner position at +6μm
NX-Wafer | Approach position at -5μm
NX-Wafer | Approach position at +5μm
NX-Wafer | * 1800nm Step Height Sample (10 Repeats)
NX-Wafer | Full range of Z Detector raw
NX-Wafer | Approach position at -3μm
NX-Wafer | Approach position at +3μm
NX-Wafer | Under 50nm Step Height Sample (10 Repeats)
NX-Wafer | 50~200nm Step Height Sample (10 Repeats)
NX-Wafer | 200~500nm Step Height Sample (10 Repeats)
NX-Wafer | PosMin
NX-Wafer | PosMax
NX-Wafer | MaxLength
NX-Wafer | FullRecoveryProfileLength
NX-Wafer | Full range of Y Detector raw
NX-Wafer | Full range of X Detector raw
NX-Wafer | EFM Phase : For the positive bias
NX-Wafer | EFM Phase : For the negative bias
NX-Wafer | Z Linearity by 5 different sample height
NX-Wafer | Laser intensity with AC160 cantilever
NX-Wafer | A－B difference during 3 times run test
NX-Wafer | A＋B difference during 3 times run test
NX-Wafer | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Wafer | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Wafer | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Wafer | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Wafer | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Wafer | Allowable max warpage, w/ SiC sample chuck
NX-Wafer | RMS of Z Height of Wafer Zeroscan (Top)
NX-Wafer | RMS of Z Height of Wafer Zeroscan (Right)
NX-Wafer | RMS of Z Height of Wafer Zeroscan (Left)
NX-Wafer | RMS of Z Height of Wafer ZeroScan (Center)
NX-Wafer | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Wafer | PosMin
NX-Wafer | PosMax
NX-Wafer | MaxLength
NX-Wafer | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Wafer | Check sample surface mapping
NX-Wafer | Z Direction, Accuracy Test (Bias not significant)
NX-Wafer | Y Direction, Accuracy Test (Bias not significant)
NX-Wafer | X Direction, Accuracy Test (Bias not significant)
NX-Wafer | Roughness Reproducibility Test (SR Panel)
NX-Wafer | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Wafer | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Wafer | Roughness Reproducibility Test (SiNx Panel)
NX-Wafer | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Wafer | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Wafer | Roughness Reproducibility Test (ABF Panel)
NX-Wafer | Roughness Repeatability Test (SR Panel)
NX-Wafer | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Wafer | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Wafer | Roughness Repeatability Test (SiNx Panel)
NX-Wafer | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Wafer | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Wafer | Roughness Repeatability Test (ABF Panel)
NX-Wafer | Lift Pin Stroke Min Check in DB
NX-Wafer | Lift Pin Stroke Max Check in DB
NX-Wafer | Limit range
NX20 300mm | Z Linearity by Z scanner position
NX20 300mm | Full range of Z Detector during Z Scan sweep
NX20 300mm | Z Height calibration error
NX20 300mm | Z Drive calibration error
NX20 300mm | Z Linearity by 4 different sample height
NX20 300mm | Full range of Z Height during Z Scan sweep
NX20 300mm | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 300mm | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 300mm | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 300mm | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 300mm | Z Height noise
NX20 300mm | Wafer placement(Loading) error on the load port1 (X offset)
NX20 300mm | Wafer placement(Loading) error on the load port1 (Y offset)
NX20 300mm | Wafer placement(Loading) error on the load port2 (X offset)
NX20 300mm | Wafer placement(Loading) error on the load port2 (Y offset)
NX20 300mm | OpticalAxisShiftPerFocusX
NX20 300mm | OpticalAxisShiftPerFocusY
NX20 300mm | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX20 300mm | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX20 300mm | Lift pins synchronization offset check (X offset)
NX20 300mm | Lift pins synchronization offset check (Y offset)
NX20 300mm | Profiler noise @200um
NX20 300mm | Roll motion to X+ full scan
NX20 300mm | Position repeatability(10repeats, Position Based)
NX20 300mm | OPM Repeatability @5mm
NX20 300mm | OPM Repeatability @1mm
NX20 300mm | Max OPM @5mm
NX20 300mm | Max OPM @1mm
NX20 300mm | Torque Profile (10mm, 200um/s)
NX20 300mm | Torque Profile (10mm, 20um/s)
NX20 300mm | Lateral(C-D) sensitivity
NX20 300mm | Lateral(C-D) noise
NX20 300mm | Min. Vibration Transmissbility (~100Hz)
NX20 300mm | RMS of ZeroScan
NX20 300mm | LightStrengthGain
NX20 300mm | Full range of X Detector during X Scan sweep
NX20 300mm | Full range of X1 Detector during X Scan sweep
NX20 300mm | Full range of X2 Detector during X Scan sweep
NX20 300mm | Full range of Y Detector during Y Scan sweep
NX20 300mm | Full range of Y1 Detector during Y Scan sweep
NX20 300mm | Full range of Y2 Detector during Y Scan sweep
NX20 300mm | Full range of X1 Detector Raw during X Scan sweep
NX20 300mm | Full range of X2 Detector Raw during X Scan sweep
NX20 300mm | Full range of Y1 Detector Raw during Y Scan sweep
NX20 300mm | Full range of Y2 Detector Raw during Y Scan sweep
NX20 300mm | X Detector stroke calibration error (nominal scale error X)
NX20 300mm | Y Detector stroke calibration error (nominal scale error Y)
NX20 300mm | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX20 300mm | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX20 300mm | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX20 300mm | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX20 300mm | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX20 300mm | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX20 300mm | X Non-linearity error
NX20 300mm | Y Non-linearity error
NX20 300mm | X axis OPM
NX20 300mm | Y axis OPM
NX20 300mm | X Ringing(Rpv)
NX20 300mm | Y Ringing(Rpv)
NX20 300mm | X Detector noise (XY Servo on)
NX20 300mm | Y Detector noise (XY Servo on)
NX20 300mm | X1 Detector noise (XY Servo on)
NX20 300mm | Y1 Detector noise (XY Servo on)
NX20 300mm | X2 Detector noise (XY Servo on)
NX20 300mm | Y2 Detector noise (XY Servo on)
NX20 300mm | RMS error after Stage Calibration
NX20 300mm | Image offset after moving from -X to +X (X scan)
NX20 300mm | Image offset after moving from -Y to +Y (Y scan)
NX20 300mm | Image offset after moving from -X to +X (Y scan)
NX20 300mm | Image offset after moving from -Y to +Y (X scan)
NX20 300mm | Image offset after moving from +X to -X (X scan)
NX20 300mm | Image offset after moving from +Y to -Y (Y scan)
NX20 300mm | Image offset after moving from +X to -X (Y scan)
NX20 300mm | Image offset after moving from +Y to -Y (X scan)
NX20 300mm | XY Stage positioning X offset (Vision Pattern Matching)
NX20 300mm | XY Stage positioning Y offset (Vision Pattern Matching)
NX20 300mm | XY Stage positioning X 1σ (Vision Pattern Matching)
NX20 300mm | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX20 300mm | XY Stage positioning X offset (In-Die Align)
NX20 300mm | XY Stage positioning Y offset (In-Die Align)
NX20 300mm | XY Stage positioning X 1σ (In-Die Align)
NX20 300mm | XY Stage positioning Y 1σ (In-Die Align)
NX20 300mm | XY Stage positioning X offset (LLC Translation)
NX20 300mm | XY Stage positioning Y offset (LLC Translation)
NX20 300mm | XY Stage positioning X 1σ (LLC Translation)
NX20 300mm | XY Stage positioning Y 1σ (LLC Translation)
NX20 300mm | XY Stage positioning X offset (Global Align)
NX20 300mm | XY Stage positioning Y offset (Global Align)
NX20 300mm | XY Stage positioning X 1σ (Global Align)
NX20 300mm | XY Stage positioning Y 1σ (Global Align)
NX20 300mm | Z Drift
NX20 300mm | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | Cycle time for 15-point measurement
NX20 300mm | Short-term stability using Step Height Sample (3σ)
NX20 300mm | Roughness Repeatability 1x1um(RSD%)
NX20 300mm | Roughness Repeatability 1x1um(1σ)
NX20 300mm | Roughness Repeatability 80x80um(RSD%)
NX20 300mm | Roughness Repeatability 80x80um(1σ)
NX20 300mm | Z Height calibrtaion error
NX20 300mm | Z Drive calibrtaion error
NX20 300mm | Z Linearity by 3 different sample height
NX20 300mm | Full range in voltage during Z Scan sweep
NX20 300mm | Full range in um during Z Scan sweep
NX20 300mm | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 300mm | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 300mm | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 300mm | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 300mm | Lateral(C-D) sensitivity
NX20 300mm | Lateral(C-D) noise
NX20 300mm | A-B noise
NX20 300mm | NCM Amplitude noise
NX20 300mm | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX20 300mm | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX20 300mm | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX20 300mm | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(0,120000)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(120000,0)
NX20 300mm | RMS of Z Height of Wafer ZeroScan (0,0)
NX20 300mm | Max Wafer Zeroscan-Min Wafer Zeroscan
NX20 300mm | XY Orthogonality
NX20 300mm | MicrometerPerAdc of Z Height (Z Height calibration)
NX20 300mm | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX20 300mm | Z linearity by sample height
NX20 300mm | 1) Z Scanner position at +3μm
NX20 300mm | 3) Z Scanner position at -3μm
NX20 300mm | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX20 300mm | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX20 300mm | X Scan Stroke
NX20 300mm | Y Scan Stroke
NX20 300mm | Full range of X Detector Raw during X Scan sweep
NX20 300mm | Full range of Y Detector Raw during Y Scan sweep
NX20 300mm | X Detector noise (XY Servo off)
NX20 300mm | Y Detector noise (XY Servo off)
NX20 300mm | X axis OPM
NX20 300mm | Y axis OPM
NX20 300mm | X Non-linearity
NX20 300mm | Y Non-linearity
NX20 300mm | OPM Repeatability @25mm
NX20 300mm | OPM Repeatability @10mm
NX20 300mm | Max OPM @25mm
NX20 300mm | Max OPM @10mm
NX20 300mm | Torque Profile (50mm, 200um/s)
NX20 300mm | Torque Profile (50mm, 20um/s)
NX20 300mm | SLD current limit check
NX20 300mm | SLD beam power at the end of the lens
NX20 300mm | 2) of 20nm
NX20 300mm | X Detector noise (XYServo off)
NX20 300mm | Y Detector noise (XYServo off)
NX20 300mm | X axis OPM raw Rpv
NX20 300mm | Y axis OPM raw Rpv
NX20 300mm | StageSafetyXMax
NX20 300mm | StageSafetyXMin
NX20 300mm | StageSafetyY
NX20 300mm | StageSafetyYMax
NX20 300mm | StageSafetyYMin
NX20 300mm | Check X axis upper limit sensor function
NX20 300mm | Check X axis lower limit sensor function
NX20 300mm | Check Y axis upper limit sensor function
NX20 300mm | Check Y axis lower limit sensor function
NX20 300mm | X axis Angle Range
NX20 300mm | Short-term stability w/ 450nm Step Height Sample (3σ)
NX20 300mm | X detector noise (XY Servo on)
NX20 300mm | Y detector noise (XY Servo on)
NX20 300mm | Full range in um during X Scan sweep (XY Servo off)
NX20 300mm | Full range in um during Y Scan sweep (XY Servo off)
NX20 300mm | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX20 300mm | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX20 300mm | X axis, Repeatability @1st position
NX20 300mm | X axis, Repeatability @ 2nd position
NX20 300mm | X axis, Repeatability @ 3rd position
NX20 300mm | X axis, Repeatability @ 4th position
NX20 300mm | X axis, Accuracy @1st position
NX20 300mm | X axis, Accuracy @ 2nd position
NX20 300mm | X axis, Accuracy @ 3rd position
NX20 300mm | X axis, Accuracy @ 4th position
NX20 300mm | Y axis, Repeatability @1st position
NX20 300mm | Y axis, Repeatability @ 2nd position
NX20 300mm | Y axis, Repeatability @ 3rd position
NX20 300mm | Y axis, Repeatability @ 4th position
NX20 300mm | Y axis, Accuracy @1st position
NX20 300mm | Y axis, Accuracy @ 2nd position
NX20 300mm | Y axis, Accuracy @ 3rd position
NX20 300mm | Y axis, Accuracy @ 4th position
NX20 300mm | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX20 300mm | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX20 300mm | Z-XY Orthogonality (X direction-Left)
NX20 300mm | Z-XY Orthogonality (X direction-Right)
NX20 300mm | Z-XY Orthogonality (Y direction-Left)
NX20 300mm | Z-XY Orthogonality (Y direction-Right)
NX20 300mm | X-Scan Repeatability (10 repeats, RSD%)
NX20 300mm | Y-Scan Repeatability (10 repeats, RSD%)
NX20 300mm | Mechanical Align Accuracy
NX20 300mm | Check the Multi Peattern Upload
NX20 300mm | X Detector noise (XY servo off)
NX20 300mm | Y Detector noise (XY servo off)
NX20 300mm | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 300mm | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 300mm | Z Detector Max
NX20 300mm | Z Detector Min
NX20 300mm | WLI_Scale(Z Calibration)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(0,90000)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(90000,0)
NX20 300mm | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | Short-term stability w/ 180nm Step Height Sample (3σ)
NX20 300mm | Laser intensity with AC160TS cantilever
NX20 300mm | Check if the driving % changes proportionally as the NCM amplitude changes
NX20 300mm | NCM driving % at an amplitude of 1 nm
NX20 300mm | NCM driving % at an amplitude of 20 nm
NX20 300mm | NCM driving % at an amplitude of 100 nm
NX20 300mm | - 940nm Step Height Sample (10 Repeats)
NX20 300mm | - 200~500nm Step Height Sample (10 Repeats)
NX20 300mm | - 50~200nm Step Height Sample (10 Repeats)
NX20 300mm | - Under 50nm Step Height Sample (10 Repeats)
NX20 300mm | A-B after Auto PSPD Alignment
NX20 300mm | NCM driving% with AC160TS cantilever
NX20 300mm | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 300mm | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 300mm | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 300mm | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 300mm | X Detector stroke calibration error (nominal scale error X)
NX20 300mm | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | Resonance Frequency in DB
NX20 300mm | Servo Cut-off Frequency in DB
NX20 300mm | Check chuck surface no damage
NX20 300mm | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX20 300mm | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX20 300mm | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX20 300mm | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX20 300mm | X1 Detector noise (XY servo off)
NX20 300mm | X2 Detector noise (XY servo off)
NX20 300mm | Y1 Detector noise (XY servo off)
NX20 300mm | Y2 Detector noise (XY servo off)
NX20 300mm | X Detector noise (XY servo on)
NX20 300mm | X1 Detector noise (XY servo on)
NX20 300mm | X2 Detector noise (XY servo on)
NX20 300mm | Y Detector noise (XY servo on)
NX20 300mm | Y1 Detector noise (XY servo on)
NX20 300mm | Y2 Detector noise (XY servo on)
NX20 300mm | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 300mm | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 300mm | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 300mm | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 300mm | Z Detector offset
NX20 300mm | Check the full stroke image(servo on)
NX20 300mm | ResonanceFrequencyHz
NX20 300mm | ServoCutoffFrequencyHz
NX20 300mm | FullPixelHeight
NX20 300mm | Serial Port#
NX20 300mm | Off-axis Camera X Position(ΔX)
NX20 300mm | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX20 300mm | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX20 300mm | Z Scanner position at -3μm
NX20 300mm | Z Scanner position at +3μm
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 300mm | Scan offset accuracy (Precision)
NX20 300mm | Aligned stage accuracy (X offset-Point Align1)
NX20 300mm | Aligned stage accuracy (Y offset-Point Align1)
NX20 300mm | Aligned stage accuracy (X offset-Point Align2)
NX20 300mm | Aligned stage accuracy (Y offset-Point Align2)
NX20 300mm | Aligned stage accuracy (X offset-Point Align3)
NX20 300mm | Aligned stage accuracy (Y offset-Point Align3)
NX20 300mm | Aligned stage accuracy (X offset)
NX20 300mm | Aligned stage accuracy (Y offset)
NX20 300mm | Imag offset after EFEM loading (X scan)
NX20 300mm | Imag offset after EFEM loading (Y scan)
NX20 300mm | Imag offset after Lock stage (X scan)
NX20 300mm | Imag offset after Lock stage (Y scan)
NX20 300mm | Mask Roughness Repeatability (10Repeats, RSD%)
NX20 300mm | Mask Roughness Repeatability (10Repeats, σ )
NX20 300mm | Mask exchange time
NX20 300mm | WLI_Scale(C Calibration)
NX20 300mm | WLI_Scale(B Calibration)
NX20 300mm | WLI_Scale(A Calibration)
NX20 300mm | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 300mm | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 300mm | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX20 300mm | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX20 300mm | Port 1 lens(to OnAxis) Y Offset
NX20 300mm | Port 1 lens(to OnAxis) X Offset
NX20 300mm | P Gain
NX20 300mm | I Gain
NX20 300mm | D Gain
NX20 300mm | Check the Port 2 lens light strength
NX20 300mm | Check the Port 1 lens light strength
NX20 300mm | Check the focus deviation between the lenses(Δf)
NX20 300mm | Check the 25mm Moving – Y Axis Offset Check
NX20 300mm | C Scanner Detector Min
NX20 300mm | C Scanner Detector Max
NX20 300mm | B Scanner Detector Min
NX20 300mm | B Scanner Detector Max
NX20 300mm | Angle align between Port 2 lens and VLSI SHS: Y
NX20 300mm | Angle align between Port 2 lens and VLSI SHS: X
NX20 300mm | Angle align between Port 2 lens and Sample chuck: Y
NX20 300mm | Angle align between Port 2 lens and Sample chuck: X
NX20 300mm | Angle align between Port 1 lens and Sample chuck: Y
NX20 300mm | Angle align between Port 1 lens and Sample chuck: X
NX20 300mm | A Scanner Detector Min
NX20 300mm | A Scanner Detector Max
NX20 300mm | Y2 Detector noise (XY Servo off)
NX20 300mm | X2 Detector noise (XY Servo off)
NX20 300mm | Y axis OPM (-52000,-52000)
NX20 300mm | Y axis OPM (-52000,52000)
NX20 300mm | Y axis OPM (52000,-52000)
NX20 300mm | Y axis OPM (52000,52000)
NX20 300mm | Y axis OPM (0,0)
NX20 300mm | XY Orthogonality (-52000,-52000)
NX20 300mm | XY Orthogonality (-52000,52000)
NX20 300mm | XY Orthogonality (52000,-52000)
NX20 300mm | XY Orthogonality (52000,52000)
NX20 300mm | XY Orthogonality (0,0)
NX20 300mm | X axis OPM (-52000,-52000)
NX20 300mm | X axis OPM (-52000,52000)
NX20 300mm | X axis OPM (52000,-52000)
NX20 300mm | X axis OPM (52000,52000)
NX20 300mm | X axis OPM (0,0)
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(60000,0)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX20 300mm | RMS of Z Height of Wafer Zeroscan(0,60000)
NX20 300mm | XY stage calibration
NX20 300mm | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX20 300mm | LD beam power at the end of the lens
NX20 300mm | Y Non-lineairty
NX20 300mm | X Non-lineairty
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX20 300mm | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX20 300mm | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX20 300mm | UpperLimit
NX20 300mm | Under 50nm Step Height Sample (10 Repeats)
NX20 300mm | Overlay Camera X Position(ΔX)
NX20 300mm | LowerLimit
NX20 300mm | Initial Z
NX20 300mm | 940nm Step Height Sample (10 Repeats)
NX20 300mm | 50~200nm Step Height Sample (10 Repeats)
NX20 300mm | RMS of Z Height of C122 Zeroscan(0,-90000)
NX20 300mm | StrokeYMin
NX20 300mm | StrokeYMax
NX20 300mm | StrokeY
NX20 300mm | StrokeXMin
NX20 300mm | StrokeXMax
NX20 300mm | StrokeX
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX20 300mm | Z Drift (Short Traveling)
NX20 300mm | 200~500nm Step Height Sample (10 Repeats)
NX20 300mm | XY orthogonality
NX20 300mm | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX20 300mm | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX20 300mm | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX20 300mm | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX20 300mm | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX20 300mm | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX20 300mm | Y1 Detector noise (XY Servo off)
NX20 300mm | X1 Detector noise (XY Servo off)
NX20 300mm | Maximum Y Scan range (Full Stroke)
NX20 300mm | Maximum X Scan range (Full Stroke)
NX20 300mm | Full range of Y2 Detector raw
NX20 300mm | Full range of Y1 Detector raw
NX20 300mm | Full range of X2 Detector raw
NX20 300mm | Full range of X1 Detector raw
NX20 300mm | Z Linearity by Z Scanner position
NX20 300mm | -Y Direction software Limit
NX20 300mm | -X Direction software Limit
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 300mm | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 300mm | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX20 300mm | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX20 300mm | %GRR w/ roughness of non-patterened wafer
NX20 300mm | Z Scanner position at -6μm
NX20 300mm | Z Scanner position at +6μm
NX20 300mm | Approach position at -5μm
NX20 300mm | Approach position at +5μm
NX20 300mm | * 1800nm Step Height Sample (10 Repeats)
NX20 300mm | Full range of Z Detector raw
NX20 300mm | Approach position at -3μm
NX20 300mm | Approach position at +3μm
NX20 300mm | Under 50nm Step Height Sample (10 Repeats)
NX20 300mm | 50~200nm Step Height Sample (10 Repeats)
NX20 300mm | 200~500nm Step Height Sample (10 Repeats)
NX20 300mm | PosMin
NX20 300mm | PosMax
NX20 300mm | MaxLength
NX20 300mm | FullRecoveryProfileLength
NX20 300mm | Full range of Y Detector raw
NX20 300mm | Full range of X Detector raw
NX20 300mm | EFM Phase : For the positive bias
NX20 300mm | EFM Phase : For the negative bias
NX20 300mm | Z Linearity by 5 different sample height
NX20 300mm | Laser intensity with AC160 cantilever
NX20 300mm | A－B difference during 3 times run test
NX20 300mm | A＋B difference during 3 times run test
NX20 300mm | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX20 300mm | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX20 300mm | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX20 300mm | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX20 300mm | RMS of Z Drive of Wafer ZeroScan (0,0)
NX20 300mm | Allowable max warpage, w/ SiC sample chuck
NX20 300mm | RMS of Z Height of Wafer Zeroscan (Top)
NX20 300mm | RMS of Z Height of Wafer Zeroscan (Right)
NX20 300mm | RMS of Z Height of Wafer Zeroscan (Left)
NX20 300mm | RMS of Z Height of Wafer ZeroScan (Center)
NX20 300mm | RMS of Z Height of Wafer Zeroscan (Bottom)
NX20 300mm | PosMin
NX20 300mm | PosMax
NX20 300mm | MaxLength
NX20 300mm | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX20 300mm | Check sample surface mapping
NX20 300mm | Z Direction, Accuracy Test (Bias not significant)
NX20 300mm | Y Direction, Accuracy Test (Bias not significant)
NX20 300mm | X Direction, Accuracy Test (Bias not significant)
NX20 300mm | Roughness Reproducibility Test (SR Panel)
NX20 300mm | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX20 300mm | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX20 300mm | Roughness Reproducibility Test (SiNx Panel)
NX20 300mm | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX20 300mm | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX20 300mm | Roughness Reproducibility Test (ABF Panel)
NX20 300mm | Roughness Repeatability Test (SR Panel)
NX20 300mm | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX20 300mm | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX20 300mm | Roughness Repeatability Test (SiNx Panel)
NX20 300mm | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX20 300mm | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX20 300mm | Roughness Repeatability Test (ABF Panel)
NX20 300mm | Lift Pin Stroke Min Check in DB
NX20 300mm | Lift Pin Stroke Max Check in DB
NX20 300mm | Limit range
NX-PSS | Z Linearity by Z scanner position
NX-PSS | Full range of Z Detector during Z Scan sweep
NX-PSS | Z Height calibration error
NX-PSS | Z Drive calibration error
NX-PSS | Z Linearity by 4 different sample height
NX-PSS | Full range of Z Height during Z Scan sweep
NX-PSS | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-PSS | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-PSS | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-PSS | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-PSS | Z Height noise
NX-PSS | Wafer placement(Loading) error on the load port1 (X offset)
NX-PSS | Wafer placement(Loading) error on the load port1 (Y offset)
NX-PSS | Wafer placement(Loading) error on the load port2 (X offset)
NX-PSS | Wafer placement(Loading) error on the load port2 (Y offset)
NX-PSS | OpticalAxisShiftPerFocusX
NX-PSS | OpticalAxisShiftPerFocusY
NX-PSS | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-PSS | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-PSS | Lift pins synchronization offset check (X offset)
NX-PSS | Lift pins synchronization offset check (Y offset)
NX-PSS | Profiler noise @200um
NX-PSS | Roll motion to X+ full scan
NX-PSS | Position repeatability(10repeats, Position Based)
NX-PSS | OPM Repeatability @5mm
NX-PSS | OPM Repeatability @1mm
NX-PSS | Max OPM @5mm
NX-PSS | Max OPM @1mm
NX-PSS | Torque Profile (10mm, 200um/s)
NX-PSS | Torque Profile (10mm, 20um/s)
NX-PSS | Lateral(C-D) sensitivity
NX-PSS | Lateral(C-D) noise
NX-PSS | Min. Vibration Transmissbility (~100Hz)
NX-PSS | RMS of ZeroScan
NX-PSS | LightStrengthGain
NX-PSS | Full range of X Detector during X Scan sweep
NX-PSS | Full range of X1 Detector during X Scan sweep
NX-PSS | Full range of X2 Detector during X Scan sweep
NX-PSS | Full range of Y Detector during Y Scan sweep
NX-PSS | Full range of Y1 Detector during Y Scan sweep
NX-PSS | Full range of Y2 Detector during Y Scan sweep
NX-PSS | Full range of X1 Detector Raw during X Scan sweep
NX-PSS | Full range of X2 Detector Raw during X Scan sweep
NX-PSS | Full range of Y1 Detector Raw during Y Scan sweep
NX-PSS | Full range of Y2 Detector Raw during Y Scan sweep
NX-PSS | X Detector stroke calibration error (nominal scale error X)
NX-PSS | Y Detector stroke calibration error (nominal scale error Y)
NX-PSS | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-PSS | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-PSS | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-PSS | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-PSS | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-PSS | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-PSS | X Non-linearity error
NX-PSS | Y Non-linearity error
NX-PSS | X axis OPM
NX-PSS | Y axis OPM
NX-PSS | X Ringing(Rpv)
NX-PSS | Y Ringing(Rpv)
NX-PSS | X Detector noise (XY Servo on)
NX-PSS | Y Detector noise (XY Servo on)
NX-PSS | X1 Detector noise (XY Servo on)
NX-PSS | Y1 Detector noise (XY Servo on)
NX-PSS | X2 Detector noise (XY Servo on)
NX-PSS | Y2 Detector noise (XY Servo on)
NX-PSS | RMS error after Stage Calibration
NX-PSS | Image offset after moving from -X to +X (X scan)
NX-PSS | Image offset after moving from -Y to +Y (Y scan)
NX-PSS | Image offset after moving from -X to +X (Y scan)
NX-PSS | Image offset after moving from -Y to +Y (X scan)
NX-PSS | Image offset after moving from +X to -X (X scan)
NX-PSS | Image offset after moving from +Y to -Y (Y scan)
NX-PSS | Image offset after moving from +X to -X (Y scan)
NX-PSS | Image offset after moving from +Y to -Y (X scan)
NX-PSS | XY Stage positioning X offset (Vision Pattern Matching)
NX-PSS | XY Stage positioning Y offset (Vision Pattern Matching)
NX-PSS | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-PSS | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-PSS | XY Stage positioning X offset (In-Die Align)
NX-PSS | XY Stage positioning Y offset (In-Die Align)
NX-PSS | XY Stage positioning X 1σ (In-Die Align)
NX-PSS | XY Stage positioning Y 1σ (In-Die Align)
NX-PSS | XY Stage positioning X offset (LLC Translation)
NX-PSS | XY Stage positioning Y offset (LLC Translation)
NX-PSS | XY Stage positioning X 1σ (LLC Translation)
NX-PSS | XY Stage positioning Y 1σ (LLC Translation)
NX-PSS | XY Stage positioning X offset (Global Align)
NX-PSS | XY Stage positioning Y offset (Global Align)
NX-PSS | XY Stage positioning X 1σ (Global Align)
NX-PSS | XY Stage positioning Y 1σ (Global Align)
NX-PSS | Z Drift
NX-PSS | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | Cycle time for 15-point measurement
NX-PSS | Short-term stability using Step Height Sample (3σ)
NX-PSS | Roughness Repeatability 1x1um(RSD%)
NX-PSS | Roughness Repeatability 1x1um(1σ)
NX-PSS | Roughness Repeatability 80x80um(RSD%)
NX-PSS | Roughness Repeatability 80x80um(1σ)
NX-PSS | Z Height calibrtaion error
NX-PSS | Z Drive calibrtaion error
NX-PSS | Z Linearity by 3 different sample height
NX-PSS | Full range in voltage during Z Scan sweep
NX-PSS | Full range in um during Z Scan sweep
NX-PSS | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-PSS | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-PSS | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-PSS | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-PSS | Lateral(C-D) sensitivity
NX-PSS | Lateral(C-D) noise
NX-PSS | A-B noise
NX-PSS | NCM Amplitude noise
NX-PSS | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-PSS | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-PSS | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-PSS | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-PSS | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-PSS | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-PSS | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-PSS | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-PSS | RMS of Z Height of Wafer ZeroScan (0,0)
NX-PSS | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-PSS | XY Orthogonality
NX-PSS | MicrometerPerAdc of Z Height (Z Height calibration)
NX-PSS | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-PSS | Z linearity by sample height
NX-PSS | 1) Z Scanner position at +3μm
NX-PSS | 3) Z Scanner position at -3μm
NX-PSS | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-PSS | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-PSS | X Scan Stroke
NX-PSS | Y Scan Stroke
NX-PSS | Full range of X Detector Raw during X Scan sweep
NX-PSS | Full range of Y Detector Raw during Y Scan sweep
NX-PSS | X Detector noise (XY Servo off)
NX-PSS | Y Detector noise (XY Servo off)
NX-PSS | X axis OPM
NX-PSS | Y axis OPM
NX-PSS | X Non-linearity
NX-PSS | Y Non-linearity
NX-PSS | OPM Repeatability @25mm
NX-PSS | OPM Repeatability @10mm
NX-PSS | Max OPM @25mm
NX-PSS | Max OPM @10mm
NX-PSS | Torque Profile (50mm, 200um/s)
NX-PSS | Torque Profile (50mm, 20um/s)
NX-PSS | SLD current limit check
NX-PSS | SLD beam power at the end of the lens
NX-PSS | 2) of 20nm
NX-PSS | X Detector noise (XYServo off)
NX-PSS | Y Detector noise (XYServo off)
NX-PSS | X axis OPM raw Rpv
NX-PSS | Y axis OPM raw Rpv
NX-PSS | StageSafetyXMax
NX-PSS | StageSafetyXMin
NX-PSS | StageSafetyY
NX-PSS | StageSafetyYMax
NX-PSS | StageSafetyYMin
NX-PSS | Check X axis upper limit sensor function
NX-PSS | Check X axis lower limit sensor function
NX-PSS | Check Y axis upper limit sensor function
NX-PSS | Check Y axis lower limit sensor function
NX-PSS | X axis Angle Range
NX-PSS | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-PSS | X detector noise (XY Servo on)
NX-PSS | Y detector noise (XY Servo on)
NX-PSS | Full range in um during X Scan sweep (XY Servo off)
NX-PSS | Full range in um during Y Scan sweep (XY Servo off)
NX-PSS | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-PSS | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-PSS | X axis, Repeatability @1st position
NX-PSS | X axis, Repeatability @ 2nd position
NX-PSS | X axis, Repeatability @ 3rd position
NX-PSS | X axis, Repeatability @ 4th position
NX-PSS | X axis, Accuracy @1st position
NX-PSS | X axis, Accuracy @ 2nd position
NX-PSS | X axis, Accuracy @ 3rd position
NX-PSS | X axis, Accuracy @ 4th position
NX-PSS | Y axis, Repeatability @1st position
NX-PSS | Y axis, Repeatability @ 2nd position
NX-PSS | Y axis, Repeatability @ 3rd position
NX-PSS | Y axis, Repeatability @ 4th position
NX-PSS | Y axis, Accuracy @1st position
NX-PSS | Y axis, Accuracy @ 2nd position
NX-PSS | Y axis, Accuracy @ 3rd position
NX-PSS | Y axis, Accuracy @ 4th position
NX-PSS | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-PSS | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-PSS | Z-XY Orthogonality (X direction-Left)
NX-PSS | Z-XY Orthogonality (X direction-Right)
NX-PSS | Z-XY Orthogonality (Y direction-Left)
NX-PSS | Z-XY Orthogonality (Y direction-Right)
NX-PSS | X-Scan Repeatability (10 repeats, RSD%)
NX-PSS | Y-Scan Repeatability (10 repeats, RSD%)
NX-PSS | Mechanical Align Accuracy
NX-PSS | Check the Multi Peattern Upload
NX-PSS | X Detector noise (XY servo off)
NX-PSS | Y Detector noise (XY servo off)
NX-PSS | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-PSS | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-PSS | Z Detector Max
NX-PSS | Z Detector Min
NX-PSS | WLI_Scale(Z Calibration)
NX-PSS | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-PSS | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-PSS | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-PSS | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-PSS | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-PSS | Laser intensity with AC160TS cantilever
NX-PSS | Check if the driving % changes proportionally as the NCM amplitude changes
NX-PSS | NCM driving % at an amplitude of 1 nm
NX-PSS | NCM driving % at an amplitude of 20 nm
NX-PSS | NCM driving % at an amplitude of 100 nm
NX-PSS | - 940nm Step Height Sample (10 Repeats)
NX-PSS | - 200~500nm Step Height Sample (10 Repeats)
NX-PSS | - 50~200nm Step Height Sample (10 Repeats)
NX-PSS | - Under 50nm Step Height Sample (10 Repeats)
NX-PSS | A-B after Auto PSPD Alignment
NX-PSS | NCM driving% with AC160TS cantilever
NX-PSS | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-PSS | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-PSS | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-PSS | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-PSS | X Detector stroke calibration error (nominal scale error X)
NX-PSS | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | Resonance Frequency in DB
NX-PSS | Servo Cut-off Frequency in DB
NX-PSS | Check chuck surface no damage
NX-PSS | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-PSS | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-PSS | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-PSS | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-PSS | X1 Detector noise (XY servo off)
NX-PSS | X2 Detector noise (XY servo off)
NX-PSS | Y1 Detector noise (XY servo off)
NX-PSS | Y2 Detector noise (XY servo off)
NX-PSS | X Detector noise (XY servo on)
NX-PSS | X1 Detector noise (XY servo on)
NX-PSS | X2 Detector noise (XY servo on)
NX-PSS | Y Detector noise (XY servo on)
NX-PSS | Y1 Detector noise (XY servo on)
NX-PSS | Y2 Detector noise (XY servo on)
NX-PSS | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PSS | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PSS | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-PSS | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-PSS | Z Detector offset
NX-PSS | Check the full stroke image(servo on)
NX-PSS | ResonanceFrequencyHz
NX-PSS | ServoCutoffFrequencyHz
NX-PSS | FullPixelHeight
NX-PSS | Serial Port#
NX-PSS | Off-axis Camera X Position(ΔX)
NX-PSS | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-PSS | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-PSS | Z Scanner position at -3μm
NX-PSS | Z Scanner position at +3μm
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-PSS | Scan offset accuracy (Precision)
NX-PSS | Aligned stage accuracy (X offset-Point Align1)
NX-PSS | Aligned stage accuracy (Y offset-Point Align1)
NX-PSS | Aligned stage accuracy (X offset-Point Align2)
NX-PSS | Aligned stage accuracy (Y offset-Point Align2)
NX-PSS | Aligned stage accuracy (X offset-Point Align3)
NX-PSS | Aligned stage accuracy (Y offset-Point Align3)
NX-PSS | Aligned stage accuracy (X offset)
NX-PSS | Aligned stage accuracy (Y offset)
NX-PSS | Imag offset after EFEM loading (X scan)
NX-PSS | Imag offset after EFEM loading (Y scan)
NX-PSS | Imag offset after Lock stage (X scan)
NX-PSS | Imag offset after Lock stage (Y scan)
NX-PSS | Mask Roughness Repeatability (10Repeats, RSD%)
NX-PSS | Mask Roughness Repeatability (10Repeats, σ )
NX-PSS | Mask exchange time
NX-PSS | WLI_Scale(C Calibration)
NX-PSS | WLI_Scale(B Calibration)
NX-PSS | WLI_Scale(A Calibration)
NX-PSS | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-PSS | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-PSS | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-PSS | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-PSS | Port 1 lens(to OnAxis) Y Offset
NX-PSS | Port 1 lens(to OnAxis) X Offset
NX-PSS | P Gain
NX-PSS | I Gain
NX-PSS | D Gain
NX-PSS | Check the Port 2 lens light strength
NX-PSS | Check the Port 1 lens light strength
NX-PSS | Check the focus deviation between the lenses(Δf)
NX-PSS | Check the 25mm Moving – Y Axis Offset Check
NX-PSS | C Scanner Detector Min
NX-PSS | C Scanner Detector Max
NX-PSS | B Scanner Detector Min
NX-PSS | B Scanner Detector Max
NX-PSS | Angle align between Port 2 lens and VLSI SHS: Y
NX-PSS | Angle align between Port 2 lens and VLSI SHS: X
NX-PSS | Angle align between Port 2 lens and Sample chuck: Y
NX-PSS | Angle align between Port 2 lens and Sample chuck: X
NX-PSS | Angle align between Port 1 lens and Sample chuck: Y
NX-PSS | Angle align between Port 1 lens and Sample chuck: X
NX-PSS | A Scanner Detector Min
NX-PSS | A Scanner Detector Max
NX-PSS | Y2 Detector noise (XY Servo off)
NX-PSS | X2 Detector noise (XY Servo off)
NX-PSS | Y axis OPM (-52000,-52000)
NX-PSS | Y axis OPM (-52000,52000)
NX-PSS | Y axis OPM (52000,-52000)
NX-PSS | Y axis OPM (52000,52000)
NX-PSS | Y axis OPM (0,0)
NX-PSS | XY Orthogonality (-52000,-52000)
NX-PSS | XY Orthogonality (-52000,52000)
NX-PSS | XY Orthogonality (52000,-52000)
NX-PSS | XY Orthogonality (52000,52000)
NX-PSS | XY Orthogonality (0,0)
NX-PSS | X axis OPM (-52000,-52000)
NX-PSS | X axis OPM (-52000,52000)
NX-PSS | X axis OPM (52000,-52000)
NX-PSS | X axis OPM (52000,52000)
NX-PSS | X axis OPM (0,0)
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-PSS | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-PSS | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-PSS | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-PSS | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-PSS | XY stage calibration
NX-PSS | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-PSS | LD beam power at the end of the lens
NX-PSS | Y Non-lineairty
NX-PSS | X Non-lineairty
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-PSS | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-PSS | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-PSS | UpperLimit
NX-PSS | Under 50nm Step Height Sample (10 Repeats)
NX-PSS | Overlay Camera X Position(ΔX)
NX-PSS | LowerLimit
NX-PSS | Initial Z
NX-PSS | 940nm Step Height Sample (10 Repeats)
NX-PSS | 50~200nm Step Height Sample (10 Repeats)
NX-PSS | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-PSS | StrokeYMin
NX-PSS | StrokeYMax
NX-PSS | StrokeY
NX-PSS | StrokeXMin
NX-PSS | StrokeXMax
NX-PSS | StrokeX
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-PSS | Z Drift (Short Traveling)
NX-PSS | 200~500nm Step Height Sample (10 Repeats)
NX-PSS | XY orthogonality
NX-PSS | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-PSS | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-PSS | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-PSS | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-PSS | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-PSS | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-PSS | Y1 Detector noise (XY Servo off)
NX-PSS | X1 Detector noise (XY Servo off)
NX-PSS | Maximum Y Scan range (Full Stroke)
NX-PSS | Maximum X Scan range (Full Stroke)
NX-PSS | Full range of Y2 Detector raw
NX-PSS | Full range of Y1 Detector raw
NX-PSS | Full range of X2 Detector raw
NX-PSS | Full range of X1 Detector raw
NX-PSS | Z Linearity by Z Scanner position
NX-PSS | -Y Direction software Limit
NX-PSS | -X Direction software Limit
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-PSS | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-PSS | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-PSS | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-PSS | %GRR w/ roughness of non-patterened wafer
NX-PSS | Z Scanner position at -6μm
NX-PSS | Z Scanner position at +6μm
NX-PSS | Approach position at -5μm
NX-PSS | Approach position at +5μm
NX-PSS | * 1800nm Step Height Sample (10 Repeats)
NX-PSS | Full range of Z Detector raw
NX-PSS | Approach position at -3μm
NX-PSS | Approach position at +3μm
NX-PSS | Under 50nm Step Height Sample (10 Repeats)
NX-PSS | 50~200nm Step Height Sample (10 Repeats)
NX-PSS | 200~500nm Step Height Sample (10 Repeats)
NX-PSS | PosMin
NX-PSS | PosMax
NX-PSS | MaxLength
NX-PSS | FullRecoveryProfileLength
NX-PSS | Full range of Y Detector raw
NX-PSS | Full range of X Detector raw
NX-PSS | EFM Phase : For the positive bias
NX-PSS | EFM Phase : For the negative bias
NX-PSS | Z Linearity by 5 different sample height
NX-PSS | Laser intensity with AC160 cantilever
NX-PSS | A－B difference during 3 times run test
NX-PSS | A＋B difference during 3 times run test
NX-PSS | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-PSS | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-PSS | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-PSS | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-PSS | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-PSS | Allowable max warpage, w/ SiC sample chuck
NX-PSS | RMS of Z Height of Wafer Zeroscan (Top)
NX-PSS | RMS of Z Height of Wafer Zeroscan (Right)
NX-PSS | RMS of Z Height of Wafer Zeroscan (Left)
NX-PSS | RMS of Z Height of Wafer ZeroScan (Center)
NX-PSS | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-PSS | PosMin
NX-PSS | PosMax
NX-PSS | MaxLength
NX-PSS | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-PSS | Check sample surface mapping
NX-PSS | Z Direction, Accuracy Test (Bias not significant)
NX-PSS | Y Direction, Accuracy Test (Bias not significant)
NX-PSS | X Direction, Accuracy Test (Bias not significant)
NX-PSS | Roughness Reproducibility Test (SR Panel)
NX-PSS | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-PSS | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-PSS | Roughness Reproducibility Test (SiNx Panel)
NX-PSS | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-PSS | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-PSS | Roughness Reproducibility Test (ABF Panel)
NX-PSS | Roughness Repeatability Test (SR Panel)
NX-PSS | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-PSS | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-PSS | Roughness Repeatability Test (SiNx Panel)
NX-PSS | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-PSS | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-PSS | Roughness Repeatability Test (ABF Panel)
NX-PSS | Lift Pin Stroke Min Check in DB
NX-PSS | Lift Pin Stroke Max Check in DB
NX-PSS | Limit range
NX10 | Z Linearity by Z scanner position
NX10 | Full range of Z Detector during Z Scan sweep
NX10 | Z Height calibration error
NX10 | Z Drive calibration error
NX10 | Z Linearity by 4 different sample height
NX10 | Full range of Z Height during Z Scan sweep
NX10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX10 | Z Height noise
NX10 | Wafer placement(Loading) error on the load port1 (X offset)
NX10 | Wafer placement(Loading) error on the load port1 (Y offset)
NX10 | Wafer placement(Loading) error on the load port2 (X offset)
NX10 | Wafer placement(Loading) error on the load port2 (Y offset)
NX10 | OpticalAxisShiftPerFocusX
NX10 | OpticalAxisShiftPerFocusY
NX10 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX10 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX10 | Lift pins synchronization offset check (X offset)
NX10 | Lift pins synchronization offset check (Y offset)
NX10 | Profiler noise @200um
NX10 | Roll motion to X+ full scan
NX10 | Position repeatability(10repeats, Position Based)
NX10 | OPM Repeatability @5mm
NX10 | OPM Repeatability @1mm
NX10 | Max OPM @5mm
NX10 | Max OPM @1mm
NX10 | Torque Profile (10mm, 200um/s)
NX10 | Torque Profile (10mm, 20um/s)
NX10 | Lateral(C-D) sensitivity
NX10 | Lateral(C-D) noise
NX10 | Min. Vibration Transmissbility (~100Hz)
NX10 | RMS of ZeroScan
NX10 | LightStrengthGain
NX10 | Full range of X Detector during X Scan sweep
NX10 | Full range of X1 Detector during X Scan sweep
NX10 | Full range of X2 Detector during X Scan sweep
NX10 | Full range of Y Detector during Y Scan sweep
NX10 | Full range of Y1 Detector during Y Scan sweep
NX10 | Full range of Y2 Detector during Y Scan sweep
NX10 | Full range of X1 Detector Raw during X Scan sweep
NX10 | Full range of X2 Detector Raw during X Scan sweep
NX10 | Full range of Y1 Detector Raw during Y Scan sweep
NX10 | Full range of Y2 Detector Raw during Y Scan sweep
NX10 | X Detector stroke calibration error (nominal scale error X)
NX10 | Y Detector stroke calibration error (nominal scale error Y)
NX10 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX10 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX10 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX10 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX10 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX10 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX10 | X Non-linearity error
NX10 | Y Non-linearity error
NX10 | X axis OPM
NX10 | Y axis OPM
NX10 | X Ringing(Rpv)
NX10 | Y Ringing(Rpv)
NX10 | X Detector noise (XY Servo on)
NX10 | Y Detector noise (XY Servo on)
NX10 | X1 Detector noise (XY Servo on)
NX10 | Y1 Detector noise (XY Servo on)
NX10 | X2 Detector noise (XY Servo on)
NX10 | Y2 Detector noise (XY Servo on)
NX10 | RMS error after Stage Calibration
NX10 | Image offset after moving from -X to +X (X scan)
NX10 | Image offset after moving from -Y to +Y (Y scan)
NX10 | Image offset after moving from -X to +X (Y scan)
NX10 | Image offset after moving from -Y to +Y (X scan)
NX10 | Image offset after moving from +X to -X (X scan)
NX10 | Image offset after moving from +Y to -Y (Y scan)
NX10 | Image offset after moving from +X to -X (Y scan)
NX10 | Image offset after moving from +Y to -Y (X scan)
NX10 | XY Stage positioning X offset (Vision Pattern Matching)
NX10 | XY Stage positioning Y offset (Vision Pattern Matching)
NX10 | XY Stage positioning X 1σ (Vision Pattern Matching)
NX10 | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX10 | XY Stage positioning X offset (In-Die Align)
NX10 | XY Stage positioning Y offset (In-Die Align)
NX10 | XY Stage positioning X 1σ (In-Die Align)
NX10 | XY Stage positioning Y 1σ (In-Die Align)
NX10 | XY Stage positioning X offset (LLC Translation)
NX10 | XY Stage positioning Y offset (LLC Translation)
NX10 | XY Stage positioning X 1σ (LLC Translation)
NX10 | XY Stage positioning Y 1σ (LLC Translation)
NX10 | XY Stage positioning X offset (Global Align)
NX10 | XY Stage positioning Y offset (Global Align)
NX10 | XY Stage positioning X 1σ (Global Align)
NX10 | XY Stage positioning Y 1σ (Global Align)
NX10 | Z Drift
NX10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | Cycle time for 15-point measurement
NX10 | Short-term stability using Step Height Sample (3σ)
NX10 | Roughness Repeatability 1x1um(RSD%)
NX10 | Roughness Repeatability 1x1um(1σ)
NX10 | Roughness Repeatability 80x80um(RSD%)
NX10 | Roughness Repeatability 80x80um(1σ)
NX10 | Z Height calibrtaion error
NX10 | Z Drive calibrtaion error
NX10 | Z Linearity by 3 different sample height
NX10 | Full range in voltage during Z Scan sweep
NX10 | Full range in um during Z Scan sweep
NX10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX10 | Lateral(C-D) sensitivity
NX10 | Lateral(C-D) noise
NX10 | A-B noise
NX10 | NCM Amplitude noise
NX10 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX10 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX10 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX10 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX10 | RMS of Z Height of Wafer Zeroscan(0,120000)
NX10 | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX10 | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX10 | RMS of Z Height of Wafer Zeroscan(120000,0)
NX10 | RMS of Z Height of Wafer ZeroScan (0,0)
NX10 | Max Wafer Zeroscan-Min Wafer Zeroscan
NX10 | XY Orthogonality
NX10 | MicrometerPerAdc of Z Height (Z Height calibration)
NX10 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX10 | Z linearity by sample height
NX10 | 1) Z Scanner position at +3μm
NX10 | 3) Z Scanner position at -3μm
NX10 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX10 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX10 | X Scan Stroke
NX10 | Y Scan Stroke
NX10 | Full range of X Detector Raw during X Scan sweep
NX10 | Full range of Y Detector Raw during Y Scan sweep
NX10 | X Detector noise (XY Servo off)
NX10 | Y Detector noise (XY Servo off)
NX10 | X axis OPM
NX10 | Y axis OPM
NX10 | X Non-linearity
NX10 | Y Non-linearity
NX10 | OPM Repeatability @25mm
NX10 | OPM Repeatability @10mm
NX10 | Max OPM @25mm
NX10 | Max OPM @10mm
NX10 | Torque Profile (50mm, 200um/s)
NX10 | Torque Profile (50mm, 20um/s)
NX10 | SLD current limit check
NX10 | SLD beam power at the end of the lens
NX10 | 2) of 20nm
NX10 | X Detector noise (XYServo off)
NX10 | Y Detector noise (XYServo off)
NX10 | X axis OPM raw Rpv
NX10 | Y axis OPM raw Rpv
NX10 | StageSafetyXMax
NX10 | StageSafetyXMin
NX10 | StageSafetyY
NX10 | StageSafetyYMax
NX10 | StageSafetyYMin
NX10 | Check X axis upper limit sensor function
NX10 | Check X axis lower limit sensor function
NX10 | Check Y axis upper limit sensor function
NX10 | Check Y axis lower limit sensor function
NX10 | X axis Angle Range
NX10 | Short-term stability w/ 450nm Step Height Sample (3σ)
NX10 | X detector noise (XY Servo on)
NX10 | Y detector noise (XY Servo on)
NX10 | Full range in um during X Scan sweep (XY Servo off)
NX10 | Full range in um during Y Scan sweep (XY Servo off)
NX10 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX10 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX10 | X axis, Repeatability @1st position
NX10 | X axis, Repeatability @ 2nd position
NX10 | X axis, Repeatability @ 3rd position
NX10 | X axis, Repeatability @ 4th position
NX10 | X axis, Accuracy @1st position
NX10 | X axis, Accuracy @ 2nd position
NX10 | X axis, Accuracy @ 3rd position
NX10 | X axis, Accuracy @ 4th position
NX10 | Y axis, Repeatability @1st position
NX10 | Y axis, Repeatability @ 2nd position
NX10 | Y axis, Repeatability @ 3rd position
NX10 | Y axis, Repeatability @ 4th position
NX10 | Y axis, Accuracy @1st position
NX10 | Y axis, Accuracy @ 2nd position
NX10 | Y axis, Accuracy @ 3rd position
NX10 | Y axis, Accuracy @ 4th position
NX10 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX10 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX10 | Z-XY Orthogonality (X direction-Left)
NX10 | Z-XY Orthogonality (X direction-Right)
NX10 | Z-XY Orthogonality (Y direction-Left)
NX10 | Z-XY Orthogonality (Y direction-Right)
NX10 | X-Scan Repeatability (10 repeats, RSD%)
NX10 | Y-Scan Repeatability (10 repeats, RSD%)
NX10 | Mechanical Align Accuracy
NX10 | Check the Multi Peattern Upload
NX10 | X Detector noise (XY servo off)
NX10 | Y Detector noise (XY servo off)
NX10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX10 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX10 | Z Detector Max
NX10 | Z Detector Min
NX10 | WLI_Scale(Z Calibration)
NX10 | RMS of Z Height of Wafer Zeroscan(0,90000)
NX10 | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX10 | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX10 | RMS of Z Height of Wafer Zeroscan(90000,0)
NX10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | Short-term stability w/ 180nm Step Height Sample (3σ)
NX10 | Laser intensity with AC160TS cantilever
NX10 | Check if the driving % changes proportionally as the NCM amplitude changes
NX10 | NCM driving % at an amplitude of 1 nm
NX10 | NCM driving % at an amplitude of 20 nm
NX10 | NCM driving % at an amplitude of 100 nm
NX10 | - 940nm Step Height Sample (10 Repeats)
NX10 | - 200~500nm Step Height Sample (10 Repeats)
NX10 | - 50~200nm Step Height Sample (10 Repeats)
NX10 | - Under 50nm Step Height Sample (10 Repeats)
NX10 | A-B after Auto PSPD Alignment
NX10 | NCM driving% with AC160TS cantilever
NX10 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX10 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX10 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX10 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX10 | X Detector stroke calibration error (nominal scale error X)
NX10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | Resonance Frequency in DB
NX10 | Servo Cut-off Frequency in DB
NX10 | Check chuck surface no damage
NX10 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX10 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX10 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX10 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX10 | X1 Detector noise (XY servo off)
NX10 | X2 Detector noise (XY servo off)
NX10 | Y1 Detector noise (XY servo off)
NX10 | Y2 Detector noise (XY servo off)
NX10 | X Detector noise (XY servo on)
NX10 | X1 Detector noise (XY servo on)
NX10 | X2 Detector noise (XY servo on)
NX10 | Y Detector noise (XY servo on)
NX10 | Y1 Detector noise (XY servo on)
NX10 | Y2 Detector noise (XY servo on)
NX10 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX10 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX10 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX10 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX10 | Z Detector offset
NX10 | Check the full stroke image(servo on)
NX10 | ResonanceFrequencyHz
NX10 | ServoCutoffFrequencyHz
NX10 | FullPixelHeight
NX10 | Serial Port#
NX10 | Off-axis Camera X Position(ΔX)
NX10 | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX10 | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX10 | Z Scanner position at -3μm
NX10 | Z Scanner position at +3μm
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX10 | Scan offset accuracy (Precision)
NX10 | Aligned stage accuracy (X offset-Point Align1)
NX10 | Aligned stage accuracy (Y offset-Point Align1)
NX10 | Aligned stage accuracy (X offset-Point Align2)
NX10 | Aligned stage accuracy (Y offset-Point Align2)
NX10 | Aligned stage accuracy (X offset-Point Align3)
NX10 | Aligned stage accuracy (Y offset-Point Align3)
NX10 | Aligned stage accuracy (X offset)
NX10 | Aligned stage accuracy (Y offset)
NX10 | Imag offset after EFEM loading (X scan)
NX10 | Imag offset after EFEM loading (Y scan)
NX10 | Imag offset after Lock stage (X scan)
NX10 | Imag offset after Lock stage (Y scan)
NX10 | Mask Roughness Repeatability (10Repeats, RSD%)
NX10 | Mask Roughness Repeatability (10Repeats, σ )
NX10 | Mask exchange time
NX10 | WLI_Scale(C Calibration)
NX10 | WLI_Scale(B Calibration)
NX10 | WLI_Scale(A Calibration)
NX10 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX10 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX10 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX10 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX10 | Port 1 lens(to OnAxis) Y Offset
NX10 | Port 1 lens(to OnAxis) X Offset
NX10 | P Gain
NX10 | I Gain
NX10 | D Gain
NX10 | Check the Port 2 lens light strength
NX10 | Check the Port 1 lens light strength
NX10 | Check the focus deviation between the lenses(Δf)
NX10 | Check the 25mm Moving – Y Axis Offset Check
NX10 | C Scanner Detector Min
NX10 | C Scanner Detector Max
NX10 | B Scanner Detector Min
NX10 | B Scanner Detector Max
NX10 | Angle align between Port 2 lens and VLSI SHS: Y
NX10 | Angle align between Port 2 lens and VLSI SHS: X
NX10 | Angle align between Port 2 lens and Sample chuck: Y
NX10 | Angle align between Port 2 lens and Sample chuck: X
NX10 | Angle align between Port 1 lens and Sample chuck: Y
NX10 | Angle align between Port 1 lens and Sample chuck: X
NX10 | A Scanner Detector Min
NX10 | A Scanner Detector Max
NX10 | Y2 Detector noise (XY Servo off)
NX10 | X2 Detector noise (XY Servo off)
NX10 | Y axis OPM (-52000,-52000)
NX10 | Y axis OPM (-52000,52000)
NX10 | Y axis OPM (52000,-52000)
NX10 | Y axis OPM (52000,52000)
NX10 | Y axis OPM (0,0)
NX10 | XY Orthogonality (-52000,-52000)
NX10 | XY Orthogonality (-52000,52000)
NX10 | XY Orthogonality (52000,-52000)
NX10 | XY Orthogonality (52000,52000)
NX10 | XY Orthogonality (0,0)
NX10 | X axis OPM (-52000,-52000)
NX10 | X axis OPM (-52000,52000)
NX10 | X axis OPM (52000,-52000)
NX10 | X axis OPM (52000,52000)
NX10 | X axis OPM (0,0)
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX10 | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX10 | RMS of Z Height of Wafer Zeroscan(60000,0)
NX10 | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX10 | RMS of Z Height of Wafer Zeroscan(0,60000)
NX10 | XY stage calibration
NX10 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX10 | LD beam power at the end of the lens
NX10 | Y Non-lineairty
NX10 | X Non-lineairty
NX10 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX10 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX10 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX10 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX10 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX10 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX10 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX10 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX10 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX10 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX10 | UpperLimit
NX10 | Under 50nm Step Height Sample (10 Repeats)
NX10 | Overlay Camera X Position(ΔX)
NX10 | LowerLimit
NX10 | Initial Z
NX10 | 940nm Step Height Sample (10 Repeats)
NX10 | 50~200nm Step Height Sample (10 Repeats)
NX10 | RMS of Z Height of C122 Zeroscan(0,-90000)
NX10 | StrokeYMin
NX10 | StrokeYMax
NX10 | StrokeY
NX10 | StrokeXMin
NX10 | StrokeXMax
NX10 | StrokeX
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX10 | Z Drift (Short Traveling)
NX10 | 200~500nm Step Height Sample (10 Repeats)
NX10 | XY orthogonality
NX10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX10 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX10 | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX10 | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX10 | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX10 | Y1 Detector noise (XY Servo off)
NX10 | X1 Detector noise (XY Servo off)
NX10 | Maximum Y Scan range (Full Stroke)
NX10 | Maximum X Scan range (Full Stroke)
NX10 | Full range of Y2 Detector raw
NX10 | Full range of Y1 Detector raw
NX10 | Full range of X2 Detector raw
NX10 | Full range of X1 Detector raw
NX10 | Z Linearity by Z Scanner position
NX10 | -Y Direction software Limit
NX10 | -X Direction software Limit
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX10 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX10 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX10 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX10 | %GRR w/ roughness of non-patterened wafer
NX10 | Z Scanner position at -6μm
NX10 | Z Scanner position at +6μm
NX10 | Approach position at -5μm
NX10 | Approach position at +5μm
NX10 | * 1800nm Step Height Sample (10 Repeats)
NX10 | Full range of Z Detector raw
NX10 | Approach position at -3μm
NX10 | Approach position at +3μm
NX10 | Under 50nm Step Height Sample (10 Repeats)
NX10 | 50~200nm Step Height Sample (10 Repeats)
NX10 | 200~500nm Step Height Sample (10 Repeats)
NX10 | PosMin
NX10 | PosMax
NX10 | MaxLength
NX10 | FullRecoveryProfileLength
NX10 | Full range of Y Detector raw
NX10 | Full range of X Detector raw
NX10 | EFM Phase : For the positive bias
NX10 | EFM Phase : For the negative bias
NX10 | Z Linearity by 5 different sample height
NX10 | Laser intensity with AC160 cantilever
NX10 | A－B difference during 3 times run test
NX10 | A＋B difference during 3 times run test
NX10 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX10 | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX10 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX10 | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX10 | RMS of Z Drive of Wafer ZeroScan (0,0)
NX10 | Allowable max warpage, w/ SiC sample chuck
NX10 | RMS of Z Height of Wafer Zeroscan (Top)
NX10 | RMS of Z Height of Wafer Zeroscan (Right)
NX10 | RMS of Z Height of Wafer Zeroscan (Left)
NX10 | RMS of Z Height of Wafer ZeroScan (Center)
NX10 | RMS of Z Height of Wafer Zeroscan (Bottom)
NX10 | PosMin
NX10 | PosMax
NX10 | MaxLength
NX10 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX10 | Check sample surface mapping
NX10 | Z Direction, Accuracy Test (Bias not significant)
NX10 | Y Direction, Accuracy Test (Bias not significant)
NX10 | X Direction, Accuracy Test (Bias not significant)
NX10 | Roughness Reproducibility Test (SR Panel)
NX10 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX10 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX10 | Roughness Reproducibility Test (SiNx Panel)
NX10 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX10 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX10 | Roughness Reproducibility Test (ABF Panel)
NX10 | Roughness Repeatability Test (SR Panel)
NX10 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX10 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX10 | Roughness Repeatability Test (SiNx Panel)
NX10 | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX10 | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX10 | Roughness Repeatability Test (ABF Panel)
NX10 | Lift Pin Stroke Min Check in DB
NX10 | Lift Pin Stroke Max Check in DB
NX10 | Limit range
FX40 | Z Linearity by Z scanner position
FX40 | Full range of Z Detector during Z Scan sweep
FX40 | Z Height calibration error
FX40 | Z Drive calibration error
FX40 | Z Linearity by 4 different sample height
FX40 | Full range of Z Height during Z Scan sweep
FX40 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
FX40 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
FX40 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
FX40 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
FX40 | Z Height noise
FX40 | Wafer placement(Loading) error on the load port1 (X offset)
FX40 | Wafer placement(Loading) error on the load port1 (Y offset)
FX40 | Wafer placement(Loading) error on the load port2 (X offset)
FX40 | Wafer placement(Loading) error on the load port2 (Y offset)
FX40 | OpticalAxisShiftPerFocusX
FX40 | OpticalAxisShiftPerFocusY
FX40 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
FX40 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
FX40 | Lift pins synchronization offset check (X offset)
FX40 | Lift pins synchronization offset check (Y offset)
FX40 | Profiler noise @200um
FX40 | Roll motion to X+ full scan
FX40 | Position repeatability(10repeats, Position Based)
FX40 | OPM Repeatability @5mm
FX40 | OPM Repeatability @1mm
FX40 | Max OPM @5mm
FX40 | Max OPM @1mm
FX40 | Torque Profile (10mm, 200um/s)
FX40 | Torque Profile (10mm, 20um/s)
FX40 | Lateral(C-D) sensitivity
FX40 | Lateral(C-D) noise
FX40 | Min. Vibration Transmissbility (~100Hz)
FX40 | RMS of ZeroScan
FX40 | LightStrengthGain
FX40 | Full range of X Detector during X Scan sweep
FX40 | Full range of X1 Detector during X Scan sweep
FX40 | Full range of X2 Detector during X Scan sweep
FX40 | Full range of Y Detector during Y Scan sweep
FX40 | Full range of Y1 Detector during Y Scan sweep
FX40 | Full range of Y2 Detector during Y Scan sweep
FX40 | Full range of X1 Detector Raw during X Scan sweep
FX40 | Full range of X2 Detector Raw during X Scan sweep
FX40 | Full range of Y1 Detector Raw during Y Scan sweep
FX40 | Full range of Y2 Detector Raw during Y Scan sweep
FX40 | X Detector stroke calibration error (nominal scale error X)
FX40 | Y Detector stroke calibration error (nominal scale error Y)
FX40 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
FX40 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
FX40 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
FX40 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
FX40 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
FX40 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
FX40 | X Non-linearity error
FX40 | Y Non-linearity error
FX40 | X axis OPM
FX40 | Y axis OPM
FX40 | X Ringing(Rpv)
FX40 | Y Ringing(Rpv)
FX40 | X Detector noise (XY Servo on)
FX40 | Y Detector noise (XY Servo on)
FX40 | X1 Detector noise (XY Servo on)
FX40 | Y1 Detector noise (XY Servo on)
FX40 | X2 Detector noise (XY Servo on)
FX40 | Y2 Detector noise (XY Servo on)
FX40 | RMS error after Stage Calibration
FX40 | Image offset after moving from -X to +X (X scan)
FX40 | Image offset after moving from -Y to +Y (Y scan)
FX40 | Image offset after moving from -X to +X (Y scan)
FX40 | Image offset after moving from -Y to +Y (X scan)
FX40 | Image offset after moving from +X to -X (X scan)
FX40 | Image offset after moving from +Y to -Y (Y scan)
FX40 | Image offset after moving from +X to -X (Y scan)
FX40 | Image offset after moving from +Y to -Y (X scan)
FX40 | XY Stage positioning X offset (Vision Pattern Matching)
FX40 | XY Stage positioning Y offset (Vision Pattern Matching)
FX40 | XY Stage positioning X 1σ (Vision Pattern Matching)
FX40 | XY Stage positioning Y 1σ (Vision Pattern Matching)
FX40 | XY Stage positioning X offset (In-Die Align)
FX40 | XY Stage positioning Y offset (In-Die Align)
FX40 | XY Stage positioning X 1σ (In-Die Align)
FX40 | XY Stage positioning Y 1σ (In-Die Align)
FX40 | XY Stage positioning X offset (LLC Translation)
FX40 | XY Stage positioning Y offset (LLC Translation)
FX40 | XY Stage positioning X 1σ (LLC Translation)
FX40 | XY Stage positioning Y 1σ (LLC Translation)
FX40 | XY Stage positioning X offset (Global Align)
FX40 | XY Stage positioning Y offset (Global Align)
FX40 | XY Stage positioning X 1σ (Global Align)
FX40 | XY Stage positioning Y 1σ (Global Align)
FX40 | Z Drift
FX40 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | Cycle time for 15-point measurement
FX40 | Short-term stability using Step Height Sample (3σ)
FX40 | Roughness Repeatability 1x1um(RSD%)
FX40 | Roughness Repeatability 1x1um(1σ)
FX40 | Roughness Repeatability 80x80um(RSD%)
FX40 | Roughness Repeatability 80x80um(1σ)
FX40 | Z Height calibrtaion error
FX40 | Z Drive calibrtaion error
FX40 | Z Linearity by 3 different sample height
FX40 | Full range in voltage during Z Scan sweep
FX40 | Full range in um during Z Scan sweep
FX40 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
FX40 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
FX40 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
FX40 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
FX40 | Lateral(C-D) sensitivity
FX40 | Lateral(C-D) noise
FX40 | A-B noise
FX40 | NCM Amplitude noise
FX40 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
FX40 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
FX40 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
FX40 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
FX40 | RMS of Z Height of Wafer Zeroscan(0,120000)
FX40 | RMS of Z Height of Wafer Zeroscan(-120000,0)
FX40 | RMS of Z Height of Wafer Zeroscan(0,-120000)
FX40 | RMS of Z Height of Wafer Zeroscan(120000,0)
FX40 | RMS of Z Height of Wafer ZeroScan (0,0)
FX40 | Max Wafer Zeroscan-Min Wafer Zeroscan
FX40 | XY Orthogonality
FX40 | MicrometerPerAdc of Z Height (Z Height calibration)
FX40 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
FX40 | Z linearity by sample height
FX40 | 1) Z Scanner position at +3μm
FX40 | 3) Z Scanner position at -3μm
FX40 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
FX40 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
FX40 | X Scan Stroke
FX40 | Y Scan Stroke
FX40 | Full range of X Detector Raw during X Scan sweep
FX40 | Full range of Y Detector Raw during Y Scan sweep
FX40 | X Detector noise (XY Servo off)
FX40 | Y Detector noise (XY Servo off)
FX40 | X axis OPM
FX40 | Y axis OPM
FX40 | X Non-linearity
FX40 | Y Non-linearity
FX40 | OPM Repeatability @25mm
FX40 | OPM Repeatability @10mm
FX40 | Max OPM @25mm
FX40 | Max OPM @10mm
FX40 | Torque Profile (50mm, 200um/s)
FX40 | Torque Profile (50mm, 20um/s)
FX40 | SLD current limit check
FX40 | SLD beam power at the end of the lens
FX40 | 2) of 20nm
FX40 | X Detector noise (XYServo off)
FX40 | Y Detector noise (XYServo off)
FX40 | X axis OPM raw Rpv
FX40 | Y axis OPM raw Rpv
FX40 | StageSafetyXMax
FX40 | StageSafetyXMin
FX40 | StageSafetyY
FX40 | StageSafetyYMax
FX40 | StageSafetyYMin
FX40 | Check X axis upper limit sensor function
FX40 | Check X axis lower limit sensor function
FX40 | Check Y axis upper limit sensor function
FX40 | Check Y axis lower limit sensor function
FX40 | X axis Angle Range
FX40 | Short-term stability w/ 450nm Step Height Sample (3σ)
FX40 | X detector noise (XY Servo on)
FX40 | Y detector noise (XY Servo on)
FX40 | Full range in um during X Scan sweep (XY Servo off)
FX40 | Full range in um during Y Scan sweep (XY Servo off)
FX40 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
FX40 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
FX40 | X axis, Repeatability @1st position
FX40 | X axis, Repeatability @ 2nd position
FX40 | X axis, Repeatability @ 3rd position
FX40 | X axis, Repeatability @ 4th position
FX40 | X axis, Accuracy @1st position
FX40 | X axis, Accuracy @ 2nd position
FX40 | X axis, Accuracy @ 3rd position
FX40 | X axis, Accuracy @ 4th position
FX40 | Y axis, Repeatability @1st position
FX40 | Y axis, Repeatability @ 2nd position
FX40 | Y axis, Repeatability @ 3rd position
FX40 | Y axis, Repeatability @ 4th position
FX40 | Y axis, Accuracy @1st position
FX40 | Y axis, Accuracy @ 2nd position
FX40 | Y axis, Accuracy @ 3rd position
FX40 | Y axis, Accuracy @ 4th position
FX40 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
FX40 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
FX40 | Z-XY Orthogonality (X direction-Left)
FX40 | Z-XY Orthogonality (X direction-Right)
FX40 | Z-XY Orthogonality (Y direction-Left)
FX40 | Z-XY Orthogonality (Y direction-Right)
FX40 | X-Scan Repeatability (10 repeats, RSD%)
FX40 | Y-Scan Repeatability (10 repeats, RSD%)
FX40 | Mechanical Align Accuracy
FX40 | Check the Multi Peattern Upload
FX40 | X Detector noise (XY servo off)
FX40 | Y Detector noise (XY servo off)
FX40 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
FX40 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
FX40 | Z Detector Max
FX40 | Z Detector Min
FX40 | WLI_Scale(Z Calibration)
FX40 | RMS of Z Height of Wafer Zeroscan(0,90000)
FX40 | RMS of Z Height of Wafer Zeroscan(-90000,0)
FX40 | RMS of Z Height of Wafer Zeroscan(0,-90000)
FX40 | RMS of Z Height of Wafer Zeroscan(90000,0)
FX40 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | Short-term stability w/ 180nm Step Height Sample (3σ)
FX40 | Laser intensity with AC160TS cantilever
FX40 | Check if the driving % changes proportionally as the NCM amplitude changes
FX40 | NCM driving % at an amplitude of 1 nm
FX40 | NCM driving % at an amplitude of 20 nm
FX40 | NCM driving % at an amplitude of 100 nm
FX40 | - 940nm Step Height Sample (10 Repeats)
FX40 | - 200~500nm Step Height Sample (10 Repeats)
FX40 | - 50~200nm Step Height Sample (10 Repeats)
FX40 | - Under 50nm Step Height Sample (10 Repeats)
FX40 | A-B after Auto PSPD Alignment
FX40 | NCM driving% with AC160TS cantilever
FX40 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
FX40 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
FX40 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
FX40 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
FX40 | X Detector stroke calibration error (nominal scale error X)
FX40 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | Resonance Frequency in DB
FX40 | Servo Cut-off Frequency in DB
FX40 | Check chuck surface no damage
FX40 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
FX40 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
FX40 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
FX40 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
FX40 | X1 Detector noise (XY servo off)
FX40 | X2 Detector noise (XY servo off)
FX40 | Y1 Detector noise (XY servo off)
FX40 | Y2 Detector noise (XY servo off)
FX40 | X Detector noise (XY servo on)
FX40 | X1 Detector noise (XY servo on)
FX40 | X2 Detector noise (XY servo on)
FX40 | Y Detector noise (XY servo on)
FX40 | Y1 Detector noise (XY servo on)
FX40 | Y2 Detector noise (XY servo on)
FX40 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX40 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX40 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
FX40 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
FX40 | Z Detector offset
FX40 | Check the full stroke image(servo on)
FX40 | ResonanceFrequencyHz
FX40 | ServoCutoffFrequencyHz
FX40 | FullPixelHeight
FX40 | Serial Port#
FX40 | Off-axis Camera X Position(ΔX)
FX40 | Roughness Repeatability 1x1um(10Repeats, RSD%)
FX40 | Roughness Repeatability 80x80um(10Repeats, RSD%)
FX40 | Z Scanner position at -3μm
FX40 | Z Scanner position at +3μm
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
FX40 | Scan offset accuracy (Precision)
FX40 | Aligned stage accuracy (X offset-Point Align1)
FX40 | Aligned stage accuracy (Y offset-Point Align1)
FX40 | Aligned stage accuracy (X offset-Point Align2)
FX40 | Aligned stage accuracy (Y offset-Point Align2)
FX40 | Aligned stage accuracy (X offset-Point Align3)
FX40 | Aligned stage accuracy (Y offset-Point Align3)
FX40 | Aligned stage accuracy (X offset)
FX40 | Aligned stage accuracy (Y offset)
FX40 | Imag offset after EFEM loading (X scan)
FX40 | Imag offset after EFEM loading (Y scan)
FX40 | Imag offset after Lock stage (X scan)
FX40 | Imag offset after Lock stage (Y scan)
FX40 | Mask Roughness Repeatability (10Repeats, RSD%)
FX40 | Mask Roughness Repeatability (10Repeats, σ )
FX40 | Mask exchange time
FX40 | WLI_Scale(C Calibration)
FX40 | WLI_Scale(B Calibration)
FX40 | WLI_Scale(A Calibration)
FX40 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
FX40 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
FX40 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
FX40 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
FX40 | Port 1 lens(to OnAxis) Y Offset
FX40 | Port 1 lens(to OnAxis) X Offset
FX40 | P Gain
FX40 | I Gain
FX40 | D Gain
FX40 | Check the Port 2 lens light strength
FX40 | Check the Port 1 lens light strength
FX40 | Check the focus deviation between the lenses(Δf)
FX40 | Check the 25mm Moving – Y Axis Offset Check
FX40 | C Scanner Detector Min
FX40 | C Scanner Detector Max
FX40 | B Scanner Detector Min
FX40 | B Scanner Detector Max
FX40 | Angle align between Port 2 lens and VLSI SHS: Y
FX40 | Angle align between Port 2 lens and VLSI SHS: X
FX40 | Angle align between Port 2 lens and Sample chuck: Y
FX40 | Angle align between Port 2 lens and Sample chuck: X
FX40 | Angle align between Port 1 lens and Sample chuck: Y
FX40 | Angle align between Port 1 lens and Sample chuck: X
FX40 | A Scanner Detector Min
FX40 | A Scanner Detector Max
FX40 | Y2 Detector noise (XY Servo off)
FX40 | X2 Detector noise (XY Servo off)
FX40 | Y axis OPM (-52000,-52000)
FX40 | Y axis OPM (-52000,52000)
FX40 | Y axis OPM (52000,-52000)
FX40 | Y axis OPM (52000,52000)
FX40 | Y axis OPM (0,0)
FX40 | XY Orthogonality (-52000,-52000)
FX40 | XY Orthogonality (-52000,52000)
FX40 | XY Orthogonality (52000,-52000)
FX40 | XY Orthogonality (52000,52000)
FX40 | XY Orthogonality (0,0)
FX40 | X axis OPM (-52000,-52000)
FX40 | X axis OPM (-52000,52000)
FX40 | X axis OPM (52000,-52000)
FX40 | X axis OPM (52000,52000)
FX40 | X axis OPM (0,0)
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
FX40 | RMS of Z Height of Wafer Zeroscan(-60000,0)
FX40 | RMS of Z Height of Wafer Zeroscan(60000,0)
FX40 | RMS of Z Height of Wafer Zeroscan(0,-60000)
FX40 | RMS of Z Height of Wafer Zeroscan(0,60000)
FX40 | XY stage calibration
FX40 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
FX40 | LD beam power at the end of the lens
FX40 | Y Non-lineairty
FX40 | X Non-lineairty
FX40 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
FX40 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
FX40 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
FX40 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
FX40 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
FX40 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
FX40 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
FX40 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
FX40 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
FX40 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
FX40 | UpperLimit
FX40 | Under 50nm Step Height Sample (10 Repeats)
FX40 | Overlay Camera X Position(ΔX)
FX40 | LowerLimit
FX40 | Initial Z
FX40 | 940nm Step Height Sample (10 Repeats)
FX40 | 50~200nm Step Height Sample (10 Repeats)
FX40 | RMS of Z Height of C122 Zeroscan(0,-90000)
FX40 | StrokeYMin
FX40 | StrokeYMax
FX40 | StrokeY
FX40 | StrokeXMin
FX40 | StrokeXMax
FX40 | StrokeX
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
FX40 | Z Drift (Short Traveling)
FX40 | 200~500nm Step Height Sample (10 Repeats)
FX40 | XY orthogonality
FX40 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
FX40 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
FX40 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
FX40 | Spring constant calibration by Sader Method with 3XC-NA-500DC
FX40 | Spring constant calibration by Sader Method with 3XC-NA-240AC
FX40 | Spring constant calibration by Sader Method with 3XC-NA-200AC
FX40 | Y1 Detector noise (XY Servo off)
FX40 | X1 Detector noise (XY Servo off)
FX40 | Maximum Y Scan range (Full Stroke)
FX40 | Maximum X Scan range (Full Stroke)
FX40 | Full range of Y2 Detector raw
FX40 | Full range of Y1 Detector raw
FX40 | Full range of X2 Detector raw
FX40 | Full range of X1 Detector raw
FX40 | Z Linearity by Z Scanner position
FX40 | -Y Direction software Limit
FX40 | -X Direction software Limit
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
FX40 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
FX40 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
FX40 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
FX40 | %GRR w/ roughness of non-patterened wafer
FX40 | Z Scanner position at -6μm
FX40 | Z Scanner position at +6μm
FX40 | Approach position at -5μm
FX40 | Approach position at +5μm
FX40 | * 1800nm Step Height Sample (10 Repeats)
FX40 | Full range of Z Detector raw
FX40 | Approach position at -3μm
FX40 | Approach position at +3μm
FX40 | Under 50nm Step Height Sample (10 Repeats)
FX40 | 50~200nm Step Height Sample (10 Repeats)
FX40 | 200~500nm Step Height Sample (10 Repeats)
FX40 | PosMin
FX40 | PosMax
FX40 | MaxLength
FX40 | FullRecoveryProfileLength
FX40 | Full range of Y Detector raw
FX40 | Full range of X Detector raw
FX40 | EFM Phase : For the positive bias
FX40 | EFM Phase : For the negative bias
FX40 | Z Linearity by 5 different sample height
FX40 | Laser intensity with AC160 cantilever
FX40 | A－B difference during 3 times run test
FX40 | A＋B difference during 3 times run test
FX40 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
FX40 | RMS of Z Drive of Wafer Zeroscan(60000,0)
FX40 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
FX40 | RMS of Z Drive of Wafer Zeroscan(0,60000)
FX40 | RMS of Z Drive of Wafer ZeroScan (0,0)
FX40 | Allowable max warpage, w/ SiC sample chuck
FX40 | RMS of Z Height of Wafer Zeroscan (Top)
FX40 | RMS of Z Height of Wafer Zeroscan (Right)
FX40 | RMS of Z Height of Wafer Zeroscan (Left)
FX40 | RMS of Z Height of Wafer ZeroScan (Center)
FX40 | RMS of Z Height of Wafer Zeroscan (Bottom)
FX40 | PosMin
FX40 | PosMax
FX40 | MaxLength
FX40 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
FX40 | Check sample surface mapping
FX40 | Z Direction, Accuracy Test (Bias not significant)
FX40 | Y Direction, Accuracy Test (Bias not significant)
FX40 | X Direction, Accuracy Test (Bias not significant)
FX40 | Roughness Reproducibility Test (SR Panel)
FX40 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
FX40 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
FX40 | Roughness Reproducibility Test (SiNx Panel)
FX40 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
FX40 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
FX40 | Roughness Reproducibility Test (ABF Panel)
FX40 | Roughness Repeatability Test (SR Panel)
FX40 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
FX40 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
FX40 | Roughness Repeatability Test (SiNx Panel)
FX40 | Roughness Repeatability Test (Rough Cu Panel, Trace)
FX40 | Roughness Repeatability Test (Rough Cu Panel, Plane)
FX40 | Roughness Repeatability Test (ABF Panel)
FX40 | Lift Pin Stroke Min Check in DB
FX40 | Lift Pin Stroke Max Check in DB
FX40 | Limit range
NX-PTR | Z Linearity by Z scanner position
NX-PTR | Full range of Z Detector during Z Scan sweep
NX-PTR | Z Height calibration error
NX-PTR | Z Drive calibration error
NX-PTR | Z Linearity by 4 different sample height
NX-PTR | Full range of Z Height during Z Scan sweep
NX-PTR | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-PTR | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-PTR | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-PTR | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-PTR | Z Height noise
NX-PTR | Wafer placement(Loading) error on the load port1 (X offset)
NX-PTR | Wafer placement(Loading) error on the load port1 (Y offset)
NX-PTR | Wafer placement(Loading) error on the load port2 (X offset)
NX-PTR | Wafer placement(Loading) error on the load port2 (Y offset)
NX-PTR | OpticalAxisShiftPerFocusX
NX-PTR | OpticalAxisShiftPerFocusY
NX-PTR | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-PTR | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-PTR | Lift pins synchronization offset check (X offset)
NX-PTR | Lift pins synchronization offset check (Y offset)
NX-PTR | Profiler noise @200um
NX-PTR | Roll motion to X+ full scan
NX-PTR | Position repeatability(10repeats, Position Based)
NX-PTR | OPM Repeatability @5mm
NX-PTR | OPM Repeatability @1mm
NX-PTR | Max OPM @5mm
NX-PTR | Max OPM @1mm
NX-PTR | Torque Profile (10mm, 200um/s)
NX-PTR | Torque Profile (10mm, 20um/s)
NX-PTR | Lateral(C-D) sensitivity
NX-PTR | Lateral(C-D) noise
NX-PTR | Min. Vibration Transmissbility (~100Hz)
NX-PTR | RMS of ZeroScan
NX-PTR | LightStrengthGain
NX-PTR | Full range of X Detector during X Scan sweep
NX-PTR | Full range of X1 Detector during X Scan sweep
NX-PTR | Full range of X2 Detector during X Scan sweep
NX-PTR | Full range of Y Detector during Y Scan sweep
NX-PTR | Full range of Y1 Detector during Y Scan sweep
NX-PTR | Full range of Y2 Detector during Y Scan sweep
NX-PTR | Full range of X1 Detector Raw during X Scan sweep
NX-PTR | Full range of X2 Detector Raw during X Scan sweep
NX-PTR | Full range of Y1 Detector Raw during Y Scan sweep
NX-PTR | Full range of Y2 Detector Raw during Y Scan sweep
NX-PTR | X Detector stroke calibration error (nominal scale error X)
NX-PTR | Y Detector stroke calibration error (nominal scale error Y)
NX-PTR | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-PTR | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-PTR | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-PTR | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-PTR | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-PTR | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-PTR | X Non-linearity error
NX-PTR | Y Non-linearity error
NX-PTR | X axis OPM
NX-PTR | Y axis OPM
NX-PTR | X Ringing(Rpv)
NX-PTR | Y Ringing(Rpv)
NX-PTR | X Detector noise (XY Servo on)
NX-PTR | Y Detector noise (XY Servo on)
NX-PTR | X1 Detector noise (XY Servo on)
NX-PTR | Y1 Detector noise (XY Servo on)
NX-PTR | X2 Detector noise (XY Servo on)
NX-PTR | Y2 Detector noise (XY Servo on)
NX-PTR | RMS error after Stage Calibration
NX-PTR | Image offset after moving from -X to +X (X scan)
NX-PTR | Image offset after moving from -Y to +Y (Y scan)
NX-PTR | Image offset after moving from -X to +X (Y scan)
NX-PTR | Image offset after moving from -Y to +Y (X scan)
NX-PTR | Image offset after moving from +X to -X (X scan)
NX-PTR | Image offset after moving from +Y to -Y (Y scan)
NX-PTR | Image offset after moving from +X to -X (Y scan)
NX-PTR | Image offset after moving from +Y to -Y (X scan)
NX-PTR | XY Stage positioning X offset (Vision Pattern Matching)
NX-PTR | XY Stage positioning Y offset (Vision Pattern Matching)
NX-PTR | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-PTR | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-PTR | XY Stage positioning X offset (In-Die Align)
NX-PTR | XY Stage positioning Y offset (In-Die Align)
NX-PTR | XY Stage positioning X 1σ (In-Die Align)
NX-PTR | XY Stage positioning Y 1σ (In-Die Align)
NX-PTR | XY Stage positioning X offset (LLC Translation)
NX-PTR | XY Stage positioning Y offset (LLC Translation)
NX-PTR | XY Stage positioning X 1σ (LLC Translation)
NX-PTR | XY Stage positioning Y 1σ (LLC Translation)
NX-PTR | XY Stage positioning X offset (Global Align)
NX-PTR | XY Stage positioning Y offset (Global Align)
NX-PTR | XY Stage positioning X 1σ (Global Align)
NX-PTR | XY Stage positioning Y 1σ (Global Align)
NX-PTR | Z Drift
NX-PTR | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | Cycle time for 15-point measurement
NX-PTR | Short-term stability using Step Height Sample (3σ)
NX-PTR | Roughness Repeatability 1x1um(RSD%)
NX-PTR | Roughness Repeatability 1x1um(1σ)
NX-PTR | Roughness Repeatability 80x80um(RSD%)
NX-PTR | Roughness Repeatability 80x80um(1σ)
NX-PTR | Z Height calibrtaion error
NX-PTR | Z Drive calibrtaion error
NX-PTR | Z Linearity by 3 different sample height
NX-PTR | Full range in voltage during Z Scan sweep
NX-PTR | Full range in um during Z Scan sweep
NX-PTR | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-PTR | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-PTR | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-PTR | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-PTR | Lateral(C-D) sensitivity
NX-PTR | Lateral(C-D) noise
NX-PTR | A-B noise
NX-PTR | NCM Amplitude noise
NX-PTR | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-PTR | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-PTR | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-PTR | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-PTR | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-PTR | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-PTR | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-PTR | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-PTR | RMS of Z Height of Wafer ZeroScan (0,0)
NX-PTR | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-PTR | XY Orthogonality
NX-PTR | MicrometerPerAdc of Z Height (Z Height calibration)
NX-PTR | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-PTR | Z linearity by sample height
NX-PTR | 1) Z Scanner position at +3μm
NX-PTR | 3) Z Scanner position at -3μm
NX-PTR | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-PTR | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-PTR | X Scan Stroke
NX-PTR | Y Scan Stroke
NX-PTR | Full range of X Detector Raw during X Scan sweep
NX-PTR | Full range of Y Detector Raw during Y Scan sweep
NX-PTR | X Detector noise (XY Servo off)
NX-PTR | Y Detector noise (XY Servo off)
NX-PTR | X axis OPM
NX-PTR | Y axis OPM
NX-PTR | X Non-linearity
NX-PTR | Y Non-linearity
NX-PTR | OPM Repeatability @25mm
NX-PTR | OPM Repeatability @10mm
NX-PTR | Max OPM @25mm
NX-PTR | Max OPM @10mm
NX-PTR | Torque Profile (50mm, 200um/s)
NX-PTR | Torque Profile (50mm, 20um/s)
NX-PTR | SLD current limit check
NX-PTR | SLD beam power at the end of the lens
NX-PTR | 2) of 20nm
NX-PTR | X Detector noise (XYServo off)
NX-PTR | Y Detector noise (XYServo off)
NX-PTR | X axis OPM raw Rpv
NX-PTR | Y axis OPM raw Rpv
NX-PTR | StageSafetyXMax
NX-PTR | StageSafetyXMin
NX-PTR | StageSafetyY
NX-PTR | StageSafetyYMax
NX-PTR | StageSafetyYMin
NX-PTR | Check X axis upper limit sensor function
NX-PTR | Check X axis lower limit sensor function
NX-PTR | Check Y axis upper limit sensor function
NX-PTR | Check Y axis lower limit sensor function
NX-PTR | X axis Angle Range
NX-PTR | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-PTR | X detector noise (XY Servo on)
NX-PTR | Y detector noise (XY Servo on)
NX-PTR | Full range in um during X Scan sweep (XY Servo off)
NX-PTR | Full range in um during Y Scan sweep (XY Servo off)
NX-PTR | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-PTR | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-PTR | X axis, Repeatability @1st position
NX-PTR | X axis, Repeatability @ 2nd position
NX-PTR | X axis, Repeatability @ 3rd position
NX-PTR | X axis, Repeatability @ 4th position
NX-PTR | X axis, Accuracy @1st position
NX-PTR | X axis, Accuracy @ 2nd position
NX-PTR | X axis, Accuracy @ 3rd position
NX-PTR | X axis, Accuracy @ 4th position
NX-PTR | Y axis, Repeatability @1st position
NX-PTR | Y axis, Repeatability @ 2nd position
NX-PTR | Y axis, Repeatability @ 3rd position
NX-PTR | Y axis, Repeatability @ 4th position
NX-PTR | Y axis, Accuracy @1st position
NX-PTR | Y axis, Accuracy @ 2nd position
NX-PTR | Y axis, Accuracy @ 3rd position
NX-PTR | Y axis, Accuracy @ 4th position
NX-PTR | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-PTR | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-PTR | Z-XY Orthogonality (X direction-Left)
NX-PTR | Z-XY Orthogonality (X direction-Right)
NX-PTR | Z-XY Orthogonality (Y direction-Left)
NX-PTR | Z-XY Orthogonality (Y direction-Right)
NX-PTR | X-Scan Repeatability (10 repeats, RSD%)
NX-PTR | Y-Scan Repeatability (10 repeats, RSD%)
NX-PTR | Mechanical Align Accuracy
NX-PTR | Check the Multi Peattern Upload
NX-PTR | X Detector noise (XY servo off)
NX-PTR | Y Detector noise (XY servo off)
NX-PTR | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-PTR | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-PTR | Z Detector Max
NX-PTR | Z Detector Min
NX-PTR | WLI_Scale(Z Calibration)
NX-PTR | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-PTR | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-PTR | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-PTR | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-PTR | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-PTR | Laser intensity with AC160TS cantilever
NX-PTR | Check if the driving % changes proportionally as the NCM amplitude changes
NX-PTR | NCM driving % at an amplitude of 1 nm
NX-PTR | NCM driving % at an amplitude of 20 nm
NX-PTR | NCM driving % at an amplitude of 100 nm
NX-PTR | - 940nm Step Height Sample (10 Repeats)
NX-PTR | - 200~500nm Step Height Sample (10 Repeats)
NX-PTR | - 50~200nm Step Height Sample (10 Repeats)
NX-PTR | - Under 50nm Step Height Sample (10 Repeats)
NX-PTR | A-B after Auto PSPD Alignment
NX-PTR | NCM driving% with AC160TS cantilever
NX-PTR | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-PTR | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-PTR | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-PTR | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-PTR | X Detector stroke calibration error (nominal scale error X)
NX-PTR | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | Resonance Frequency in DB
NX-PTR | Servo Cut-off Frequency in DB
NX-PTR | Check chuck surface no damage
NX-PTR | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-PTR | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-PTR | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-PTR | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-PTR | X1 Detector noise (XY servo off)
NX-PTR | X2 Detector noise (XY servo off)
NX-PTR | Y1 Detector noise (XY servo off)
NX-PTR | Y2 Detector noise (XY servo off)
NX-PTR | X Detector noise (XY servo on)
NX-PTR | X1 Detector noise (XY servo on)
NX-PTR | X2 Detector noise (XY servo on)
NX-PTR | Y Detector noise (XY servo on)
NX-PTR | Y1 Detector noise (XY servo on)
NX-PTR | Y2 Detector noise (XY servo on)
NX-PTR | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-PTR | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-PTR | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-PTR | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-PTR | Z Detector offset
NX-PTR | Check the full stroke image(servo on)
NX-PTR | ResonanceFrequencyHz
NX-PTR | ServoCutoffFrequencyHz
NX-PTR | FullPixelHeight
NX-PTR | Serial Port#
NX-PTR | Off-axis Camera X Position(ΔX)
NX-PTR | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-PTR | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-PTR | Z Scanner position at -3μm
NX-PTR | Z Scanner position at +3μm
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-PTR | Scan offset accuracy (Precision)
NX-PTR | Aligned stage accuracy (X offset-Point Align1)
NX-PTR | Aligned stage accuracy (Y offset-Point Align1)
NX-PTR | Aligned stage accuracy (X offset-Point Align2)
NX-PTR | Aligned stage accuracy (Y offset-Point Align2)
NX-PTR | Aligned stage accuracy (X offset-Point Align3)
NX-PTR | Aligned stage accuracy (Y offset-Point Align3)
NX-PTR | Aligned stage accuracy (X offset)
NX-PTR | Aligned stage accuracy (Y offset)
NX-PTR | Imag offset after EFEM loading (X scan)
NX-PTR | Imag offset after EFEM loading (Y scan)
NX-PTR | Imag offset after Lock stage (X scan)
NX-PTR | Imag offset after Lock stage (Y scan)
NX-PTR | Mask Roughness Repeatability (10Repeats, RSD%)
NX-PTR | Mask Roughness Repeatability (10Repeats, σ )
NX-PTR | Mask exchange time
NX-PTR | WLI_Scale(C Calibration)
NX-PTR | WLI_Scale(B Calibration)
NX-PTR | WLI_Scale(A Calibration)
NX-PTR | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-PTR | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-PTR | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-PTR | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-PTR | Port 1 lens(to OnAxis) Y Offset
NX-PTR | Port 1 lens(to OnAxis) X Offset
NX-PTR | P Gain
NX-PTR | I Gain
NX-PTR | D Gain
NX-PTR | Check the Port 2 lens light strength
NX-PTR | Check the Port 1 lens light strength
NX-PTR | Check the focus deviation between the lenses(Δf)
NX-PTR | Check the 25mm Moving – Y Axis Offset Check
NX-PTR | C Scanner Detector Min
NX-PTR | C Scanner Detector Max
NX-PTR | B Scanner Detector Min
NX-PTR | B Scanner Detector Max
NX-PTR | Angle align between Port 2 lens and VLSI SHS: Y
NX-PTR | Angle align between Port 2 lens and VLSI SHS: X
NX-PTR | Angle align between Port 2 lens and Sample chuck: Y
NX-PTR | Angle align between Port 2 lens and Sample chuck: X
NX-PTR | Angle align between Port 1 lens and Sample chuck: Y
NX-PTR | Angle align between Port 1 lens and Sample chuck: X
NX-PTR | A Scanner Detector Min
NX-PTR | A Scanner Detector Max
NX-PTR | Y2 Detector noise (XY Servo off)
NX-PTR | X2 Detector noise (XY Servo off)
NX-PTR | Y axis OPM (-52000,-52000)
NX-PTR | Y axis OPM (-52000,52000)
NX-PTR | Y axis OPM (52000,-52000)
NX-PTR | Y axis OPM (52000,52000)
NX-PTR | Y axis OPM (0,0)
NX-PTR | XY Orthogonality (-52000,-52000)
NX-PTR | XY Orthogonality (-52000,52000)
NX-PTR | XY Orthogonality (52000,-52000)
NX-PTR | XY Orthogonality (52000,52000)
NX-PTR | XY Orthogonality (0,0)
NX-PTR | X axis OPM (-52000,-52000)
NX-PTR | X axis OPM (-52000,52000)
NX-PTR | X axis OPM (52000,-52000)
NX-PTR | X axis OPM (52000,52000)
NX-PTR | X axis OPM (0,0)
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-PTR | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-PTR | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-PTR | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-PTR | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-PTR | XY stage calibration
NX-PTR | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-PTR | LD beam power at the end of the lens
NX-PTR | Y Non-lineairty
NX-PTR | X Non-lineairty
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-PTR | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-PTR | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-PTR | UpperLimit
NX-PTR | Under 50nm Step Height Sample (10 Repeats)
NX-PTR | Overlay Camera X Position(ΔX)
NX-PTR | LowerLimit
NX-PTR | Initial Z
NX-PTR | 940nm Step Height Sample (10 Repeats)
NX-PTR | 50~200nm Step Height Sample (10 Repeats)
NX-PTR | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-PTR | StrokeYMin
NX-PTR | StrokeYMax
NX-PTR | StrokeY
NX-PTR | StrokeXMin
NX-PTR | StrokeXMax
NX-PTR | StrokeX
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-PTR | Z Drift (Short Traveling)
NX-PTR | 200~500nm Step Height Sample (10 Repeats)
NX-PTR | XY orthogonality
NX-PTR | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-PTR | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-PTR | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-PTR | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-PTR | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-PTR | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-PTR | Y1 Detector noise (XY Servo off)
NX-PTR | X1 Detector noise (XY Servo off)
NX-PTR | Maximum Y Scan range (Full Stroke)
NX-PTR | Maximum X Scan range (Full Stroke)
NX-PTR | Full range of Y2 Detector raw
NX-PTR | Full range of Y1 Detector raw
NX-PTR | Full range of X2 Detector raw
NX-PTR | Full range of X1 Detector raw
NX-PTR | Z Linearity by Z Scanner position
NX-PTR | -Y Direction software Limit
NX-PTR | -X Direction software Limit
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-PTR | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-PTR | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-PTR | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-PTR | %GRR w/ roughness of non-patterened wafer
NX-PTR | Z Scanner position at -6μm
NX-PTR | Z Scanner position at +6μm
NX-PTR | Approach position at -5μm
NX-PTR | Approach position at +5μm
NX-PTR | * 1800nm Step Height Sample (10 Repeats)
NX-PTR | Full range of Z Detector raw
NX-PTR | Approach position at -3μm
NX-PTR | Approach position at +3μm
NX-PTR | Under 50nm Step Height Sample (10 Repeats)
NX-PTR | 50~200nm Step Height Sample (10 Repeats)
NX-PTR | 200~500nm Step Height Sample (10 Repeats)
NX-PTR | PosMin
NX-PTR | PosMax
NX-PTR | MaxLength
NX-PTR | FullRecoveryProfileLength
NX-PTR | Full range of Y Detector raw
NX-PTR | Full range of X Detector raw
NX-PTR | EFM Phase : For the positive bias
NX-PTR | EFM Phase : For the negative bias
NX-PTR | Z Linearity by 5 different sample height
NX-PTR | Laser intensity with AC160 cantilever
NX-PTR | A－B difference during 3 times run test
NX-PTR | A＋B difference during 3 times run test
NX-PTR | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-PTR | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-PTR | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-PTR | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-PTR | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-PTR | Allowable max warpage, w/ SiC sample chuck
NX-PTR | RMS of Z Height of Wafer Zeroscan (Top)
NX-PTR | RMS of Z Height of Wafer Zeroscan (Right)
NX-PTR | RMS of Z Height of Wafer Zeroscan (Left)
NX-PTR | RMS of Z Height of Wafer ZeroScan (Center)
NX-PTR | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-PTR | PosMin
NX-PTR | PosMax
NX-PTR | MaxLength
NX-PTR | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-PTR | Check sample surface mapping
NX-PTR | Z Direction, Accuracy Test (Bias not significant)
NX-PTR | Y Direction, Accuracy Test (Bias not significant)
NX-PTR | X Direction, Accuracy Test (Bias not significant)
NX-PTR | Roughness Reproducibility Test (SR Panel)
NX-PTR | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-PTR | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-PTR | Roughness Reproducibility Test (SiNx Panel)
NX-PTR | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-PTR | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-PTR | Roughness Reproducibility Test (ABF Panel)
NX-PTR | Roughness Repeatability Test (SR Panel)
NX-PTR | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-PTR | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-PTR | Roughness Repeatability Test (SiNx Panel)
NX-PTR | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-PTR | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-PTR | Roughness Repeatability Test (ABF Panel)
NX-PTR | Lift Pin Stroke Min Check in DB
NX-PTR | Lift Pin Stroke Max Check in DB
NX-PTR | Limit range
NX-TSH1518 | Z Linearity by Z scanner position
NX-TSH1518 | Full range of Z Detector during Z Scan sweep
NX-TSH1518 | Z Height calibration error
NX-TSH1518 | Z Drive calibration error
NX-TSH1518 | Z Linearity by 4 different sample height
NX-TSH1518 | Full range of Z Height during Z Scan sweep
NX-TSH1518 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH1518 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH1518 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH1518 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH1518 | Z Height noise
NX-TSH1518 | Wafer placement(Loading) error on the load port1 (X offset)
NX-TSH1518 | Wafer placement(Loading) error on the load port1 (Y offset)
NX-TSH1518 | Wafer placement(Loading) error on the load port2 (X offset)
NX-TSH1518 | Wafer placement(Loading) error on the load port2 (Y offset)
NX-TSH1518 | OpticalAxisShiftPerFocusX
NX-TSH1518 | OpticalAxisShiftPerFocusY
NX-TSH1518 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-TSH1518 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-TSH1518 | Lift pins synchronization offset check (X offset)
NX-TSH1518 | Lift pins synchronization offset check (Y offset)
NX-TSH1518 | Profiler noise @200um
NX-TSH1518 | Roll motion to X+ full scan
NX-TSH1518 | Position repeatability(10repeats, Position Based)
NX-TSH1518 | OPM Repeatability @5mm
NX-TSH1518 | OPM Repeatability @1mm
NX-TSH1518 | Max OPM @5mm
NX-TSH1518 | Max OPM @1mm
NX-TSH1518 | Torque Profile (10mm, 200um/s)
NX-TSH1518 | Torque Profile (10mm, 20um/s)
NX-TSH1518 | Lateral(C-D) sensitivity
NX-TSH1518 | Lateral(C-D) noise
NX-TSH1518 | Min. Vibration Transmissbility (~100Hz)
NX-TSH1518 | RMS of ZeroScan
NX-TSH1518 | LightStrengthGain
NX-TSH1518 | Full range of X Detector during X Scan sweep
NX-TSH1518 | Full range of X1 Detector during X Scan sweep
NX-TSH1518 | Full range of X2 Detector during X Scan sweep
NX-TSH1518 | Full range of Y Detector during Y Scan sweep
NX-TSH1518 | Full range of Y1 Detector during Y Scan sweep
NX-TSH1518 | Full range of Y2 Detector during Y Scan sweep
NX-TSH1518 | Full range of X1 Detector Raw during X Scan sweep
NX-TSH1518 | Full range of X2 Detector Raw during X Scan sweep
NX-TSH1518 | Full range of Y1 Detector Raw during Y Scan sweep
NX-TSH1518 | Full range of Y2 Detector Raw during Y Scan sweep
NX-TSH1518 | X Detector stroke calibration error (nominal scale error X)
NX-TSH1518 | Y Detector stroke calibration error (nominal scale error Y)
NX-TSH1518 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-TSH1518 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-TSH1518 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-TSH1518 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-TSH1518 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-TSH1518 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-TSH1518 | X Non-linearity error
NX-TSH1518 | Y Non-linearity error
NX-TSH1518 | X axis OPM
NX-TSH1518 | Y axis OPM
NX-TSH1518 | X Ringing(Rpv)
NX-TSH1518 | Y Ringing(Rpv)
NX-TSH1518 | X Detector noise (XY Servo on)
NX-TSH1518 | Y Detector noise (XY Servo on)
NX-TSH1518 | X1 Detector noise (XY Servo on)
NX-TSH1518 | Y1 Detector noise (XY Servo on)
NX-TSH1518 | X2 Detector noise (XY Servo on)
NX-TSH1518 | Y2 Detector noise (XY Servo on)
NX-TSH1518 | RMS error after Stage Calibration
NX-TSH1518 | Image offset after moving from -X to +X (X scan)
NX-TSH1518 | Image offset after moving from -Y to +Y (Y scan)
NX-TSH1518 | Image offset after moving from -X to +X (Y scan)
NX-TSH1518 | Image offset after moving from -Y to +Y (X scan)
NX-TSH1518 | Image offset after moving from +X to -X (X scan)
NX-TSH1518 | Image offset after moving from +Y to -Y (Y scan)
NX-TSH1518 | Image offset after moving from +X to -X (Y scan)
NX-TSH1518 | Image offset after moving from +Y to -Y (X scan)
NX-TSH1518 | XY Stage positioning X offset (Vision Pattern Matching)
NX-TSH1518 | XY Stage positioning Y offset (Vision Pattern Matching)
NX-TSH1518 | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-TSH1518 | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-TSH1518 | XY Stage positioning X offset (In-Die Align)
NX-TSH1518 | XY Stage positioning Y offset (In-Die Align)
NX-TSH1518 | XY Stage positioning X 1σ (In-Die Align)
NX-TSH1518 | XY Stage positioning Y 1σ (In-Die Align)
NX-TSH1518 | XY Stage positioning X offset (LLC Translation)
NX-TSH1518 | XY Stage positioning Y offset (LLC Translation)
NX-TSH1518 | XY Stage positioning X 1σ (LLC Translation)
NX-TSH1518 | XY Stage positioning Y 1σ (LLC Translation)
NX-TSH1518 | XY Stage positioning X offset (Global Align)
NX-TSH1518 | XY Stage positioning Y offset (Global Align)
NX-TSH1518 | XY Stage positioning X 1σ (Global Align)
NX-TSH1518 | XY Stage positioning Y 1σ (Global Align)
NX-TSH1518 | Z Drift
NX-TSH1518 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | Cycle time for 15-point measurement
NX-TSH1518 | Short-term stability using Step Height Sample (3σ)
NX-TSH1518 | Roughness Repeatability 1x1um(RSD%)
NX-TSH1518 | Roughness Repeatability 1x1um(1σ)
NX-TSH1518 | Roughness Repeatability 80x80um(RSD%)
NX-TSH1518 | Roughness Repeatability 80x80um(1σ)
NX-TSH1518 | Z Height calibrtaion error
NX-TSH1518 | Z Drive calibrtaion error
NX-TSH1518 | Z Linearity by 3 different sample height
NX-TSH1518 | Full range in voltage during Z Scan sweep
NX-TSH1518 | Full range in um during Z Scan sweep
NX-TSH1518 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH1518 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH1518 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH1518 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH1518 | Lateral(C-D) sensitivity
NX-TSH1518 | Lateral(C-D) noise
NX-TSH1518 | A-B noise
NX-TSH1518 | NCM Amplitude noise
NX-TSH1518 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-TSH1518 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-TSH1518 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-TSH1518 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-TSH1518 | RMS of Z Height of Wafer ZeroScan (0,0)
NX-TSH1518 | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-TSH1518 | XY Orthogonality
NX-TSH1518 | MicrometerPerAdc of Z Height (Z Height calibration)
NX-TSH1518 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-TSH1518 | Z linearity by sample height
NX-TSH1518 | 1) Z Scanner position at +3μm
NX-TSH1518 | 3) Z Scanner position at -3μm
NX-TSH1518 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-TSH1518 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-TSH1518 | X Scan Stroke
NX-TSH1518 | Y Scan Stroke
NX-TSH1518 | Full range of X Detector Raw during X Scan sweep
NX-TSH1518 | Full range of Y Detector Raw during Y Scan sweep
NX-TSH1518 | X Detector noise (XY Servo off)
NX-TSH1518 | Y Detector noise (XY Servo off)
NX-TSH1518 | X axis OPM
NX-TSH1518 | Y axis OPM
NX-TSH1518 | X Non-linearity
NX-TSH1518 | Y Non-linearity
NX-TSH1518 | OPM Repeatability @25mm
NX-TSH1518 | OPM Repeatability @10mm
NX-TSH1518 | Max OPM @25mm
NX-TSH1518 | Max OPM @10mm
NX-TSH1518 | Torque Profile (50mm, 200um/s)
NX-TSH1518 | Torque Profile (50mm, 20um/s)
NX-TSH1518 | SLD current limit check
NX-TSH1518 | SLD beam power at the end of the lens
NX-TSH1518 | 2) of 20nm
NX-TSH1518 | X Detector noise (XYServo off)
NX-TSH1518 | Y Detector noise (XYServo off)
NX-TSH1518 | X axis OPM raw Rpv
NX-TSH1518 | Y axis OPM raw Rpv
NX-TSH1518 | StageSafetyXMax
NX-TSH1518 | StageSafetyXMin
NX-TSH1518 | StageSafetyY
NX-TSH1518 | StageSafetyYMax
NX-TSH1518 | StageSafetyYMin
NX-TSH1518 | Check X axis upper limit sensor function
NX-TSH1518 | Check X axis lower limit sensor function
NX-TSH1518 | Check Y axis upper limit sensor function
NX-TSH1518 | Check Y axis lower limit sensor function
NX-TSH1518 | X axis Angle Range
NX-TSH1518 | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-TSH1518 | X detector noise (XY Servo on)
NX-TSH1518 | Y detector noise (XY Servo on)
NX-TSH1518 | Full range in um during X Scan sweep (XY Servo off)
NX-TSH1518 | Full range in um during Y Scan sweep (XY Servo off)
NX-TSH1518 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-TSH1518 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-TSH1518 | X axis, Repeatability @1st position
NX-TSH1518 | X axis, Repeatability @ 2nd position
NX-TSH1518 | X axis, Repeatability @ 3rd position
NX-TSH1518 | X axis, Repeatability @ 4th position
NX-TSH1518 | X axis, Accuracy @1st position
NX-TSH1518 | X axis, Accuracy @ 2nd position
NX-TSH1518 | X axis, Accuracy @ 3rd position
NX-TSH1518 | X axis, Accuracy @ 4th position
NX-TSH1518 | Y axis, Repeatability @1st position
NX-TSH1518 | Y axis, Repeatability @ 2nd position
NX-TSH1518 | Y axis, Repeatability @ 3rd position
NX-TSH1518 | Y axis, Repeatability @ 4th position
NX-TSH1518 | Y axis, Accuracy @1st position
NX-TSH1518 | Y axis, Accuracy @ 2nd position
NX-TSH1518 | Y axis, Accuracy @ 3rd position
NX-TSH1518 | Y axis, Accuracy @ 4th position
NX-TSH1518 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-TSH1518 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-TSH1518 | Z-XY Orthogonality (X direction-Left)
NX-TSH1518 | Z-XY Orthogonality (X direction-Right)
NX-TSH1518 | Z-XY Orthogonality (Y direction-Left)
NX-TSH1518 | Z-XY Orthogonality (Y direction-Right)
NX-TSH1518 | X-Scan Repeatability (10 repeats, RSD%)
NX-TSH1518 | Y-Scan Repeatability (10 repeats, RSD%)
NX-TSH1518 | Mechanical Align Accuracy
NX-TSH1518 | Check the Multi Peattern Upload
NX-TSH1518 | X Detector noise (XY servo off)
NX-TSH1518 | Y Detector noise (XY servo off)
NX-TSH1518 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH1518 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH1518 | Z Detector Max
NX-TSH1518 | Z Detector Min
NX-TSH1518 | WLI_Scale(Z Calibration)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-TSH1518 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-TSH1518 | Laser intensity with AC160TS cantilever
NX-TSH1518 | Check if the driving % changes proportionally as the NCM amplitude changes
NX-TSH1518 | NCM driving % at an amplitude of 1 nm
NX-TSH1518 | NCM driving % at an amplitude of 20 nm
NX-TSH1518 | NCM driving % at an amplitude of 100 nm
NX-TSH1518 | - 940nm Step Height Sample (10 Repeats)
NX-TSH1518 | - 200~500nm Step Height Sample (10 Repeats)
NX-TSH1518 | - 50~200nm Step Height Sample (10 Repeats)
NX-TSH1518 | - Under 50nm Step Height Sample (10 Repeats)
NX-TSH1518 | A-B after Auto PSPD Alignment
NX-TSH1518 | NCM driving% with AC160TS cantilever
NX-TSH1518 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH1518 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH1518 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH1518 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH1518 | X Detector stroke calibration error (nominal scale error X)
NX-TSH1518 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | Resonance Frequency in DB
NX-TSH1518 | Servo Cut-off Frequency in DB
NX-TSH1518 | Check chuck surface no damage
NX-TSH1518 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-TSH1518 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-TSH1518 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-TSH1518 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-TSH1518 | X1 Detector noise (XY servo off)
NX-TSH1518 | X2 Detector noise (XY servo off)
NX-TSH1518 | Y1 Detector noise (XY servo off)
NX-TSH1518 | Y2 Detector noise (XY servo off)
NX-TSH1518 | X Detector noise (XY servo on)
NX-TSH1518 | X1 Detector noise (XY servo on)
NX-TSH1518 | X2 Detector noise (XY servo on)
NX-TSH1518 | Y Detector noise (XY servo on)
NX-TSH1518 | Y1 Detector noise (XY servo on)
NX-TSH1518 | Y2 Detector noise (XY servo on)
NX-TSH1518 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH1518 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH1518 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH1518 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH1518 | Z Detector offset
NX-TSH1518 | Check the full stroke image(servo on)
NX-TSH1518 | ResonanceFrequencyHz
NX-TSH1518 | ServoCutoffFrequencyHz
NX-TSH1518 | FullPixelHeight
NX-TSH1518 | Serial Port#
NX-TSH1518 | Off-axis Camera X Position(ΔX)
NX-TSH1518 | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-TSH1518 | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-TSH1518 | Z Scanner position at -3μm
NX-TSH1518 | Z Scanner position at +3μm
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH1518 | Scan offset accuracy (Precision)
NX-TSH1518 | Aligned stage accuracy (X offset-Point Align1)
NX-TSH1518 | Aligned stage accuracy (Y offset-Point Align1)
NX-TSH1518 | Aligned stage accuracy (X offset-Point Align2)
NX-TSH1518 | Aligned stage accuracy (Y offset-Point Align2)
NX-TSH1518 | Aligned stage accuracy (X offset-Point Align3)
NX-TSH1518 | Aligned stage accuracy (Y offset-Point Align3)
NX-TSH1518 | Aligned stage accuracy (X offset)
NX-TSH1518 | Aligned stage accuracy (Y offset)
NX-TSH1518 | Imag offset after EFEM loading (X scan)
NX-TSH1518 | Imag offset after EFEM loading (Y scan)
NX-TSH1518 | Imag offset after Lock stage (X scan)
NX-TSH1518 | Imag offset after Lock stage (Y scan)
NX-TSH1518 | Mask Roughness Repeatability (10Repeats, RSD%)
NX-TSH1518 | Mask Roughness Repeatability (10Repeats, σ )
NX-TSH1518 | Mask exchange time
NX-TSH1518 | WLI_Scale(C Calibration)
NX-TSH1518 | WLI_Scale(B Calibration)
NX-TSH1518 | WLI_Scale(A Calibration)
NX-TSH1518 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH1518 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH1518 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-TSH1518 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-TSH1518 | Port 1 lens(to OnAxis) Y Offset
NX-TSH1518 | Port 1 lens(to OnAxis) X Offset
NX-TSH1518 | P Gain
NX-TSH1518 | I Gain
NX-TSH1518 | D Gain
NX-TSH1518 | Check the Port 2 lens light strength
NX-TSH1518 | Check the Port 1 lens light strength
NX-TSH1518 | Check the focus deviation between the lenses(Δf)
NX-TSH1518 | Check the 25mm Moving – Y Axis Offset Check
NX-TSH1518 | C Scanner Detector Min
NX-TSH1518 | C Scanner Detector Max
NX-TSH1518 | B Scanner Detector Min
NX-TSH1518 | B Scanner Detector Max
NX-TSH1518 | Angle align between Port 2 lens and VLSI SHS: Y
NX-TSH1518 | Angle align between Port 2 lens and VLSI SHS: X
NX-TSH1518 | Angle align between Port 2 lens and Sample chuck: Y
NX-TSH1518 | Angle align between Port 2 lens and Sample chuck: X
NX-TSH1518 | Angle align between Port 1 lens and Sample chuck: Y
NX-TSH1518 | Angle align between Port 1 lens and Sample chuck: X
NX-TSH1518 | A Scanner Detector Min
NX-TSH1518 | A Scanner Detector Max
NX-TSH1518 | Y2 Detector noise (XY Servo off)
NX-TSH1518 | X2 Detector noise (XY Servo off)
NX-TSH1518 | Y axis OPM (-52000,-52000)
NX-TSH1518 | Y axis OPM (-52000,52000)
NX-TSH1518 | Y axis OPM (52000,-52000)
NX-TSH1518 | Y axis OPM (52000,52000)
NX-TSH1518 | Y axis OPM (0,0)
NX-TSH1518 | XY Orthogonality (-52000,-52000)
NX-TSH1518 | XY Orthogonality (-52000,52000)
NX-TSH1518 | XY Orthogonality (52000,-52000)
NX-TSH1518 | XY Orthogonality (52000,52000)
NX-TSH1518 | XY Orthogonality (0,0)
NX-TSH1518 | X axis OPM (-52000,-52000)
NX-TSH1518 | X axis OPM (-52000,52000)
NX-TSH1518 | X axis OPM (52000,-52000)
NX-TSH1518 | X axis OPM (52000,52000)
NX-TSH1518 | X axis OPM (0,0)
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-TSH1518 | XY stage calibration
NX-TSH1518 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-TSH1518 | LD beam power at the end of the lens
NX-TSH1518 | Y Non-lineairty
NX-TSH1518 | X Non-lineairty
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-TSH1518 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-TSH1518 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-TSH1518 | UpperLimit
NX-TSH1518 | Under 50nm Step Height Sample (10 Repeats)
NX-TSH1518 | Overlay Camera X Position(ΔX)
NX-TSH1518 | LowerLimit
NX-TSH1518 | Initial Z
NX-TSH1518 | 940nm Step Height Sample (10 Repeats)
NX-TSH1518 | 50~200nm Step Height Sample (10 Repeats)
NX-TSH1518 | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-TSH1518 | StrokeYMin
NX-TSH1518 | StrokeYMax
NX-TSH1518 | StrokeY
NX-TSH1518 | StrokeXMin
NX-TSH1518 | StrokeXMax
NX-TSH1518 | StrokeX
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-TSH1518 | Z Drift (Short Traveling)
NX-TSH1518 | 200~500nm Step Height Sample (10 Repeats)
NX-TSH1518 | XY orthogonality
NX-TSH1518 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-TSH1518 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-TSH1518 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-TSH1518 | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-TSH1518 | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-TSH1518 | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-TSH1518 | Y1 Detector noise (XY Servo off)
NX-TSH1518 | X1 Detector noise (XY Servo off)
NX-TSH1518 | Maximum Y Scan range (Full Stroke)
NX-TSH1518 | Maximum X Scan range (Full Stroke)
NX-TSH1518 | Full range of Y2 Detector raw
NX-TSH1518 | Full range of Y1 Detector raw
NX-TSH1518 | Full range of X2 Detector raw
NX-TSH1518 | Full range of X1 Detector raw
NX-TSH1518 | Z Linearity by Z Scanner position
NX-TSH1518 | -Y Direction software Limit
NX-TSH1518 | -X Direction software Limit
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH1518 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH1518 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-TSH1518 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-TSH1518 | %GRR w/ roughness of non-patterened wafer
NX-TSH1518 | Z Scanner position at -6μm
NX-TSH1518 | Z Scanner position at +6μm
NX-TSH1518 | Approach position at -5μm
NX-TSH1518 | Approach position at +5μm
NX-TSH1518 | * 1800nm Step Height Sample (10 Repeats)
NX-TSH1518 | Full range of Z Detector raw
NX-TSH1518 | Approach position at -3μm
NX-TSH1518 | Approach position at +3μm
NX-TSH1518 | Under 50nm Step Height Sample (10 Repeats)
NX-TSH1518 | 50~200nm Step Height Sample (10 Repeats)
NX-TSH1518 | 200~500nm Step Height Sample (10 Repeats)
NX-TSH1518 | PosMin
NX-TSH1518 | PosMax
NX-TSH1518 | MaxLength
NX-TSH1518 | FullRecoveryProfileLength
NX-TSH1518 | Full range of Y Detector raw
NX-TSH1518 | Full range of X Detector raw
NX-TSH1518 | EFM Phase : For the positive bias
NX-TSH1518 | EFM Phase : For the negative bias
NX-TSH1518 | Z Linearity by 5 different sample height
NX-TSH1518 | Laser intensity with AC160 cantilever
NX-TSH1518 | A－B difference during 3 times run test
NX-TSH1518 | A＋B difference during 3 times run test
NX-TSH1518 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-TSH1518 | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-TSH1518 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-TSH1518 | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-TSH1518 | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-TSH1518 | Allowable max warpage, w/ SiC sample chuck
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan (Top)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan (Right)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan (Left)
NX-TSH1518 | RMS of Z Height of Wafer ZeroScan (Center)
NX-TSH1518 | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-TSH1518 | PosMin
NX-TSH1518 | PosMax
NX-TSH1518 | MaxLength
NX-TSH1518 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-TSH1518 | Check sample surface mapping
NX-TSH1518 | Z Direction, Accuracy Test (Bias not significant)
NX-TSH1518 | Y Direction, Accuracy Test (Bias not significant)
NX-TSH1518 | X Direction, Accuracy Test (Bias not significant)
NX-TSH1518 | Roughness Reproducibility Test (SR Panel)
NX-TSH1518 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-TSH1518 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-TSH1518 | Roughness Reproducibility Test (SiNx Panel)
NX-TSH1518 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-TSH1518 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-TSH1518 | Roughness Reproducibility Test (ABF Panel)
NX-TSH1518 | Roughness Repeatability Test (SR Panel)
NX-TSH1518 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-TSH1518 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-TSH1518 | Roughness Repeatability Test (SiNx Panel)
NX-TSH1518 | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-TSH1518 | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-TSH1518 | Roughness Repeatability Test (ABF Panel)
NX-TSH1518 | Lift Pin Stroke Min Check in DB
NX-TSH1518 | Lift Pin Stroke Max Check in DB
NX-TSH1518 | Limit range
NX12 | Z Linearity by Z scanner position
NX12 | Full range of Z Detector during Z Scan sweep
NX12 | Z Height calibration error
NX12 | Z Drive calibration error
NX12 | Z Linearity by 4 different sample height
NX12 | Full range of Z Height during Z Scan sweep
NX12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX12 | Z Height noise
NX12 | Wafer placement(Loading) error on the load port1 (X offset)
NX12 | Wafer placement(Loading) error on the load port1 (Y offset)
NX12 | Wafer placement(Loading) error on the load port2 (X offset)
NX12 | Wafer placement(Loading) error on the load port2 (Y offset)
NX12 | OpticalAxisShiftPerFocusX
NX12 | OpticalAxisShiftPerFocusY
NX12 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX12 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX12 | Lift pins synchronization offset check (X offset)
NX12 | Lift pins synchronization offset check (Y offset)
NX12 | Profiler noise @200um
NX12 | Roll motion to X+ full scan
NX12 | Position repeatability(10repeats, Position Based)
NX12 | OPM Repeatability @5mm
NX12 | OPM Repeatability @1mm
NX12 | Max OPM @5mm
NX12 | Max OPM @1mm
NX12 | Torque Profile (10mm, 200um/s)
NX12 | Torque Profile (10mm, 20um/s)
NX12 | Lateral(C-D) sensitivity
NX12 | Lateral(C-D) noise
NX12 | Min. Vibration Transmissbility (~100Hz)
NX12 | RMS of ZeroScan
NX12 | LightStrengthGain
NX12 | Full range of X Detector during X Scan sweep
NX12 | Full range of X1 Detector during X Scan sweep
NX12 | Full range of X2 Detector during X Scan sweep
NX12 | Full range of Y Detector during Y Scan sweep
NX12 | Full range of Y1 Detector during Y Scan sweep
NX12 | Full range of Y2 Detector during Y Scan sweep
NX12 | Full range of X1 Detector Raw during X Scan sweep
NX12 | Full range of X2 Detector Raw during X Scan sweep
NX12 | Full range of Y1 Detector Raw during Y Scan sweep
NX12 | Full range of Y2 Detector Raw during Y Scan sweep
NX12 | X Detector stroke calibration error (nominal scale error X)
NX12 | Y Detector stroke calibration error (nominal scale error Y)
NX12 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX12 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX12 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX12 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX12 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX12 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX12 | X Non-linearity error
NX12 | Y Non-linearity error
NX12 | X axis OPM
NX12 | Y axis OPM
NX12 | X Ringing(Rpv)
NX12 | Y Ringing(Rpv)
NX12 | X Detector noise (XY Servo on)
NX12 | Y Detector noise (XY Servo on)
NX12 | X1 Detector noise (XY Servo on)
NX12 | Y1 Detector noise (XY Servo on)
NX12 | X2 Detector noise (XY Servo on)
NX12 | Y2 Detector noise (XY Servo on)
NX12 | RMS error after Stage Calibration
NX12 | Image offset after moving from -X to +X (X scan)
NX12 | Image offset after moving from -Y to +Y (Y scan)
NX12 | Image offset after moving from -X to +X (Y scan)
NX12 | Image offset after moving from -Y to +Y (X scan)
NX12 | Image offset after moving from +X to -X (X scan)
NX12 | Image offset after moving from +Y to -Y (Y scan)
NX12 | Image offset after moving from +X to -X (Y scan)
NX12 | Image offset after moving from +Y to -Y (X scan)
NX12 | XY Stage positioning X offset (Vision Pattern Matching)
NX12 | XY Stage positioning Y offset (Vision Pattern Matching)
NX12 | XY Stage positioning X 1σ (Vision Pattern Matching)
NX12 | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX12 | XY Stage positioning X offset (In-Die Align)
NX12 | XY Stage positioning Y offset (In-Die Align)
NX12 | XY Stage positioning X 1σ (In-Die Align)
NX12 | XY Stage positioning Y 1σ (In-Die Align)
NX12 | XY Stage positioning X offset (LLC Translation)
NX12 | XY Stage positioning Y offset (LLC Translation)
NX12 | XY Stage positioning X 1σ (LLC Translation)
NX12 | XY Stage positioning Y 1σ (LLC Translation)
NX12 | XY Stage positioning X offset (Global Align)
NX12 | XY Stage positioning Y offset (Global Align)
NX12 | XY Stage positioning X 1σ (Global Align)
NX12 | XY Stage positioning Y 1σ (Global Align)
NX12 | Z Drift
NX12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | Cycle time for 15-point measurement
NX12 | Short-term stability using Step Height Sample (3σ)
NX12 | Roughness Repeatability 1x1um(RSD%)
NX12 | Roughness Repeatability 1x1um(1σ)
NX12 | Roughness Repeatability 80x80um(RSD%)
NX12 | Roughness Repeatability 80x80um(1σ)
NX12 | Z Height calibrtaion error
NX12 | Z Drive calibrtaion error
NX12 | Z Linearity by 3 different sample height
NX12 | Full range in voltage during Z Scan sweep
NX12 | Full range in um during Z Scan sweep
NX12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX12 | Lateral(C-D) sensitivity
NX12 | Lateral(C-D) noise
NX12 | A-B noise
NX12 | NCM Amplitude noise
NX12 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX12 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX12 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX12 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX12 | RMS of Z Height of Wafer Zeroscan(0,120000)
NX12 | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX12 | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX12 | RMS of Z Height of Wafer Zeroscan(120000,0)
NX12 | RMS of Z Height of Wafer ZeroScan (0,0)
NX12 | Max Wafer Zeroscan-Min Wafer Zeroscan
NX12 | XY Orthogonality
NX12 | MicrometerPerAdc of Z Height (Z Height calibration)
NX12 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX12 | Z linearity by sample height
NX12 | 1) Z Scanner position at +3μm
NX12 | 3) Z Scanner position at -3μm
NX12 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX12 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX12 | X Scan Stroke
NX12 | Y Scan Stroke
NX12 | Full range of X Detector Raw during X Scan sweep
NX12 | Full range of Y Detector Raw during Y Scan sweep
NX12 | X Detector noise (XY Servo off)
NX12 | Y Detector noise (XY Servo off)
NX12 | X axis OPM
NX12 | Y axis OPM
NX12 | X Non-linearity
NX12 | Y Non-linearity
NX12 | OPM Repeatability @25mm
NX12 | OPM Repeatability @10mm
NX12 | Max OPM @25mm
NX12 | Max OPM @10mm
NX12 | Torque Profile (50mm, 200um/s)
NX12 | Torque Profile (50mm, 20um/s)
NX12 | SLD current limit check
NX12 | SLD beam power at the end of the lens
NX12 | 2) of 20nm
NX12 | X Detector noise (XYServo off)
NX12 | Y Detector noise (XYServo off)
NX12 | X axis OPM raw Rpv
NX12 | Y axis OPM raw Rpv
NX12 | StageSafetyXMax
NX12 | StageSafetyXMin
NX12 | StageSafetyY
NX12 | StageSafetyYMax
NX12 | StageSafetyYMin
NX12 | Check X axis upper limit sensor function
NX12 | Check X axis lower limit sensor function
NX12 | Check Y axis upper limit sensor function
NX12 | Check Y axis lower limit sensor function
NX12 | X axis Angle Range
NX12 | Short-term stability w/ 450nm Step Height Sample (3σ)
NX12 | X detector noise (XY Servo on)
NX12 | Y detector noise (XY Servo on)
NX12 | Full range in um during X Scan sweep (XY Servo off)
NX12 | Full range in um during Y Scan sweep (XY Servo off)
NX12 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX12 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX12 | X axis, Repeatability @1st position
NX12 | X axis, Repeatability @ 2nd position
NX12 | X axis, Repeatability @ 3rd position
NX12 | X axis, Repeatability @ 4th position
NX12 | X axis, Accuracy @1st position
NX12 | X axis, Accuracy @ 2nd position
NX12 | X axis, Accuracy @ 3rd position
NX12 | X axis, Accuracy @ 4th position
NX12 | Y axis, Repeatability @1st position
NX12 | Y axis, Repeatability @ 2nd position
NX12 | Y axis, Repeatability @ 3rd position
NX12 | Y axis, Repeatability @ 4th position
NX12 | Y axis, Accuracy @1st position
NX12 | Y axis, Accuracy @ 2nd position
NX12 | Y axis, Accuracy @ 3rd position
NX12 | Y axis, Accuracy @ 4th position
NX12 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX12 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX12 | Z-XY Orthogonality (X direction-Left)
NX12 | Z-XY Orthogonality (X direction-Right)
NX12 | Z-XY Orthogonality (Y direction-Left)
NX12 | Z-XY Orthogonality (Y direction-Right)
NX12 | X-Scan Repeatability (10 repeats, RSD%)
NX12 | Y-Scan Repeatability (10 repeats, RSD%)
NX12 | Mechanical Align Accuracy
NX12 | Check the Multi Peattern Upload
NX12 | X Detector noise (XY servo off)
NX12 | Y Detector noise (XY servo off)
NX12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX12 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX12 | Z Detector Max
NX12 | Z Detector Min
NX12 | WLI_Scale(Z Calibration)
NX12 | RMS of Z Height of Wafer Zeroscan(0,90000)
NX12 | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX12 | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX12 | RMS of Z Height of Wafer Zeroscan(90000,0)
NX12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | Short-term stability w/ 180nm Step Height Sample (3σ)
NX12 | Laser intensity with AC160TS cantilever
NX12 | Check if the driving % changes proportionally as the NCM amplitude changes
NX12 | NCM driving % at an amplitude of 1 nm
NX12 | NCM driving % at an amplitude of 20 nm
NX12 | NCM driving % at an amplitude of 100 nm
NX12 | - 940nm Step Height Sample (10 Repeats)
NX12 | - 200~500nm Step Height Sample (10 Repeats)
NX12 | - 50~200nm Step Height Sample (10 Repeats)
NX12 | - Under 50nm Step Height Sample (10 Repeats)
NX12 | A-B after Auto PSPD Alignment
NX12 | NCM driving% with AC160TS cantilever
NX12 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX12 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX12 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX12 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX12 | X Detector stroke calibration error (nominal scale error X)
NX12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | Resonance Frequency in DB
NX12 | Servo Cut-off Frequency in DB
NX12 | Check chuck surface no damage
NX12 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX12 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX12 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX12 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX12 | X1 Detector noise (XY servo off)
NX12 | X2 Detector noise (XY servo off)
NX12 | Y1 Detector noise (XY servo off)
NX12 | Y2 Detector noise (XY servo off)
NX12 | X Detector noise (XY servo on)
NX12 | X1 Detector noise (XY servo on)
NX12 | X2 Detector noise (XY servo on)
NX12 | Y Detector noise (XY servo on)
NX12 | Y1 Detector noise (XY servo on)
NX12 | Y2 Detector noise (XY servo on)
NX12 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX12 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX12 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX12 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX12 | Z Detector offset
NX12 | Check the full stroke image(servo on)
NX12 | ResonanceFrequencyHz
NX12 | ServoCutoffFrequencyHz
NX12 | FullPixelHeight
NX12 | Serial Port#
NX12 | Off-axis Camera X Position(ΔX)
NX12 | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX12 | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX12 | Z Scanner position at -3μm
NX12 | Z Scanner position at +3μm
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX12 | Scan offset accuracy (Precision)
NX12 | Aligned stage accuracy (X offset-Point Align1)
NX12 | Aligned stage accuracy (Y offset-Point Align1)
NX12 | Aligned stage accuracy (X offset-Point Align2)
NX12 | Aligned stage accuracy (Y offset-Point Align2)
NX12 | Aligned stage accuracy (X offset-Point Align3)
NX12 | Aligned stage accuracy (Y offset-Point Align3)
NX12 | Aligned stage accuracy (X offset)
NX12 | Aligned stage accuracy (Y offset)
NX12 | Imag offset after EFEM loading (X scan)
NX12 | Imag offset after EFEM loading (Y scan)
NX12 | Imag offset after Lock stage (X scan)
NX12 | Imag offset after Lock stage (Y scan)
NX12 | Mask Roughness Repeatability (10Repeats, RSD%)
NX12 | Mask Roughness Repeatability (10Repeats, σ )
NX12 | Mask exchange time
NX12 | WLI_Scale(C Calibration)
NX12 | WLI_Scale(B Calibration)
NX12 | WLI_Scale(A Calibration)
NX12 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX12 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX12 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX12 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX12 | Port 1 lens(to OnAxis) Y Offset
NX12 | Port 1 lens(to OnAxis) X Offset
NX12 | P Gain
NX12 | I Gain
NX12 | D Gain
NX12 | Check the Port 2 lens light strength
NX12 | Check the Port 1 lens light strength
NX12 | Check the focus deviation between the lenses(Δf)
NX12 | Check the 25mm Moving – Y Axis Offset Check
NX12 | C Scanner Detector Min
NX12 | C Scanner Detector Max
NX12 | B Scanner Detector Min
NX12 | B Scanner Detector Max
NX12 | Angle align between Port 2 lens and VLSI SHS: Y
NX12 | Angle align between Port 2 lens and VLSI SHS: X
NX12 | Angle align between Port 2 lens and Sample chuck: Y
NX12 | Angle align between Port 2 lens and Sample chuck: X
NX12 | Angle align between Port 1 lens and Sample chuck: Y
NX12 | Angle align between Port 1 lens and Sample chuck: X
NX12 | A Scanner Detector Min
NX12 | A Scanner Detector Max
NX12 | Y2 Detector noise (XY Servo off)
NX12 | X2 Detector noise (XY Servo off)
NX12 | Y axis OPM (-52000,-52000)
NX12 | Y axis OPM (-52000,52000)
NX12 | Y axis OPM (52000,-52000)
NX12 | Y axis OPM (52000,52000)
NX12 | Y axis OPM (0,0)
NX12 | XY Orthogonality (-52000,-52000)
NX12 | XY Orthogonality (-52000,52000)
NX12 | XY Orthogonality (52000,-52000)
NX12 | XY Orthogonality (52000,52000)
NX12 | XY Orthogonality (0,0)
NX12 | X axis OPM (-52000,-52000)
NX12 | X axis OPM (-52000,52000)
NX12 | X axis OPM (52000,-52000)
NX12 | X axis OPM (52000,52000)
NX12 | X axis OPM (0,0)
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX12 | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX12 | RMS of Z Height of Wafer Zeroscan(60000,0)
NX12 | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX12 | RMS of Z Height of Wafer Zeroscan(0,60000)
NX12 | XY stage calibration
NX12 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX12 | LD beam power at the end of the lens
NX12 | Y Non-lineairty
NX12 | X Non-lineairty
NX12 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX12 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX12 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX12 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX12 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX12 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX12 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX12 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX12 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX12 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX12 | UpperLimit
NX12 | Under 50nm Step Height Sample (10 Repeats)
NX12 | Overlay Camera X Position(ΔX)
NX12 | LowerLimit
NX12 | Initial Z
NX12 | 940nm Step Height Sample (10 Repeats)
NX12 | 50~200nm Step Height Sample (10 Repeats)
NX12 | RMS of Z Height of C122 Zeroscan(0,-90000)
NX12 | StrokeYMin
NX12 | StrokeYMax
NX12 | StrokeY
NX12 | StrokeXMin
NX12 | StrokeXMax
NX12 | StrokeX
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX12 | Z Drift (Short Traveling)
NX12 | 200~500nm Step Height Sample (10 Repeats)
NX12 | XY orthogonality
NX12 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX12 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX12 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX12 | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX12 | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX12 | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX12 | Y1 Detector noise (XY Servo off)
NX12 | X1 Detector noise (XY Servo off)
NX12 | Maximum Y Scan range (Full Stroke)
NX12 | Maximum X Scan range (Full Stroke)
NX12 | Full range of Y2 Detector raw
NX12 | Full range of Y1 Detector raw
NX12 | Full range of X2 Detector raw
NX12 | Full range of X1 Detector raw
NX12 | Z Linearity by Z Scanner position
NX12 | -Y Direction software Limit
NX12 | -X Direction software Limit
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX12 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX12 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX12 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX12 | %GRR w/ roughness of non-patterened wafer
NX12 | Z Scanner position at -6μm
NX12 | Z Scanner position at +6μm
NX12 | Approach position at -5μm
NX12 | Approach position at +5μm
NX12 | * 1800nm Step Height Sample (10 Repeats)
NX12 | Full range of Z Detector raw
NX12 | Approach position at -3μm
NX12 | Approach position at +3μm
NX12 | Under 50nm Step Height Sample (10 Repeats)
NX12 | 50~200nm Step Height Sample (10 Repeats)
NX12 | 200~500nm Step Height Sample (10 Repeats)
NX12 | PosMin
NX12 | PosMax
NX12 | MaxLength
NX12 | FullRecoveryProfileLength
NX12 | Full range of Y Detector raw
NX12 | Full range of X Detector raw
NX12 | EFM Phase : For the positive bias
NX12 | EFM Phase : For the negative bias
NX12 | Z Linearity by 5 different sample height
NX12 | Laser intensity with AC160 cantilever
NX12 | A－B difference during 3 times run test
NX12 | A＋B difference during 3 times run test
NX12 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX12 | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX12 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX12 | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX12 | RMS of Z Drive of Wafer ZeroScan (0,0)
NX12 | Allowable max warpage, w/ SiC sample chuck
NX12 | RMS of Z Height of Wafer Zeroscan (Top)
NX12 | RMS of Z Height of Wafer Zeroscan (Right)
NX12 | RMS of Z Height of Wafer Zeroscan (Left)
NX12 | RMS of Z Height of Wafer ZeroScan (Center)
NX12 | RMS of Z Height of Wafer Zeroscan (Bottom)
NX12 | PosMin
NX12 | PosMax
NX12 | MaxLength
NX12 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX12 | Check sample surface mapping
NX12 | Z Direction, Accuracy Test (Bias not significant)
NX12 | Y Direction, Accuracy Test (Bias not significant)
NX12 | X Direction, Accuracy Test (Bias not significant)
NX12 | Roughness Reproducibility Test (SR Panel)
NX12 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX12 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX12 | Roughness Reproducibility Test (SiNx Panel)
NX12 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX12 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX12 | Roughness Reproducibility Test (ABF Panel)
NX12 | Roughness Repeatability Test (SR Panel)
NX12 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX12 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX12 | Roughness Repeatability Test (SiNx Panel)
NX12 | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX12 | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX12 | Roughness Repeatability Test (ABF Panel)
NX12 | Lift Pin Stroke Min Check in DB
NX12 | Lift Pin Stroke Max Check in DB
NX12 | Limit range
NX-Hybrid WLI | Z Linearity by Z scanner position
NX-Hybrid WLI | Full range of Z Detector during Z Scan sweep
NX-Hybrid WLI | Z Height calibration error
NX-Hybrid WLI | Z Drive calibration error
NX-Hybrid WLI | Z Linearity by 4 different sample height
NX-Hybrid WLI | Full range of Z Height during Z Scan sweep
NX-Hybrid WLI | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hybrid WLI | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hybrid WLI | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hybrid WLI | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hybrid WLI | Z Height noise
NX-Hybrid WLI | Wafer placement(Loading) error on the load port1 (X offset)
NX-Hybrid WLI | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Hybrid WLI | Wafer placement(Loading) error on the load port2 (X offset)
NX-Hybrid WLI | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Hybrid WLI | OpticalAxisShiftPerFocusX
NX-Hybrid WLI | OpticalAxisShiftPerFocusY
NX-Hybrid WLI | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Hybrid WLI | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Hybrid WLI | Lift pins synchronization offset check (X offset)
NX-Hybrid WLI | Lift pins synchronization offset check (Y offset)
NX-Hybrid WLI | Profiler noise @200um
NX-Hybrid WLI | Roll motion to X+ full scan
NX-Hybrid WLI | Position repeatability(10repeats, Position Based)
NX-Hybrid WLI | OPM Repeatability @5mm
NX-Hybrid WLI | OPM Repeatability @1mm
NX-Hybrid WLI | Max OPM @5mm
NX-Hybrid WLI | Max OPM @1mm
NX-Hybrid WLI | Torque Profile (10mm, 200um/s)
NX-Hybrid WLI | Torque Profile (10mm, 20um/s)
NX-Hybrid WLI | Lateral(C-D) sensitivity
NX-Hybrid WLI | Lateral(C-D) noise
NX-Hybrid WLI | Min. Vibration Transmissbility (~100Hz)
NX-Hybrid WLI | RMS of ZeroScan
NX-Hybrid WLI | LightStrengthGain
NX-Hybrid WLI | Full range of X Detector during X Scan sweep
NX-Hybrid WLI | Full range of X1 Detector during X Scan sweep
NX-Hybrid WLI | Full range of X2 Detector during X Scan sweep
NX-Hybrid WLI | Full range of Y Detector during Y Scan sweep
NX-Hybrid WLI | Full range of Y1 Detector during Y Scan sweep
NX-Hybrid WLI | Full range of Y2 Detector during Y Scan sweep
NX-Hybrid WLI | Full range of X1 Detector Raw during X Scan sweep
NX-Hybrid WLI | Full range of X2 Detector Raw during X Scan sweep
NX-Hybrid WLI | Full range of Y1 Detector Raw during Y Scan sweep
NX-Hybrid WLI | Full range of Y2 Detector Raw during Y Scan sweep
NX-Hybrid WLI | X Detector stroke calibration error (nominal scale error X)
NX-Hybrid WLI | Y Detector stroke calibration error (nominal scale error Y)
NX-Hybrid WLI | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Hybrid WLI | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Hybrid WLI | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Hybrid WLI | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Hybrid WLI | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Hybrid WLI | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Hybrid WLI | X Non-linearity error
NX-Hybrid WLI | Y Non-linearity error
NX-Hybrid WLI | X axis OPM
NX-Hybrid WLI | Y axis OPM
NX-Hybrid WLI | X Ringing(Rpv)
NX-Hybrid WLI | Y Ringing(Rpv)
NX-Hybrid WLI | X Detector noise (XY Servo on)
NX-Hybrid WLI | Y Detector noise (XY Servo on)
NX-Hybrid WLI | X1 Detector noise (XY Servo on)
NX-Hybrid WLI | Y1 Detector noise (XY Servo on)
NX-Hybrid WLI | X2 Detector noise (XY Servo on)
NX-Hybrid WLI | Y2 Detector noise (XY Servo on)
NX-Hybrid WLI | RMS error after Stage Calibration
NX-Hybrid WLI | Image offset after moving from -X to +X (X scan)
NX-Hybrid WLI | Image offset after moving from -Y to +Y (Y scan)
NX-Hybrid WLI | Image offset after moving from -X to +X (Y scan)
NX-Hybrid WLI | Image offset after moving from -Y to +Y (X scan)
NX-Hybrid WLI | Image offset after moving from +X to -X (X scan)
NX-Hybrid WLI | Image offset after moving from +Y to -Y (Y scan)
NX-Hybrid WLI | Image offset after moving from +X to -X (Y scan)
NX-Hybrid WLI | Image offset after moving from +Y to -Y (X scan)
NX-Hybrid WLI | XY Stage positioning X offset (Vision Pattern Matching)
NX-Hybrid WLI | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Hybrid WLI | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Hybrid WLI | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Hybrid WLI | XY Stage positioning X offset (In-Die Align)
NX-Hybrid WLI | XY Stage positioning Y offset (In-Die Align)
NX-Hybrid WLI | XY Stage positioning X 1σ (In-Die Align)
NX-Hybrid WLI | XY Stage positioning Y 1σ (In-Die Align)
NX-Hybrid WLI | XY Stage positioning X offset (LLC Translation)
NX-Hybrid WLI | XY Stage positioning Y offset (LLC Translation)
NX-Hybrid WLI | XY Stage positioning X 1σ (LLC Translation)
NX-Hybrid WLI | XY Stage positioning Y 1σ (LLC Translation)
NX-Hybrid WLI | XY Stage positioning X offset (Global Align)
NX-Hybrid WLI | XY Stage positioning Y offset (Global Align)
NX-Hybrid WLI | XY Stage positioning X 1σ (Global Align)
NX-Hybrid WLI | XY Stage positioning Y 1σ (Global Align)
NX-Hybrid WLI | Z Drift
NX-Hybrid WLI | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | Cycle time for 15-point measurement
NX-Hybrid WLI | Short-term stability using Step Height Sample (3σ)
NX-Hybrid WLI | Roughness Repeatability 1x1um(RSD%)
NX-Hybrid WLI | Roughness Repeatability 1x1um(1σ)
NX-Hybrid WLI | Roughness Repeatability 80x80um(RSD%)
NX-Hybrid WLI | Roughness Repeatability 80x80um(1σ)
NX-Hybrid WLI | Z Height calibrtaion error
NX-Hybrid WLI | Z Drive calibrtaion error
NX-Hybrid WLI | Z Linearity by 3 different sample height
NX-Hybrid WLI | Full range in voltage during Z Scan sweep
NX-Hybrid WLI | Full range in um during Z Scan sweep
NX-Hybrid WLI | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hybrid WLI | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hybrid WLI | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hybrid WLI | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hybrid WLI | Lateral(C-D) sensitivity
NX-Hybrid WLI | Lateral(C-D) noise
NX-Hybrid WLI | A-B noise
NX-Hybrid WLI | NCM Amplitude noise
NX-Hybrid WLI | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Hybrid WLI | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Hybrid WLI | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Hybrid WLI | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Hybrid WLI | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Hybrid WLI | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Hybrid WLI | XY Orthogonality
NX-Hybrid WLI | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Hybrid WLI | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Hybrid WLI | Z linearity by sample height
NX-Hybrid WLI | 1) Z Scanner position at +3μm
NX-Hybrid WLI | 3) Z Scanner position at -3μm
NX-Hybrid WLI | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Hybrid WLI | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Hybrid WLI | X Scan Stroke
NX-Hybrid WLI | Y Scan Stroke
NX-Hybrid WLI | Full range of X Detector Raw during X Scan sweep
NX-Hybrid WLI | Full range of Y Detector Raw during Y Scan sweep
NX-Hybrid WLI | X Detector noise (XY Servo off)
NX-Hybrid WLI | Y Detector noise (XY Servo off)
NX-Hybrid WLI | X axis OPM
NX-Hybrid WLI | Y axis OPM
NX-Hybrid WLI | X Non-linearity
NX-Hybrid WLI | Y Non-linearity
NX-Hybrid WLI | OPM Repeatability @25mm
NX-Hybrid WLI | OPM Repeatability @10mm
NX-Hybrid WLI | Max OPM @25mm
NX-Hybrid WLI | Max OPM @10mm
NX-Hybrid WLI | Torque Profile (50mm, 200um/s)
NX-Hybrid WLI | Torque Profile (50mm, 20um/s)
NX-Hybrid WLI | SLD current limit check
NX-Hybrid WLI | SLD beam power at the end of the lens
NX-Hybrid WLI | 2) of 20nm
NX-Hybrid WLI | X Detector noise (XYServo off)
NX-Hybrid WLI | Y Detector noise (XYServo off)
NX-Hybrid WLI | X axis OPM raw Rpv
NX-Hybrid WLI | Y axis OPM raw Rpv
NX-Hybrid WLI | StageSafetyXMax
NX-Hybrid WLI | StageSafetyXMin
NX-Hybrid WLI | StageSafetyY
NX-Hybrid WLI | StageSafetyYMax
NX-Hybrid WLI | StageSafetyYMin
NX-Hybrid WLI | Check X axis upper limit sensor function
NX-Hybrid WLI | Check X axis lower limit sensor function
NX-Hybrid WLI | Check Y axis upper limit sensor function
NX-Hybrid WLI | Check Y axis lower limit sensor function
NX-Hybrid WLI | X axis Angle Range
NX-Hybrid WLI | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Hybrid WLI | X detector noise (XY Servo on)
NX-Hybrid WLI | Y detector noise (XY Servo on)
NX-Hybrid WLI | Full range in um during X Scan sweep (XY Servo off)
NX-Hybrid WLI | Full range in um during Y Scan sweep (XY Servo off)
NX-Hybrid WLI | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Hybrid WLI | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Hybrid WLI | X axis, Repeatability @1st position
NX-Hybrid WLI | X axis, Repeatability @ 2nd position
NX-Hybrid WLI | X axis, Repeatability @ 3rd position
NX-Hybrid WLI | X axis, Repeatability @ 4th position
NX-Hybrid WLI | X axis, Accuracy @1st position
NX-Hybrid WLI | X axis, Accuracy @ 2nd position
NX-Hybrid WLI | X axis, Accuracy @ 3rd position
NX-Hybrid WLI | X axis, Accuracy @ 4th position
NX-Hybrid WLI | Y axis, Repeatability @1st position
NX-Hybrid WLI | Y axis, Repeatability @ 2nd position
NX-Hybrid WLI | Y axis, Repeatability @ 3rd position
NX-Hybrid WLI | Y axis, Repeatability @ 4th position
NX-Hybrid WLI | Y axis, Accuracy @1st position
NX-Hybrid WLI | Y axis, Accuracy @ 2nd position
NX-Hybrid WLI | Y axis, Accuracy @ 3rd position
NX-Hybrid WLI | Y axis, Accuracy @ 4th position
NX-Hybrid WLI | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Hybrid WLI | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Hybrid WLI | Z-XY Orthogonality (X direction-Left)
NX-Hybrid WLI | Z-XY Orthogonality (X direction-Right)
NX-Hybrid WLI | Z-XY Orthogonality (Y direction-Left)
NX-Hybrid WLI | Z-XY Orthogonality (Y direction-Right)
NX-Hybrid WLI | X-Scan Repeatability (10 repeats, RSD%)
NX-Hybrid WLI | Y-Scan Repeatability (10 repeats, RSD%)
NX-Hybrid WLI | Mechanical Align Accuracy
NX-Hybrid WLI | Check the Multi Peattern Upload
NX-Hybrid WLI | X Detector noise (XY servo off)
NX-Hybrid WLI | Y Detector noise (XY servo off)
NX-Hybrid WLI | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hybrid WLI | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hybrid WLI | Z Detector Max
NX-Hybrid WLI | Z Detector Min
NX-Hybrid WLI | WLI_Scale(Z Calibration)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Hybrid WLI | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Hybrid WLI | Laser intensity with AC160TS cantilever
NX-Hybrid WLI | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Hybrid WLI | NCM driving % at an amplitude of 1 nm
NX-Hybrid WLI | NCM driving % at an amplitude of 20 nm
NX-Hybrid WLI | NCM driving % at an amplitude of 100 nm
NX-Hybrid WLI | - 940nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | - 200~500nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | - 50~200nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | - Under 50nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | A-B after Auto PSPD Alignment
NX-Hybrid WLI | NCM driving% with AC160TS cantilever
NX-Hybrid WLI | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hybrid WLI | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hybrid WLI | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hybrid WLI | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hybrid WLI | X Detector stroke calibration error (nominal scale error X)
NX-Hybrid WLI | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | Resonance Frequency in DB
NX-Hybrid WLI | Servo Cut-off Frequency in DB
NX-Hybrid WLI | Check chuck surface no damage
NX-Hybrid WLI | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Hybrid WLI | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Hybrid WLI | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Hybrid WLI | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Hybrid WLI | X1 Detector noise (XY servo off)
NX-Hybrid WLI | X2 Detector noise (XY servo off)
NX-Hybrid WLI | Y1 Detector noise (XY servo off)
NX-Hybrid WLI | Y2 Detector noise (XY servo off)
NX-Hybrid WLI | X Detector noise (XY servo on)
NX-Hybrid WLI | X1 Detector noise (XY servo on)
NX-Hybrid WLI | X2 Detector noise (XY servo on)
NX-Hybrid WLI | Y Detector noise (XY servo on)
NX-Hybrid WLI | Y1 Detector noise (XY servo on)
NX-Hybrid WLI | Y2 Detector noise (XY servo on)
NX-Hybrid WLI | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hybrid WLI | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hybrid WLI | Z Detector offset
NX-Hybrid WLI | Check the full stroke image(servo on)
NX-Hybrid WLI | ResonanceFrequencyHz
NX-Hybrid WLI | ServoCutoffFrequencyHz
NX-Hybrid WLI | FullPixelHeight
NX-Hybrid WLI | Serial Port#
NX-Hybrid WLI | Off-axis Camera X Position(ΔX)
NX-Hybrid WLI | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Hybrid WLI | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Hybrid WLI | Z Scanner position at -3μm
NX-Hybrid WLI | Z Scanner position at +3μm
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hybrid WLI | Scan offset accuracy (Precision)
NX-Hybrid WLI | Aligned stage accuracy (X offset-Point Align1)
NX-Hybrid WLI | Aligned stage accuracy (Y offset-Point Align1)
NX-Hybrid WLI | Aligned stage accuracy (X offset-Point Align2)
NX-Hybrid WLI | Aligned stage accuracy (Y offset-Point Align2)
NX-Hybrid WLI | Aligned stage accuracy (X offset-Point Align3)
NX-Hybrid WLI | Aligned stage accuracy (Y offset-Point Align3)
NX-Hybrid WLI | Aligned stage accuracy (X offset)
NX-Hybrid WLI | Aligned stage accuracy (Y offset)
NX-Hybrid WLI | Imag offset after EFEM loading (X scan)
NX-Hybrid WLI | Imag offset after EFEM loading (Y scan)
NX-Hybrid WLI | Imag offset after Lock stage (X scan)
NX-Hybrid WLI | Imag offset after Lock stage (Y scan)
NX-Hybrid WLI | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Hybrid WLI | Mask Roughness Repeatability (10Repeats, σ )
NX-Hybrid WLI | Mask exchange time
NX-Hybrid WLI | WLI_Scale(C Calibration)
NX-Hybrid WLI | WLI_Scale(B Calibration)
NX-Hybrid WLI | WLI_Scale(A Calibration)
NX-Hybrid WLI | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hybrid WLI | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hybrid WLI | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Hybrid WLI | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Hybrid WLI | Port 1 lens(to OnAxis) Y Offset
NX-Hybrid WLI | Port 1 lens(to OnAxis) X Offset
NX-Hybrid WLI | P Gain
NX-Hybrid WLI | I Gain
NX-Hybrid WLI | D Gain
NX-Hybrid WLI | Check the Port 2 lens light strength
NX-Hybrid WLI | Check the Port 1 lens light strength
NX-Hybrid WLI | Check the focus deviation between the lenses(Δf)
NX-Hybrid WLI | Check the 25mm Moving – Y Axis Offset Check
NX-Hybrid WLI | C Scanner Detector Min
NX-Hybrid WLI | C Scanner Detector Max
NX-Hybrid WLI | B Scanner Detector Min
NX-Hybrid WLI | B Scanner Detector Max
NX-Hybrid WLI | Angle align between Port 2 lens and VLSI SHS: Y
NX-Hybrid WLI | Angle align between Port 2 lens and VLSI SHS: X
NX-Hybrid WLI | Angle align between Port 2 lens and Sample chuck: Y
NX-Hybrid WLI | Angle align between Port 2 lens and Sample chuck: X
NX-Hybrid WLI | Angle align between Port 1 lens and Sample chuck: Y
NX-Hybrid WLI | Angle align between Port 1 lens and Sample chuck: X
NX-Hybrid WLI | A Scanner Detector Min
NX-Hybrid WLI | A Scanner Detector Max
NX-Hybrid WLI | Y2 Detector noise (XY Servo off)
NX-Hybrid WLI | X2 Detector noise (XY Servo off)
NX-Hybrid WLI | Y axis OPM (-52000,-52000)
NX-Hybrid WLI | Y axis OPM (-52000,52000)
NX-Hybrid WLI | Y axis OPM (52000,-52000)
NX-Hybrid WLI | Y axis OPM (52000,52000)
NX-Hybrid WLI | Y axis OPM (0,0)
NX-Hybrid WLI | XY Orthogonality (-52000,-52000)
NX-Hybrid WLI | XY Orthogonality (-52000,52000)
NX-Hybrid WLI | XY Orthogonality (52000,-52000)
NX-Hybrid WLI | XY Orthogonality (52000,52000)
NX-Hybrid WLI | XY Orthogonality (0,0)
NX-Hybrid WLI | X axis OPM (-52000,-52000)
NX-Hybrid WLI | X axis OPM (-52000,52000)
NX-Hybrid WLI | X axis OPM (52000,-52000)
NX-Hybrid WLI | X axis OPM (52000,52000)
NX-Hybrid WLI | X axis OPM (0,0)
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Hybrid WLI | XY stage calibration
NX-Hybrid WLI | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Hybrid WLI | LD beam power at the end of the lens
NX-Hybrid WLI | Y Non-lineairty
NX-Hybrid WLI | X Non-lineairty
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Hybrid WLI | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Hybrid WLI | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Hybrid WLI | UpperLimit
NX-Hybrid WLI | Under 50nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | Overlay Camera X Position(ΔX)
NX-Hybrid WLI | LowerLimit
NX-Hybrid WLI | Initial Z
NX-Hybrid WLI | 940nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | 50~200nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Hybrid WLI | StrokeYMin
NX-Hybrid WLI | StrokeYMax
NX-Hybrid WLI | StrokeY
NX-Hybrid WLI | StrokeXMin
NX-Hybrid WLI | StrokeXMax
NX-Hybrid WLI | StrokeX
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Hybrid WLI | Z Drift (Short Traveling)
NX-Hybrid WLI | 200~500nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | XY orthogonality
NX-Hybrid WLI | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Hybrid WLI | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Hybrid WLI | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Hybrid WLI | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Hybrid WLI | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Hybrid WLI | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Hybrid WLI | Y1 Detector noise (XY Servo off)
NX-Hybrid WLI | X1 Detector noise (XY Servo off)
NX-Hybrid WLI | Maximum Y Scan range (Full Stroke)
NX-Hybrid WLI | Maximum X Scan range (Full Stroke)
NX-Hybrid WLI | Full range of Y2 Detector raw
NX-Hybrid WLI | Full range of Y1 Detector raw
NX-Hybrid WLI | Full range of X2 Detector raw
NX-Hybrid WLI | Full range of X1 Detector raw
NX-Hybrid WLI | Z Linearity by Z Scanner position
NX-Hybrid WLI | -Y Direction software Limit
NX-Hybrid WLI | -X Direction software Limit
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hybrid WLI | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hybrid WLI | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Hybrid WLI | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Hybrid WLI | %GRR w/ roughness of non-patterened wafer
NX-Hybrid WLI | Z Scanner position at -6μm
NX-Hybrid WLI | Z Scanner position at +6μm
NX-Hybrid WLI | Approach position at -5μm
NX-Hybrid WLI | Approach position at +5μm
NX-Hybrid WLI | * 1800nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | Full range of Z Detector raw
NX-Hybrid WLI | Approach position at -3μm
NX-Hybrid WLI | Approach position at +3μm
NX-Hybrid WLI | Under 50nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | 50~200nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | 200~500nm Step Height Sample (10 Repeats)
NX-Hybrid WLI | PosMin
NX-Hybrid WLI | PosMax
NX-Hybrid WLI | MaxLength
NX-Hybrid WLI | FullRecoveryProfileLength
NX-Hybrid WLI | Full range of Y Detector raw
NX-Hybrid WLI | Full range of X Detector raw
NX-Hybrid WLI | EFM Phase : For the positive bias
NX-Hybrid WLI | EFM Phase : For the negative bias
NX-Hybrid WLI | Z Linearity by 5 different sample height
NX-Hybrid WLI | Laser intensity with AC160 cantilever
NX-Hybrid WLI | A－B difference during 3 times run test
NX-Hybrid WLI | A＋B difference during 3 times run test
NX-Hybrid WLI | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Hybrid WLI | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Hybrid WLI | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Hybrid WLI | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Hybrid WLI | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Hybrid WLI | Allowable max warpage, w/ SiC sample chuck
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan (Top)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan (Right)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan (Left)
NX-Hybrid WLI | RMS of Z Height of Wafer ZeroScan (Center)
NX-Hybrid WLI | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Hybrid WLI | PosMin
NX-Hybrid WLI | PosMax
NX-Hybrid WLI | MaxLength
NX-Hybrid WLI | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Hybrid WLI | Check sample surface mapping
NX-Hybrid WLI | Z Direction, Accuracy Test (Bias not significant)
NX-Hybrid WLI | Y Direction, Accuracy Test (Bias not significant)
NX-Hybrid WLI | X Direction, Accuracy Test (Bias not significant)
NX-Hybrid WLI | Roughness Reproducibility Test (SR Panel)
NX-Hybrid WLI | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Hybrid WLI | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Hybrid WLI | Roughness Reproducibility Test (SiNx Panel)
NX-Hybrid WLI | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Hybrid WLI | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Hybrid WLI | Roughness Reproducibility Test (ABF Panel)
NX-Hybrid WLI | Roughness Repeatability Test (SR Panel)
NX-Hybrid WLI | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Hybrid WLI | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Hybrid WLI | Roughness Repeatability Test (SiNx Panel)
NX-Hybrid WLI | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Hybrid WLI | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Hybrid WLI | Roughness Repeatability Test (ABF Panel)
NX-Hybrid WLI | Lift Pin Stroke Min Check in DB
NX-Hybrid WLI | Lift Pin Stroke Max Check in DB
NX-Hybrid WLI | Limit range
NX7 | Z Linearity by Z scanner position
NX7 | Full range of Z Detector during Z Scan sweep
NX7 | Z Height calibration error
NX7 | Z Drive calibration error
NX7 | Z Linearity by 4 different sample height
NX7 | Full range of Z Height during Z Scan sweep
NX7 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX7 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX7 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX7 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX7 | Z Height noise
NX7 | Wafer placement(Loading) error on the load port1 (X offset)
NX7 | Wafer placement(Loading) error on the load port1 (Y offset)
NX7 | Wafer placement(Loading) error on the load port2 (X offset)
NX7 | Wafer placement(Loading) error on the load port2 (Y offset)
NX7 | OpticalAxisShiftPerFocusX
NX7 | OpticalAxisShiftPerFocusY
NX7 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX7 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX7 | Lift pins synchronization offset check (X offset)
NX7 | Lift pins synchronization offset check (Y offset)
NX7 | Profiler noise @200um
NX7 | Roll motion to X+ full scan
NX7 | Position repeatability(10repeats, Position Based)
NX7 | OPM Repeatability @5mm
NX7 | OPM Repeatability @1mm
NX7 | Max OPM @5mm
NX7 | Max OPM @1mm
NX7 | Torque Profile (10mm, 200um/s)
NX7 | Torque Profile (10mm, 20um/s)
NX7 | Lateral(C-D) sensitivity
NX7 | Lateral(C-D) noise
NX7 | Min. Vibration Transmissbility (~100Hz)
NX7 | RMS of ZeroScan
NX7 | LightStrengthGain
NX7 | Full range of X Detector during X Scan sweep
NX7 | Full range of X1 Detector during X Scan sweep
NX7 | Full range of X2 Detector during X Scan sweep
NX7 | Full range of Y Detector during Y Scan sweep
NX7 | Full range of Y1 Detector during Y Scan sweep
NX7 | Full range of Y2 Detector during Y Scan sweep
NX7 | Full range of X1 Detector Raw during X Scan sweep
NX7 | Full range of X2 Detector Raw during X Scan sweep
NX7 | Full range of Y1 Detector Raw during Y Scan sweep
NX7 | Full range of Y2 Detector Raw during Y Scan sweep
NX7 | X Detector stroke calibration error (nominal scale error X)
NX7 | Y Detector stroke calibration error (nominal scale error Y)
NX7 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX7 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX7 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX7 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX7 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX7 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX7 | X Non-linearity error
NX7 | Y Non-linearity error
NX7 | X axis OPM
NX7 | Y axis OPM
NX7 | X Ringing(Rpv)
NX7 | Y Ringing(Rpv)
NX7 | X Detector noise (XY Servo on)
NX7 | Y Detector noise (XY Servo on)
NX7 | X1 Detector noise (XY Servo on)
NX7 | Y1 Detector noise (XY Servo on)
NX7 | X2 Detector noise (XY Servo on)
NX7 | Y2 Detector noise (XY Servo on)
NX7 | RMS error after Stage Calibration
NX7 | Image offset after moving from -X to +X (X scan)
NX7 | Image offset after moving from -Y to +Y (Y scan)
NX7 | Image offset after moving from -X to +X (Y scan)
NX7 | Image offset after moving from -Y to +Y (X scan)
NX7 | Image offset after moving from +X to -X (X scan)
NX7 | Image offset after moving from +Y to -Y (Y scan)
NX7 | Image offset after moving from +X to -X (Y scan)
NX7 | Image offset after moving from +Y to -Y (X scan)
NX7 | XY Stage positioning X offset (Vision Pattern Matching)
NX7 | XY Stage positioning Y offset (Vision Pattern Matching)
NX7 | XY Stage positioning X 1σ (Vision Pattern Matching)
NX7 | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX7 | XY Stage positioning X offset (In-Die Align)
NX7 | XY Stage positioning Y offset (In-Die Align)
NX7 | XY Stage positioning X 1σ (In-Die Align)
NX7 | XY Stage positioning Y 1σ (In-Die Align)
NX7 | XY Stage positioning X offset (LLC Translation)
NX7 | XY Stage positioning Y offset (LLC Translation)
NX7 | XY Stage positioning X 1σ (LLC Translation)
NX7 | XY Stage positioning Y 1σ (LLC Translation)
NX7 | XY Stage positioning X offset (Global Align)
NX7 | XY Stage positioning Y offset (Global Align)
NX7 | XY Stage positioning X 1σ (Global Align)
NX7 | XY Stage positioning Y 1σ (Global Align)
NX7 | Z Drift
NX7 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | Cycle time for 15-point measurement
NX7 | Short-term stability using Step Height Sample (3σ)
NX7 | Roughness Repeatability 1x1um(RSD%)
NX7 | Roughness Repeatability 1x1um(1σ)
NX7 | Roughness Repeatability 80x80um(RSD%)
NX7 | Roughness Repeatability 80x80um(1σ)
NX7 | Z Height calibrtaion error
NX7 | Z Drive calibrtaion error
NX7 | Z Linearity by 3 different sample height
NX7 | Full range in voltage during Z Scan sweep
NX7 | Full range in um during Z Scan sweep
NX7 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX7 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX7 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX7 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX7 | Lateral(C-D) sensitivity
NX7 | Lateral(C-D) noise
NX7 | A-B noise
NX7 | NCM Amplitude noise
NX7 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX7 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX7 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX7 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX7 | RMS of Z Height of Wafer Zeroscan(0,120000)
NX7 | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX7 | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX7 | RMS of Z Height of Wafer Zeroscan(120000,0)
NX7 | RMS of Z Height of Wafer ZeroScan (0,0)
NX7 | Max Wafer Zeroscan-Min Wafer Zeroscan
NX7 | XY Orthogonality
NX7 | MicrometerPerAdc of Z Height (Z Height calibration)
NX7 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX7 | Z linearity by sample height
NX7 | 1) Z Scanner position at +3μm
NX7 | 3) Z Scanner position at -3μm
NX7 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX7 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX7 | X Scan Stroke
NX7 | Y Scan Stroke
NX7 | Full range of X Detector Raw during X Scan sweep
NX7 | Full range of Y Detector Raw during Y Scan sweep
NX7 | X Detector noise (XY Servo off)
NX7 | Y Detector noise (XY Servo off)
NX7 | X axis OPM
NX7 | Y axis OPM
NX7 | X Non-linearity
NX7 | Y Non-linearity
NX7 | OPM Repeatability @25mm
NX7 | OPM Repeatability @10mm
NX7 | Max OPM @25mm
NX7 | Max OPM @10mm
NX7 | Torque Profile (50mm, 200um/s)
NX7 | Torque Profile (50mm, 20um/s)
NX7 | SLD current limit check
NX7 | SLD beam power at the end of the lens
NX7 | 2) of 20nm
NX7 | X Detector noise (XYServo off)
NX7 | Y Detector noise (XYServo off)
NX7 | X axis OPM raw Rpv
NX7 | Y axis OPM raw Rpv
NX7 | StageSafetyXMax
NX7 | StageSafetyXMin
NX7 | StageSafetyY
NX7 | StageSafetyYMax
NX7 | StageSafetyYMin
NX7 | Check X axis upper limit sensor function
NX7 | Check X axis lower limit sensor function
NX7 | Check Y axis upper limit sensor function
NX7 | Check Y axis lower limit sensor function
NX7 | X axis Angle Range
NX7 | Short-term stability w/ 450nm Step Height Sample (3σ)
NX7 | X detector noise (XY Servo on)
NX7 | Y detector noise (XY Servo on)
NX7 | Full range in um during X Scan sweep (XY Servo off)
NX7 | Full range in um during Y Scan sweep (XY Servo off)
NX7 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX7 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX7 | X axis, Repeatability @1st position
NX7 | X axis, Repeatability @ 2nd position
NX7 | X axis, Repeatability @ 3rd position
NX7 | X axis, Repeatability @ 4th position
NX7 | X axis, Accuracy @1st position
NX7 | X axis, Accuracy @ 2nd position
NX7 | X axis, Accuracy @ 3rd position
NX7 | X axis, Accuracy @ 4th position
NX7 | Y axis, Repeatability @1st position
NX7 | Y axis, Repeatability @ 2nd position
NX7 | Y axis, Repeatability @ 3rd position
NX7 | Y axis, Repeatability @ 4th position
NX7 | Y axis, Accuracy @1st position
NX7 | Y axis, Accuracy @ 2nd position
NX7 | Y axis, Accuracy @ 3rd position
NX7 | Y axis, Accuracy @ 4th position
NX7 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX7 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX7 | Z-XY Orthogonality (X direction-Left)
NX7 | Z-XY Orthogonality (X direction-Right)
NX7 | Z-XY Orthogonality (Y direction-Left)
NX7 | Z-XY Orthogonality (Y direction-Right)
NX7 | X-Scan Repeatability (10 repeats, RSD%)
NX7 | Y-Scan Repeatability (10 repeats, RSD%)
NX7 | Mechanical Align Accuracy
NX7 | Check the Multi Peattern Upload
NX7 | X Detector noise (XY servo off)
NX7 | Y Detector noise (XY servo off)
NX7 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX7 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX7 | Z Detector Max
NX7 | Z Detector Min
NX7 | WLI_Scale(Z Calibration)
NX7 | RMS of Z Height of Wafer Zeroscan(0,90000)
NX7 | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX7 | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX7 | RMS of Z Height of Wafer Zeroscan(90000,0)
NX7 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | Short-term stability w/ 180nm Step Height Sample (3σ)
NX7 | Laser intensity with AC160TS cantilever
NX7 | Check if the driving % changes proportionally as the NCM amplitude changes
NX7 | NCM driving % at an amplitude of 1 nm
NX7 | NCM driving % at an amplitude of 20 nm
NX7 | NCM driving % at an amplitude of 100 nm
NX7 | - 940nm Step Height Sample (10 Repeats)
NX7 | - 200~500nm Step Height Sample (10 Repeats)
NX7 | - 50~200nm Step Height Sample (10 Repeats)
NX7 | - Under 50nm Step Height Sample (10 Repeats)
NX7 | A-B after Auto PSPD Alignment
NX7 | NCM driving% with AC160TS cantilever
NX7 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX7 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX7 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX7 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX7 | X Detector stroke calibration error (nominal scale error X)
NX7 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | Resonance Frequency in DB
NX7 | Servo Cut-off Frequency in DB
NX7 | Check chuck surface no damage
NX7 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX7 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX7 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX7 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX7 | X1 Detector noise (XY servo off)
NX7 | X2 Detector noise (XY servo off)
NX7 | Y1 Detector noise (XY servo off)
NX7 | Y2 Detector noise (XY servo off)
NX7 | X Detector noise (XY servo on)
NX7 | X1 Detector noise (XY servo on)
NX7 | X2 Detector noise (XY servo on)
NX7 | Y Detector noise (XY servo on)
NX7 | Y1 Detector noise (XY servo on)
NX7 | Y2 Detector noise (XY servo on)
NX7 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX7 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX7 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX7 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX7 | Z Detector offset
NX7 | Check the full stroke image(servo on)
NX7 | ResonanceFrequencyHz
NX7 | ServoCutoffFrequencyHz
NX7 | FullPixelHeight
NX7 | Serial Port#
NX7 | Off-axis Camera X Position(ΔX)
NX7 | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX7 | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX7 | Z Scanner position at -3μm
NX7 | Z Scanner position at +3μm
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX7 | Scan offset accuracy (Precision)
NX7 | Aligned stage accuracy (X offset-Point Align1)
NX7 | Aligned stage accuracy (Y offset-Point Align1)
NX7 | Aligned stage accuracy (X offset-Point Align2)
NX7 | Aligned stage accuracy (Y offset-Point Align2)
NX7 | Aligned stage accuracy (X offset-Point Align3)
NX7 | Aligned stage accuracy (Y offset-Point Align3)
NX7 | Aligned stage accuracy (X offset)
NX7 | Aligned stage accuracy (Y offset)
NX7 | Imag offset after EFEM loading (X scan)
NX7 | Imag offset after EFEM loading (Y scan)
NX7 | Imag offset after Lock stage (X scan)
NX7 | Imag offset after Lock stage (Y scan)
NX7 | Mask Roughness Repeatability (10Repeats, RSD%)
NX7 | Mask Roughness Repeatability (10Repeats, σ )
NX7 | Mask exchange time
NX7 | WLI_Scale(C Calibration)
NX7 | WLI_Scale(B Calibration)
NX7 | WLI_Scale(A Calibration)
NX7 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX7 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX7 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX7 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX7 | Port 1 lens(to OnAxis) Y Offset
NX7 | Port 1 lens(to OnAxis) X Offset
NX7 | P Gain
NX7 | I Gain
NX7 | D Gain
NX7 | Check the Port 2 lens light strength
NX7 | Check the Port 1 lens light strength
NX7 | Check the focus deviation between the lenses(Δf)
NX7 | Check the 25mm Moving – Y Axis Offset Check
NX7 | C Scanner Detector Min
NX7 | C Scanner Detector Max
NX7 | B Scanner Detector Min
NX7 | B Scanner Detector Max
NX7 | Angle align between Port 2 lens and VLSI SHS: Y
NX7 | Angle align between Port 2 lens and VLSI SHS: X
NX7 | Angle align between Port 2 lens and Sample chuck: Y
NX7 | Angle align between Port 2 lens and Sample chuck: X
NX7 | Angle align between Port 1 lens and Sample chuck: Y
NX7 | Angle align between Port 1 lens and Sample chuck: X
NX7 | A Scanner Detector Min
NX7 | A Scanner Detector Max
NX7 | Y2 Detector noise (XY Servo off)
NX7 | X2 Detector noise (XY Servo off)
NX7 | Y axis OPM (-52000,-52000)
NX7 | Y axis OPM (-52000,52000)
NX7 | Y axis OPM (52000,-52000)
NX7 | Y axis OPM (52000,52000)
NX7 | Y axis OPM (0,0)
NX7 | XY Orthogonality (-52000,-52000)
NX7 | XY Orthogonality (-52000,52000)
NX7 | XY Orthogonality (52000,-52000)
NX7 | XY Orthogonality (52000,52000)
NX7 | XY Orthogonality (0,0)
NX7 | X axis OPM (-52000,-52000)
NX7 | X axis OPM (-52000,52000)
NX7 | X axis OPM (52000,-52000)
NX7 | X axis OPM (52000,52000)
NX7 | X axis OPM (0,0)
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX7 | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX7 | RMS of Z Height of Wafer Zeroscan(60000,0)
NX7 | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX7 | RMS of Z Height of Wafer Zeroscan(0,60000)
NX7 | XY stage calibration
NX7 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX7 | LD beam power at the end of the lens
NX7 | Y Non-lineairty
NX7 | X Non-lineairty
NX7 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX7 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX7 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX7 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX7 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX7 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX7 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX7 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX7 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX7 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX7 | UpperLimit
NX7 | Under 50nm Step Height Sample (10 Repeats)
NX7 | Overlay Camera X Position(ΔX)
NX7 | LowerLimit
NX7 | Initial Z
NX7 | 940nm Step Height Sample (10 Repeats)
NX7 | 50~200nm Step Height Sample (10 Repeats)
NX7 | RMS of Z Height of C122 Zeroscan(0,-90000)
NX7 | StrokeYMin
NX7 | StrokeYMax
NX7 | StrokeY
NX7 | StrokeXMin
NX7 | StrokeXMax
NX7 | StrokeX
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX7 | Z Drift (Short Traveling)
NX7 | 200~500nm Step Height Sample (10 Repeats)
NX7 | XY orthogonality
NX7 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX7 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX7 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX7 | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX7 | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX7 | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX7 | Y1 Detector noise (XY Servo off)
NX7 | X1 Detector noise (XY Servo off)
NX7 | Maximum Y Scan range (Full Stroke)
NX7 | Maximum X Scan range (Full Stroke)
NX7 | Full range of Y2 Detector raw
NX7 | Full range of Y1 Detector raw
NX7 | Full range of X2 Detector raw
NX7 | Full range of X1 Detector raw
NX7 | Z Linearity by Z Scanner position
NX7 | -Y Direction software Limit
NX7 | -X Direction software Limit
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX7 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX7 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX7 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX7 | %GRR w/ roughness of non-patterened wafer
NX7 | Z Scanner position at -6μm
NX7 | Z Scanner position at +6μm
NX7 | Approach position at -5μm
NX7 | Approach position at +5μm
NX7 | * 1800nm Step Height Sample (10 Repeats)
NX7 | Full range of Z Detector raw
NX7 | Approach position at -3μm
NX7 | Approach position at +3μm
NX7 | Under 50nm Step Height Sample (10 Repeats)
NX7 | 50~200nm Step Height Sample (10 Repeats)
NX7 | 200~500nm Step Height Sample (10 Repeats)
NX7 | PosMin
NX7 | PosMax
NX7 | MaxLength
NX7 | FullRecoveryProfileLength
NX7 | Full range of Y Detector raw
NX7 | Full range of X Detector raw
NX7 | EFM Phase : For the positive bias
NX7 | EFM Phase : For the negative bias
NX7 | Z Linearity by 5 different sample height
NX7 | Laser intensity with AC160 cantilever
NX7 | A－B difference during 3 times run test
NX7 | A＋B difference during 3 times run test
NX7 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX7 | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX7 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX7 | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX7 | RMS of Z Drive of Wafer ZeroScan (0,0)
NX7 | Allowable max warpage, w/ SiC sample chuck
NX7 | RMS of Z Height of Wafer Zeroscan (Top)
NX7 | RMS of Z Height of Wafer Zeroscan (Right)
NX7 | RMS of Z Height of Wafer Zeroscan (Left)
NX7 | RMS of Z Height of Wafer ZeroScan (Center)
NX7 | RMS of Z Height of Wafer Zeroscan (Bottom)
NX7 | PosMin
NX7 | PosMax
NX7 | MaxLength
NX7 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX7 | Check sample surface mapping
NX7 | Z Direction, Accuracy Test (Bias not significant)
NX7 | Y Direction, Accuracy Test (Bias not significant)
NX7 | X Direction, Accuracy Test (Bias not significant)
NX7 | Roughness Reproducibility Test (SR Panel)
NX7 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX7 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX7 | Roughness Reproducibility Test (SiNx Panel)
NX7 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX7 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX7 | Roughness Reproducibility Test (ABF Panel)
NX7 | Roughness Repeatability Test (SR Panel)
NX7 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX7 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX7 | Roughness Repeatability Test (SiNx Panel)
NX7 | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX7 | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX7 | Roughness Repeatability Test (ABF Panel)
NX7 | Lift Pin Stroke Min Check in DB
NX7 | Lift Pin Stroke Max Check in DB
NX7 | Limit range
NX-Hivac | Z Linearity by Z scanner position
NX-Hivac | Full range of Z Detector during Z Scan sweep
NX-Hivac | Z Height calibration error
NX-Hivac | Z Drive calibration error
NX-Hivac | Z Linearity by 4 different sample height
NX-Hivac | Full range of Z Height during Z Scan sweep
NX-Hivac | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hivac | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hivac | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hivac | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hivac | Z Height noise
NX-Hivac | Wafer placement(Loading) error on the load port1 (X offset)
NX-Hivac | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Hivac | Wafer placement(Loading) error on the load port2 (X offset)
NX-Hivac | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Hivac | OpticalAxisShiftPerFocusX
NX-Hivac | OpticalAxisShiftPerFocusY
NX-Hivac | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Hivac | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Hivac | Lift pins synchronization offset check (X offset)
NX-Hivac | Lift pins synchronization offset check (Y offset)
NX-Hivac | Profiler noise @200um
NX-Hivac | Roll motion to X+ full scan
NX-Hivac | Position repeatability(10repeats, Position Based)
NX-Hivac | OPM Repeatability @5mm
NX-Hivac | OPM Repeatability @1mm
NX-Hivac | Max OPM @5mm
NX-Hivac | Max OPM @1mm
NX-Hivac | Torque Profile (10mm, 200um/s)
NX-Hivac | Torque Profile (10mm, 20um/s)
NX-Hivac | Lateral(C-D) sensitivity
NX-Hivac | Lateral(C-D) noise
NX-Hivac | Min. Vibration Transmissbility (~100Hz)
NX-Hivac | RMS of ZeroScan
NX-Hivac | LightStrengthGain
NX-Hivac | Full range of X Detector during X Scan sweep
NX-Hivac | Full range of X1 Detector during X Scan sweep
NX-Hivac | Full range of X2 Detector during X Scan sweep
NX-Hivac | Full range of Y Detector during Y Scan sweep
NX-Hivac | Full range of Y1 Detector during Y Scan sweep
NX-Hivac | Full range of Y2 Detector during Y Scan sweep
NX-Hivac | Full range of X1 Detector Raw during X Scan sweep
NX-Hivac | Full range of X2 Detector Raw during X Scan sweep
NX-Hivac | Full range of Y1 Detector Raw during Y Scan sweep
NX-Hivac | Full range of Y2 Detector Raw during Y Scan sweep
NX-Hivac | X Detector stroke calibration error (nominal scale error X)
NX-Hivac | Y Detector stroke calibration error (nominal scale error Y)
NX-Hivac | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Hivac | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Hivac | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Hivac | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Hivac | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Hivac | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Hivac | X Non-linearity error
NX-Hivac | Y Non-linearity error
NX-Hivac | X axis OPM
NX-Hivac | Y axis OPM
NX-Hivac | X Ringing(Rpv)
NX-Hivac | Y Ringing(Rpv)
NX-Hivac | X Detector noise (XY Servo on)
NX-Hivac | Y Detector noise (XY Servo on)
NX-Hivac | X1 Detector noise (XY Servo on)
NX-Hivac | Y1 Detector noise (XY Servo on)
NX-Hivac | X2 Detector noise (XY Servo on)
NX-Hivac | Y2 Detector noise (XY Servo on)
NX-Hivac | RMS error after Stage Calibration
NX-Hivac | Image offset after moving from -X to +X (X scan)
NX-Hivac | Image offset after moving from -Y to +Y (Y scan)
NX-Hivac | Image offset after moving from -X to +X (Y scan)
NX-Hivac | Image offset after moving from -Y to +Y (X scan)
NX-Hivac | Image offset after moving from +X to -X (X scan)
NX-Hivac | Image offset after moving from +Y to -Y (Y scan)
NX-Hivac | Image offset after moving from +X to -X (Y scan)
NX-Hivac | Image offset after moving from +Y to -Y (X scan)
NX-Hivac | XY Stage positioning X offset (Vision Pattern Matching)
NX-Hivac | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Hivac | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Hivac | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Hivac | XY Stage positioning X offset (In-Die Align)
NX-Hivac | XY Stage positioning Y offset (In-Die Align)
NX-Hivac | XY Stage positioning X 1σ (In-Die Align)
NX-Hivac | XY Stage positioning Y 1σ (In-Die Align)
NX-Hivac | XY Stage positioning X offset (LLC Translation)
NX-Hivac | XY Stage positioning Y offset (LLC Translation)
NX-Hivac | XY Stage positioning X 1σ (LLC Translation)
NX-Hivac | XY Stage positioning Y 1σ (LLC Translation)
NX-Hivac | XY Stage positioning X offset (Global Align)
NX-Hivac | XY Stage positioning Y offset (Global Align)
NX-Hivac | XY Stage positioning X 1σ (Global Align)
NX-Hivac | XY Stage positioning Y 1σ (Global Align)
NX-Hivac | Z Drift
NX-Hivac | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | Cycle time for 15-point measurement
NX-Hivac | Short-term stability using Step Height Sample (3σ)
NX-Hivac | Roughness Repeatability 1x1um(RSD%)
NX-Hivac | Roughness Repeatability 1x1um(1σ)
NX-Hivac | Roughness Repeatability 80x80um(RSD%)
NX-Hivac | Roughness Repeatability 80x80um(1σ)
NX-Hivac | Z Height calibrtaion error
NX-Hivac | Z Drive calibrtaion error
NX-Hivac | Z Linearity by 3 different sample height
NX-Hivac | Full range in voltage during Z Scan sweep
NX-Hivac | Full range in um during Z Scan sweep
NX-Hivac | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hivac | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hivac | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hivac | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hivac | Lateral(C-D) sensitivity
NX-Hivac | Lateral(C-D) noise
NX-Hivac | A-B noise
NX-Hivac | NCM Amplitude noise
NX-Hivac | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Hivac | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Hivac | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Hivac | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Hivac | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Hivac | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Hivac | XY Orthogonality
NX-Hivac | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Hivac | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Hivac | Z linearity by sample height
NX-Hivac | 1) Z Scanner position at +3μm
NX-Hivac | 3) Z Scanner position at -3μm
NX-Hivac | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Hivac | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Hivac | X Scan Stroke
NX-Hivac | Y Scan Stroke
NX-Hivac | Full range of X Detector Raw during X Scan sweep
NX-Hivac | Full range of Y Detector Raw during Y Scan sweep
NX-Hivac | X Detector noise (XY Servo off)
NX-Hivac | Y Detector noise (XY Servo off)
NX-Hivac | X axis OPM
NX-Hivac | Y axis OPM
NX-Hivac | X Non-linearity
NX-Hivac | Y Non-linearity
NX-Hivac | OPM Repeatability @25mm
NX-Hivac | OPM Repeatability @10mm
NX-Hivac | Max OPM @25mm
NX-Hivac | Max OPM @10mm
NX-Hivac | Torque Profile (50mm, 200um/s)
NX-Hivac | Torque Profile (50mm, 20um/s)
NX-Hivac | SLD current limit check
NX-Hivac | SLD beam power at the end of the lens
NX-Hivac | 2) of 20nm
NX-Hivac | X Detector noise (XYServo off)
NX-Hivac | Y Detector noise (XYServo off)
NX-Hivac | X axis OPM raw Rpv
NX-Hivac | Y axis OPM raw Rpv
NX-Hivac | StageSafetyXMax
NX-Hivac | StageSafetyXMin
NX-Hivac | StageSafetyY
NX-Hivac | StageSafetyYMax
NX-Hivac | StageSafetyYMin
NX-Hivac | Check X axis upper limit sensor function
NX-Hivac | Check X axis lower limit sensor function
NX-Hivac | Check Y axis upper limit sensor function
NX-Hivac | Check Y axis lower limit sensor function
NX-Hivac | X axis Angle Range
NX-Hivac | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Hivac | X detector noise (XY Servo on)
NX-Hivac | Y detector noise (XY Servo on)
NX-Hivac | Full range in um during X Scan sweep (XY Servo off)
NX-Hivac | Full range in um during Y Scan sweep (XY Servo off)
NX-Hivac | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Hivac | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Hivac | X axis, Repeatability @1st position
NX-Hivac | X axis, Repeatability @ 2nd position
NX-Hivac | X axis, Repeatability @ 3rd position
NX-Hivac | X axis, Repeatability @ 4th position
NX-Hivac | X axis, Accuracy @1st position
NX-Hivac | X axis, Accuracy @ 2nd position
NX-Hivac | X axis, Accuracy @ 3rd position
NX-Hivac | X axis, Accuracy @ 4th position
NX-Hivac | Y axis, Repeatability @1st position
NX-Hivac | Y axis, Repeatability @ 2nd position
NX-Hivac | Y axis, Repeatability @ 3rd position
NX-Hivac | Y axis, Repeatability @ 4th position
NX-Hivac | Y axis, Accuracy @1st position
NX-Hivac | Y axis, Accuracy @ 2nd position
NX-Hivac | Y axis, Accuracy @ 3rd position
NX-Hivac | Y axis, Accuracy @ 4th position
NX-Hivac | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Hivac | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Hivac | Z-XY Orthogonality (X direction-Left)
NX-Hivac | Z-XY Orthogonality (X direction-Right)
NX-Hivac | Z-XY Orthogonality (Y direction-Left)
NX-Hivac | Z-XY Orthogonality (Y direction-Right)
NX-Hivac | X-Scan Repeatability (10 repeats, RSD%)
NX-Hivac | Y-Scan Repeatability (10 repeats, RSD%)
NX-Hivac | Mechanical Align Accuracy
NX-Hivac | Check the Multi Peattern Upload
NX-Hivac | X Detector noise (XY servo off)
NX-Hivac | Y Detector noise (XY servo off)
NX-Hivac | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hivac | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hivac | Z Detector Max
NX-Hivac | Z Detector Min
NX-Hivac | WLI_Scale(Z Calibration)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Hivac | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Hivac | Laser intensity with AC160TS cantilever
NX-Hivac | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Hivac | NCM driving % at an amplitude of 1 nm
NX-Hivac | NCM driving % at an amplitude of 20 nm
NX-Hivac | NCM driving % at an amplitude of 100 nm
NX-Hivac | - 940nm Step Height Sample (10 Repeats)
NX-Hivac | - 200~500nm Step Height Sample (10 Repeats)
NX-Hivac | - 50~200nm Step Height Sample (10 Repeats)
NX-Hivac | - Under 50nm Step Height Sample (10 Repeats)
NX-Hivac | A-B after Auto PSPD Alignment
NX-Hivac | NCM driving% with AC160TS cantilever
NX-Hivac | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hivac | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hivac | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hivac | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hivac | X Detector stroke calibration error (nominal scale error X)
NX-Hivac | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | Resonance Frequency in DB
NX-Hivac | Servo Cut-off Frequency in DB
NX-Hivac | Check chuck surface no damage
NX-Hivac | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Hivac | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Hivac | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Hivac | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Hivac | X1 Detector noise (XY servo off)
NX-Hivac | X2 Detector noise (XY servo off)
NX-Hivac | Y1 Detector noise (XY servo off)
NX-Hivac | Y2 Detector noise (XY servo off)
NX-Hivac | X Detector noise (XY servo on)
NX-Hivac | X1 Detector noise (XY servo on)
NX-Hivac | X2 Detector noise (XY servo on)
NX-Hivac | Y Detector noise (XY servo on)
NX-Hivac | Y1 Detector noise (XY servo on)
NX-Hivac | Y2 Detector noise (XY servo on)
NX-Hivac | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hivac | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hivac | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hivac | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hivac | Z Detector offset
NX-Hivac | Check the full stroke image(servo on)
NX-Hivac | ResonanceFrequencyHz
NX-Hivac | ServoCutoffFrequencyHz
NX-Hivac | FullPixelHeight
NX-Hivac | Serial Port#
NX-Hivac | Off-axis Camera X Position(ΔX)
NX-Hivac | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Hivac | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Hivac | Z Scanner position at -3μm
NX-Hivac | Z Scanner position at +3μm
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hivac | Scan offset accuracy (Precision)
NX-Hivac | Aligned stage accuracy (X offset-Point Align1)
NX-Hivac | Aligned stage accuracy (Y offset-Point Align1)
NX-Hivac | Aligned stage accuracy (X offset-Point Align2)
NX-Hivac | Aligned stage accuracy (Y offset-Point Align2)
NX-Hivac | Aligned stage accuracy (X offset-Point Align3)
NX-Hivac | Aligned stage accuracy (Y offset-Point Align3)
NX-Hivac | Aligned stage accuracy (X offset)
NX-Hivac | Aligned stage accuracy (Y offset)
NX-Hivac | Imag offset after EFEM loading (X scan)
NX-Hivac | Imag offset after EFEM loading (Y scan)
NX-Hivac | Imag offset after Lock stage (X scan)
NX-Hivac | Imag offset after Lock stage (Y scan)
NX-Hivac | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Hivac | Mask Roughness Repeatability (10Repeats, σ )
NX-Hivac | Mask exchange time
NX-Hivac | WLI_Scale(C Calibration)
NX-Hivac | WLI_Scale(B Calibration)
NX-Hivac | WLI_Scale(A Calibration)
NX-Hivac | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hivac | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hivac | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Hivac | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Hivac | Port 1 lens(to OnAxis) Y Offset
NX-Hivac | Port 1 lens(to OnAxis) X Offset
NX-Hivac | P Gain
NX-Hivac | I Gain
NX-Hivac | D Gain
NX-Hivac | Check the Port 2 lens light strength
NX-Hivac | Check the Port 1 lens light strength
NX-Hivac | Check the focus deviation between the lenses(Δf)
NX-Hivac | Check the 25mm Moving – Y Axis Offset Check
NX-Hivac | C Scanner Detector Min
NX-Hivac | C Scanner Detector Max
NX-Hivac | B Scanner Detector Min
NX-Hivac | B Scanner Detector Max
NX-Hivac | Angle align between Port 2 lens and VLSI SHS: Y
NX-Hivac | Angle align between Port 2 lens and VLSI SHS: X
NX-Hivac | Angle align between Port 2 lens and Sample chuck: Y
NX-Hivac | Angle align between Port 2 lens and Sample chuck: X
NX-Hivac | Angle align between Port 1 lens and Sample chuck: Y
NX-Hivac | Angle align between Port 1 lens and Sample chuck: X
NX-Hivac | A Scanner Detector Min
NX-Hivac | A Scanner Detector Max
NX-Hivac | Y2 Detector noise (XY Servo off)
NX-Hivac | X2 Detector noise (XY Servo off)
NX-Hivac | Y axis OPM (-52000,-52000)
NX-Hivac | Y axis OPM (-52000,52000)
NX-Hivac | Y axis OPM (52000,-52000)
NX-Hivac | Y axis OPM (52000,52000)
NX-Hivac | Y axis OPM (0,0)
NX-Hivac | XY Orthogonality (-52000,-52000)
NX-Hivac | XY Orthogonality (-52000,52000)
NX-Hivac | XY Orthogonality (52000,-52000)
NX-Hivac | XY Orthogonality (52000,52000)
NX-Hivac | XY Orthogonality (0,0)
NX-Hivac | X axis OPM (-52000,-52000)
NX-Hivac | X axis OPM (-52000,52000)
NX-Hivac | X axis OPM (52000,-52000)
NX-Hivac | X axis OPM (52000,52000)
NX-Hivac | X axis OPM (0,0)
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Hivac | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Hivac | XY stage calibration
NX-Hivac | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Hivac | LD beam power at the end of the lens
NX-Hivac | Y Non-lineairty
NX-Hivac | X Non-lineairty
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Hivac | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Hivac | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Hivac | UpperLimit
NX-Hivac | Under 50nm Step Height Sample (10 Repeats)
NX-Hivac | Overlay Camera X Position(ΔX)
NX-Hivac | LowerLimit
NX-Hivac | Initial Z
NX-Hivac | 940nm Step Height Sample (10 Repeats)
NX-Hivac | 50~200nm Step Height Sample (10 Repeats)
NX-Hivac | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Hivac | StrokeYMin
NX-Hivac | StrokeYMax
NX-Hivac | StrokeY
NX-Hivac | StrokeXMin
NX-Hivac | StrokeXMax
NX-Hivac | StrokeX
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Hivac | Z Drift (Short Traveling)
NX-Hivac | 200~500nm Step Height Sample (10 Repeats)
NX-Hivac | XY orthogonality
NX-Hivac | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Hivac | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Hivac | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Hivac | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Hivac | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Hivac | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Hivac | Y1 Detector noise (XY Servo off)
NX-Hivac | X1 Detector noise (XY Servo off)
NX-Hivac | Maximum Y Scan range (Full Stroke)
NX-Hivac | Maximum X Scan range (Full Stroke)
NX-Hivac | Full range of Y2 Detector raw
NX-Hivac | Full range of Y1 Detector raw
NX-Hivac | Full range of X2 Detector raw
NX-Hivac | Full range of X1 Detector raw
NX-Hivac | Z Linearity by Z Scanner position
NX-Hivac | -Y Direction software Limit
NX-Hivac | -X Direction software Limit
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hivac | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hivac | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Hivac | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Hivac | %GRR w/ roughness of non-patterened wafer
NX-Hivac | Z Scanner position at -6μm
NX-Hivac | Z Scanner position at +6μm
NX-Hivac | Approach position at -5μm
NX-Hivac | Approach position at +5μm
NX-Hivac | * 1800nm Step Height Sample (10 Repeats)
NX-Hivac | Full range of Z Detector raw
NX-Hivac | Approach position at -3μm
NX-Hivac | Approach position at +3μm
NX-Hivac | Under 50nm Step Height Sample (10 Repeats)
NX-Hivac | 50~200nm Step Height Sample (10 Repeats)
NX-Hivac | 200~500nm Step Height Sample (10 Repeats)
NX-Hivac | PosMin
NX-Hivac | PosMax
NX-Hivac | MaxLength
NX-Hivac | FullRecoveryProfileLength
NX-Hivac | Full range of Y Detector raw
NX-Hivac | Full range of X Detector raw
NX-Hivac | EFM Phase : For the positive bias
NX-Hivac | EFM Phase : For the negative bias
NX-Hivac | Z Linearity by 5 different sample height
NX-Hivac | Laser intensity with AC160 cantilever
NX-Hivac | A－B difference during 3 times run test
NX-Hivac | A＋B difference during 3 times run test
NX-Hivac | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Hivac | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Hivac | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Hivac | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Hivac | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Hivac | Allowable max warpage, w/ SiC sample chuck
NX-Hivac | RMS of Z Height of Wafer Zeroscan (Top)
NX-Hivac | RMS of Z Height of Wafer Zeroscan (Right)
NX-Hivac | RMS of Z Height of Wafer Zeroscan (Left)
NX-Hivac | RMS of Z Height of Wafer ZeroScan (Center)
NX-Hivac | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Hivac | PosMin
NX-Hivac | PosMax
NX-Hivac | MaxLength
NX-Hivac | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Hivac | Check sample surface mapping
NX-Hivac | Z Direction, Accuracy Test (Bias not significant)
NX-Hivac | Y Direction, Accuracy Test (Bias not significant)
NX-Hivac | X Direction, Accuracy Test (Bias not significant)
NX-Hivac | Roughness Reproducibility Test (SR Panel)
NX-Hivac | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Hivac | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Hivac | Roughness Reproducibility Test (SiNx Panel)
NX-Hivac | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Hivac | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Hivac | Roughness Reproducibility Test (ABF Panel)
NX-Hivac | Roughness Repeatability Test (SR Panel)
NX-Hivac | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Hivac | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Hivac | Roughness Repeatability Test (SiNx Panel)
NX-Hivac | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Hivac | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Hivac | Roughness Repeatability Test (ABF Panel)
NX-Hivac | Lift Pin Stroke Min Check in DB
NX-Hivac | Lift Pin Stroke Max Check in DB
NX-Hivac | Limit range
NX-HDM | Z Linearity by Z scanner position
NX-HDM | Full range of Z Detector during Z Scan sweep
NX-HDM | Z Height calibration error
NX-HDM | Z Drive calibration error
NX-HDM | Z Linearity by 4 different sample height
NX-HDM | Full range of Z Height during Z Scan sweep
NX-HDM | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-HDM | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-HDM | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-HDM | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-HDM | Z Height noise
NX-HDM | Wafer placement(Loading) error on the load port1 (X offset)
NX-HDM | Wafer placement(Loading) error on the load port1 (Y offset)
NX-HDM | Wafer placement(Loading) error on the load port2 (X offset)
NX-HDM | Wafer placement(Loading) error on the load port2 (Y offset)
NX-HDM | OpticalAxisShiftPerFocusX
NX-HDM | OpticalAxisShiftPerFocusY
NX-HDM | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-HDM | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-HDM | Lift pins synchronization offset check (X offset)
NX-HDM | Lift pins synchronization offset check (Y offset)
NX-HDM | Profiler noise @200um
NX-HDM | Roll motion to X+ full scan
NX-HDM | Position repeatability(10repeats, Position Based)
NX-HDM | OPM Repeatability @5mm
NX-HDM | OPM Repeatability @1mm
NX-HDM | Max OPM @5mm
NX-HDM | Max OPM @1mm
NX-HDM | Torque Profile (10mm, 200um/s)
NX-HDM | Torque Profile (10mm, 20um/s)
NX-HDM | Lateral(C-D) sensitivity
NX-HDM | Lateral(C-D) noise
NX-HDM | Min. Vibration Transmissbility (~100Hz)
NX-HDM | RMS of ZeroScan
NX-HDM | LightStrengthGain
NX-HDM | Full range of X Detector during X Scan sweep
NX-HDM | Full range of X1 Detector during X Scan sweep
NX-HDM | Full range of X2 Detector during X Scan sweep
NX-HDM | Full range of Y Detector during Y Scan sweep
NX-HDM | Full range of Y1 Detector during Y Scan sweep
NX-HDM | Full range of Y2 Detector during Y Scan sweep
NX-HDM | Full range of X1 Detector Raw during X Scan sweep
NX-HDM | Full range of X2 Detector Raw during X Scan sweep
NX-HDM | Full range of Y1 Detector Raw during Y Scan sweep
NX-HDM | Full range of Y2 Detector Raw during Y Scan sweep
NX-HDM | X Detector stroke calibration error (nominal scale error X)
NX-HDM | Y Detector stroke calibration error (nominal scale error Y)
NX-HDM | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-HDM | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-HDM | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-HDM | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-HDM | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-HDM | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-HDM | X Non-linearity error
NX-HDM | Y Non-linearity error
NX-HDM | X axis OPM
NX-HDM | Y axis OPM
NX-HDM | X Ringing(Rpv)
NX-HDM | Y Ringing(Rpv)
NX-HDM | X Detector noise (XY Servo on)
NX-HDM | Y Detector noise (XY Servo on)
NX-HDM | X1 Detector noise (XY Servo on)
NX-HDM | Y1 Detector noise (XY Servo on)
NX-HDM | X2 Detector noise (XY Servo on)
NX-HDM | Y2 Detector noise (XY Servo on)
NX-HDM | RMS error after Stage Calibration
NX-HDM | Image offset after moving from -X to +X (X scan)
NX-HDM | Image offset after moving from -Y to +Y (Y scan)
NX-HDM | Image offset after moving from -X to +X (Y scan)
NX-HDM | Image offset after moving from -Y to +Y (X scan)
NX-HDM | Image offset after moving from +X to -X (X scan)
NX-HDM | Image offset after moving from +Y to -Y (Y scan)
NX-HDM | Image offset after moving from +X to -X (Y scan)
NX-HDM | Image offset after moving from +Y to -Y (X scan)
NX-HDM | XY Stage positioning X offset (Vision Pattern Matching)
NX-HDM | XY Stage positioning Y offset (Vision Pattern Matching)
NX-HDM | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-HDM | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-HDM | XY Stage positioning X offset (In-Die Align)
NX-HDM | XY Stage positioning Y offset (In-Die Align)
NX-HDM | XY Stage positioning X 1σ (In-Die Align)
NX-HDM | XY Stage positioning Y 1σ (In-Die Align)
NX-HDM | XY Stage positioning X offset (LLC Translation)
NX-HDM | XY Stage positioning Y offset (LLC Translation)
NX-HDM | XY Stage positioning X 1σ (LLC Translation)
NX-HDM | XY Stage positioning Y 1σ (LLC Translation)
NX-HDM | XY Stage positioning X offset (Global Align)
NX-HDM | XY Stage positioning Y offset (Global Align)
NX-HDM | XY Stage positioning X 1σ (Global Align)
NX-HDM | XY Stage positioning Y 1σ (Global Align)
NX-HDM | Z Drift
NX-HDM | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | Cycle time for 15-point measurement
NX-HDM | Short-term stability using Step Height Sample (3σ)
NX-HDM | Roughness Repeatability 1x1um(RSD%)
NX-HDM | Roughness Repeatability 1x1um(1σ)
NX-HDM | Roughness Repeatability 80x80um(RSD%)
NX-HDM | Roughness Repeatability 80x80um(1σ)
NX-HDM | Z Height calibrtaion error
NX-HDM | Z Drive calibrtaion error
NX-HDM | Z Linearity by 3 different sample height
NX-HDM | Full range in voltage during Z Scan sweep
NX-HDM | Full range in um during Z Scan sweep
NX-HDM | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-HDM | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-HDM | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-HDM | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-HDM | Lateral(C-D) sensitivity
NX-HDM | Lateral(C-D) noise
NX-HDM | A-B noise
NX-HDM | NCM Amplitude noise
NX-HDM | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-HDM | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-HDM | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-HDM | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-HDM | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-HDM | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-HDM | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-HDM | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-HDM | RMS of Z Height of Wafer ZeroScan (0,0)
NX-HDM | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-HDM | XY Orthogonality
NX-HDM | MicrometerPerAdc of Z Height (Z Height calibration)
NX-HDM | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-HDM | Z linearity by sample height
NX-HDM | 1) Z Scanner position at +3μm
NX-HDM | 3) Z Scanner position at -3μm
NX-HDM | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-HDM | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-HDM | X Scan Stroke
NX-HDM | Y Scan Stroke
NX-HDM | Full range of X Detector Raw during X Scan sweep
NX-HDM | Full range of Y Detector Raw during Y Scan sweep
NX-HDM | X Detector noise (XY Servo off)
NX-HDM | Y Detector noise (XY Servo off)
NX-HDM | X axis OPM
NX-HDM | Y axis OPM
NX-HDM | X Non-linearity
NX-HDM | Y Non-linearity
NX-HDM | OPM Repeatability @25mm
NX-HDM | OPM Repeatability @10mm
NX-HDM | Max OPM @25mm
NX-HDM | Max OPM @10mm
NX-HDM | Torque Profile (50mm, 200um/s)
NX-HDM | Torque Profile (50mm, 20um/s)
NX-HDM | SLD current limit check
NX-HDM | SLD beam power at the end of the lens
NX-HDM | 2) of 20nm
NX-HDM | X Detector noise (XYServo off)
NX-HDM | Y Detector noise (XYServo off)
NX-HDM | X axis OPM raw Rpv
NX-HDM | Y axis OPM raw Rpv
NX-HDM | StageSafetyXMax
NX-HDM | StageSafetyXMin
NX-HDM | StageSafetyY
NX-HDM | StageSafetyYMax
NX-HDM | StageSafetyYMin
NX-HDM | Check X axis upper limit sensor function
NX-HDM | Check X axis lower limit sensor function
NX-HDM | Check Y axis upper limit sensor function
NX-HDM | Check Y axis lower limit sensor function
NX-HDM | X axis Angle Range
NX-HDM | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-HDM | X detector noise (XY Servo on)
NX-HDM | Y detector noise (XY Servo on)
NX-HDM | Full range in um during X Scan sweep (XY Servo off)
NX-HDM | Full range in um during Y Scan sweep (XY Servo off)
NX-HDM | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-HDM | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-HDM | X axis, Repeatability @1st position
NX-HDM | X axis, Repeatability @ 2nd position
NX-HDM | X axis, Repeatability @ 3rd position
NX-HDM | X axis, Repeatability @ 4th position
NX-HDM | X axis, Accuracy @1st position
NX-HDM | X axis, Accuracy @ 2nd position
NX-HDM | X axis, Accuracy @ 3rd position
NX-HDM | X axis, Accuracy @ 4th position
NX-HDM | Y axis, Repeatability @1st position
NX-HDM | Y axis, Repeatability @ 2nd position
NX-HDM | Y axis, Repeatability @ 3rd position
NX-HDM | Y axis, Repeatability @ 4th position
NX-HDM | Y axis, Accuracy @1st position
NX-HDM | Y axis, Accuracy @ 2nd position
NX-HDM | Y axis, Accuracy @ 3rd position
NX-HDM | Y axis, Accuracy @ 4th position
NX-HDM | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-HDM | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-HDM | Z-XY Orthogonality (X direction-Left)
NX-HDM | Z-XY Orthogonality (X direction-Right)
NX-HDM | Z-XY Orthogonality (Y direction-Left)
NX-HDM | Z-XY Orthogonality (Y direction-Right)
NX-HDM | X-Scan Repeatability (10 repeats, RSD%)
NX-HDM | Y-Scan Repeatability (10 repeats, RSD%)
NX-HDM | Mechanical Align Accuracy
NX-HDM | Check the Multi Peattern Upload
NX-HDM | X Detector noise (XY servo off)
NX-HDM | Y Detector noise (XY servo off)
NX-HDM | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-HDM | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-HDM | Z Detector Max
NX-HDM | Z Detector Min
NX-HDM | WLI_Scale(Z Calibration)
NX-HDM | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-HDM | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-HDM | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-HDM | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-HDM | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-HDM | Laser intensity with AC160TS cantilever
NX-HDM | Check if the driving % changes proportionally as the NCM amplitude changes
NX-HDM | NCM driving % at an amplitude of 1 nm
NX-HDM | NCM driving % at an amplitude of 20 nm
NX-HDM | NCM driving % at an amplitude of 100 nm
NX-HDM | - 940nm Step Height Sample (10 Repeats)
NX-HDM | - 200~500nm Step Height Sample (10 Repeats)
NX-HDM | - 50~200nm Step Height Sample (10 Repeats)
NX-HDM | - Under 50nm Step Height Sample (10 Repeats)
NX-HDM | A-B after Auto PSPD Alignment
NX-HDM | NCM driving% with AC160TS cantilever
NX-HDM | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-HDM | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-HDM | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-HDM | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-HDM | X Detector stroke calibration error (nominal scale error X)
NX-HDM | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | Resonance Frequency in DB
NX-HDM | Servo Cut-off Frequency in DB
NX-HDM | Check chuck surface no damage
NX-HDM | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-HDM | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-HDM | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-HDM | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-HDM | X1 Detector noise (XY servo off)
NX-HDM | X2 Detector noise (XY servo off)
NX-HDM | Y1 Detector noise (XY servo off)
NX-HDM | Y2 Detector noise (XY servo off)
NX-HDM | X Detector noise (XY servo on)
NX-HDM | X1 Detector noise (XY servo on)
NX-HDM | X2 Detector noise (XY servo on)
NX-HDM | Y Detector noise (XY servo on)
NX-HDM | Y1 Detector noise (XY servo on)
NX-HDM | Y2 Detector noise (XY servo on)
NX-HDM | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-HDM | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-HDM | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-HDM | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-HDM | Z Detector offset
NX-HDM | Check the full stroke image(servo on)
NX-HDM | ResonanceFrequencyHz
NX-HDM | ServoCutoffFrequencyHz
NX-HDM | FullPixelHeight
NX-HDM | Serial Port#
NX-HDM | Off-axis Camera X Position(ΔX)
NX-HDM | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-HDM | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-HDM | Z Scanner position at -3μm
NX-HDM | Z Scanner position at +3μm
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-HDM | Scan offset accuracy (Precision)
NX-HDM | Aligned stage accuracy (X offset-Point Align1)
NX-HDM | Aligned stage accuracy (Y offset-Point Align1)
NX-HDM | Aligned stage accuracy (X offset-Point Align2)
NX-HDM | Aligned stage accuracy (Y offset-Point Align2)
NX-HDM | Aligned stage accuracy (X offset-Point Align3)
NX-HDM | Aligned stage accuracy (Y offset-Point Align3)
NX-HDM | Aligned stage accuracy (X offset)
NX-HDM | Aligned stage accuracy (Y offset)
NX-HDM | Imag offset after EFEM loading (X scan)
NX-HDM | Imag offset after EFEM loading (Y scan)
NX-HDM | Imag offset after Lock stage (X scan)
NX-HDM | Imag offset after Lock stage (Y scan)
NX-HDM | Mask Roughness Repeatability (10Repeats, RSD%)
NX-HDM | Mask Roughness Repeatability (10Repeats, σ )
NX-HDM | Mask exchange time
NX-HDM | WLI_Scale(C Calibration)
NX-HDM | WLI_Scale(B Calibration)
NX-HDM | WLI_Scale(A Calibration)
NX-HDM | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-HDM | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-HDM | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-HDM | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-HDM | Port 1 lens(to OnAxis) Y Offset
NX-HDM | Port 1 lens(to OnAxis) X Offset
NX-HDM | P Gain
NX-HDM | I Gain
NX-HDM | D Gain
NX-HDM | Check the Port 2 lens light strength
NX-HDM | Check the Port 1 lens light strength
NX-HDM | Check the focus deviation between the lenses(Δf)
NX-HDM | Check the 25mm Moving – Y Axis Offset Check
NX-HDM | C Scanner Detector Min
NX-HDM | C Scanner Detector Max
NX-HDM | B Scanner Detector Min
NX-HDM | B Scanner Detector Max
NX-HDM | Angle align between Port 2 lens and VLSI SHS: Y
NX-HDM | Angle align between Port 2 lens and VLSI SHS: X
NX-HDM | Angle align between Port 2 lens and Sample chuck: Y
NX-HDM | Angle align between Port 2 lens and Sample chuck: X
NX-HDM | Angle align between Port 1 lens and Sample chuck: Y
NX-HDM | Angle align between Port 1 lens and Sample chuck: X
NX-HDM | A Scanner Detector Min
NX-HDM | A Scanner Detector Max
NX-HDM | Y2 Detector noise (XY Servo off)
NX-HDM | X2 Detector noise (XY Servo off)
NX-HDM | Y axis OPM (-52000,-52000)
NX-HDM | Y axis OPM (-52000,52000)
NX-HDM | Y axis OPM (52000,-52000)
NX-HDM | Y axis OPM (52000,52000)
NX-HDM | Y axis OPM (0,0)
NX-HDM | XY Orthogonality (-52000,-52000)
NX-HDM | XY Orthogonality (-52000,52000)
NX-HDM | XY Orthogonality (52000,-52000)
NX-HDM | XY Orthogonality (52000,52000)
NX-HDM | XY Orthogonality (0,0)
NX-HDM | X axis OPM (-52000,-52000)
NX-HDM | X axis OPM (-52000,52000)
NX-HDM | X axis OPM (52000,-52000)
NX-HDM | X axis OPM (52000,52000)
NX-HDM | X axis OPM (0,0)
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-HDM | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-HDM | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-HDM | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-HDM | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-HDM | XY stage calibration
NX-HDM | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-HDM | LD beam power at the end of the lens
NX-HDM | Y Non-lineairty
NX-HDM | X Non-lineairty
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-HDM | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-HDM | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-HDM | UpperLimit
NX-HDM | Under 50nm Step Height Sample (10 Repeats)
NX-HDM | Overlay Camera X Position(ΔX)
NX-HDM | LowerLimit
NX-HDM | Initial Z
NX-HDM | 940nm Step Height Sample (10 Repeats)
NX-HDM | 50~200nm Step Height Sample (10 Repeats)
NX-HDM | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-HDM | StrokeYMin
NX-HDM | StrokeYMax
NX-HDM | StrokeY
NX-HDM | StrokeXMin
NX-HDM | StrokeXMax
NX-HDM | StrokeX
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-HDM | Z Drift (Short Traveling)
NX-HDM | 200~500nm Step Height Sample (10 Repeats)
NX-HDM | XY orthogonality
NX-HDM | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-HDM | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-HDM | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-HDM | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-HDM | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-HDM | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-HDM | Y1 Detector noise (XY Servo off)
NX-HDM | X1 Detector noise (XY Servo off)
NX-HDM | Maximum Y Scan range (Full Stroke)
NX-HDM | Maximum X Scan range (Full Stroke)
NX-HDM | Full range of Y2 Detector raw
NX-HDM | Full range of Y1 Detector raw
NX-HDM | Full range of X2 Detector raw
NX-HDM | Full range of X1 Detector raw
NX-HDM | Z Linearity by Z Scanner position
NX-HDM | -Y Direction software Limit
NX-HDM | -X Direction software Limit
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-HDM | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-HDM | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-HDM | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-HDM | %GRR w/ roughness of non-patterened wafer
NX-HDM | Z Scanner position at -6μm
NX-HDM | Z Scanner position at +6μm
NX-HDM | Approach position at -5μm
NX-HDM | Approach position at +5μm
NX-HDM | * 1800nm Step Height Sample (10 Repeats)
NX-HDM | Full range of Z Detector raw
NX-HDM | Approach position at -3μm
NX-HDM | Approach position at +3μm
NX-HDM | Under 50nm Step Height Sample (10 Repeats)
NX-HDM | 50~200nm Step Height Sample (10 Repeats)
NX-HDM | 200~500nm Step Height Sample (10 Repeats)
NX-HDM | PosMin
NX-HDM | PosMax
NX-HDM | MaxLength
NX-HDM | FullRecoveryProfileLength
NX-HDM | Full range of Y Detector raw
NX-HDM | Full range of X Detector raw
NX-HDM | EFM Phase : For the positive bias
NX-HDM | EFM Phase : For the negative bias
NX-HDM | Z Linearity by 5 different sample height
NX-HDM | Laser intensity with AC160 cantilever
NX-HDM | A－B difference during 3 times run test
NX-HDM | A＋B difference during 3 times run test
NX-HDM | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-HDM | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-HDM | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-HDM | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-HDM | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-HDM | Allowable max warpage, w/ SiC sample chuck
NX-HDM | RMS of Z Height of Wafer Zeroscan (Top)
NX-HDM | RMS of Z Height of Wafer Zeroscan (Right)
NX-HDM | RMS of Z Height of Wafer Zeroscan (Left)
NX-HDM | RMS of Z Height of Wafer ZeroScan (Center)
NX-HDM | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-HDM | PosMin
NX-HDM | PosMax
NX-HDM | MaxLength
NX-HDM | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-HDM | Check sample surface mapping
NX-HDM | Z Direction, Accuracy Test (Bias not significant)
NX-HDM | Y Direction, Accuracy Test (Bias not significant)
NX-HDM | X Direction, Accuracy Test (Bias not significant)
NX-HDM | Roughness Reproducibility Test (SR Panel)
NX-HDM | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-HDM | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-HDM | Roughness Reproducibility Test (SiNx Panel)
NX-HDM | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-HDM | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-HDM | Roughness Reproducibility Test (ABF Panel)
NX-HDM | Roughness Repeatability Test (SR Panel)
NX-HDM | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-HDM | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-HDM | Roughness Repeatability Test (SiNx Panel)
NX-HDM | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-HDM | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-HDM | Roughness Repeatability Test (ABF Panel)
NX-HDM | Lift Pin Stroke Min Check in DB
NX-HDM | Lift Pin Stroke Max Check in DB
NX-HDM | Limit range
NX20 | Z Linearity by Z scanner position
NX20 | Full range of Z Detector during Z Scan sweep
NX20 | Z Height calibration error
NX20 | Z Drive calibration error
NX20 | Z Linearity by 4 different sample height
NX20 | Full range of Z Height during Z Scan sweep
NX20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 | Z Height noise
NX20 | Wafer placement(Loading) error on the load port1 (X offset)
NX20 | Wafer placement(Loading) error on the load port1 (Y offset)
NX20 | Wafer placement(Loading) error on the load port2 (X offset)
NX20 | Wafer placement(Loading) error on the load port2 (Y offset)
NX20 | OpticalAxisShiftPerFocusX
NX20 | OpticalAxisShiftPerFocusY
NX20 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX20 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX20 | Lift pins synchronization offset check (X offset)
NX20 | Lift pins synchronization offset check (Y offset)
NX20 | Profiler noise @200um
NX20 | Roll motion to X+ full scan
NX20 | Position repeatability(10repeats, Position Based)
NX20 | OPM Repeatability @5mm
NX20 | OPM Repeatability @1mm
NX20 | Max OPM @5mm
NX20 | Max OPM @1mm
NX20 | Torque Profile (10mm, 200um/s)
NX20 | Torque Profile (10mm, 20um/s)
NX20 | Lateral(C-D) sensitivity
NX20 | Lateral(C-D) noise
NX20 | Min. Vibration Transmissbility (~100Hz)
NX20 | RMS of ZeroScan
NX20 | LightStrengthGain
NX20 | Full range of X Detector during X Scan sweep
NX20 | Full range of X1 Detector during X Scan sweep
NX20 | Full range of X2 Detector during X Scan sweep
NX20 | Full range of Y Detector during Y Scan sweep
NX20 | Full range of Y1 Detector during Y Scan sweep
NX20 | Full range of Y2 Detector during Y Scan sweep
NX20 | Full range of X1 Detector Raw during X Scan sweep
NX20 | Full range of X2 Detector Raw during X Scan sweep
NX20 | Full range of Y1 Detector Raw during Y Scan sweep
NX20 | Full range of Y2 Detector Raw during Y Scan sweep
NX20 | X Detector stroke calibration error (nominal scale error X)
NX20 | Y Detector stroke calibration error (nominal scale error Y)
NX20 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX20 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX20 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX20 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX20 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX20 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX20 | X Non-linearity error
NX20 | Y Non-linearity error
NX20 | X axis OPM
NX20 | Y axis OPM
NX20 | X Ringing(Rpv)
NX20 | Y Ringing(Rpv)
NX20 | X Detector noise (XY Servo on)
NX20 | Y Detector noise (XY Servo on)
NX20 | X1 Detector noise (XY Servo on)
NX20 | Y1 Detector noise (XY Servo on)
NX20 | X2 Detector noise (XY Servo on)
NX20 | Y2 Detector noise (XY Servo on)
NX20 | RMS error after Stage Calibration
NX20 | Image offset after moving from -X to +X (X scan)
NX20 | Image offset after moving from -Y to +Y (Y scan)
NX20 | Image offset after moving from -X to +X (Y scan)
NX20 | Image offset after moving from -Y to +Y (X scan)
NX20 | Image offset after moving from +X to -X (X scan)
NX20 | Image offset after moving from +Y to -Y (Y scan)
NX20 | Image offset after moving from +X to -X (Y scan)
NX20 | Image offset after moving from +Y to -Y (X scan)
NX20 | XY Stage positioning X offset (Vision Pattern Matching)
NX20 | XY Stage positioning Y offset (Vision Pattern Matching)
NX20 | XY Stage positioning X 1σ (Vision Pattern Matching)
NX20 | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX20 | XY Stage positioning X offset (In-Die Align)
NX20 | XY Stage positioning Y offset (In-Die Align)
NX20 | XY Stage positioning X 1σ (In-Die Align)
NX20 | XY Stage positioning Y 1σ (In-Die Align)
NX20 | XY Stage positioning X offset (LLC Translation)
NX20 | XY Stage positioning Y offset (LLC Translation)
NX20 | XY Stage positioning X 1σ (LLC Translation)
NX20 | XY Stage positioning Y 1σ (LLC Translation)
NX20 | XY Stage positioning X offset (Global Align)
NX20 | XY Stage positioning Y offset (Global Align)
NX20 | XY Stage positioning X 1σ (Global Align)
NX20 | XY Stage positioning Y 1σ (Global Align)
NX20 | Z Drift
NX20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | Cycle time for 15-point measurement
NX20 | Short-term stability using Step Height Sample (3σ)
NX20 | Roughness Repeatability 1x1um(RSD%)
NX20 | Roughness Repeatability 1x1um(1σ)
NX20 | Roughness Repeatability 80x80um(RSD%)
NX20 | Roughness Repeatability 80x80um(1σ)
NX20 | Z Height calibrtaion error
NX20 | Z Drive calibrtaion error
NX20 | Z Linearity by 3 different sample height
NX20 | Full range in voltage during Z Scan sweep
NX20 | Full range in um during Z Scan sweep
NX20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 | Lateral(C-D) sensitivity
NX20 | Lateral(C-D) noise
NX20 | A-B noise
NX20 | NCM Amplitude noise
NX20 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX20 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX20 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX20 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX20 | RMS of Z Height of Wafer Zeroscan(0,120000)
NX20 | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX20 | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX20 | RMS of Z Height of Wafer Zeroscan(120000,0)
NX20 | RMS of Z Height of Wafer ZeroScan (0,0)
NX20 | Max Wafer Zeroscan-Min Wafer Zeroscan
NX20 | XY Orthogonality
NX20 | MicrometerPerAdc of Z Height (Z Height calibration)
NX20 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX20 | Z linearity by sample height
NX20 | 1) Z Scanner position at +3μm
NX20 | 3) Z Scanner position at -3μm
NX20 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX20 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX20 | X Scan Stroke
NX20 | Y Scan Stroke
NX20 | Full range of X Detector Raw during X Scan sweep
NX20 | Full range of Y Detector Raw during Y Scan sweep
NX20 | X Detector noise (XY Servo off)
NX20 | Y Detector noise (XY Servo off)
NX20 | X axis OPM
NX20 | Y axis OPM
NX20 | X Non-linearity
NX20 | Y Non-linearity
NX20 | OPM Repeatability @25mm
NX20 | OPM Repeatability @10mm
NX20 | Max OPM @25mm
NX20 | Max OPM @10mm
NX20 | Torque Profile (50mm, 200um/s)
NX20 | Torque Profile (50mm, 20um/s)
NX20 | SLD current limit check
NX20 | SLD beam power at the end of the lens
NX20 | 2) of 20nm
NX20 | X Detector noise (XYServo off)
NX20 | Y Detector noise (XYServo off)
NX20 | X axis OPM raw Rpv
NX20 | Y axis OPM raw Rpv
NX20 | StageSafetyXMax
NX20 | StageSafetyXMin
NX20 | StageSafetyY
NX20 | StageSafetyYMax
NX20 | StageSafetyYMin
NX20 | Check X axis upper limit sensor function
NX20 | Check X axis lower limit sensor function
NX20 | Check Y axis upper limit sensor function
NX20 | Check Y axis lower limit sensor function
NX20 | X axis Angle Range
NX20 | Short-term stability w/ 450nm Step Height Sample (3σ)
NX20 | X detector noise (XY Servo on)
NX20 | Y detector noise (XY Servo on)
NX20 | Full range in um during X Scan sweep (XY Servo off)
NX20 | Full range in um during Y Scan sweep (XY Servo off)
NX20 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX20 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX20 | X axis, Repeatability @1st position
NX20 | X axis, Repeatability @ 2nd position
NX20 | X axis, Repeatability @ 3rd position
NX20 | X axis, Repeatability @ 4th position
NX20 | X axis, Accuracy @1st position
NX20 | X axis, Accuracy @ 2nd position
NX20 | X axis, Accuracy @ 3rd position
NX20 | X axis, Accuracy @ 4th position
NX20 | Y axis, Repeatability @1st position
NX20 | Y axis, Repeatability @ 2nd position
NX20 | Y axis, Repeatability @ 3rd position
NX20 | Y axis, Repeatability @ 4th position
NX20 | Y axis, Accuracy @1st position
NX20 | Y axis, Accuracy @ 2nd position
NX20 | Y axis, Accuracy @ 3rd position
NX20 | Y axis, Accuracy @ 4th position
NX20 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX20 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX20 | Z-XY Orthogonality (X direction-Left)
NX20 | Z-XY Orthogonality (X direction-Right)
NX20 | Z-XY Orthogonality (Y direction-Left)
NX20 | Z-XY Orthogonality (Y direction-Right)
NX20 | X-Scan Repeatability (10 repeats, RSD%)
NX20 | Y-Scan Repeatability (10 repeats, RSD%)
NX20 | Mechanical Align Accuracy
NX20 | Check the Multi Peattern Upload
NX20 | X Detector noise (XY servo off)
NX20 | Y Detector noise (XY servo off)
NX20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 | Z Detector Max
NX20 | Z Detector Min
NX20 | WLI_Scale(Z Calibration)
NX20 | RMS of Z Height of Wafer Zeroscan(0,90000)
NX20 | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX20 | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX20 | RMS of Z Height of Wafer Zeroscan(90000,0)
NX20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | Short-term stability w/ 180nm Step Height Sample (3σ)
NX20 | Laser intensity with AC160TS cantilever
NX20 | Check if the driving % changes proportionally as the NCM amplitude changes
NX20 | NCM driving % at an amplitude of 1 nm
NX20 | NCM driving % at an amplitude of 20 nm
NX20 | NCM driving % at an amplitude of 100 nm
NX20 | - 940nm Step Height Sample (10 Repeats)
NX20 | - 200~500nm Step Height Sample (10 Repeats)
NX20 | - 50~200nm Step Height Sample (10 Repeats)
NX20 | - Under 50nm Step Height Sample (10 Repeats)
NX20 | A-B after Auto PSPD Alignment
NX20 | NCM driving% with AC160TS cantilever
NX20 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 | X Detector stroke calibration error (nominal scale error X)
NX20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | Resonance Frequency in DB
NX20 | Servo Cut-off Frequency in DB
NX20 | Check chuck surface no damage
NX20 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX20 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX20 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX20 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX20 | X1 Detector noise (XY servo off)
NX20 | X2 Detector noise (XY servo off)
NX20 | Y1 Detector noise (XY servo off)
NX20 | Y2 Detector noise (XY servo off)
NX20 | X Detector noise (XY servo on)
NX20 | X1 Detector noise (XY servo on)
NX20 | X2 Detector noise (XY servo on)
NX20 | Y Detector noise (XY servo on)
NX20 | Y1 Detector noise (XY servo on)
NX20 | Y2 Detector noise (XY servo on)
NX20 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 | Z Detector offset
NX20 | Check the full stroke image(servo on)
NX20 | ResonanceFrequencyHz
NX20 | ServoCutoffFrequencyHz
NX20 | FullPixelHeight
NX20 | Serial Port#
NX20 | Off-axis Camera X Position(ΔX)
NX20 | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX20 | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX20 | Z Scanner position at -3μm
NX20 | Z Scanner position at +3μm
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 | Scan offset accuracy (Precision)
NX20 | Aligned stage accuracy (X offset-Point Align1)
NX20 | Aligned stage accuracy (Y offset-Point Align1)
NX20 | Aligned stage accuracy (X offset-Point Align2)
NX20 | Aligned stage accuracy (Y offset-Point Align2)
NX20 | Aligned stage accuracy (X offset-Point Align3)
NX20 | Aligned stage accuracy (Y offset-Point Align3)
NX20 | Aligned stage accuracy (X offset)
NX20 | Aligned stage accuracy (Y offset)
NX20 | Imag offset after EFEM loading (X scan)
NX20 | Imag offset after EFEM loading (Y scan)
NX20 | Imag offset after Lock stage (X scan)
NX20 | Imag offset after Lock stage (Y scan)
NX20 | Mask Roughness Repeatability (10Repeats, RSD%)
NX20 | Mask Roughness Repeatability (10Repeats, σ )
NX20 | Mask exchange time
NX20 | WLI_Scale(C Calibration)
NX20 | WLI_Scale(B Calibration)
NX20 | WLI_Scale(A Calibration)
NX20 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX20 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX20 | Port 1 lens(to OnAxis) Y Offset
NX20 | Port 1 lens(to OnAxis) X Offset
NX20 | P Gain
NX20 | I Gain
NX20 | D Gain
NX20 | Check the Port 2 lens light strength
NX20 | Check the Port 1 lens light strength
NX20 | Check the focus deviation between the lenses(Δf)
NX20 | Check the 25mm Moving – Y Axis Offset Check
NX20 | C Scanner Detector Min
NX20 | C Scanner Detector Max
NX20 | B Scanner Detector Min
NX20 | B Scanner Detector Max
NX20 | Angle align between Port 2 lens and VLSI SHS: Y
NX20 | Angle align between Port 2 lens and VLSI SHS: X
NX20 | Angle align between Port 2 lens and Sample chuck: Y
NX20 | Angle align between Port 2 lens and Sample chuck: X
NX20 | Angle align between Port 1 lens and Sample chuck: Y
NX20 | Angle align between Port 1 lens and Sample chuck: X
NX20 | A Scanner Detector Min
NX20 | A Scanner Detector Max
NX20 | Y2 Detector noise (XY Servo off)
NX20 | X2 Detector noise (XY Servo off)
NX20 | Y axis OPM (-52000,-52000)
NX20 | Y axis OPM (-52000,52000)
NX20 | Y axis OPM (52000,-52000)
NX20 | Y axis OPM (52000,52000)
NX20 | Y axis OPM (0,0)
NX20 | XY Orthogonality (-52000,-52000)
NX20 | XY Orthogonality (-52000,52000)
NX20 | XY Orthogonality (52000,-52000)
NX20 | XY Orthogonality (52000,52000)
NX20 | XY Orthogonality (0,0)
NX20 | X axis OPM (-52000,-52000)
NX20 | X axis OPM (-52000,52000)
NX20 | X axis OPM (52000,-52000)
NX20 | X axis OPM (52000,52000)
NX20 | X axis OPM (0,0)
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX20 | RMS of Z Height of Wafer Zeroscan(60000,0)
NX20 | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX20 | RMS of Z Height of Wafer Zeroscan(0,60000)
NX20 | XY stage calibration
NX20 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX20 | LD beam power at the end of the lens
NX20 | Y Non-lineairty
NX20 | X Non-lineairty
NX20 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX20 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX20 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX20 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX20 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX20 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX20 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX20 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX20 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX20 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX20 | UpperLimit
NX20 | Under 50nm Step Height Sample (10 Repeats)
NX20 | Overlay Camera X Position(ΔX)
NX20 | LowerLimit
NX20 | Initial Z
NX20 | 940nm Step Height Sample (10 Repeats)
NX20 | 50~200nm Step Height Sample (10 Repeats)
NX20 | RMS of Z Height of C122 Zeroscan(0,-90000)
NX20 | StrokeYMin
NX20 | StrokeYMax
NX20 | StrokeY
NX20 | StrokeXMin
NX20 | StrokeXMax
NX20 | StrokeX
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX20 | Z Drift (Short Traveling)
NX20 | 200~500nm Step Height Sample (10 Repeats)
NX20 | XY orthogonality
NX20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX20 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX20 | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX20 | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX20 | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX20 | Y1 Detector noise (XY Servo off)
NX20 | X1 Detector noise (XY Servo off)
NX20 | Maximum Y Scan range (Full Stroke)
NX20 | Maximum X Scan range (Full Stroke)
NX20 | Full range of Y2 Detector raw
NX20 | Full range of Y1 Detector raw
NX20 | Full range of X2 Detector raw
NX20 | Full range of X1 Detector raw
NX20 | Z Linearity by Z Scanner position
NX20 | -Y Direction software Limit
NX20 | -X Direction software Limit
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX20 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX20 | %GRR w/ roughness of non-patterened wafer
NX20 | Z Scanner position at -6μm
NX20 | Z Scanner position at +6μm
NX20 | Approach position at -5μm
NX20 | Approach position at +5μm
NX20 | * 1800nm Step Height Sample (10 Repeats)
NX20 | Full range of Z Detector raw
NX20 | Approach position at -3μm
NX20 | Approach position at +3μm
NX20 | Under 50nm Step Height Sample (10 Repeats)
NX20 | 50~200nm Step Height Sample (10 Repeats)
NX20 | 200~500nm Step Height Sample (10 Repeats)
NX20 | PosMin
NX20 | PosMax
NX20 | MaxLength
NX20 | FullRecoveryProfileLength
NX20 | Full range of Y Detector raw
NX20 | Full range of X Detector raw
NX20 | EFM Phase : For the positive bias
NX20 | EFM Phase : For the negative bias
NX20 | Z Linearity by 5 different sample height
NX20 | Laser intensity with AC160 cantilever
NX20 | A－B difference during 3 times run test
NX20 | A＋B difference during 3 times run test
NX20 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX20 | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX20 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX20 | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX20 | RMS of Z Drive of Wafer ZeroScan (0,0)
NX20 | Allowable max warpage, w/ SiC sample chuck
NX20 | RMS of Z Height of Wafer Zeroscan (Top)
NX20 | RMS of Z Height of Wafer Zeroscan (Right)
NX20 | RMS of Z Height of Wafer Zeroscan (Left)
NX20 | RMS of Z Height of Wafer ZeroScan (Center)
NX20 | RMS of Z Height of Wafer Zeroscan (Bottom)
NX20 | PosMin
NX20 | PosMax
NX20 | MaxLength
NX20 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX20 | Check sample surface mapping
NX20 | Z Direction, Accuracy Test (Bias not significant)
NX20 | Y Direction, Accuracy Test (Bias not significant)
NX20 | X Direction, Accuracy Test (Bias not significant)
NX20 | Roughness Reproducibility Test (SR Panel)
NX20 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX20 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX20 | Roughness Reproducibility Test (SiNx Panel)
NX20 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX20 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX20 | Roughness Reproducibility Test (ABF Panel)
NX20 | Roughness Repeatability Test (SR Panel)
NX20 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX20 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX20 | Roughness Repeatability Test (SiNx Panel)
NX20 | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX20 | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX20 | Roughness Repeatability Test (ABF Panel)
NX20 | Lift Pin Stroke Min Check in DB
NX20 | Lift Pin Stroke Max Check in DB
NX20 | Limit range
NX20 Lite | Z Linearity by Z scanner position
NX20 Lite | Full range of Z Detector during Z Scan sweep
NX20 Lite | Z Height calibration error
NX20 Lite | Z Drive calibration error
NX20 Lite | Z Linearity by 4 different sample height
NX20 Lite | Full range of Z Height during Z Scan sweep
NX20 Lite | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 Lite | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 Lite | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 Lite | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 Lite | Z Height noise
NX20 Lite | Wafer placement(Loading) error on the load port1 (X offset)
NX20 Lite | Wafer placement(Loading) error on the load port1 (Y offset)
NX20 Lite | Wafer placement(Loading) error on the load port2 (X offset)
NX20 Lite | Wafer placement(Loading) error on the load port2 (Y offset)
NX20 Lite | OpticalAxisShiftPerFocusX
NX20 Lite | OpticalAxisShiftPerFocusY
NX20 Lite | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX20 Lite | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX20 Lite | Lift pins synchronization offset check (X offset)
NX20 Lite | Lift pins synchronization offset check (Y offset)
NX20 Lite | Profiler noise @200um
NX20 Lite | Roll motion to X+ full scan
NX20 Lite | Position repeatability(10repeats, Position Based)
NX20 Lite | OPM Repeatability @5mm
NX20 Lite | OPM Repeatability @1mm
NX20 Lite | Max OPM @5mm
NX20 Lite | Max OPM @1mm
NX20 Lite | Torque Profile (10mm, 200um/s)
NX20 Lite | Torque Profile (10mm, 20um/s)
NX20 Lite | Lateral(C-D) sensitivity
NX20 Lite | Lateral(C-D) noise
NX20 Lite | Min. Vibration Transmissbility (~100Hz)
NX20 Lite | RMS of ZeroScan
NX20 Lite | LightStrengthGain
NX20 Lite | Full range of X Detector during X Scan sweep
NX20 Lite | Full range of X1 Detector during X Scan sweep
NX20 Lite | Full range of X2 Detector during X Scan sweep
NX20 Lite | Full range of Y Detector during Y Scan sweep
NX20 Lite | Full range of Y1 Detector during Y Scan sweep
NX20 Lite | Full range of Y2 Detector during Y Scan sweep
NX20 Lite | Full range of X1 Detector Raw during X Scan sweep
NX20 Lite | Full range of X2 Detector Raw during X Scan sweep
NX20 Lite | Full range of Y1 Detector Raw during Y Scan sweep
NX20 Lite | Full range of Y2 Detector Raw during Y Scan sweep
NX20 Lite | X Detector stroke calibration error (nominal scale error X)
NX20 Lite | Y Detector stroke calibration error (nominal scale error Y)
NX20 Lite | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX20 Lite | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX20 Lite | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX20 Lite | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX20 Lite | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX20 Lite | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX20 Lite | X Non-linearity error
NX20 Lite | Y Non-linearity error
NX20 Lite | X axis OPM
NX20 Lite | Y axis OPM
NX20 Lite | X Ringing(Rpv)
NX20 Lite | Y Ringing(Rpv)
NX20 Lite | X Detector noise (XY Servo on)
NX20 Lite | Y Detector noise (XY Servo on)
NX20 Lite | X1 Detector noise (XY Servo on)
NX20 Lite | Y1 Detector noise (XY Servo on)
NX20 Lite | X2 Detector noise (XY Servo on)
NX20 Lite | Y2 Detector noise (XY Servo on)
NX20 Lite | RMS error after Stage Calibration
NX20 Lite | Image offset after moving from -X to +X (X scan)
NX20 Lite | Image offset after moving from -Y to +Y (Y scan)
NX20 Lite | Image offset after moving from -X to +X (Y scan)
NX20 Lite | Image offset after moving from -Y to +Y (X scan)
NX20 Lite | Image offset after moving from +X to -X (X scan)
NX20 Lite | Image offset after moving from +Y to -Y (Y scan)
NX20 Lite | Image offset after moving from +X to -X (Y scan)
NX20 Lite | Image offset after moving from +Y to -Y (X scan)
NX20 Lite | XY Stage positioning X offset (Vision Pattern Matching)
NX20 Lite | XY Stage positioning Y offset (Vision Pattern Matching)
NX20 Lite | XY Stage positioning X 1σ (Vision Pattern Matching)
NX20 Lite | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX20 Lite | XY Stage positioning X offset (In-Die Align)
NX20 Lite | XY Stage positioning Y offset (In-Die Align)
NX20 Lite | XY Stage positioning X 1σ (In-Die Align)
NX20 Lite | XY Stage positioning Y 1σ (In-Die Align)
NX20 Lite | XY Stage positioning X offset (LLC Translation)
NX20 Lite | XY Stage positioning Y offset (LLC Translation)
NX20 Lite | XY Stage positioning X 1σ (LLC Translation)
NX20 Lite | XY Stage positioning Y 1σ (LLC Translation)
NX20 Lite | XY Stage positioning X offset (Global Align)
NX20 Lite | XY Stage positioning Y offset (Global Align)
NX20 Lite | XY Stage positioning X 1σ (Global Align)
NX20 Lite | XY Stage positioning Y 1σ (Global Align)
NX20 Lite | Z Drift
NX20 Lite | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | Cycle time for 15-point measurement
NX20 Lite | Short-term stability using Step Height Sample (3σ)
NX20 Lite | Roughness Repeatability 1x1um(RSD%)
NX20 Lite | Roughness Repeatability 1x1um(1σ)
NX20 Lite | Roughness Repeatability 80x80um(RSD%)
NX20 Lite | Roughness Repeatability 80x80um(1σ)
NX20 Lite | Z Height calibrtaion error
NX20 Lite | Z Drive calibrtaion error
NX20 Lite | Z Linearity by 3 different sample height
NX20 Lite | Full range in voltage during Z Scan sweep
NX20 Lite | Full range in um during Z Scan sweep
NX20 Lite | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 Lite | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 Lite | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 Lite | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 Lite | Lateral(C-D) sensitivity
NX20 Lite | Lateral(C-D) noise
NX20 Lite | A-B noise
NX20 Lite | NCM Amplitude noise
NX20 Lite | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX20 Lite | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX20 Lite | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX20 Lite | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(0,120000)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(120000,0)
NX20 Lite | RMS of Z Height of Wafer ZeroScan (0,0)
NX20 Lite | Max Wafer Zeroscan-Min Wafer Zeroscan
NX20 Lite | XY Orthogonality
NX20 Lite | MicrometerPerAdc of Z Height (Z Height calibration)
NX20 Lite | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX20 Lite | Z linearity by sample height
NX20 Lite | 1) Z Scanner position at +3μm
NX20 Lite | 3) Z Scanner position at -3μm
NX20 Lite | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX20 Lite | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX20 Lite | X Scan Stroke
NX20 Lite | Y Scan Stroke
NX20 Lite | Full range of X Detector Raw during X Scan sweep
NX20 Lite | Full range of Y Detector Raw during Y Scan sweep
NX20 Lite | X Detector noise (XY Servo off)
NX20 Lite | Y Detector noise (XY Servo off)
NX20 Lite | X axis OPM
NX20 Lite | Y axis OPM
NX20 Lite | X Non-linearity
NX20 Lite | Y Non-linearity
NX20 Lite | OPM Repeatability @25mm
NX20 Lite | OPM Repeatability @10mm
NX20 Lite | Max OPM @25mm
NX20 Lite | Max OPM @10mm
NX20 Lite | Torque Profile (50mm, 200um/s)
NX20 Lite | Torque Profile (50mm, 20um/s)
NX20 Lite | SLD current limit check
NX20 Lite | SLD beam power at the end of the lens
NX20 Lite | 2) of 20nm
NX20 Lite | X Detector noise (XYServo off)
NX20 Lite | Y Detector noise (XYServo off)
NX20 Lite | X axis OPM raw Rpv
NX20 Lite | Y axis OPM raw Rpv
NX20 Lite | StageSafetyXMax
NX20 Lite | StageSafetyXMin
NX20 Lite | StageSafetyY
NX20 Lite | StageSafetyYMax
NX20 Lite | StageSafetyYMin
NX20 Lite | Check X axis upper limit sensor function
NX20 Lite | Check X axis lower limit sensor function
NX20 Lite | Check Y axis upper limit sensor function
NX20 Lite | Check Y axis lower limit sensor function
NX20 Lite | X axis Angle Range
NX20 Lite | Short-term stability w/ 450nm Step Height Sample (3σ)
NX20 Lite | X detector noise (XY Servo on)
NX20 Lite | Y detector noise (XY Servo on)
NX20 Lite | Full range in um during X Scan sweep (XY Servo off)
NX20 Lite | Full range in um during Y Scan sweep (XY Servo off)
NX20 Lite | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX20 Lite | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX20 Lite | X axis, Repeatability @1st position
NX20 Lite | X axis, Repeatability @ 2nd position
NX20 Lite | X axis, Repeatability @ 3rd position
NX20 Lite | X axis, Repeatability @ 4th position
NX20 Lite | X axis, Accuracy @1st position
NX20 Lite | X axis, Accuracy @ 2nd position
NX20 Lite | X axis, Accuracy @ 3rd position
NX20 Lite | X axis, Accuracy @ 4th position
NX20 Lite | Y axis, Repeatability @1st position
NX20 Lite | Y axis, Repeatability @ 2nd position
NX20 Lite | Y axis, Repeatability @ 3rd position
NX20 Lite | Y axis, Repeatability @ 4th position
NX20 Lite | Y axis, Accuracy @1st position
NX20 Lite | Y axis, Accuracy @ 2nd position
NX20 Lite | Y axis, Accuracy @ 3rd position
NX20 Lite | Y axis, Accuracy @ 4th position
NX20 Lite | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX20 Lite | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX20 Lite | Z-XY Orthogonality (X direction-Left)
NX20 Lite | Z-XY Orthogonality (X direction-Right)
NX20 Lite | Z-XY Orthogonality (Y direction-Left)
NX20 Lite | Z-XY Orthogonality (Y direction-Right)
NX20 Lite | X-Scan Repeatability (10 repeats, RSD%)
NX20 Lite | Y-Scan Repeatability (10 repeats, RSD%)
NX20 Lite | Mechanical Align Accuracy
NX20 Lite | Check the Multi Peattern Upload
NX20 Lite | X Detector noise (XY servo off)
NX20 Lite | Y Detector noise (XY servo off)
NX20 Lite | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 Lite | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 Lite | Z Detector Max
NX20 Lite | Z Detector Min
NX20 Lite | WLI_Scale(Z Calibration)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(0,90000)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(90000,0)
NX20 Lite | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | Short-term stability w/ 180nm Step Height Sample (3σ)
NX20 Lite | Laser intensity with AC160TS cantilever
NX20 Lite | Check if the driving % changes proportionally as the NCM amplitude changes
NX20 Lite | NCM driving % at an amplitude of 1 nm
NX20 Lite | NCM driving % at an amplitude of 20 nm
NX20 Lite | NCM driving % at an amplitude of 100 nm
NX20 Lite | - 940nm Step Height Sample (10 Repeats)
NX20 Lite | - 200~500nm Step Height Sample (10 Repeats)
NX20 Lite | - 50~200nm Step Height Sample (10 Repeats)
NX20 Lite | - Under 50nm Step Height Sample (10 Repeats)
NX20 Lite | A-B after Auto PSPD Alignment
NX20 Lite | NCM driving% with AC160TS cantilever
NX20 Lite | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX20 Lite | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX20 Lite | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX20 Lite | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX20 Lite | X Detector stroke calibration error (nominal scale error X)
NX20 Lite | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | Resonance Frequency in DB
NX20 Lite | Servo Cut-off Frequency in DB
NX20 Lite | Check chuck surface no damage
NX20 Lite | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX20 Lite | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX20 Lite | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX20 Lite | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX20 Lite | X1 Detector noise (XY servo off)
NX20 Lite | X2 Detector noise (XY servo off)
NX20 Lite | Y1 Detector noise (XY servo off)
NX20 Lite | Y2 Detector noise (XY servo off)
NX20 Lite | X Detector noise (XY servo on)
NX20 Lite | X1 Detector noise (XY servo on)
NX20 Lite | X2 Detector noise (XY servo on)
NX20 Lite | Y Detector noise (XY servo on)
NX20 Lite | Y1 Detector noise (XY servo on)
NX20 Lite | Y2 Detector noise (XY servo on)
NX20 Lite | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX20 Lite | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX20 Lite | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 Lite | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 Lite | Z Detector offset
NX20 Lite | Check the full stroke image(servo on)
NX20 Lite | ResonanceFrequencyHz
NX20 Lite | ServoCutoffFrequencyHz
NX20 Lite | FullPixelHeight
NX20 Lite | Serial Port#
NX20 Lite | Off-axis Camera X Position(ΔX)
NX20 Lite | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX20 Lite | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX20 Lite | Z Scanner position at -3μm
NX20 Lite | Z Scanner position at +3μm
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 Lite | Scan offset accuracy (Precision)
NX20 Lite | Aligned stage accuracy (X offset-Point Align1)
NX20 Lite | Aligned stage accuracy (Y offset-Point Align1)
NX20 Lite | Aligned stage accuracy (X offset-Point Align2)
NX20 Lite | Aligned stage accuracy (Y offset-Point Align2)
NX20 Lite | Aligned stage accuracy (X offset-Point Align3)
NX20 Lite | Aligned stage accuracy (Y offset-Point Align3)
NX20 Lite | Aligned stage accuracy (X offset)
NX20 Lite | Aligned stage accuracy (Y offset)
NX20 Lite | Imag offset after EFEM loading (X scan)
NX20 Lite | Imag offset after EFEM loading (Y scan)
NX20 Lite | Imag offset after Lock stage (X scan)
NX20 Lite | Imag offset after Lock stage (Y scan)
NX20 Lite | Mask Roughness Repeatability (10Repeats, RSD%)
NX20 Lite | Mask Roughness Repeatability (10Repeats, σ )
NX20 Lite | Mask exchange time
NX20 Lite | WLI_Scale(C Calibration)
NX20 Lite | WLI_Scale(B Calibration)
NX20 Lite | WLI_Scale(A Calibration)
NX20 Lite | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX20 Lite | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX20 Lite | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX20 Lite | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX20 Lite | Port 1 lens(to OnAxis) Y Offset
NX20 Lite | Port 1 lens(to OnAxis) X Offset
NX20 Lite | P Gain
NX20 Lite | I Gain
NX20 Lite | D Gain
NX20 Lite | Check the Port 2 lens light strength
NX20 Lite | Check the Port 1 lens light strength
NX20 Lite | Check the focus deviation between the lenses(Δf)
NX20 Lite | Check the 25mm Moving – Y Axis Offset Check
NX20 Lite | C Scanner Detector Min
NX20 Lite | C Scanner Detector Max
NX20 Lite | B Scanner Detector Min
NX20 Lite | B Scanner Detector Max
NX20 Lite | Angle align between Port 2 lens and VLSI SHS: Y
NX20 Lite | Angle align between Port 2 lens and VLSI SHS: X
NX20 Lite | Angle align between Port 2 lens and Sample chuck: Y
NX20 Lite | Angle align between Port 2 lens and Sample chuck: X
NX20 Lite | Angle align between Port 1 lens and Sample chuck: Y
NX20 Lite | Angle align between Port 1 lens and Sample chuck: X
NX20 Lite | A Scanner Detector Min
NX20 Lite | A Scanner Detector Max
NX20 Lite | Y2 Detector noise (XY Servo off)
NX20 Lite | X2 Detector noise (XY Servo off)
NX20 Lite | Y axis OPM (-52000,-52000)
NX20 Lite | Y axis OPM (-52000,52000)
NX20 Lite | Y axis OPM (52000,-52000)
NX20 Lite | Y axis OPM (52000,52000)
NX20 Lite | Y axis OPM (0,0)
NX20 Lite | XY Orthogonality (-52000,-52000)
NX20 Lite | XY Orthogonality (-52000,52000)
NX20 Lite | XY Orthogonality (52000,-52000)
NX20 Lite | XY Orthogonality (52000,52000)
NX20 Lite | XY Orthogonality (0,0)
NX20 Lite | X axis OPM (-52000,-52000)
NX20 Lite | X axis OPM (-52000,52000)
NX20 Lite | X axis OPM (52000,-52000)
NX20 Lite | X axis OPM (52000,52000)
NX20 Lite | X axis OPM (0,0)
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(60000,0)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX20 Lite | RMS of Z Height of Wafer Zeroscan(0,60000)
NX20 Lite | XY stage calibration
NX20 Lite | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX20 Lite | LD beam power at the end of the lens
NX20 Lite | Y Non-lineairty
NX20 Lite | X Non-lineairty
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX20 Lite | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX20 Lite | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX20 Lite | UpperLimit
NX20 Lite | Under 50nm Step Height Sample (10 Repeats)
NX20 Lite | Overlay Camera X Position(ΔX)
NX20 Lite | LowerLimit
NX20 Lite | Initial Z
NX20 Lite | 940nm Step Height Sample (10 Repeats)
NX20 Lite | 50~200nm Step Height Sample (10 Repeats)
NX20 Lite | RMS of Z Height of C122 Zeroscan(0,-90000)
NX20 Lite | StrokeYMin
NX20 Lite | StrokeYMax
NX20 Lite | StrokeY
NX20 Lite | StrokeXMin
NX20 Lite | StrokeXMax
NX20 Lite | StrokeX
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX20 Lite | Z Drift (Short Traveling)
NX20 Lite | 200~500nm Step Height Sample (10 Repeats)
NX20 Lite | XY orthogonality
NX20 Lite | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX20 Lite | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX20 Lite | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX20 Lite | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX20 Lite | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX20 Lite | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX20 Lite | Y1 Detector noise (XY Servo off)
NX20 Lite | X1 Detector noise (XY Servo off)
NX20 Lite | Maximum Y Scan range (Full Stroke)
NX20 Lite | Maximum X Scan range (Full Stroke)
NX20 Lite | Full range of Y2 Detector raw
NX20 Lite | Full range of Y1 Detector raw
NX20 Lite | Full range of X2 Detector raw
NX20 Lite | Full range of X1 Detector raw
NX20 Lite | Z Linearity by Z Scanner position
NX20 Lite | -Y Direction software Limit
NX20 Lite | -X Direction software Limit
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX20 Lite | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX20 Lite | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX20 Lite | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX20 Lite | %GRR w/ roughness of non-patterened wafer
NX20 Lite | Z Scanner position at -6μm
NX20 Lite | Z Scanner position at +6μm
NX20 Lite | Approach position at -5μm
NX20 Lite | Approach position at +5μm
NX20 Lite | * 1800nm Step Height Sample (10 Repeats)
NX20 Lite | Full range of Z Detector raw
NX20 Lite | Approach position at -3μm
NX20 Lite | Approach position at +3μm
NX20 Lite | Under 50nm Step Height Sample (10 Repeats)
NX20 Lite | 50~200nm Step Height Sample (10 Repeats)
NX20 Lite | 200~500nm Step Height Sample (10 Repeats)
NX20 Lite | PosMin
NX20 Lite | PosMax
NX20 Lite | MaxLength
NX20 Lite | FullRecoveryProfileLength
NX20 Lite | Full range of Y Detector raw
NX20 Lite | Full range of X Detector raw
NX20 Lite | EFM Phase : For the positive bias
NX20 Lite | EFM Phase : For the negative bias
NX20 Lite | Z Linearity by 5 different sample height
NX20 Lite | Laser intensity with AC160 cantilever
NX20 Lite | A－B difference during 3 times run test
NX20 Lite | A＋B difference during 3 times run test
NX20 Lite | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX20 Lite | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX20 Lite | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX20 Lite | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX20 Lite | RMS of Z Drive of Wafer ZeroScan (0,0)
NX20 Lite | Allowable max warpage, w/ SiC sample chuck
NX20 Lite | RMS of Z Height of Wafer Zeroscan (Top)
NX20 Lite | RMS of Z Height of Wafer Zeroscan (Right)
NX20 Lite | RMS of Z Height of Wafer Zeroscan (Left)
NX20 Lite | RMS of Z Height of Wafer ZeroScan (Center)
NX20 Lite | RMS of Z Height of Wafer Zeroscan (Bottom)
NX20 Lite | PosMin
NX20 Lite | PosMax
NX20 Lite | MaxLength
NX20 Lite | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX20 Lite | Check sample surface mapping
NX20 Lite | Z Direction, Accuracy Test (Bias not significant)
NX20 Lite | Y Direction, Accuracy Test (Bias not significant)
NX20 Lite | X Direction, Accuracy Test (Bias not significant)
NX20 Lite | Roughness Reproducibility Test (SR Panel)
NX20 Lite | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX20 Lite | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX20 Lite | Roughness Reproducibility Test (SiNx Panel)
NX20 Lite | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX20 Lite | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX20 Lite | Roughness Reproducibility Test (ABF Panel)
NX20 Lite | Roughness Repeatability Test (SR Panel)
NX20 Lite | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX20 Lite | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX20 Lite | Roughness Repeatability Test (SiNx Panel)
NX20 Lite | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX20 Lite | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX20 Lite | Roughness Repeatability Test (ABF Panel)
NX20 Lite | Lift Pin Stroke Min Check in DB
NX20 Lite | Lift Pin Stroke Max Check in DB
NX20 Lite | Limit range
NX-TSH400TF | Z Linearity by Z scanner position
NX-TSH400TF | Full range of Z Detector during Z Scan sweep
NX-TSH400TF | Z Height calibration error
NX-TSH400TF | Z Drive calibration error
NX-TSH400TF | Z Linearity by 4 different sample height
NX-TSH400TF | Full range of Z Height during Z Scan sweep
NX-TSH400TF | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH400TF | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH400TF | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH400TF | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH400TF | Z Height noise
NX-TSH400TF | Wafer placement(Loading) error on the load port1 (X offset)
NX-TSH400TF | Wafer placement(Loading) error on the load port1 (Y offset)
NX-TSH400TF | Wafer placement(Loading) error on the load port2 (X offset)
NX-TSH400TF | Wafer placement(Loading) error on the load port2 (Y offset)
NX-TSH400TF | OpticalAxisShiftPerFocusX
NX-TSH400TF | OpticalAxisShiftPerFocusY
NX-TSH400TF | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-TSH400TF | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-TSH400TF | Lift pins synchronization offset check (X offset)
NX-TSH400TF | Lift pins synchronization offset check (Y offset)
NX-TSH400TF | Profiler noise @200um
NX-TSH400TF | Roll motion to X+ full scan
NX-TSH400TF | Position repeatability(10repeats, Position Based)
NX-TSH400TF | OPM Repeatability @5mm
NX-TSH400TF | OPM Repeatability @1mm
NX-TSH400TF | Max OPM @5mm
NX-TSH400TF | Max OPM @1mm
NX-TSH400TF | Torque Profile (10mm, 200um/s)
NX-TSH400TF | Torque Profile (10mm, 20um/s)
NX-TSH400TF | Lateral(C-D) sensitivity
NX-TSH400TF | Lateral(C-D) noise
NX-TSH400TF | Min. Vibration Transmissbility (~100Hz)
NX-TSH400TF | RMS of ZeroScan
NX-TSH400TF | LightStrengthGain
NX-TSH400TF | Full range of X Detector during X Scan sweep
NX-TSH400TF | Full range of X1 Detector during X Scan sweep
NX-TSH400TF | Full range of X2 Detector during X Scan sweep
NX-TSH400TF | Full range of Y Detector during Y Scan sweep
NX-TSH400TF | Full range of Y1 Detector during Y Scan sweep
NX-TSH400TF | Full range of Y2 Detector during Y Scan sweep
NX-TSH400TF | Full range of X1 Detector Raw during X Scan sweep
NX-TSH400TF | Full range of X2 Detector Raw during X Scan sweep
NX-TSH400TF | Full range of Y1 Detector Raw during Y Scan sweep
NX-TSH400TF | Full range of Y2 Detector Raw during Y Scan sweep
NX-TSH400TF | X Detector stroke calibration error (nominal scale error X)
NX-TSH400TF | Y Detector stroke calibration error (nominal scale error Y)
NX-TSH400TF | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-TSH400TF | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-TSH400TF | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-TSH400TF | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-TSH400TF | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-TSH400TF | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-TSH400TF | X Non-linearity error
NX-TSH400TF | Y Non-linearity error
NX-TSH400TF | X axis OPM
NX-TSH400TF | Y axis OPM
NX-TSH400TF | X Ringing(Rpv)
NX-TSH400TF | Y Ringing(Rpv)
NX-TSH400TF | X Detector noise (XY Servo on)
NX-TSH400TF | Y Detector noise (XY Servo on)
NX-TSH400TF | X1 Detector noise (XY Servo on)
NX-TSH400TF | Y1 Detector noise (XY Servo on)
NX-TSH400TF | X2 Detector noise (XY Servo on)
NX-TSH400TF | Y2 Detector noise (XY Servo on)
NX-TSH400TF | RMS error after Stage Calibration
NX-TSH400TF | Image offset after moving from -X to +X (X scan)
NX-TSH400TF | Image offset after moving from -Y to +Y (Y scan)
NX-TSH400TF | Image offset after moving from -X to +X (Y scan)
NX-TSH400TF | Image offset after moving from -Y to +Y (X scan)
NX-TSH400TF | Image offset after moving from +X to -X (X scan)
NX-TSH400TF | Image offset after moving from +Y to -Y (Y scan)
NX-TSH400TF | Image offset after moving from +X to -X (Y scan)
NX-TSH400TF | Image offset after moving from +Y to -Y (X scan)
NX-TSH400TF | XY Stage positioning X offset (Vision Pattern Matching)
NX-TSH400TF | XY Stage positioning Y offset (Vision Pattern Matching)
NX-TSH400TF | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-TSH400TF | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-TSH400TF | XY Stage positioning X offset (In-Die Align)
NX-TSH400TF | XY Stage positioning Y offset (In-Die Align)
NX-TSH400TF | XY Stage positioning X 1σ (In-Die Align)
NX-TSH400TF | XY Stage positioning Y 1σ (In-Die Align)
NX-TSH400TF | XY Stage positioning X offset (LLC Translation)
NX-TSH400TF | XY Stage positioning Y offset (LLC Translation)
NX-TSH400TF | XY Stage positioning X 1σ (LLC Translation)
NX-TSH400TF | XY Stage positioning Y 1σ (LLC Translation)
NX-TSH400TF | XY Stage positioning X offset (Global Align)
NX-TSH400TF | XY Stage positioning Y offset (Global Align)
NX-TSH400TF | XY Stage positioning X 1σ (Global Align)
NX-TSH400TF | XY Stage positioning Y 1σ (Global Align)
NX-TSH400TF | Z Drift
NX-TSH400TF | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | Cycle time for 15-point measurement
NX-TSH400TF | Short-term stability using Step Height Sample (3σ)
NX-TSH400TF | Roughness Repeatability 1x1um(RSD%)
NX-TSH400TF | Roughness Repeatability 1x1um(1σ)
NX-TSH400TF | Roughness Repeatability 80x80um(RSD%)
NX-TSH400TF | Roughness Repeatability 80x80um(1σ)
NX-TSH400TF | Z Height calibrtaion error
NX-TSH400TF | Z Drive calibrtaion error
NX-TSH400TF | Z Linearity by 3 different sample height
NX-TSH400TF | Full range in voltage during Z Scan sweep
NX-TSH400TF | Full range in um during Z Scan sweep
NX-TSH400TF | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH400TF | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH400TF | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH400TF | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH400TF | Lateral(C-D) sensitivity
NX-TSH400TF | Lateral(C-D) noise
NX-TSH400TF | A-B noise
NX-TSH400TF | NCM Amplitude noise
NX-TSH400TF | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-TSH400TF | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-TSH400TF | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-TSH400TF | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-TSH400TF | RMS of Z Height of Wafer ZeroScan (0,0)
NX-TSH400TF | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-TSH400TF | XY Orthogonality
NX-TSH400TF | MicrometerPerAdc of Z Height (Z Height calibration)
NX-TSH400TF | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-TSH400TF | Z linearity by sample height
NX-TSH400TF | 1) Z Scanner position at +3μm
NX-TSH400TF | 3) Z Scanner position at -3μm
NX-TSH400TF | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-TSH400TF | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-TSH400TF | X Scan Stroke
NX-TSH400TF | Y Scan Stroke
NX-TSH400TF | Full range of X Detector Raw during X Scan sweep
NX-TSH400TF | Full range of Y Detector Raw during Y Scan sweep
NX-TSH400TF | X Detector noise (XY Servo off)
NX-TSH400TF | Y Detector noise (XY Servo off)
NX-TSH400TF | X axis OPM
NX-TSH400TF | Y axis OPM
NX-TSH400TF | X Non-linearity
NX-TSH400TF | Y Non-linearity
NX-TSH400TF | OPM Repeatability @25mm
NX-TSH400TF | OPM Repeatability @10mm
NX-TSH400TF | Max OPM @25mm
NX-TSH400TF | Max OPM @10mm
NX-TSH400TF | Torque Profile (50mm, 200um/s)
NX-TSH400TF | Torque Profile (50mm, 20um/s)
NX-TSH400TF | SLD current limit check
NX-TSH400TF | SLD beam power at the end of the lens
NX-TSH400TF | 2) of 20nm
NX-TSH400TF | X Detector noise (XYServo off)
NX-TSH400TF | Y Detector noise (XYServo off)
NX-TSH400TF | X axis OPM raw Rpv
NX-TSH400TF | Y axis OPM raw Rpv
NX-TSH400TF | StageSafetyXMax
NX-TSH400TF | StageSafetyXMin
NX-TSH400TF | StageSafetyY
NX-TSH400TF | StageSafetyYMax
NX-TSH400TF | StageSafetyYMin
NX-TSH400TF | Check X axis upper limit sensor function
NX-TSH400TF | Check X axis lower limit sensor function
NX-TSH400TF | Check Y axis upper limit sensor function
NX-TSH400TF | Check Y axis lower limit sensor function
NX-TSH400TF | X axis Angle Range
NX-TSH400TF | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-TSH400TF | X detector noise (XY Servo on)
NX-TSH400TF | Y detector noise (XY Servo on)
NX-TSH400TF | Full range in um during X Scan sweep (XY Servo off)
NX-TSH400TF | Full range in um during Y Scan sweep (XY Servo off)
NX-TSH400TF | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-TSH400TF | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-TSH400TF | X axis, Repeatability @1st position
NX-TSH400TF | X axis, Repeatability @ 2nd position
NX-TSH400TF | X axis, Repeatability @ 3rd position
NX-TSH400TF | X axis, Repeatability @ 4th position
NX-TSH400TF | X axis, Accuracy @1st position
NX-TSH400TF | X axis, Accuracy @ 2nd position
NX-TSH400TF | X axis, Accuracy @ 3rd position
NX-TSH400TF | X axis, Accuracy @ 4th position
NX-TSH400TF | Y axis, Repeatability @1st position
NX-TSH400TF | Y axis, Repeatability @ 2nd position
NX-TSH400TF | Y axis, Repeatability @ 3rd position
NX-TSH400TF | Y axis, Repeatability @ 4th position
NX-TSH400TF | Y axis, Accuracy @1st position
NX-TSH400TF | Y axis, Accuracy @ 2nd position
NX-TSH400TF | Y axis, Accuracy @ 3rd position
NX-TSH400TF | Y axis, Accuracy @ 4th position
NX-TSH400TF | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-TSH400TF | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-TSH400TF | Z-XY Orthogonality (X direction-Left)
NX-TSH400TF | Z-XY Orthogonality (X direction-Right)
NX-TSH400TF | Z-XY Orthogonality (Y direction-Left)
NX-TSH400TF | Z-XY Orthogonality (Y direction-Right)
NX-TSH400TF | X-Scan Repeatability (10 repeats, RSD%)
NX-TSH400TF | Y-Scan Repeatability (10 repeats, RSD%)
NX-TSH400TF | Mechanical Align Accuracy
NX-TSH400TF | Check the Multi Peattern Upload
NX-TSH400TF | X Detector noise (XY servo off)
NX-TSH400TF | Y Detector noise (XY servo off)
NX-TSH400TF | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH400TF | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH400TF | Z Detector Max
NX-TSH400TF | Z Detector Min
NX-TSH400TF | WLI_Scale(Z Calibration)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-TSH400TF | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-TSH400TF | Laser intensity with AC160TS cantilever
NX-TSH400TF | Check if the driving % changes proportionally as the NCM amplitude changes
NX-TSH400TF | NCM driving % at an amplitude of 1 nm
NX-TSH400TF | NCM driving % at an amplitude of 20 nm
NX-TSH400TF | NCM driving % at an amplitude of 100 nm
NX-TSH400TF | - 940nm Step Height Sample (10 Repeats)
NX-TSH400TF | - 200~500nm Step Height Sample (10 Repeats)
NX-TSH400TF | - 50~200nm Step Height Sample (10 Repeats)
NX-TSH400TF | - Under 50nm Step Height Sample (10 Repeats)
NX-TSH400TF | A-B after Auto PSPD Alignment
NX-TSH400TF | NCM driving% with AC160TS cantilever
NX-TSH400TF | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH400TF | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH400TF | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH400TF | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH400TF | X Detector stroke calibration error (nominal scale error X)
NX-TSH400TF | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | Resonance Frequency in DB
NX-TSH400TF | Servo Cut-off Frequency in DB
NX-TSH400TF | Check chuck surface no damage
NX-TSH400TF | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-TSH400TF | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-TSH400TF | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-TSH400TF | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-TSH400TF | X1 Detector noise (XY servo off)
NX-TSH400TF | X2 Detector noise (XY servo off)
NX-TSH400TF | Y1 Detector noise (XY servo off)
NX-TSH400TF | Y2 Detector noise (XY servo off)
NX-TSH400TF | X Detector noise (XY servo on)
NX-TSH400TF | X1 Detector noise (XY servo on)
NX-TSH400TF | X2 Detector noise (XY servo on)
NX-TSH400TF | Y Detector noise (XY servo on)
NX-TSH400TF | Y1 Detector noise (XY servo on)
NX-TSH400TF | Y2 Detector noise (XY servo on)
NX-TSH400TF | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH400TF | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH400TF | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH400TF | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH400TF | Z Detector offset
NX-TSH400TF | Check the full stroke image(servo on)
NX-TSH400TF | ResonanceFrequencyHz
NX-TSH400TF | ServoCutoffFrequencyHz
NX-TSH400TF | FullPixelHeight
NX-TSH400TF | Serial Port#
NX-TSH400TF | Off-axis Camera X Position(ΔX)
NX-TSH400TF | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-TSH400TF | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-TSH400TF | Z Scanner position at -3μm
NX-TSH400TF | Z Scanner position at +3μm
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH400TF | Scan offset accuracy (Precision)
NX-TSH400TF | Aligned stage accuracy (X offset-Point Align1)
NX-TSH400TF | Aligned stage accuracy (Y offset-Point Align1)
NX-TSH400TF | Aligned stage accuracy (X offset-Point Align2)
NX-TSH400TF | Aligned stage accuracy (Y offset-Point Align2)
NX-TSH400TF | Aligned stage accuracy (X offset-Point Align3)
NX-TSH400TF | Aligned stage accuracy (Y offset-Point Align3)
NX-TSH400TF | Aligned stage accuracy (X offset)
NX-TSH400TF | Aligned stage accuracy (Y offset)
NX-TSH400TF | Imag offset after EFEM loading (X scan)
NX-TSH400TF | Imag offset after EFEM loading (Y scan)
NX-TSH400TF | Imag offset after Lock stage (X scan)
NX-TSH400TF | Imag offset after Lock stage (Y scan)
NX-TSH400TF | Mask Roughness Repeatability (10Repeats, RSD%)
NX-TSH400TF | Mask Roughness Repeatability (10Repeats, σ )
NX-TSH400TF | Mask exchange time
NX-TSH400TF | WLI_Scale(C Calibration)
NX-TSH400TF | WLI_Scale(B Calibration)
NX-TSH400TF | WLI_Scale(A Calibration)
NX-TSH400TF | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH400TF | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH400TF | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-TSH400TF | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-TSH400TF | Port 1 lens(to OnAxis) Y Offset
NX-TSH400TF | Port 1 lens(to OnAxis) X Offset
NX-TSH400TF | P Gain
NX-TSH400TF | I Gain
NX-TSH400TF | D Gain
NX-TSH400TF | Check the Port 2 lens light strength
NX-TSH400TF | Check the Port 1 lens light strength
NX-TSH400TF | Check the focus deviation between the lenses(Δf)
NX-TSH400TF | Check the 25mm Moving – Y Axis Offset Check
NX-TSH400TF | C Scanner Detector Min
NX-TSH400TF | C Scanner Detector Max
NX-TSH400TF | B Scanner Detector Min
NX-TSH400TF | B Scanner Detector Max
NX-TSH400TF | Angle align between Port 2 lens and VLSI SHS: Y
NX-TSH400TF | Angle align between Port 2 lens and VLSI SHS: X
NX-TSH400TF | Angle align between Port 2 lens and Sample chuck: Y
NX-TSH400TF | Angle align between Port 2 lens and Sample chuck: X
NX-TSH400TF | Angle align between Port 1 lens and Sample chuck: Y
NX-TSH400TF | Angle align between Port 1 lens and Sample chuck: X
NX-TSH400TF | A Scanner Detector Min
NX-TSH400TF | A Scanner Detector Max
NX-TSH400TF | Y2 Detector noise (XY Servo off)
NX-TSH400TF | X2 Detector noise (XY Servo off)
NX-TSH400TF | Y axis OPM (-52000,-52000)
NX-TSH400TF | Y axis OPM (-52000,52000)
NX-TSH400TF | Y axis OPM (52000,-52000)
NX-TSH400TF | Y axis OPM (52000,52000)
NX-TSH400TF | Y axis OPM (0,0)
NX-TSH400TF | XY Orthogonality (-52000,-52000)
NX-TSH400TF | XY Orthogonality (-52000,52000)
NX-TSH400TF | XY Orthogonality (52000,-52000)
NX-TSH400TF | XY Orthogonality (52000,52000)
NX-TSH400TF | XY Orthogonality (0,0)
NX-TSH400TF | X axis OPM (-52000,-52000)
NX-TSH400TF | X axis OPM (-52000,52000)
NX-TSH400TF | X axis OPM (52000,-52000)
NX-TSH400TF | X axis OPM (52000,52000)
NX-TSH400TF | X axis OPM (0,0)
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-TSH400TF | XY stage calibration
NX-TSH400TF | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-TSH400TF | LD beam power at the end of the lens
NX-TSH400TF | Y Non-lineairty
NX-TSH400TF | X Non-lineairty
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-TSH400TF | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-TSH400TF | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-TSH400TF | UpperLimit
NX-TSH400TF | Under 50nm Step Height Sample (10 Repeats)
NX-TSH400TF | Overlay Camera X Position(ΔX)
NX-TSH400TF | LowerLimit
NX-TSH400TF | Initial Z
NX-TSH400TF | 940nm Step Height Sample (10 Repeats)
NX-TSH400TF | 50~200nm Step Height Sample (10 Repeats)
NX-TSH400TF | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-TSH400TF | StrokeYMin
NX-TSH400TF | StrokeYMax
NX-TSH400TF | StrokeY
NX-TSH400TF | StrokeXMin
NX-TSH400TF | StrokeXMax
NX-TSH400TF | StrokeX
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-TSH400TF | Z Drift (Short Traveling)
NX-TSH400TF | 200~500nm Step Height Sample (10 Repeats)
NX-TSH400TF | XY orthogonality
NX-TSH400TF | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-TSH400TF | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-TSH400TF | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-TSH400TF | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-TSH400TF | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-TSH400TF | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-TSH400TF | Y1 Detector noise (XY Servo off)
NX-TSH400TF | X1 Detector noise (XY Servo off)
NX-TSH400TF | Maximum Y Scan range (Full Stroke)
NX-TSH400TF | Maximum X Scan range (Full Stroke)
NX-TSH400TF | Full range of Y2 Detector raw
NX-TSH400TF | Full range of Y1 Detector raw
NX-TSH400TF | Full range of X2 Detector raw
NX-TSH400TF | Full range of X1 Detector raw
NX-TSH400TF | Z Linearity by Z Scanner position
NX-TSH400TF | -Y Direction software Limit
NX-TSH400TF | -X Direction software Limit
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH400TF | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH400TF | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-TSH400TF | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-TSH400TF | %GRR w/ roughness of non-patterened wafer
NX-TSH400TF | Z Scanner position at -6μm
NX-TSH400TF | Z Scanner position at +6μm
NX-TSH400TF | Approach position at -5μm
NX-TSH400TF | Approach position at +5μm
NX-TSH400TF | * 1800nm Step Height Sample (10 Repeats)
NX-TSH400TF | Full range of Z Detector raw
NX-TSH400TF | Approach position at -3μm
NX-TSH400TF | Approach position at +3μm
NX-TSH400TF | Under 50nm Step Height Sample (10 Repeats)
NX-TSH400TF | 50~200nm Step Height Sample (10 Repeats)
NX-TSH400TF | 200~500nm Step Height Sample (10 Repeats)
NX-TSH400TF | PosMin
NX-TSH400TF | PosMax
NX-TSH400TF | MaxLength
NX-TSH400TF | FullRecoveryProfileLength
NX-TSH400TF | Full range of Y Detector raw
NX-TSH400TF | Full range of X Detector raw
NX-TSH400TF | EFM Phase : For the positive bias
NX-TSH400TF | EFM Phase : For the negative bias
NX-TSH400TF | Z Linearity by 5 different sample height
NX-TSH400TF | Laser intensity with AC160 cantilever
NX-TSH400TF | A－B difference during 3 times run test
NX-TSH400TF | A＋B difference during 3 times run test
NX-TSH400TF | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-TSH400TF | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-TSH400TF | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-TSH400TF | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-TSH400TF | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-TSH400TF | Allowable max warpage, w/ SiC sample chuck
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan (Top)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan (Right)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan (Left)
NX-TSH400TF | RMS of Z Height of Wafer ZeroScan (Center)
NX-TSH400TF | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-TSH400TF | PosMin
NX-TSH400TF | PosMax
NX-TSH400TF | MaxLength
NX-TSH400TF | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-TSH400TF | Check sample surface mapping
NX-TSH400TF | Z Direction, Accuracy Test (Bias not significant)
NX-TSH400TF | Y Direction, Accuracy Test (Bias not significant)
NX-TSH400TF | X Direction, Accuracy Test (Bias not significant)
NX-TSH400TF | Roughness Reproducibility Test (SR Panel)
NX-TSH400TF | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-TSH400TF | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-TSH400TF | Roughness Reproducibility Test (SiNx Panel)
NX-TSH400TF | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-TSH400TF | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-TSH400TF | Roughness Reproducibility Test (ABF Panel)
NX-TSH400TF | Roughness Repeatability Test (SR Panel)
NX-TSH400TF | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-TSH400TF | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-TSH400TF | Roughness Repeatability Test (SiNx Panel)
NX-TSH400TF | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-TSH400TF | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-TSH400TF | Roughness Repeatability Test (ABF Panel)
NX-TSH400TF | Lift Pin Stroke Min Check in DB
NX-TSH400TF | Lift Pin Stroke Max Check in DB
NX-TSH400TF | Limit range
NX-Wafer Plus | Z Linearity by Z scanner position
NX-Wafer Plus | Full range of Z Detector during Z Scan sweep
NX-Wafer Plus | Z Height calibration error
NX-Wafer Plus | Z Drive calibration error
NX-Wafer Plus | Z Linearity by 4 different sample height
NX-Wafer Plus | Full range of Z Height during Z Scan sweep
NX-Wafer Plus | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer Plus | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer Plus | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer Plus | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer Plus | Z Height noise
NX-Wafer Plus | Wafer placement(Loading) error on the load port1 (X offset)
NX-Wafer Plus | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Wafer Plus | Wafer placement(Loading) error on the load port2 (X offset)
NX-Wafer Plus | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Wafer Plus | OpticalAxisShiftPerFocusX
NX-Wafer Plus | OpticalAxisShiftPerFocusY
NX-Wafer Plus | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Wafer Plus | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Wafer Plus | Lift pins synchronization offset check (X offset)
NX-Wafer Plus | Lift pins synchronization offset check (Y offset)
NX-Wafer Plus | Profiler noise @200um
NX-Wafer Plus | Roll motion to X+ full scan
NX-Wafer Plus | Position repeatability(10repeats, Position Based)
NX-Wafer Plus | OPM Repeatability @5mm
NX-Wafer Plus | OPM Repeatability @1mm
NX-Wafer Plus | Max OPM @5mm
NX-Wafer Plus | Max OPM @1mm
NX-Wafer Plus | Torque Profile (10mm, 200um/s)
NX-Wafer Plus | Torque Profile (10mm, 20um/s)
NX-Wafer Plus | Lateral(C-D) sensitivity
NX-Wafer Plus | Lateral(C-D) noise
NX-Wafer Plus | Min. Vibration Transmissbility (~100Hz)
NX-Wafer Plus | RMS of ZeroScan
NX-Wafer Plus | LightStrengthGain
NX-Wafer Plus | Full range of X Detector during X Scan sweep
NX-Wafer Plus | Full range of X1 Detector during X Scan sweep
NX-Wafer Plus | Full range of X2 Detector during X Scan sweep
NX-Wafer Plus | Full range of Y Detector during Y Scan sweep
NX-Wafer Plus | Full range of Y1 Detector during Y Scan sweep
NX-Wafer Plus | Full range of Y2 Detector during Y Scan sweep
NX-Wafer Plus | Full range of X1 Detector Raw during X Scan sweep
NX-Wafer Plus | Full range of X2 Detector Raw during X Scan sweep
NX-Wafer Plus | Full range of Y1 Detector Raw during Y Scan sweep
NX-Wafer Plus | Full range of Y2 Detector Raw during Y Scan sweep
NX-Wafer Plus | X Detector stroke calibration error (nominal scale error X)
NX-Wafer Plus | Y Detector stroke calibration error (nominal scale error Y)
NX-Wafer Plus | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Wafer Plus | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Wafer Plus | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Wafer Plus | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Wafer Plus | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Wafer Plus | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Wafer Plus | X Non-linearity error
NX-Wafer Plus | Y Non-linearity error
NX-Wafer Plus | X axis OPM
NX-Wafer Plus | Y axis OPM
NX-Wafer Plus | X Ringing(Rpv)
NX-Wafer Plus | Y Ringing(Rpv)
NX-Wafer Plus | X Detector noise (XY Servo on)
NX-Wafer Plus | Y Detector noise (XY Servo on)
NX-Wafer Plus | X1 Detector noise (XY Servo on)
NX-Wafer Plus | Y1 Detector noise (XY Servo on)
NX-Wafer Plus | X2 Detector noise (XY Servo on)
NX-Wafer Plus | Y2 Detector noise (XY Servo on)
NX-Wafer Plus | RMS error after Stage Calibration
NX-Wafer Plus | Image offset after moving from -X to +X (X scan)
NX-Wafer Plus | Image offset after moving from -Y to +Y (Y scan)
NX-Wafer Plus | Image offset after moving from -X to +X (Y scan)
NX-Wafer Plus | Image offset after moving from -Y to +Y (X scan)
NX-Wafer Plus | Image offset after moving from +X to -X (X scan)
NX-Wafer Plus | Image offset after moving from +Y to -Y (Y scan)
NX-Wafer Plus | Image offset after moving from +X to -X (Y scan)
NX-Wafer Plus | Image offset after moving from +Y to -Y (X scan)
NX-Wafer Plus | XY Stage positioning X offset (Vision Pattern Matching)
NX-Wafer Plus | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Wafer Plus | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Wafer Plus | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Wafer Plus | XY Stage positioning X offset (In-Die Align)
NX-Wafer Plus | XY Stage positioning Y offset (In-Die Align)
NX-Wafer Plus | XY Stage positioning X 1σ (In-Die Align)
NX-Wafer Plus | XY Stage positioning Y 1σ (In-Die Align)
NX-Wafer Plus | XY Stage positioning X offset (LLC Translation)
NX-Wafer Plus | XY Stage positioning Y offset (LLC Translation)
NX-Wafer Plus | XY Stage positioning X 1σ (LLC Translation)
NX-Wafer Plus | XY Stage positioning Y 1σ (LLC Translation)
NX-Wafer Plus | XY Stage positioning X offset (Global Align)
NX-Wafer Plus | XY Stage positioning Y offset (Global Align)
NX-Wafer Plus | XY Stage positioning X 1σ (Global Align)
NX-Wafer Plus | XY Stage positioning Y 1σ (Global Align)
NX-Wafer Plus | Z Drift
NX-Wafer Plus | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | Cycle time for 15-point measurement
NX-Wafer Plus | Short-term stability using Step Height Sample (3σ)
NX-Wafer Plus | Roughness Repeatability 1x1um(RSD%)
NX-Wafer Plus | Roughness Repeatability 1x1um(1σ)
NX-Wafer Plus | Roughness Repeatability 80x80um(RSD%)
NX-Wafer Plus | Roughness Repeatability 80x80um(1σ)
NX-Wafer Plus | Z Height calibrtaion error
NX-Wafer Plus | Z Drive calibrtaion error
NX-Wafer Plus | Z Linearity by 3 different sample height
NX-Wafer Plus | Full range in voltage during Z Scan sweep
NX-Wafer Plus | Full range in um during Z Scan sweep
NX-Wafer Plus | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer Plus | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer Plus | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer Plus | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer Plus | Lateral(C-D) sensitivity
NX-Wafer Plus | Lateral(C-D) noise
NX-Wafer Plus | A-B noise
NX-Wafer Plus | NCM Amplitude noise
NX-Wafer Plus | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Wafer Plus | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Wafer Plus | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Wafer Plus | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Wafer Plus | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Wafer Plus | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Wafer Plus | XY Orthogonality
NX-Wafer Plus | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Wafer Plus | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Wafer Plus | Z linearity by sample height
NX-Wafer Plus | 1) Z Scanner position at +3μm
NX-Wafer Plus | 3) Z Scanner position at -3μm
NX-Wafer Plus | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Wafer Plus | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Wafer Plus | X Scan Stroke
NX-Wafer Plus | Y Scan Stroke
NX-Wafer Plus | Full range of X Detector Raw during X Scan sweep
NX-Wafer Plus | Full range of Y Detector Raw during Y Scan sweep
NX-Wafer Plus | X Detector noise (XY Servo off)
NX-Wafer Plus | Y Detector noise (XY Servo off)
NX-Wafer Plus | X axis OPM
NX-Wafer Plus | Y axis OPM
NX-Wafer Plus | X Non-linearity
NX-Wafer Plus | Y Non-linearity
NX-Wafer Plus | OPM Repeatability @25mm
NX-Wafer Plus | OPM Repeatability @10mm
NX-Wafer Plus | Max OPM @25mm
NX-Wafer Plus | Max OPM @10mm
NX-Wafer Plus | Torque Profile (50mm, 200um/s)
NX-Wafer Plus | Torque Profile (50mm, 20um/s)
NX-Wafer Plus | SLD current limit check
NX-Wafer Plus | SLD beam power at the end of the lens
NX-Wafer Plus | 2) of 20nm
NX-Wafer Plus | X Detector noise (XYServo off)
NX-Wafer Plus | Y Detector noise (XYServo off)
NX-Wafer Plus | X axis OPM raw Rpv
NX-Wafer Plus | Y axis OPM raw Rpv
NX-Wafer Plus | StageSafetyXMax
NX-Wafer Plus | StageSafetyXMin
NX-Wafer Plus | StageSafetyY
NX-Wafer Plus | StageSafetyYMax
NX-Wafer Plus | StageSafetyYMin
NX-Wafer Plus | Check X axis upper limit sensor function
NX-Wafer Plus | Check X axis lower limit sensor function
NX-Wafer Plus | Check Y axis upper limit sensor function
NX-Wafer Plus | Check Y axis lower limit sensor function
NX-Wafer Plus | X axis Angle Range
NX-Wafer Plus | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Wafer Plus | X detector noise (XY Servo on)
NX-Wafer Plus | Y detector noise (XY Servo on)
NX-Wafer Plus | Full range in um during X Scan sweep (XY Servo off)
NX-Wafer Plus | Full range in um during Y Scan sweep (XY Servo off)
NX-Wafer Plus | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Wafer Plus | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Wafer Plus | X axis, Repeatability @1st position
NX-Wafer Plus | X axis, Repeatability @ 2nd position
NX-Wafer Plus | X axis, Repeatability @ 3rd position
NX-Wafer Plus | X axis, Repeatability @ 4th position
NX-Wafer Plus | X axis, Accuracy @1st position
NX-Wafer Plus | X axis, Accuracy @ 2nd position
NX-Wafer Plus | X axis, Accuracy @ 3rd position
NX-Wafer Plus | X axis, Accuracy @ 4th position
NX-Wafer Plus | Y axis, Repeatability @1st position
NX-Wafer Plus | Y axis, Repeatability @ 2nd position
NX-Wafer Plus | Y axis, Repeatability @ 3rd position
NX-Wafer Plus | Y axis, Repeatability @ 4th position
NX-Wafer Plus | Y axis, Accuracy @1st position
NX-Wafer Plus | Y axis, Accuracy @ 2nd position
NX-Wafer Plus | Y axis, Accuracy @ 3rd position
NX-Wafer Plus | Y axis, Accuracy @ 4th position
NX-Wafer Plus | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Wafer Plus | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Wafer Plus | Z-XY Orthogonality (X direction-Left)
NX-Wafer Plus | Z-XY Orthogonality (X direction-Right)
NX-Wafer Plus | Z-XY Orthogonality (Y direction-Left)
NX-Wafer Plus | Z-XY Orthogonality (Y direction-Right)
NX-Wafer Plus | X-Scan Repeatability (10 repeats, RSD%)
NX-Wafer Plus | Y-Scan Repeatability (10 repeats, RSD%)
NX-Wafer Plus | Mechanical Align Accuracy
NX-Wafer Plus | Check the Multi Peattern Upload
NX-Wafer Plus | X Detector noise (XY servo off)
NX-Wafer Plus | Y Detector noise (XY servo off)
NX-Wafer Plus | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer Plus | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer Plus | Z Detector Max
NX-Wafer Plus | Z Detector Min
NX-Wafer Plus | WLI_Scale(Z Calibration)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Wafer Plus | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Wafer Plus | Laser intensity with AC160TS cantilever
NX-Wafer Plus | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Wafer Plus | NCM driving % at an amplitude of 1 nm
NX-Wafer Plus | NCM driving % at an amplitude of 20 nm
NX-Wafer Plus | NCM driving % at an amplitude of 100 nm
NX-Wafer Plus | - 940nm Step Height Sample (10 Repeats)
NX-Wafer Plus | - 200~500nm Step Height Sample (10 Repeats)
NX-Wafer Plus | - 50~200nm Step Height Sample (10 Repeats)
NX-Wafer Plus | - Under 50nm Step Height Sample (10 Repeats)
NX-Wafer Plus | A-B after Auto PSPD Alignment
NX-Wafer Plus | NCM driving% with AC160TS cantilever
NX-Wafer Plus | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer Plus | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer Plus | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer Plus | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer Plus | X Detector stroke calibration error (nominal scale error X)
NX-Wafer Plus | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | Resonance Frequency in DB
NX-Wafer Plus | Servo Cut-off Frequency in DB
NX-Wafer Plus | Check chuck surface no damage
NX-Wafer Plus | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Wafer Plus | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Wafer Plus | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Wafer Plus | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Wafer Plus | X1 Detector noise (XY servo off)
NX-Wafer Plus | X2 Detector noise (XY servo off)
NX-Wafer Plus | Y1 Detector noise (XY servo off)
NX-Wafer Plus | Y2 Detector noise (XY servo off)
NX-Wafer Plus | X Detector noise (XY servo on)
NX-Wafer Plus | X1 Detector noise (XY servo on)
NX-Wafer Plus | X2 Detector noise (XY servo on)
NX-Wafer Plus | Y Detector noise (XY servo on)
NX-Wafer Plus | Y1 Detector noise (XY servo on)
NX-Wafer Plus | Y2 Detector noise (XY servo on)
NX-Wafer Plus | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Plus | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Plus | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer Plus | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer Plus | Z Detector offset
NX-Wafer Plus | Check the full stroke image(servo on)
NX-Wafer Plus | ResonanceFrequencyHz
NX-Wafer Plus | ServoCutoffFrequencyHz
NX-Wafer Plus | FullPixelHeight
NX-Wafer Plus | Serial Port#
NX-Wafer Plus | Off-axis Camera X Position(ΔX)
NX-Wafer Plus | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Wafer Plus | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Wafer Plus | Z Scanner position at -3μm
NX-Wafer Plus | Z Scanner position at +3μm
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer Plus | Scan offset accuracy (Precision)
NX-Wafer Plus | Aligned stage accuracy (X offset-Point Align1)
NX-Wafer Plus | Aligned stage accuracy (Y offset-Point Align1)
NX-Wafer Plus | Aligned stage accuracy (X offset-Point Align2)
NX-Wafer Plus | Aligned stage accuracy (Y offset-Point Align2)
NX-Wafer Plus | Aligned stage accuracy (X offset-Point Align3)
NX-Wafer Plus | Aligned stage accuracy (Y offset-Point Align3)
NX-Wafer Plus | Aligned stage accuracy (X offset)
NX-Wafer Plus | Aligned stage accuracy (Y offset)
NX-Wafer Plus | Imag offset after EFEM loading (X scan)
NX-Wafer Plus | Imag offset after EFEM loading (Y scan)
NX-Wafer Plus | Imag offset after Lock stage (X scan)
NX-Wafer Plus | Imag offset after Lock stage (Y scan)
NX-Wafer Plus | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Wafer Plus | Mask Roughness Repeatability (10Repeats, σ )
NX-Wafer Plus | Mask exchange time
NX-Wafer Plus | WLI_Scale(C Calibration)
NX-Wafer Plus | WLI_Scale(B Calibration)
NX-Wafer Plus | WLI_Scale(A Calibration)
NX-Wafer Plus | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer Plus | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer Plus | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Wafer Plus | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Wafer Plus | Port 1 lens(to OnAxis) Y Offset
NX-Wafer Plus | Port 1 lens(to OnAxis) X Offset
NX-Wafer Plus | P Gain
NX-Wafer Plus | I Gain
NX-Wafer Plus | D Gain
NX-Wafer Plus | Check the Port 2 lens light strength
NX-Wafer Plus | Check the Port 1 lens light strength
NX-Wafer Plus | Check the focus deviation between the lenses(Δf)
NX-Wafer Plus | Check the 25mm Moving – Y Axis Offset Check
NX-Wafer Plus | C Scanner Detector Min
NX-Wafer Plus | C Scanner Detector Max
NX-Wafer Plus | B Scanner Detector Min
NX-Wafer Plus | B Scanner Detector Max
NX-Wafer Plus | Angle align between Port 2 lens and VLSI SHS: Y
NX-Wafer Plus | Angle align between Port 2 lens and VLSI SHS: X
NX-Wafer Plus | Angle align between Port 2 lens and Sample chuck: Y
NX-Wafer Plus | Angle align between Port 2 lens and Sample chuck: X
NX-Wafer Plus | Angle align between Port 1 lens and Sample chuck: Y
NX-Wafer Plus | Angle align between Port 1 lens and Sample chuck: X
NX-Wafer Plus | A Scanner Detector Min
NX-Wafer Plus | A Scanner Detector Max
NX-Wafer Plus | Y2 Detector noise (XY Servo off)
NX-Wafer Plus | X2 Detector noise (XY Servo off)
NX-Wafer Plus | Y axis OPM (-52000,-52000)
NX-Wafer Plus | Y axis OPM (-52000,52000)
NX-Wafer Plus | Y axis OPM (52000,-52000)
NX-Wafer Plus | Y axis OPM (52000,52000)
NX-Wafer Plus | Y axis OPM (0,0)
NX-Wafer Plus | XY Orthogonality (-52000,-52000)
NX-Wafer Plus | XY Orthogonality (-52000,52000)
NX-Wafer Plus | XY Orthogonality (52000,-52000)
NX-Wafer Plus | XY Orthogonality (52000,52000)
NX-Wafer Plus | XY Orthogonality (0,0)
NX-Wafer Plus | X axis OPM (-52000,-52000)
NX-Wafer Plus | X axis OPM (-52000,52000)
NX-Wafer Plus | X axis OPM (52000,-52000)
NX-Wafer Plus | X axis OPM (52000,52000)
NX-Wafer Plus | X axis OPM (0,0)
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Wafer Plus | XY stage calibration
NX-Wafer Plus | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Wafer Plus | LD beam power at the end of the lens
NX-Wafer Plus | Y Non-lineairty
NX-Wafer Plus | X Non-lineairty
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Wafer Plus | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Wafer Plus | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Wafer Plus | UpperLimit
NX-Wafer Plus | Under 50nm Step Height Sample (10 Repeats)
NX-Wafer Plus | Overlay Camera X Position(ΔX)
NX-Wafer Plus | LowerLimit
NX-Wafer Plus | Initial Z
NX-Wafer Plus | 940nm Step Height Sample (10 Repeats)
NX-Wafer Plus | 50~200nm Step Height Sample (10 Repeats)
NX-Wafer Plus | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Wafer Plus | StrokeYMin
NX-Wafer Plus | StrokeYMax
NX-Wafer Plus | StrokeY
NX-Wafer Plus | StrokeXMin
NX-Wafer Plus | StrokeXMax
NX-Wafer Plus | StrokeX
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Wafer Plus | Z Drift (Short Traveling)
NX-Wafer Plus | 200~500nm Step Height Sample (10 Repeats)
NX-Wafer Plus | XY orthogonality
NX-Wafer Plus | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Wafer Plus | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Wafer Plus | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Wafer Plus | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Wafer Plus | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Wafer Plus | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Wafer Plus | Y1 Detector noise (XY Servo off)
NX-Wafer Plus | X1 Detector noise (XY Servo off)
NX-Wafer Plus | Maximum Y Scan range (Full Stroke)
NX-Wafer Plus | Maximum X Scan range (Full Stroke)
NX-Wafer Plus | Full range of Y2 Detector raw
NX-Wafer Plus | Full range of Y1 Detector raw
NX-Wafer Plus | Full range of X2 Detector raw
NX-Wafer Plus | Full range of X1 Detector raw
NX-Wafer Plus | Z Linearity by Z Scanner position
NX-Wafer Plus | -Y Direction software Limit
NX-Wafer Plus | -X Direction software Limit
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer Plus | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer Plus | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Wafer Plus | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Wafer Plus | %GRR w/ roughness of non-patterened wafer
NX-Wafer Plus | Z Scanner position at -6μm
NX-Wafer Plus | Z Scanner position at +6μm
NX-Wafer Plus | Approach position at -5μm
NX-Wafer Plus | Approach position at +5μm
NX-Wafer Plus | * 1800nm Step Height Sample (10 Repeats)
NX-Wafer Plus | Full range of Z Detector raw
NX-Wafer Plus | Approach position at -3μm
NX-Wafer Plus | Approach position at +3μm
NX-Wafer Plus | Under 50nm Step Height Sample (10 Repeats)
NX-Wafer Plus | 50~200nm Step Height Sample (10 Repeats)
NX-Wafer Plus | 200~500nm Step Height Sample (10 Repeats)
NX-Wafer Plus | PosMin
NX-Wafer Plus | PosMax
NX-Wafer Plus | MaxLength
NX-Wafer Plus | FullRecoveryProfileLength
NX-Wafer Plus | Full range of Y Detector raw
NX-Wafer Plus | Full range of X Detector raw
NX-Wafer Plus | EFM Phase : For the positive bias
NX-Wafer Plus | EFM Phase : For the negative bias
NX-Wafer Plus | Z Linearity by 5 different sample height
NX-Wafer Plus | Laser intensity with AC160 cantilever
NX-Wafer Plus | A－B difference during 3 times run test
NX-Wafer Plus | A＋B difference during 3 times run test
NX-Wafer Plus | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Wafer Plus | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Wafer Plus | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Wafer Plus | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Wafer Plus | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Wafer Plus | Allowable max warpage, w/ SiC sample chuck
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan (Top)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan (Right)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan (Left)
NX-Wafer Plus | RMS of Z Height of Wafer ZeroScan (Center)
NX-Wafer Plus | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Wafer Plus | PosMin
NX-Wafer Plus | PosMax
NX-Wafer Plus | MaxLength
NX-Wafer Plus | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Wafer Plus | Check sample surface mapping
NX-Wafer Plus | Z Direction, Accuracy Test (Bias not significant)
NX-Wafer Plus | Y Direction, Accuracy Test (Bias not significant)
NX-Wafer Plus | X Direction, Accuracy Test (Bias not significant)
NX-Wafer Plus | Roughness Reproducibility Test (SR Panel)
NX-Wafer Plus | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Wafer Plus | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Wafer Plus | Roughness Reproducibility Test (SiNx Panel)
NX-Wafer Plus | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Wafer Plus | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Wafer Plus | Roughness Reproducibility Test (ABF Panel)
NX-Wafer Plus | Roughness Repeatability Test (SR Panel)
NX-Wafer Plus | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Wafer Plus | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Wafer Plus | Roughness Repeatability Test (SiNx Panel)
NX-Wafer Plus | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Wafer Plus | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Wafer Plus | Roughness Repeatability Test (ABF Panel)
NX-Wafer Plus | Lift Pin Stroke Min Check in DB
NX-Wafer Plus | Lift Pin Stroke Max Check in DB
NX-Wafer Plus | Limit range
NX-Mask | Z Linearity by Z scanner position
NX-Mask | Full range of Z Detector during Z Scan sweep
NX-Mask | Z Height calibration error
NX-Mask | Z Drive calibration error
NX-Mask | Z Linearity by 4 different sample height
NX-Mask | Full range of Z Height during Z Scan sweep
NX-Mask | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Mask | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Mask | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Mask | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Mask | Z Height noise
NX-Mask | Wafer placement(Loading) error on the load port1 (X offset)
NX-Mask | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Mask | Wafer placement(Loading) error on the load port2 (X offset)
NX-Mask | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Mask | OpticalAxisShiftPerFocusX
NX-Mask | OpticalAxisShiftPerFocusY
NX-Mask | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Mask | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Mask | Lift pins synchronization offset check (X offset)
NX-Mask | Lift pins synchronization offset check (Y offset)
NX-Mask | Profiler noise @200um
NX-Mask | Roll motion to X+ full scan
NX-Mask | Position repeatability(10repeats, Position Based)
NX-Mask | OPM Repeatability @5mm
NX-Mask | OPM Repeatability @1mm
NX-Mask | Max OPM @5mm
NX-Mask | Max OPM @1mm
NX-Mask | Torque Profile (10mm, 200um/s)
NX-Mask | Torque Profile (10mm, 20um/s)
NX-Mask | Lateral(C-D) sensitivity
NX-Mask | Lateral(C-D) noise
NX-Mask | Min. Vibration Transmissbility (~100Hz)
NX-Mask | RMS of ZeroScan
NX-Mask | LightStrengthGain
NX-Mask | Full range of X Detector during X Scan sweep
NX-Mask | Full range of X1 Detector during X Scan sweep
NX-Mask | Full range of X2 Detector during X Scan sweep
NX-Mask | Full range of Y Detector during Y Scan sweep
NX-Mask | Full range of Y1 Detector during Y Scan sweep
NX-Mask | Full range of Y2 Detector during Y Scan sweep
NX-Mask | Full range of X1 Detector Raw during X Scan sweep
NX-Mask | Full range of X2 Detector Raw during X Scan sweep
NX-Mask | Full range of Y1 Detector Raw during Y Scan sweep
NX-Mask | Full range of Y2 Detector Raw during Y Scan sweep
NX-Mask | X Detector stroke calibration error (nominal scale error X)
NX-Mask | Y Detector stroke calibration error (nominal scale error Y)
NX-Mask | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Mask | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Mask | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Mask | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Mask | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Mask | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Mask | X Non-linearity error
NX-Mask | Y Non-linearity error
NX-Mask | X axis OPM
NX-Mask | Y axis OPM
NX-Mask | X Ringing(Rpv)
NX-Mask | Y Ringing(Rpv)
NX-Mask | X Detector noise (XY Servo on)
NX-Mask | Y Detector noise (XY Servo on)
NX-Mask | X1 Detector noise (XY Servo on)
NX-Mask | Y1 Detector noise (XY Servo on)
NX-Mask | X2 Detector noise (XY Servo on)
NX-Mask | Y2 Detector noise (XY Servo on)
NX-Mask | RMS error after Stage Calibration
NX-Mask | Image offset after moving from -X to +X (X scan)
NX-Mask | Image offset after moving from -Y to +Y (Y scan)
NX-Mask | Image offset after moving from -X to +X (Y scan)
NX-Mask | Image offset after moving from -Y to +Y (X scan)
NX-Mask | Image offset after moving from +X to -X (X scan)
NX-Mask | Image offset after moving from +Y to -Y (Y scan)
NX-Mask | Image offset after moving from +X to -X (Y scan)
NX-Mask | Image offset after moving from +Y to -Y (X scan)
NX-Mask | XY Stage positioning X offset (Vision Pattern Matching)
NX-Mask | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Mask | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Mask | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Mask | XY Stage positioning X offset (In-Die Align)
NX-Mask | XY Stage positioning Y offset (In-Die Align)
NX-Mask | XY Stage positioning X 1σ (In-Die Align)
NX-Mask | XY Stage positioning Y 1σ (In-Die Align)
NX-Mask | XY Stage positioning X offset (LLC Translation)
NX-Mask | XY Stage positioning Y offset (LLC Translation)
NX-Mask | XY Stage positioning X 1σ (LLC Translation)
NX-Mask | XY Stage positioning Y 1σ (LLC Translation)
NX-Mask | XY Stage positioning X offset (Global Align)
NX-Mask | XY Stage positioning Y offset (Global Align)
NX-Mask | XY Stage positioning X 1σ (Global Align)
NX-Mask | XY Stage positioning Y 1σ (Global Align)
NX-Mask | Z Drift
NX-Mask | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | Cycle time for 15-point measurement
NX-Mask | Short-term stability using Step Height Sample (3σ)
NX-Mask | Roughness Repeatability 1x1um(RSD%)
NX-Mask | Roughness Repeatability 1x1um(1σ)
NX-Mask | Roughness Repeatability 80x80um(RSD%)
NX-Mask | Roughness Repeatability 80x80um(1σ)
NX-Mask | Z Height calibrtaion error
NX-Mask | Z Drive calibrtaion error
NX-Mask | Z Linearity by 3 different sample height
NX-Mask | Full range in voltage during Z Scan sweep
NX-Mask | Full range in um during Z Scan sweep
NX-Mask | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Mask | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Mask | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Mask | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Mask | Lateral(C-D) sensitivity
NX-Mask | Lateral(C-D) noise
NX-Mask | A-B noise
NX-Mask | NCM Amplitude noise
NX-Mask | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Mask | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Mask | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Mask | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Mask | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Mask | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Mask | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Mask | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Mask | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Mask | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Mask | XY Orthogonality
NX-Mask | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Mask | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Mask | Z linearity by sample height
NX-Mask | 1) Z Scanner position at +3μm
NX-Mask | 3) Z Scanner position at -3μm
NX-Mask | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Mask | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Mask | X Scan Stroke
NX-Mask | Y Scan Stroke
NX-Mask | Full range of X Detector Raw during X Scan sweep
NX-Mask | Full range of Y Detector Raw during Y Scan sweep
NX-Mask | X Detector noise (XY Servo off)
NX-Mask | Y Detector noise (XY Servo off)
NX-Mask | X axis OPM
NX-Mask | Y axis OPM
NX-Mask | X Non-linearity
NX-Mask | Y Non-linearity
NX-Mask | OPM Repeatability @25mm
NX-Mask | OPM Repeatability @10mm
NX-Mask | Max OPM @25mm
NX-Mask | Max OPM @10mm
NX-Mask | Torque Profile (50mm, 200um/s)
NX-Mask | Torque Profile (50mm, 20um/s)
NX-Mask | SLD current limit check
NX-Mask | SLD beam power at the end of the lens
NX-Mask | 2) of 20nm
NX-Mask | X Detector noise (XYServo off)
NX-Mask | Y Detector noise (XYServo off)
NX-Mask | X axis OPM raw Rpv
NX-Mask | Y axis OPM raw Rpv
NX-Mask | StageSafetyXMax
NX-Mask | StageSafetyXMin
NX-Mask | StageSafetyY
NX-Mask | StageSafetyYMax
NX-Mask | StageSafetyYMin
NX-Mask | Check X axis upper limit sensor function
NX-Mask | Check X axis lower limit sensor function
NX-Mask | Check Y axis upper limit sensor function
NX-Mask | Check Y axis lower limit sensor function
NX-Mask | X axis Angle Range
NX-Mask | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Mask | X detector noise (XY Servo on)
NX-Mask | Y detector noise (XY Servo on)
NX-Mask | Full range in um during X Scan sweep (XY Servo off)
NX-Mask | Full range in um during Y Scan sweep (XY Servo off)
NX-Mask | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Mask | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Mask | X axis, Repeatability @1st position
NX-Mask | X axis, Repeatability @ 2nd position
NX-Mask | X axis, Repeatability @ 3rd position
NX-Mask | X axis, Repeatability @ 4th position
NX-Mask | X axis, Accuracy @1st position
NX-Mask | X axis, Accuracy @ 2nd position
NX-Mask | X axis, Accuracy @ 3rd position
NX-Mask | X axis, Accuracy @ 4th position
NX-Mask | Y axis, Repeatability @1st position
NX-Mask | Y axis, Repeatability @ 2nd position
NX-Mask | Y axis, Repeatability @ 3rd position
NX-Mask | Y axis, Repeatability @ 4th position
NX-Mask | Y axis, Accuracy @1st position
NX-Mask | Y axis, Accuracy @ 2nd position
NX-Mask | Y axis, Accuracy @ 3rd position
NX-Mask | Y axis, Accuracy @ 4th position
NX-Mask | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Mask | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Mask | Z-XY Orthogonality (X direction-Left)
NX-Mask | Z-XY Orthogonality (X direction-Right)
NX-Mask | Z-XY Orthogonality (Y direction-Left)
NX-Mask | Z-XY Orthogonality (Y direction-Right)
NX-Mask | X-Scan Repeatability (10 repeats, RSD%)
NX-Mask | Y-Scan Repeatability (10 repeats, RSD%)
NX-Mask | Mechanical Align Accuracy
NX-Mask | Check the Multi Peattern Upload
NX-Mask | X Detector noise (XY servo off)
NX-Mask | Y Detector noise (XY servo off)
NX-Mask | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Mask | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Mask | Z Detector Max
NX-Mask | Z Detector Min
NX-Mask | WLI_Scale(Z Calibration)
NX-Mask | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Mask | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Mask | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Mask | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Mask | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Mask | Laser intensity with AC160TS cantilever
NX-Mask | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Mask | NCM driving % at an amplitude of 1 nm
NX-Mask | NCM driving % at an amplitude of 20 nm
NX-Mask | NCM driving % at an amplitude of 100 nm
NX-Mask | - 940nm Step Height Sample (10 Repeats)
NX-Mask | - 200~500nm Step Height Sample (10 Repeats)
NX-Mask | - 50~200nm Step Height Sample (10 Repeats)
NX-Mask | - Under 50nm Step Height Sample (10 Repeats)
NX-Mask | A-B after Auto PSPD Alignment
NX-Mask | NCM driving% with AC160TS cantilever
NX-Mask | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Mask | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Mask | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Mask | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Mask | X Detector stroke calibration error (nominal scale error X)
NX-Mask | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | Resonance Frequency in DB
NX-Mask | Servo Cut-off Frequency in DB
NX-Mask | Check chuck surface no damage
NX-Mask | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Mask | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Mask | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Mask | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Mask | X1 Detector noise (XY servo off)
NX-Mask | X2 Detector noise (XY servo off)
NX-Mask | Y1 Detector noise (XY servo off)
NX-Mask | Y2 Detector noise (XY servo off)
NX-Mask | X Detector noise (XY servo on)
NX-Mask | X1 Detector noise (XY servo on)
NX-Mask | X2 Detector noise (XY servo on)
NX-Mask | Y Detector noise (XY servo on)
NX-Mask | Y1 Detector noise (XY servo on)
NX-Mask | Y2 Detector noise (XY servo on)
NX-Mask | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Mask | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Mask | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Mask | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Mask | Z Detector offset
NX-Mask | Check the full stroke image(servo on)
NX-Mask | ResonanceFrequencyHz
NX-Mask | ServoCutoffFrequencyHz
NX-Mask | FullPixelHeight
NX-Mask | Serial Port#
NX-Mask | Off-axis Camera X Position(ΔX)
NX-Mask | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Mask | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Mask | Z Scanner position at -3μm
NX-Mask | Z Scanner position at +3μm
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Mask | Scan offset accuracy (Precision)
NX-Mask | Aligned stage accuracy (X offset-Point Align1)
NX-Mask | Aligned stage accuracy (Y offset-Point Align1)
NX-Mask | Aligned stage accuracy (X offset-Point Align2)
NX-Mask | Aligned stage accuracy (Y offset-Point Align2)
NX-Mask | Aligned stage accuracy (X offset-Point Align3)
NX-Mask | Aligned stage accuracy (Y offset-Point Align3)
NX-Mask | Aligned stage accuracy (X offset)
NX-Mask | Aligned stage accuracy (Y offset)
NX-Mask | Imag offset after EFEM loading (X scan)
NX-Mask | Imag offset after EFEM loading (Y scan)
NX-Mask | Imag offset after Lock stage (X scan)
NX-Mask | Imag offset after Lock stage (Y scan)
NX-Mask | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Mask | Mask Roughness Repeatability (10Repeats, σ )
NX-Mask | Mask exchange time
NX-Mask | WLI_Scale(C Calibration)
NX-Mask | WLI_Scale(B Calibration)
NX-Mask | WLI_Scale(A Calibration)
NX-Mask | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Mask | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Mask | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Mask | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Mask | Port 1 lens(to OnAxis) Y Offset
NX-Mask | Port 1 lens(to OnAxis) X Offset
NX-Mask | P Gain
NX-Mask | I Gain
NX-Mask | D Gain
NX-Mask | Check the Port 2 lens light strength
NX-Mask | Check the Port 1 lens light strength
NX-Mask | Check the focus deviation between the lenses(Δf)
NX-Mask | Check the 25mm Moving – Y Axis Offset Check
NX-Mask | C Scanner Detector Min
NX-Mask | C Scanner Detector Max
NX-Mask | B Scanner Detector Min
NX-Mask | B Scanner Detector Max
NX-Mask | Angle align between Port 2 lens and VLSI SHS: Y
NX-Mask | Angle align between Port 2 lens and VLSI SHS: X
NX-Mask | Angle align between Port 2 lens and Sample chuck: Y
NX-Mask | Angle align between Port 2 lens and Sample chuck: X
NX-Mask | Angle align between Port 1 lens and Sample chuck: Y
NX-Mask | Angle align between Port 1 lens and Sample chuck: X
NX-Mask | A Scanner Detector Min
NX-Mask | A Scanner Detector Max
NX-Mask | Y2 Detector noise (XY Servo off)
NX-Mask | X2 Detector noise (XY Servo off)
NX-Mask | Y axis OPM (-52000,-52000)
NX-Mask | Y axis OPM (-52000,52000)
NX-Mask | Y axis OPM (52000,-52000)
NX-Mask | Y axis OPM (52000,52000)
NX-Mask | Y axis OPM (0,0)
NX-Mask | XY Orthogonality (-52000,-52000)
NX-Mask | XY Orthogonality (-52000,52000)
NX-Mask | XY Orthogonality (52000,-52000)
NX-Mask | XY Orthogonality (52000,52000)
NX-Mask | XY Orthogonality (0,0)
NX-Mask | X axis OPM (-52000,-52000)
NX-Mask | X axis OPM (-52000,52000)
NX-Mask | X axis OPM (52000,-52000)
NX-Mask | X axis OPM (52000,52000)
NX-Mask | X axis OPM (0,0)
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Mask | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Mask | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Mask | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Mask | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Mask | XY stage calibration
NX-Mask | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Mask | LD beam power at the end of the lens
NX-Mask | Y Non-lineairty
NX-Mask | X Non-lineairty
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Mask | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Mask | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Mask | UpperLimit
NX-Mask | Under 50nm Step Height Sample (10 Repeats)
NX-Mask | Overlay Camera X Position(ΔX)
NX-Mask | LowerLimit
NX-Mask | Initial Z
NX-Mask | 940nm Step Height Sample (10 Repeats)
NX-Mask | 50~200nm Step Height Sample (10 Repeats)
NX-Mask | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Mask | StrokeYMin
NX-Mask | StrokeYMax
NX-Mask | StrokeY
NX-Mask | StrokeXMin
NX-Mask | StrokeXMax
NX-Mask | StrokeX
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Mask | Z Drift (Short Traveling)
NX-Mask | 200~500nm Step Height Sample (10 Repeats)
NX-Mask | XY orthogonality
NX-Mask | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Mask | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Mask | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Mask | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Mask | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Mask | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Mask | Y1 Detector noise (XY Servo off)
NX-Mask | X1 Detector noise (XY Servo off)
NX-Mask | Maximum Y Scan range (Full Stroke)
NX-Mask | Maximum X Scan range (Full Stroke)
NX-Mask | Full range of Y2 Detector raw
NX-Mask | Full range of Y1 Detector raw
NX-Mask | Full range of X2 Detector raw
NX-Mask | Full range of X1 Detector raw
NX-Mask | Z Linearity by Z Scanner position
NX-Mask | -Y Direction software Limit
NX-Mask | -X Direction software Limit
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Mask | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Mask | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Mask | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Mask | %GRR w/ roughness of non-patterened wafer
NX-Mask | Z Scanner position at -6μm
NX-Mask | Z Scanner position at +6μm
NX-Mask | Approach position at -5μm
NX-Mask | Approach position at +5μm
NX-Mask | * 1800nm Step Height Sample (10 Repeats)
NX-Mask | Full range of Z Detector raw
NX-Mask | Approach position at -3μm
NX-Mask | Approach position at +3μm
NX-Mask | Under 50nm Step Height Sample (10 Repeats)
NX-Mask | 50~200nm Step Height Sample (10 Repeats)
NX-Mask | 200~500nm Step Height Sample (10 Repeats)
NX-Mask | PosMin
NX-Mask | PosMax
NX-Mask | MaxLength
NX-Mask | FullRecoveryProfileLength
NX-Mask | Full range of Y Detector raw
NX-Mask | Full range of X Detector raw
NX-Mask | EFM Phase : For the positive bias
NX-Mask | EFM Phase : For the negative bias
NX-Mask | Z Linearity by 5 different sample height
NX-Mask | Laser intensity with AC160 cantilever
NX-Mask | A－B difference during 3 times run test
NX-Mask | A＋B difference during 3 times run test
NX-Mask | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Mask | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Mask | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Mask | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Mask | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Mask | Allowable max warpage, w/ SiC sample chuck
NX-Mask | RMS of Z Height of Wafer Zeroscan (Top)
NX-Mask | RMS of Z Height of Wafer Zeroscan (Right)
NX-Mask | RMS of Z Height of Wafer Zeroscan (Left)
NX-Mask | RMS of Z Height of Wafer ZeroScan (Center)
NX-Mask | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Mask | PosMin
NX-Mask | PosMax
NX-Mask | MaxLength
NX-Mask | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Mask | Check sample surface mapping
NX-Mask | Z Direction, Accuracy Test (Bias not significant)
NX-Mask | Y Direction, Accuracy Test (Bias not significant)
NX-Mask | X Direction, Accuracy Test (Bias not significant)
NX-Mask | Roughness Reproducibility Test (SR Panel)
NX-Mask | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Mask | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Mask | Roughness Reproducibility Test (SiNx Panel)
NX-Mask | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Mask | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Mask | Roughness Reproducibility Test (ABF Panel)
NX-Mask | Roughness Repeatability Test (SR Panel)
NX-Mask | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Mask | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Mask | Roughness Repeatability Test (SiNx Panel)
NX-Mask | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Mask | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Mask | Roughness Repeatability Test (ABF Panel)
NX-Mask | Lift Pin Stroke Min Check in DB
NX-Mask | Lift Pin Stroke Max Check in DB
NX-Mask | Limit range
NX-Hybrid WLI(Interferom Ver.) | Z Linearity by Z scanner position
NX-Hybrid WLI(Interferom Ver.) | Full range of Z Detector during Z Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration error
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration error
NX-Hybrid WLI(Interferom Ver.) | Z Linearity by 4 different sample height
NX-Hybrid WLI(Interferom Ver.) | Full range of Z Height during Z Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Z Height noise
NX-Hybrid WLI(Interferom Ver.) | Wafer placement(Loading) error on the load port1 (X offset)
NX-Hybrid WLI(Interferom Ver.) | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Hybrid WLI(Interferom Ver.) | Wafer placement(Loading) error on the load port2 (X offset)
NX-Hybrid WLI(Interferom Ver.) | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Hybrid WLI(Interferom Ver.) | OpticalAxisShiftPerFocusX
NX-Hybrid WLI(Interferom Ver.) | OpticalAxisShiftPerFocusY
NX-Hybrid WLI(Interferom Ver.) | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Hybrid WLI(Interferom Ver.) | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Hybrid WLI(Interferom Ver.) | Lift pins synchronization offset check (X offset)
NX-Hybrid WLI(Interferom Ver.) | Lift pins synchronization offset check (Y offset)
NX-Hybrid WLI(Interferom Ver.) | Profiler noise @200um
NX-Hybrid WLI(Interferom Ver.) | Roll motion to X+ full scan
NX-Hybrid WLI(Interferom Ver.) | Position repeatability(10repeats, Position Based)
NX-Hybrid WLI(Interferom Ver.) | OPM Repeatability @5mm
NX-Hybrid WLI(Interferom Ver.) | OPM Repeatability @1mm
NX-Hybrid WLI(Interferom Ver.) | Max OPM @5mm
NX-Hybrid WLI(Interferom Ver.) | Max OPM @1mm
NX-Hybrid WLI(Interferom Ver.) | Torque Profile (10mm, 200um/s)
NX-Hybrid WLI(Interferom Ver.) | Torque Profile (10mm, 20um/s)
NX-Hybrid WLI(Interferom Ver.) | Lateral(C-D) sensitivity
NX-Hybrid WLI(Interferom Ver.) | Lateral(C-D) noise
NX-Hybrid WLI(Interferom Ver.) | Min. Vibration Transmissbility (~100Hz)
NX-Hybrid WLI(Interferom Ver.) | RMS of ZeroScan
NX-Hybrid WLI(Interferom Ver.) | LightStrengthGain
NX-Hybrid WLI(Interferom Ver.) | Full range of X Detector during X Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of X1 Detector during X Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of X2 Detector during X Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of Y Detector during Y Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of Y1 Detector during Y Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of Y2 Detector during Y Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of X1 Detector Raw during X Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of X2 Detector Raw during X Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of Y1 Detector Raw during Y Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of Y2 Detector Raw during Y Scan sweep
NX-Hybrid WLI(Interferom Ver.) | X Detector stroke calibration error (nominal scale error X)
NX-Hybrid WLI(Interferom Ver.) | Y Detector stroke calibration error (nominal scale error Y)
NX-Hybrid WLI(Interferom Ver.) | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Hybrid WLI(Interferom Ver.) | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Hybrid WLI(Interferom Ver.) | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Hybrid WLI(Interferom Ver.) | X Non-linearity error
NX-Hybrid WLI(Interferom Ver.) | Y Non-linearity error
NX-Hybrid WLI(Interferom Ver.) | X axis OPM
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM
NX-Hybrid WLI(Interferom Ver.) | X Ringing(Rpv)
NX-Hybrid WLI(Interferom Ver.) | Y Ringing(Rpv)
NX-Hybrid WLI(Interferom Ver.) | X Detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | Y Detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | X1 Detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | Y1 Detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | X2 Detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | Y2 Detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | RMS error after Stage Calibration
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from -X to +X (X scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from -Y to +Y (Y scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from -X to +X (Y scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from -Y to +Y (X scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from +X to -X (X scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from +Y to -Y (Y scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from +X to -X (Y scan)
NX-Hybrid WLI(Interferom Ver.) | Image offset after moving from +Y to -Y (X scan)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X offset (Vision Pattern Matching)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X offset (In-Die Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y offset (In-Die Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X 1σ (In-Die Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y 1σ (In-Die Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X offset (LLC Translation)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y offset (LLC Translation)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X 1σ (LLC Translation)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y 1σ (LLC Translation)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X offset (Global Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y offset (Global Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning X 1σ (Global Align)
NX-Hybrid WLI(Interferom Ver.) | XY Stage positioning Y 1σ (Global Align)
NX-Hybrid WLI(Interferom Ver.) | Z Drift
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Cycle time for 15-point measurement
NX-Hybrid WLI(Interferom Ver.) | Short-term stability using Step Height Sample (3σ)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 1x1um(RSD%)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 1x1um(1σ)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 80x80um(RSD%)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 80x80um(1σ)
NX-Hybrid WLI(Interferom Ver.) | Z Height calibrtaion error
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibrtaion error
NX-Hybrid WLI(Interferom Ver.) | Z Linearity by 3 different sample height
NX-Hybrid WLI(Interferom Ver.) | Full range in voltage during Z Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range in um during Z Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Lateral(C-D) sensitivity
NX-Hybrid WLI(Interferom Ver.) | Lateral(C-D) noise
NX-Hybrid WLI(Interferom Ver.) | A-B noise
NX-Hybrid WLI(Interferom Ver.) | NCM Amplitude noise
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Hybrid WLI(Interferom Ver.) | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Hybrid WLI(Interferom Ver.) | XY Orthogonality
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Hybrid WLI(Interferom Ver.) | Z linearity by sample height
NX-Hybrid WLI(Interferom Ver.) | 1) Z Scanner position at +3μm
NX-Hybrid WLI(Interferom Ver.) | 3) Z Scanner position at -3μm
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Hybrid WLI(Interferom Ver.) | X Scan Stroke
NX-Hybrid WLI(Interferom Ver.) | Y Scan Stroke
NX-Hybrid WLI(Interferom Ver.) | Full range of X Detector Raw during X Scan sweep
NX-Hybrid WLI(Interferom Ver.) | Full range of Y Detector Raw during Y Scan sweep
NX-Hybrid WLI(Interferom Ver.) | X Detector noise (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | Y Detector noise (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM
NX-Hybrid WLI(Interferom Ver.) | X Non-linearity
NX-Hybrid WLI(Interferom Ver.) | Y Non-linearity
NX-Hybrid WLI(Interferom Ver.) | OPM Repeatability @25mm
NX-Hybrid WLI(Interferom Ver.) | OPM Repeatability @10mm
NX-Hybrid WLI(Interferom Ver.) | Max OPM @25mm
NX-Hybrid WLI(Interferom Ver.) | Max OPM @10mm
NX-Hybrid WLI(Interferom Ver.) | Torque Profile (50mm, 200um/s)
NX-Hybrid WLI(Interferom Ver.) | Torque Profile (50mm, 20um/s)
NX-Hybrid WLI(Interferom Ver.) | SLD current limit check
NX-Hybrid WLI(Interferom Ver.) | SLD beam power at the end of the lens
NX-Hybrid WLI(Interferom Ver.) | 2) of 20nm
NX-Hybrid WLI(Interferom Ver.) | X Detector noise (XYServo off)
NX-Hybrid WLI(Interferom Ver.) | Y Detector noise (XYServo off)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM raw Rpv
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM raw Rpv
NX-Hybrid WLI(Interferom Ver.) | StageSafetyXMax
NX-Hybrid WLI(Interferom Ver.) | StageSafetyXMin
NX-Hybrid WLI(Interferom Ver.) | StageSafetyY
NX-Hybrid WLI(Interferom Ver.) | StageSafetyYMax
NX-Hybrid WLI(Interferom Ver.) | StageSafetyYMin
NX-Hybrid WLI(Interferom Ver.) | Check X axis upper limit sensor function
NX-Hybrid WLI(Interferom Ver.) | Check X axis lower limit sensor function
NX-Hybrid WLI(Interferom Ver.) | Check Y axis upper limit sensor function
NX-Hybrid WLI(Interferom Ver.) | Check Y axis lower limit sensor function
NX-Hybrid WLI(Interferom Ver.) | X axis Angle Range
NX-Hybrid WLI(Interferom Ver.) | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Hybrid WLI(Interferom Ver.) | X detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | Y detector noise (XY Servo on)
NX-Hybrid WLI(Interferom Ver.) | Full range in um during X Scan sweep (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | Full range in um during Y Scan sweep (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Hybrid WLI(Interferom Ver.) | X axis, Repeatability @1st position
NX-Hybrid WLI(Interferom Ver.) | X axis, Repeatability @ 2nd position
NX-Hybrid WLI(Interferom Ver.) | X axis, Repeatability @ 3rd position
NX-Hybrid WLI(Interferom Ver.) | X axis, Repeatability @ 4th position
NX-Hybrid WLI(Interferom Ver.) | X axis, Accuracy @1st position
NX-Hybrid WLI(Interferom Ver.) | X axis, Accuracy @ 2nd position
NX-Hybrid WLI(Interferom Ver.) | X axis, Accuracy @ 3rd position
NX-Hybrid WLI(Interferom Ver.) | X axis, Accuracy @ 4th position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Repeatability @1st position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Repeatability @ 2nd position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Repeatability @ 3rd position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Repeatability @ 4th position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Accuracy @1st position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Accuracy @ 2nd position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Accuracy @ 3rd position
NX-Hybrid WLI(Interferom Ver.) | Y axis, Accuracy @ 4th position
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Z-XY Orthogonality (X direction-Left)
NX-Hybrid WLI(Interferom Ver.) | Z-XY Orthogonality (X direction-Right)
NX-Hybrid WLI(Interferom Ver.) | Z-XY Orthogonality (Y direction-Left)
NX-Hybrid WLI(Interferom Ver.) | Z-XY Orthogonality (Y direction-Right)
NX-Hybrid WLI(Interferom Ver.) | X-Scan Repeatability (10 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Y-Scan Repeatability (10 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Mechanical Align Accuracy
NX-Hybrid WLI(Interferom Ver.) | Check the Multi Peattern Upload
NX-Hybrid WLI(Interferom Ver.) | X Detector noise (XY servo off)
NX-Hybrid WLI(Interferom Ver.) | Y Detector noise (XY servo off)
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hybrid WLI(Interferom Ver.) | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hybrid WLI(Interferom Ver.) | Z Detector Max
NX-Hybrid WLI(Interferom Ver.) | Z Detector Min
NX-Hybrid WLI(Interferom Ver.) | WLI_Scale(Z Calibration)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Hybrid WLI(Interferom Ver.) | Laser intensity with AC160TS cantilever
NX-Hybrid WLI(Interferom Ver.) | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Hybrid WLI(Interferom Ver.) | NCM driving % at an amplitude of 1 nm
NX-Hybrid WLI(Interferom Ver.) | NCM driving % at an amplitude of 20 nm
NX-Hybrid WLI(Interferom Ver.) | NCM driving % at an amplitude of 100 nm
NX-Hybrid WLI(Interferom Ver.) | - 940nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | - 200~500nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | - 50~200nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | - Under 50nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | A-B after Auto PSPD Alignment
NX-Hybrid WLI(Interferom Ver.) | NCM driving% with AC160TS cantilever
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Hybrid WLI(Interferom Ver.) | X Detector stroke calibration error (nominal scale error X)
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Resonance Frequency in DB
NX-Hybrid WLI(Interferom Ver.) | Servo Cut-off Frequency in DB
NX-Hybrid WLI(Interferom Ver.) | Check chuck surface no damage
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Hybrid WLI(Interferom Ver.) | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Hybrid WLI(Interferom Ver.) | X1 Detector noise (XY servo off)
NX-Hybrid WLI(Interferom Ver.) | X2 Detector noise (XY servo off)
NX-Hybrid WLI(Interferom Ver.) | Y1 Detector noise (XY servo off)
NX-Hybrid WLI(Interferom Ver.) | Y2 Detector noise (XY servo off)
NX-Hybrid WLI(Interferom Ver.) | X Detector noise (XY servo on)
NX-Hybrid WLI(Interferom Ver.) | X1 Detector noise (XY servo on)
NX-Hybrid WLI(Interferom Ver.) | X2 Detector noise (XY servo on)
NX-Hybrid WLI(Interferom Ver.) | Y Detector noise (XY servo on)
NX-Hybrid WLI(Interferom Ver.) | Y1 Detector noise (XY servo on)
NX-Hybrid WLI(Interferom Ver.) | Y2 Detector noise (XY servo on)
NX-Hybrid WLI(Interferom Ver.) | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Hybrid WLI(Interferom Ver.) | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hybrid WLI(Interferom Ver.) | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hybrid WLI(Interferom Ver.) | Z Detector offset
NX-Hybrid WLI(Interferom Ver.) | Check the full stroke image(servo on)
NX-Hybrid WLI(Interferom Ver.) | ResonanceFrequencyHz
NX-Hybrid WLI(Interferom Ver.) | ServoCutoffFrequencyHz
NX-Hybrid WLI(Interferom Ver.) | FullPixelHeight
NX-Hybrid WLI(Interferom Ver.) | Serial Port#
NX-Hybrid WLI(Interferom Ver.) | Off-axis Camera X Position(ΔX)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Z Scanner position at -3μm
NX-Hybrid WLI(Interferom Ver.) | Z Scanner position at +3μm
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Scan offset accuracy (Precision)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (X offset-Point Align1)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (Y offset-Point Align1)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (X offset-Point Align2)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (Y offset-Point Align2)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (X offset-Point Align3)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (Y offset-Point Align3)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (X offset)
NX-Hybrid WLI(Interferom Ver.) | Aligned stage accuracy (Y offset)
NX-Hybrid WLI(Interferom Ver.) | Imag offset after EFEM loading (X scan)
NX-Hybrid WLI(Interferom Ver.) | Imag offset after EFEM loading (Y scan)
NX-Hybrid WLI(Interferom Ver.) | Imag offset after Lock stage (X scan)
NX-Hybrid WLI(Interferom Ver.) | Imag offset after Lock stage (Y scan)
NX-Hybrid WLI(Interferom Ver.) | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Hybrid WLI(Interferom Ver.) | Mask Roughness Repeatability (10Repeats, σ )
NX-Hybrid WLI(Interferom Ver.) | Mask exchange time
NX-Hybrid WLI(Interferom Ver.) | WLI_Scale(C Calibration)
NX-Hybrid WLI(Interferom Ver.) | WLI_Scale(B Calibration)
NX-Hybrid WLI(Interferom Ver.) | WLI_Scale(A Calibration)
NX-Hybrid WLI(Interferom Ver.) | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Hybrid WLI(Interferom Ver.) | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Hybrid WLI(Interferom Ver.) | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Hybrid WLI(Interferom Ver.) | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Hybrid WLI(Interferom Ver.) | Port 1 lens(to OnAxis) Y Offset
NX-Hybrid WLI(Interferom Ver.) | Port 1 lens(to OnAxis) X Offset
NX-Hybrid WLI(Interferom Ver.) | P Gain
NX-Hybrid WLI(Interferom Ver.) | I Gain
NX-Hybrid WLI(Interferom Ver.) | D Gain
NX-Hybrid WLI(Interferom Ver.) | Check the Port 2 lens light strength
NX-Hybrid WLI(Interferom Ver.) | Check the Port 1 lens light strength
NX-Hybrid WLI(Interferom Ver.) | Check the focus deviation between the lenses(Δf)
NX-Hybrid WLI(Interferom Ver.) | Check the 25mm Moving – Y Axis Offset Check
NX-Hybrid WLI(Interferom Ver.) | C Scanner Detector Min
NX-Hybrid WLI(Interferom Ver.) | C Scanner Detector Max
NX-Hybrid WLI(Interferom Ver.) | B Scanner Detector Min
NX-Hybrid WLI(Interferom Ver.) | B Scanner Detector Max
NX-Hybrid WLI(Interferom Ver.) | Angle align between Port 2 lens and VLSI SHS: Y
NX-Hybrid WLI(Interferom Ver.) | Angle align between Port 2 lens and VLSI SHS: X
NX-Hybrid WLI(Interferom Ver.) | Angle align between Port 2 lens and Sample chuck: Y
NX-Hybrid WLI(Interferom Ver.) | Angle align between Port 2 lens and Sample chuck: X
NX-Hybrid WLI(Interferom Ver.) | Angle align between Port 1 lens and Sample chuck: Y
NX-Hybrid WLI(Interferom Ver.) | Angle align between Port 1 lens and Sample chuck: X
NX-Hybrid WLI(Interferom Ver.) | A Scanner Detector Min
NX-Hybrid WLI(Interferom Ver.) | A Scanner Detector Max
NX-Hybrid WLI(Interferom Ver.) | Y2 Detector noise (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | X2 Detector noise (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM (-52000,-52000)
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM (-52000,52000)
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM (52000,-52000)
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM (52000,52000)
NX-Hybrid WLI(Interferom Ver.) | Y axis OPM (0,0)
NX-Hybrid WLI(Interferom Ver.) | XY Orthogonality (-52000,-52000)
NX-Hybrid WLI(Interferom Ver.) | XY Orthogonality (-52000,52000)
NX-Hybrid WLI(Interferom Ver.) | XY Orthogonality (52000,-52000)
NX-Hybrid WLI(Interferom Ver.) | XY Orthogonality (52000,52000)
NX-Hybrid WLI(Interferom Ver.) | XY Orthogonality (0,0)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM (-52000,-52000)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM (-52000,52000)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM (52000,-52000)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM (52000,52000)
NX-Hybrid WLI(Interferom Ver.) | X axis OPM (0,0)
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Hybrid WLI(Interferom Ver.) | XY stage calibration
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Hybrid WLI(Interferom Ver.) | LD beam power at the end of the lens
NX-Hybrid WLI(Interferom Ver.) | Y Non-lineairty
NX-Hybrid WLI(Interferom Ver.) | X Non-lineairty
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Hybrid WLI(Interferom Ver.) | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Hybrid WLI(Interferom Ver.) | UpperLimit
NX-Hybrid WLI(Interferom Ver.) | Under 50nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | Overlay Camera X Position(ΔX)
NX-Hybrid WLI(Interferom Ver.) | LowerLimit
NX-Hybrid WLI(Interferom Ver.) | Initial Z
NX-Hybrid WLI(Interferom Ver.) | 940nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | 50~200nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Hybrid WLI(Interferom Ver.) | StrokeYMin
NX-Hybrid WLI(Interferom Ver.) | StrokeYMax
NX-Hybrid WLI(Interferom Ver.) | StrokeY
NX-Hybrid WLI(Interferom Ver.) | StrokeXMin
NX-Hybrid WLI(Interferom Ver.) | StrokeXMax
NX-Hybrid WLI(Interferom Ver.) | StrokeX
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Z Drift (Short Traveling)
NX-Hybrid WLI(Interferom Ver.) | 200~500nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | XY orthogonality
NX-Hybrid WLI(Interferom Ver.) | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Hybrid WLI(Interferom Ver.) | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Hybrid WLI(Interferom Ver.) | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Hybrid WLI(Interferom Ver.) | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Hybrid WLI(Interferom Ver.) | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Hybrid WLI(Interferom Ver.) | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Hybrid WLI(Interferom Ver.) | Y1 Detector noise (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | X1 Detector noise (XY Servo off)
NX-Hybrid WLI(Interferom Ver.) | Maximum Y Scan range (Full Stroke)
NX-Hybrid WLI(Interferom Ver.) | Maximum X Scan range (Full Stroke)
NX-Hybrid WLI(Interferom Ver.) | Full range of Y2 Detector raw
NX-Hybrid WLI(Interferom Ver.) | Full range of Y1 Detector raw
NX-Hybrid WLI(Interferom Ver.) | Full range of X2 Detector raw
NX-Hybrid WLI(Interferom Ver.) | Full range of X1 Detector raw
NX-Hybrid WLI(Interferom Ver.) | Z Linearity by Z Scanner position
NX-Hybrid WLI(Interferom Ver.) | -Y Direction software Limit
NX-Hybrid WLI(Interferom Ver.) | -X Direction software Limit
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Hybrid WLI(Interferom Ver.) | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Hybrid WLI(Interferom Ver.) | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Hybrid WLI(Interferom Ver.) | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Hybrid WLI(Interferom Ver.) | %GRR w/ roughness of non-patterened wafer
NX-Hybrid WLI(Interferom Ver.) | Z Scanner position at -6μm
NX-Hybrid WLI(Interferom Ver.) | Z Scanner position at +6μm
NX-Hybrid WLI(Interferom Ver.) | Approach position at -5μm
NX-Hybrid WLI(Interferom Ver.) | Approach position at +5μm
NX-Hybrid WLI(Interferom Ver.) | * 1800nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | Full range of Z Detector raw
NX-Hybrid WLI(Interferom Ver.) | Approach position at -3μm
NX-Hybrid WLI(Interferom Ver.) | Approach position at +3μm
NX-Hybrid WLI(Interferom Ver.) | Under 50nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | 50~200nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | 200~500nm Step Height Sample (10 Repeats)
NX-Hybrid WLI(Interferom Ver.) | PosMin
NX-Hybrid WLI(Interferom Ver.) | PosMax
NX-Hybrid WLI(Interferom Ver.) | MaxLength
NX-Hybrid WLI(Interferom Ver.) | FullRecoveryProfileLength
NX-Hybrid WLI(Interferom Ver.) | Full range of Y Detector raw
NX-Hybrid WLI(Interferom Ver.) | Full range of X Detector raw
NX-Hybrid WLI(Interferom Ver.) | EFM Phase : For the positive bias
NX-Hybrid WLI(Interferom Ver.) | EFM Phase : For the negative bias
NX-Hybrid WLI(Interferom Ver.) | Z Linearity by 5 different sample height
NX-Hybrid WLI(Interferom Ver.) | Laser intensity with AC160 cantilever
NX-Hybrid WLI(Interferom Ver.) | A－B difference during 3 times run test
NX-Hybrid WLI(Interferom Ver.) | A＋B difference during 3 times run test
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Hybrid WLI(Interferom Ver.) | Allowable max warpage, w/ SiC sample chuck
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan (Top)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan (Right)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan (Left)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer ZeroScan (Center)
NX-Hybrid WLI(Interferom Ver.) | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Hybrid WLI(Interferom Ver.) | PosMin
NX-Hybrid WLI(Interferom Ver.) | PosMax
NX-Hybrid WLI(Interferom Ver.) | MaxLength
NX-Hybrid WLI(Interferom Ver.) | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Hybrid WLI(Interferom Ver.) | Check sample surface mapping
NX-Hybrid WLI(Interferom Ver.) | Z Direction, Accuracy Test (Bias not significant)
NX-Hybrid WLI(Interferom Ver.) | Y Direction, Accuracy Test (Bias not significant)
NX-Hybrid WLI(Interferom Ver.) | X Direction, Accuracy Test (Bias not significant)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (SR Panel)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (SiNx Panel)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Hybrid WLI(Interferom Ver.) | Roughness Reproducibility Test (ABF Panel)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (SR Panel)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (SiNx Panel)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Hybrid WLI(Interferom Ver.) | Roughness Repeatability Test (ABF Panel)
NX-Hybrid WLI(Interferom Ver.) | Lift Pin Stroke Min Check in DB
NX-Hybrid WLI(Interferom Ver.) | Lift Pin Stroke Max Check in DB
NX-Hybrid WLI(Interferom Ver.) | Limit range
FX200 | Z Linearity by Z scanner position
FX200 | Full range of Z Detector during Z Scan sweep
FX200 | Z Height calibration error
FX200 | Z Drive calibration error
FX200 | Z Linearity by 4 different sample height
FX200 | Full range of Z Height during Z Scan sweep
FX200 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
FX200 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
FX200 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
FX200 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
FX200 | Z Height noise
FX200 | Wafer placement(Loading) error on the load port1 (X offset)
FX200 | Wafer placement(Loading) error on the load port1 (Y offset)
FX200 | Wafer placement(Loading) error on the load port2 (X offset)
FX200 | Wafer placement(Loading) error on the load port2 (Y offset)
FX200 | OpticalAxisShiftPerFocusX
FX200 | OpticalAxisShiftPerFocusY
FX200 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
FX200 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
FX200 | Lift pins synchronization offset check (X offset)
FX200 | Lift pins synchronization offset check (Y offset)
FX200 | Profiler noise @200um
FX200 | Roll motion to X+ full scan
FX200 | Position repeatability(10repeats, Position Based)
FX200 | OPM Repeatability @5mm
FX200 | OPM Repeatability @1mm
FX200 | Max OPM @5mm
FX200 | Max OPM @1mm
FX200 | Torque Profile (10mm, 200um/s)
FX200 | Torque Profile (10mm, 20um/s)
FX200 | Lateral(C-D) sensitivity
FX200 | Lateral(C-D) noise
FX200 | Min. Vibration Transmissbility (~100Hz)
FX200 | RMS of ZeroScan
FX200 | LightStrengthGain
FX200 | Full range of X Detector during X Scan sweep
FX200 | Full range of X1 Detector during X Scan sweep
FX200 | Full range of X2 Detector during X Scan sweep
FX200 | Full range of Y Detector during Y Scan sweep
FX200 | Full range of Y1 Detector during Y Scan sweep
FX200 | Full range of Y2 Detector during Y Scan sweep
FX200 | Full range of X1 Detector Raw during X Scan sweep
FX200 | Full range of X2 Detector Raw during X Scan sweep
FX200 | Full range of Y1 Detector Raw during Y Scan sweep
FX200 | Full range of Y2 Detector Raw during Y Scan sweep
FX200 | X Detector stroke calibration error (nominal scale error X)
FX200 | Y Detector stroke calibration error (nominal scale error Y)
FX200 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
FX200 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
FX200 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
FX200 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
FX200 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
FX200 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
FX200 | X Non-linearity error
FX200 | Y Non-linearity error
FX200 | X axis OPM
FX200 | Y axis OPM
FX200 | X Ringing(Rpv)
FX200 | Y Ringing(Rpv)
FX200 | X Detector noise (XY Servo on)
FX200 | Y Detector noise (XY Servo on)
FX200 | X1 Detector noise (XY Servo on)
FX200 | Y1 Detector noise (XY Servo on)
FX200 | X2 Detector noise (XY Servo on)
FX200 | Y2 Detector noise (XY Servo on)
FX200 | RMS error after Stage Calibration
FX200 | Image offset after moving from -X to +X (X scan)
FX200 | Image offset after moving from -Y to +Y (Y scan)
FX200 | Image offset after moving from -X to +X (Y scan)
FX200 | Image offset after moving from -Y to +Y (X scan)
FX200 | Image offset after moving from +X to -X (X scan)
FX200 | Image offset after moving from +Y to -Y (Y scan)
FX200 | Image offset after moving from +X to -X (Y scan)
FX200 | Image offset after moving from +Y to -Y (X scan)
FX200 | XY Stage positioning X offset (Vision Pattern Matching)
FX200 | XY Stage positioning Y offset (Vision Pattern Matching)
FX200 | XY Stage positioning X 1σ (Vision Pattern Matching)
FX200 | XY Stage positioning Y 1σ (Vision Pattern Matching)
FX200 | XY Stage positioning X offset (In-Die Align)
FX200 | XY Stage positioning Y offset (In-Die Align)
FX200 | XY Stage positioning X 1σ (In-Die Align)
FX200 | XY Stage positioning Y 1σ (In-Die Align)
FX200 | XY Stage positioning X offset (LLC Translation)
FX200 | XY Stage positioning Y offset (LLC Translation)
FX200 | XY Stage positioning X 1σ (LLC Translation)
FX200 | XY Stage positioning Y 1σ (LLC Translation)
FX200 | XY Stage positioning X offset (Global Align)
FX200 | XY Stage positioning Y offset (Global Align)
FX200 | XY Stage positioning X 1σ (Global Align)
FX200 | XY Stage positioning Y 1σ (Global Align)
FX200 | Z Drift
FX200 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | Cycle time for 15-point measurement
FX200 | Short-term stability using Step Height Sample (3σ)
FX200 | Roughness Repeatability 1x1um(RSD%)
FX200 | Roughness Repeatability 1x1um(1σ)
FX200 | Roughness Repeatability 80x80um(RSD%)
FX200 | Roughness Repeatability 80x80um(1σ)
FX200 | Z Height calibrtaion error
FX200 | Z Drive calibrtaion error
FX200 | Z Linearity by 3 different sample height
FX200 | Full range in voltage during Z Scan sweep
FX200 | Full range in um during Z Scan sweep
FX200 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
FX200 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
FX200 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
FX200 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
FX200 | Lateral(C-D) sensitivity
FX200 | Lateral(C-D) noise
FX200 | A-B noise
FX200 | NCM Amplitude noise
FX200 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
FX200 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
FX200 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
FX200 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
FX200 | RMS of Z Height of Wafer Zeroscan(0,120000)
FX200 | RMS of Z Height of Wafer Zeroscan(-120000,0)
FX200 | RMS of Z Height of Wafer Zeroscan(0,-120000)
FX200 | RMS of Z Height of Wafer Zeroscan(120000,0)
FX200 | RMS of Z Height of Wafer ZeroScan (0,0)
FX200 | Max Wafer Zeroscan-Min Wafer Zeroscan
FX200 | XY Orthogonality
FX200 | MicrometerPerAdc of Z Height (Z Height calibration)
FX200 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
FX200 | Z linearity by sample height
FX200 | 1) Z Scanner position at +3μm
FX200 | 3) Z Scanner position at -3μm
FX200 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
FX200 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
FX200 | X Scan Stroke
FX200 | Y Scan Stroke
FX200 | Full range of X Detector Raw during X Scan sweep
FX200 | Full range of Y Detector Raw during Y Scan sweep
FX200 | X Detector noise (XY Servo off)
FX200 | Y Detector noise (XY Servo off)
FX200 | X axis OPM
FX200 | Y axis OPM
FX200 | X Non-linearity
FX200 | Y Non-linearity
FX200 | OPM Repeatability @25mm
FX200 | OPM Repeatability @10mm
FX200 | Max OPM @25mm
FX200 | Max OPM @10mm
FX200 | Torque Profile (50mm, 200um/s)
FX200 | Torque Profile (50mm, 20um/s)
FX200 | SLD current limit check
FX200 | SLD beam power at the end of the lens
FX200 | 2) of 20nm
FX200 | X Detector noise (XYServo off)
FX200 | Y Detector noise (XYServo off)
FX200 | X axis OPM raw Rpv
FX200 | Y axis OPM raw Rpv
FX200 | StageSafetyXMax
FX200 | StageSafetyXMin
FX200 | StageSafetyY
FX200 | StageSafetyYMax
FX200 | StageSafetyYMin
FX200 | Check X axis upper limit sensor function
FX200 | Check X axis lower limit sensor function
FX200 | Check Y axis upper limit sensor function
FX200 | Check Y axis lower limit sensor function
FX200 | X axis Angle Range
FX200 | Short-term stability w/ 450nm Step Height Sample (3σ)
FX200 | X detector noise (XY Servo on)
FX200 | Y detector noise (XY Servo on)
FX200 | Full range in um during X Scan sweep (XY Servo off)
FX200 | Full range in um during Y Scan sweep (XY Servo off)
FX200 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
FX200 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
FX200 | X axis, Repeatability @1st position
FX200 | X axis, Repeatability @ 2nd position
FX200 | X axis, Repeatability @ 3rd position
FX200 | X axis, Repeatability @ 4th position
FX200 | X axis, Accuracy @1st position
FX200 | X axis, Accuracy @ 2nd position
FX200 | X axis, Accuracy @ 3rd position
FX200 | X axis, Accuracy @ 4th position
FX200 | Y axis, Repeatability @1st position
FX200 | Y axis, Repeatability @ 2nd position
FX200 | Y axis, Repeatability @ 3rd position
FX200 | Y axis, Repeatability @ 4th position
FX200 | Y axis, Accuracy @1st position
FX200 | Y axis, Accuracy @ 2nd position
FX200 | Y axis, Accuracy @ 3rd position
FX200 | Y axis, Accuracy @ 4th position
FX200 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
FX200 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
FX200 | Z-XY Orthogonality (X direction-Left)
FX200 | Z-XY Orthogonality (X direction-Right)
FX200 | Z-XY Orthogonality (Y direction-Left)
FX200 | Z-XY Orthogonality (Y direction-Right)
FX200 | X-Scan Repeatability (10 repeats, RSD%)
FX200 | Y-Scan Repeatability (10 repeats, RSD%)
FX200 | Mechanical Align Accuracy
FX200 | Check the Multi Peattern Upload
FX200 | X Detector noise (XY servo off)
FX200 | Y Detector noise (XY servo off)
FX200 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
FX200 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
FX200 | Z Detector Max
FX200 | Z Detector Min
FX200 | WLI_Scale(Z Calibration)
FX200 | RMS of Z Height of Wafer Zeroscan(0,90000)
FX200 | RMS of Z Height of Wafer Zeroscan(-90000,0)
FX200 | RMS of Z Height of Wafer Zeroscan(0,-90000)
FX200 | RMS of Z Height of Wafer Zeroscan(90000,0)
FX200 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | Short-term stability w/ 180nm Step Height Sample (3σ)
FX200 | Laser intensity with AC160TS cantilever
FX200 | Check if the driving % changes proportionally as the NCM amplitude changes
FX200 | NCM driving % at an amplitude of 1 nm
FX200 | NCM driving % at an amplitude of 20 nm
FX200 | NCM driving % at an amplitude of 100 nm
FX200 | - 940nm Step Height Sample (10 Repeats)
FX200 | - 200~500nm Step Height Sample (10 Repeats)
FX200 | - 50~200nm Step Height Sample (10 Repeats)
FX200 | - Under 50nm Step Height Sample (10 Repeats)
FX200 | A-B after Auto PSPD Alignment
FX200 | NCM driving% with AC160TS cantilever
FX200 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
FX200 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
FX200 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
FX200 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
FX200 | X Detector stroke calibration error (nominal scale error X)
FX200 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | Resonance Frequency in DB
FX200 | Servo Cut-off Frequency in DB
FX200 | Check chuck surface no damage
FX200 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
FX200 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
FX200 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
FX200 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
FX200 | X1 Detector noise (XY servo off)
FX200 | X2 Detector noise (XY servo off)
FX200 | Y1 Detector noise (XY servo off)
FX200 | Y2 Detector noise (XY servo off)
FX200 | X Detector noise (XY servo on)
FX200 | X1 Detector noise (XY servo on)
FX200 | X2 Detector noise (XY servo on)
FX200 | Y Detector noise (XY servo on)
FX200 | Y1 Detector noise (XY servo on)
FX200 | Y2 Detector noise (XY servo on)
FX200 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
FX200 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
FX200 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
FX200 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
FX200 | Z Detector offset
FX200 | Check the full stroke image(servo on)
FX200 | ResonanceFrequencyHz
FX200 | ServoCutoffFrequencyHz
FX200 | FullPixelHeight
FX200 | Serial Port#
FX200 | Off-axis Camera X Position(ΔX)
FX200 | Roughness Repeatability 1x1um(10Repeats, RSD%)
FX200 | Roughness Repeatability 80x80um(10Repeats, RSD%)
FX200 | Z Scanner position at -3μm
FX200 | Z Scanner position at +3μm
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
FX200 | Scan offset accuracy (Precision)
FX200 | Aligned stage accuracy (X offset-Point Align1)
FX200 | Aligned stage accuracy (Y offset-Point Align1)
FX200 | Aligned stage accuracy (X offset-Point Align2)
FX200 | Aligned stage accuracy (Y offset-Point Align2)
FX200 | Aligned stage accuracy (X offset-Point Align3)
FX200 | Aligned stage accuracy (Y offset-Point Align3)
FX200 | Aligned stage accuracy (X offset)
FX200 | Aligned stage accuracy (Y offset)
FX200 | Imag offset after EFEM loading (X scan)
FX200 | Imag offset after EFEM loading (Y scan)
FX200 | Imag offset after Lock stage (X scan)
FX200 | Imag offset after Lock stage (Y scan)
FX200 | Mask Roughness Repeatability (10Repeats, RSD%)
FX200 | Mask Roughness Repeatability (10Repeats, σ )
FX200 | Mask exchange time
FX200 | WLI_Scale(C Calibration)
FX200 | WLI_Scale(B Calibration)
FX200 | WLI_Scale(A Calibration)
FX200 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
FX200 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
FX200 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
FX200 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
FX200 | Port 1 lens(to OnAxis) Y Offset
FX200 | Port 1 lens(to OnAxis) X Offset
FX200 | P Gain
FX200 | I Gain
FX200 | D Gain
FX200 | Check the Port 2 lens light strength
FX200 | Check the Port 1 lens light strength
FX200 | Check the focus deviation between the lenses(Δf)
FX200 | Check the 25mm Moving – Y Axis Offset Check
FX200 | C Scanner Detector Min
FX200 | C Scanner Detector Max
FX200 | B Scanner Detector Min
FX200 | B Scanner Detector Max
FX200 | Angle align between Port 2 lens and VLSI SHS: Y
FX200 | Angle align between Port 2 lens and VLSI SHS: X
FX200 | Angle align between Port 2 lens and Sample chuck: Y
FX200 | Angle align between Port 2 lens and Sample chuck: X
FX200 | Angle align between Port 1 lens and Sample chuck: Y
FX200 | Angle align between Port 1 lens and Sample chuck: X
FX200 | A Scanner Detector Min
FX200 | A Scanner Detector Max
FX200 | Y2 Detector noise (XY Servo off)
FX200 | X2 Detector noise (XY Servo off)
FX200 | Y axis OPM (-52000,-52000)
FX200 | Y axis OPM (-52000,52000)
FX200 | Y axis OPM (52000,-52000)
FX200 | Y axis OPM (52000,52000)
FX200 | Y axis OPM (0,0)
FX200 | XY Orthogonality (-52000,-52000)
FX200 | XY Orthogonality (-52000,52000)
FX200 | XY Orthogonality (52000,-52000)
FX200 | XY Orthogonality (52000,52000)
FX200 | XY Orthogonality (0,0)
FX200 | X axis OPM (-52000,-52000)
FX200 | X axis OPM (-52000,52000)
FX200 | X axis OPM (52000,-52000)
FX200 | X axis OPM (52000,52000)
FX200 | X axis OPM (0,0)
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
FX200 | RMS of Z Height of Wafer Zeroscan(-60000,0)
FX200 | RMS of Z Height of Wafer Zeroscan(60000,0)
FX200 | RMS of Z Height of Wafer Zeroscan(0,-60000)
FX200 | RMS of Z Height of Wafer Zeroscan(0,60000)
FX200 | XY stage calibration
FX200 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
FX200 | LD beam power at the end of the lens
FX200 | Y Non-lineairty
FX200 | X Non-lineairty
FX200 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
FX200 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
FX200 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
FX200 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
FX200 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
FX200 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
FX200 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
FX200 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
FX200 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
FX200 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
FX200 | UpperLimit
FX200 | Under 50nm Step Height Sample (10 Repeats)
FX200 | Overlay Camera X Position(ΔX)
FX200 | LowerLimit
FX200 | Initial Z
FX200 | 940nm Step Height Sample (10 Repeats)
FX200 | 50~200nm Step Height Sample (10 Repeats)
FX200 | RMS of Z Height of C122 Zeroscan(0,-90000)
FX200 | StrokeYMin
FX200 | StrokeYMax
FX200 | StrokeY
FX200 | StrokeXMin
FX200 | StrokeXMax
FX200 | StrokeX
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
FX200 | Z Drift (Short Traveling)
FX200 | 200~500nm Step Height Sample (10 Repeats)
FX200 | XY orthogonality
FX200 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
FX200 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
FX200 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
FX200 | Spring constant calibration by Sader Method with 3XC-NA-500DC
FX200 | Spring constant calibration by Sader Method with 3XC-NA-240AC
FX200 | Spring constant calibration by Sader Method with 3XC-NA-200AC
FX200 | Y1 Detector noise (XY Servo off)
FX200 | X1 Detector noise (XY Servo off)
FX200 | Maximum Y Scan range (Full Stroke)
FX200 | Maximum X Scan range (Full Stroke)
FX200 | Full range of Y2 Detector raw
FX200 | Full range of Y1 Detector raw
FX200 | Full range of X2 Detector raw
FX200 | Full range of X1 Detector raw
FX200 | Z Linearity by Z Scanner position
FX200 | -Y Direction software Limit
FX200 | -X Direction software Limit
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
FX200 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
FX200 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
FX200 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
FX200 | %GRR w/ roughness of non-patterened wafer
FX200 | Z Scanner position at -6μm
FX200 | Z Scanner position at +6μm
FX200 | Approach position at -5μm
FX200 | Approach position at +5μm
FX200 | * 1800nm Step Height Sample (10 Repeats)
FX200 | Full range of Z Detector raw
FX200 | Approach position at -3μm
FX200 | Approach position at +3μm
FX200 | Under 50nm Step Height Sample (10 Repeats)
FX200 | 50~200nm Step Height Sample (10 Repeats)
FX200 | 200~500nm Step Height Sample (10 Repeats)
FX200 | PosMin
FX200 | PosMax
FX200 | MaxLength
FX200 | FullRecoveryProfileLength
FX200 | Full range of Y Detector raw
FX200 | Full range of X Detector raw
FX200 | EFM Phase : For the positive bias
FX200 | EFM Phase : For the negative bias
FX200 | Z Linearity by 5 different sample height
FX200 | Laser intensity with AC160 cantilever
FX200 | A－B difference during 3 times run test
FX200 | A＋B difference during 3 times run test
FX200 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
FX200 | RMS of Z Drive of Wafer Zeroscan(60000,0)
FX200 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
FX200 | RMS of Z Drive of Wafer Zeroscan(0,60000)
FX200 | RMS of Z Drive of Wafer ZeroScan (0,0)
FX200 | Allowable max warpage, w/ SiC sample chuck
FX200 | RMS of Z Height of Wafer Zeroscan (Top)
FX200 | RMS of Z Height of Wafer Zeroscan (Right)
FX200 | RMS of Z Height of Wafer Zeroscan (Left)
FX200 | RMS of Z Height of Wafer ZeroScan (Center)
FX200 | RMS of Z Height of Wafer Zeroscan (Bottom)
FX200 | PosMin
FX200 | PosMax
FX200 | MaxLength
FX200 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
FX200 | Check sample surface mapping
FX200 | Z Direction, Accuracy Test (Bias not significant)
FX200 | Y Direction, Accuracy Test (Bias not significant)
FX200 | X Direction, Accuracy Test (Bias not significant)
FX200 | Roughness Reproducibility Test (SR Panel)
FX200 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
FX200 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
FX200 | Roughness Reproducibility Test (SiNx Panel)
FX200 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
FX200 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
FX200 | Roughness Reproducibility Test (ABF Panel)
FX200 | Roughness Repeatability Test (SR Panel)
FX200 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
FX200 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
FX200 | Roughness Repeatability Test (SiNx Panel)
FX200 | Roughness Repeatability Test (Rough Cu Panel, Trace)
FX200 | Roughness Repeatability Test (Rough Cu Panel, Plane)
FX200 | Roughness Repeatability Test (ABF Panel)
FX200 | Lift Pin Stroke Min Check in DB
FX200 | Lift Pin Stroke Max Check in DB
FX200 | Limit range
NX15 | Z Linearity by Z scanner position
NX15 | Full range of Z Detector during Z Scan sweep
NX15 | Z Height calibration error
NX15 | Z Drive calibration error
NX15 | Z Linearity by 4 different sample height
NX15 | Full range of Z Height during Z Scan sweep
NX15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX15 | Z Height noise
NX15 | Wafer placement(Loading) error on the load port1 (X offset)
NX15 | Wafer placement(Loading) error on the load port1 (Y offset)
NX15 | Wafer placement(Loading) error on the load port2 (X offset)
NX15 | Wafer placement(Loading) error on the load port2 (Y offset)
NX15 | OpticalAxisShiftPerFocusX
NX15 | OpticalAxisShiftPerFocusY
NX15 | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX15 | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX15 | Lift pins synchronization offset check (X offset)
NX15 | Lift pins synchronization offset check (Y offset)
NX15 | Profiler noise @200um
NX15 | Roll motion to X+ full scan
NX15 | Position repeatability(10repeats, Position Based)
NX15 | OPM Repeatability @5mm
NX15 | OPM Repeatability @1mm
NX15 | Max OPM @5mm
NX15 | Max OPM @1mm
NX15 | Torque Profile (10mm, 200um/s)
NX15 | Torque Profile (10mm, 20um/s)
NX15 | Lateral(C-D) sensitivity
NX15 | Lateral(C-D) noise
NX15 | Min. Vibration Transmissbility (~100Hz)
NX15 | RMS of ZeroScan
NX15 | LightStrengthGain
NX15 | Full range of X Detector during X Scan sweep
NX15 | Full range of X1 Detector during X Scan sweep
NX15 | Full range of X2 Detector during X Scan sweep
NX15 | Full range of Y Detector during Y Scan sweep
NX15 | Full range of Y1 Detector during Y Scan sweep
NX15 | Full range of Y2 Detector during Y Scan sweep
NX15 | Full range of X1 Detector Raw during X Scan sweep
NX15 | Full range of X2 Detector Raw during X Scan sweep
NX15 | Full range of Y1 Detector Raw during Y Scan sweep
NX15 | Full range of Y2 Detector Raw during Y Scan sweep
NX15 | X Detector stroke calibration error (nominal scale error X)
NX15 | Y Detector stroke calibration error (nominal scale error Y)
NX15 | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX15 | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX15 | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX15 | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX15 | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX15 | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX15 | X Non-linearity error
NX15 | Y Non-linearity error
NX15 | X axis OPM
NX15 | Y axis OPM
NX15 | X Ringing(Rpv)
NX15 | Y Ringing(Rpv)
NX15 | X Detector noise (XY Servo on)
NX15 | Y Detector noise (XY Servo on)
NX15 | X1 Detector noise (XY Servo on)
NX15 | Y1 Detector noise (XY Servo on)
NX15 | X2 Detector noise (XY Servo on)
NX15 | Y2 Detector noise (XY Servo on)
NX15 | RMS error after Stage Calibration
NX15 | Image offset after moving from -X to +X (X scan)
NX15 | Image offset after moving from -Y to +Y (Y scan)
NX15 | Image offset after moving from -X to +X (Y scan)
NX15 | Image offset after moving from -Y to +Y (X scan)
NX15 | Image offset after moving from +X to -X (X scan)
NX15 | Image offset after moving from +Y to -Y (Y scan)
NX15 | Image offset after moving from +X to -X (Y scan)
NX15 | Image offset after moving from +Y to -Y (X scan)
NX15 | XY Stage positioning X offset (Vision Pattern Matching)
NX15 | XY Stage positioning Y offset (Vision Pattern Matching)
NX15 | XY Stage positioning X 1σ (Vision Pattern Matching)
NX15 | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX15 | XY Stage positioning X offset (In-Die Align)
NX15 | XY Stage positioning Y offset (In-Die Align)
NX15 | XY Stage positioning X 1σ (In-Die Align)
NX15 | XY Stage positioning Y 1σ (In-Die Align)
NX15 | XY Stage positioning X offset (LLC Translation)
NX15 | XY Stage positioning Y offset (LLC Translation)
NX15 | XY Stage positioning X 1σ (LLC Translation)
NX15 | XY Stage positioning Y 1σ (LLC Translation)
NX15 | XY Stage positioning X offset (Global Align)
NX15 | XY Stage positioning Y offset (Global Align)
NX15 | XY Stage positioning X 1σ (Global Align)
NX15 | XY Stage positioning Y 1σ (Global Align)
NX15 | Z Drift
NX15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | Cycle time for 15-point measurement
NX15 | Short-term stability using Step Height Sample (3σ)
NX15 | Roughness Repeatability 1x1um(RSD%)
NX15 | Roughness Repeatability 1x1um(1σ)
NX15 | Roughness Repeatability 80x80um(RSD%)
NX15 | Roughness Repeatability 80x80um(1σ)
NX15 | Z Height calibrtaion error
NX15 | Z Drive calibrtaion error
NX15 | Z Linearity by 3 different sample height
NX15 | Full range in voltage during Z Scan sweep
NX15 | Full range in um during Z Scan sweep
NX15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX15 | Lateral(C-D) sensitivity
NX15 | Lateral(C-D) noise
NX15 | A-B noise
NX15 | NCM Amplitude noise
NX15 | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX15 | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX15 | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX15 | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX15 | RMS of Z Height of Wafer Zeroscan(0,120000)
NX15 | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX15 | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX15 | RMS of Z Height of Wafer Zeroscan(120000,0)
NX15 | RMS of Z Height of Wafer ZeroScan (0,0)
NX15 | Max Wafer Zeroscan-Min Wafer Zeroscan
NX15 | XY Orthogonality
NX15 | MicrometerPerAdc of Z Height (Z Height calibration)
NX15 | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX15 | Z linearity by sample height
NX15 | 1) Z Scanner position at +3μm
NX15 | 3) Z Scanner position at -3μm
NX15 | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX15 | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX15 | X Scan Stroke
NX15 | Y Scan Stroke
NX15 | Full range of X Detector Raw during X Scan sweep
NX15 | Full range of Y Detector Raw during Y Scan sweep
NX15 | X Detector noise (XY Servo off)
NX15 | Y Detector noise (XY Servo off)
NX15 | X axis OPM
NX15 | Y axis OPM
NX15 | X Non-linearity
NX15 | Y Non-linearity
NX15 | OPM Repeatability @25mm
NX15 | OPM Repeatability @10mm
NX15 | Max OPM @25mm
NX15 | Max OPM @10mm
NX15 | Torque Profile (50mm, 200um/s)
NX15 | Torque Profile (50mm, 20um/s)
NX15 | SLD current limit check
NX15 | SLD beam power at the end of the lens
NX15 | 2) of 20nm
NX15 | X Detector noise (XYServo off)
NX15 | Y Detector noise (XYServo off)
NX15 | X axis OPM raw Rpv
NX15 | Y axis OPM raw Rpv
NX15 | StageSafetyXMax
NX15 | StageSafetyXMin
NX15 | StageSafetyY
NX15 | StageSafetyYMax
NX15 | StageSafetyYMin
NX15 | Check X axis upper limit sensor function
NX15 | Check X axis lower limit sensor function
NX15 | Check Y axis upper limit sensor function
NX15 | Check Y axis lower limit sensor function
NX15 | X axis Angle Range
NX15 | Short-term stability w/ 450nm Step Height Sample (3σ)
NX15 | X detector noise (XY Servo on)
NX15 | Y detector noise (XY Servo on)
NX15 | Full range in um during X Scan sweep (XY Servo off)
NX15 | Full range in um during Y Scan sweep (XY Servo off)
NX15 | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX15 | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX15 | X axis, Repeatability @1st position
NX15 | X axis, Repeatability @ 2nd position
NX15 | X axis, Repeatability @ 3rd position
NX15 | X axis, Repeatability @ 4th position
NX15 | X axis, Accuracy @1st position
NX15 | X axis, Accuracy @ 2nd position
NX15 | X axis, Accuracy @ 3rd position
NX15 | X axis, Accuracy @ 4th position
NX15 | Y axis, Repeatability @1st position
NX15 | Y axis, Repeatability @ 2nd position
NX15 | Y axis, Repeatability @ 3rd position
NX15 | Y axis, Repeatability @ 4th position
NX15 | Y axis, Accuracy @1st position
NX15 | Y axis, Accuracy @ 2nd position
NX15 | Y axis, Accuracy @ 3rd position
NX15 | Y axis, Accuracy @ 4th position
NX15 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX15 | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX15 | Z-XY Orthogonality (X direction-Left)
NX15 | Z-XY Orthogonality (X direction-Right)
NX15 | Z-XY Orthogonality (Y direction-Left)
NX15 | Z-XY Orthogonality (Y direction-Right)
NX15 | X-Scan Repeatability (10 repeats, RSD%)
NX15 | Y-Scan Repeatability (10 repeats, RSD%)
NX15 | Mechanical Align Accuracy
NX15 | Check the Multi Peattern Upload
NX15 | X Detector noise (XY servo off)
NX15 | Y Detector noise (XY servo off)
NX15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX15 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX15 | Z Detector Max
NX15 | Z Detector Min
NX15 | WLI_Scale(Z Calibration)
NX15 | RMS of Z Height of Wafer Zeroscan(0,90000)
NX15 | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX15 | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX15 | RMS of Z Height of Wafer Zeroscan(90000,0)
NX15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | Short-term stability w/ 180nm Step Height Sample (3σ)
NX15 | Laser intensity with AC160TS cantilever
NX15 | Check if the driving % changes proportionally as the NCM amplitude changes
NX15 | NCM driving % at an amplitude of 1 nm
NX15 | NCM driving % at an amplitude of 20 nm
NX15 | NCM driving % at an amplitude of 100 nm
NX15 | - 940nm Step Height Sample (10 Repeats)
NX15 | - 200~500nm Step Height Sample (10 Repeats)
NX15 | - 50~200nm Step Height Sample (10 Repeats)
NX15 | - Under 50nm Step Height Sample (10 Repeats)
NX15 | A-B after Auto PSPD Alignment
NX15 | NCM driving% with AC160TS cantilever
NX15 | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX15 | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX15 | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX15 | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX15 | X Detector stroke calibration error (nominal scale error X)
NX15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | Resonance Frequency in DB
NX15 | Servo Cut-off Frequency in DB
NX15 | Check chuck surface no damage
NX15 | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX15 | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX15 | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX15 | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX15 | X1 Detector noise (XY servo off)
NX15 | X2 Detector noise (XY servo off)
NX15 | Y1 Detector noise (XY servo off)
NX15 | Y2 Detector noise (XY servo off)
NX15 | X Detector noise (XY servo on)
NX15 | X1 Detector noise (XY servo on)
NX15 | X2 Detector noise (XY servo on)
NX15 | Y Detector noise (XY servo on)
NX15 | Y1 Detector noise (XY servo on)
NX15 | Y2 Detector noise (XY servo on)
NX15 | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX15 | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX15 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX15 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX15 | Z Detector offset
NX15 | Check the full stroke image(servo on)
NX15 | ResonanceFrequencyHz
NX15 | ServoCutoffFrequencyHz
NX15 | FullPixelHeight
NX15 | Serial Port#
NX15 | Off-axis Camera X Position(ΔX)
NX15 | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX15 | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX15 | Z Scanner position at -3μm
NX15 | Z Scanner position at +3μm
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX15 | Scan offset accuracy (Precision)
NX15 | Aligned stage accuracy (X offset-Point Align1)
NX15 | Aligned stage accuracy (Y offset-Point Align1)
NX15 | Aligned stage accuracy (X offset-Point Align2)
NX15 | Aligned stage accuracy (Y offset-Point Align2)
NX15 | Aligned stage accuracy (X offset-Point Align3)
NX15 | Aligned stage accuracy (Y offset-Point Align3)
NX15 | Aligned stage accuracy (X offset)
NX15 | Aligned stage accuracy (Y offset)
NX15 | Imag offset after EFEM loading (X scan)
NX15 | Imag offset after EFEM loading (Y scan)
NX15 | Imag offset after Lock stage (X scan)
NX15 | Imag offset after Lock stage (Y scan)
NX15 | Mask Roughness Repeatability (10Repeats, RSD%)
NX15 | Mask Roughness Repeatability (10Repeats, σ )
NX15 | Mask exchange time
NX15 | WLI_Scale(C Calibration)
NX15 | WLI_Scale(B Calibration)
NX15 | WLI_Scale(A Calibration)
NX15 | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX15 | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX15 | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX15 | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX15 | Port 1 lens(to OnAxis) Y Offset
NX15 | Port 1 lens(to OnAxis) X Offset
NX15 | P Gain
NX15 | I Gain
NX15 | D Gain
NX15 | Check the Port 2 lens light strength
NX15 | Check the Port 1 lens light strength
NX15 | Check the focus deviation between the lenses(Δf)
NX15 | Check the 25mm Moving – Y Axis Offset Check
NX15 | C Scanner Detector Min
NX15 | C Scanner Detector Max
NX15 | B Scanner Detector Min
NX15 | B Scanner Detector Max
NX15 | Angle align between Port 2 lens and VLSI SHS: Y
NX15 | Angle align between Port 2 lens and VLSI SHS: X
NX15 | Angle align between Port 2 lens and Sample chuck: Y
NX15 | Angle align between Port 2 lens and Sample chuck: X
NX15 | Angle align between Port 1 lens and Sample chuck: Y
NX15 | Angle align between Port 1 lens and Sample chuck: X
NX15 | A Scanner Detector Min
NX15 | A Scanner Detector Max
NX15 | Y2 Detector noise (XY Servo off)
NX15 | X2 Detector noise (XY Servo off)
NX15 | Y axis OPM (-52000,-52000)
NX15 | Y axis OPM (-52000,52000)
NX15 | Y axis OPM (52000,-52000)
NX15 | Y axis OPM (52000,52000)
NX15 | Y axis OPM (0,0)
NX15 | XY Orthogonality (-52000,-52000)
NX15 | XY Orthogonality (-52000,52000)
NX15 | XY Orthogonality (52000,-52000)
NX15 | XY Orthogonality (52000,52000)
NX15 | XY Orthogonality (0,0)
NX15 | X axis OPM (-52000,-52000)
NX15 | X axis OPM (-52000,52000)
NX15 | X axis OPM (52000,-52000)
NX15 | X axis OPM (52000,52000)
NX15 | X axis OPM (0,0)
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX15 | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX15 | RMS of Z Height of Wafer Zeroscan(60000,0)
NX15 | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX15 | RMS of Z Height of Wafer Zeroscan(0,60000)
NX15 | XY stage calibration
NX15 | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX15 | LD beam power at the end of the lens
NX15 | Y Non-lineairty
NX15 | X Non-lineairty
NX15 | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX15 | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX15 | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX15 | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX15 | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX15 | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX15 | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX15 | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX15 | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX15 | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX15 | UpperLimit
NX15 | Under 50nm Step Height Sample (10 Repeats)
NX15 | Overlay Camera X Position(ΔX)
NX15 | LowerLimit
NX15 | Initial Z
NX15 | 940nm Step Height Sample (10 Repeats)
NX15 | 50~200nm Step Height Sample (10 Repeats)
NX15 | RMS of Z Height of C122 Zeroscan(0,-90000)
NX15 | StrokeYMin
NX15 | StrokeYMax
NX15 | StrokeY
NX15 | StrokeXMin
NX15 | StrokeXMax
NX15 | StrokeX
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX15 | Z Drift (Short Traveling)
NX15 | 200~500nm Step Height Sample (10 Repeats)
NX15 | XY orthogonality
NX15 | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX15 | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX15 | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX15 | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX15 | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX15 | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX15 | Y1 Detector noise (XY Servo off)
NX15 | X1 Detector noise (XY Servo off)
NX15 | Maximum Y Scan range (Full Stroke)
NX15 | Maximum X Scan range (Full Stroke)
NX15 | Full range of Y2 Detector raw
NX15 | Full range of Y1 Detector raw
NX15 | Full range of X2 Detector raw
NX15 | Full range of X1 Detector raw
NX15 | Z Linearity by Z Scanner position
NX15 | -Y Direction software Limit
NX15 | -X Direction software Limit
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX15 | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX15 | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX15 | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX15 | %GRR w/ roughness of non-patterened wafer
NX15 | Z Scanner position at -6μm
NX15 | Z Scanner position at +6μm
NX15 | Approach position at -5μm
NX15 | Approach position at +5μm
NX15 | * 1800nm Step Height Sample (10 Repeats)
NX15 | Full range of Z Detector raw
NX15 | Approach position at -3μm
NX15 | Approach position at +3μm
NX15 | Under 50nm Step Height Sample (10 Repeats)
NX15 | 50~200nm Step Height Sample (10 Repeats)
NX15 | 200~500nm Step Height Sample (10 Repeats)
NX15 | PosMin
NX15 | PosMax
NX15 | MaxLength
NX15 | FullRecoveryProfileLength
NX15 | Full range of Y Detector raw
NX15 | Full range of X Detector raw
NX15 | EFM Phase : For the positive bias
NX15 | EFM Phase : For the negative bias
NX15 | Z Linearity by 5 different sample height
NX15 | Laser intensity with AC160 cantilever
NX15 | A－B difference during 3 times run test
NX15 | A＋B difference during 3 times run test
NX15 | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX15 | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX15 | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX15 | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX15 | RMS of Z Drive of Wafer ZeroScan (0,0)
NX15 | Allowable max warpage, w/ SiC sample chuck
NX15 | RMS of Z Height of Wafer Zeroscan (Top)
NX15 | RMS of Z Height of Wafer Zeroscan (Right)
NX15 | RMS of Z Height of Wafer Zeroscan (Left)
NX15 | RMS of Z Height of Wafer ZeroScan (Center)
NX15 | RMS of Z Height of Wafer Zeroscan (Bottom)
NX15 | PosMin
NX15 | PosMax
NX15 | MaxLength
NX15 | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX15 | Check sample surface mapping
NX15 | Z Direction, Accuracy Test (Bias not significant)
NX15 | Y Direction, Accuracy Test (Bias not significant)
NX15 | X Direction, Accuracy Test (Bias not significant)
NX15 | Roughness Reproducibility Test (SR Panel)
NX15 | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX15 | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX15 | Roughness Reproducibility Test (SiNx Panel)
NX15 | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX15 | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX15 | Roughness Reproducibility Test (ABF Panel)
NX15 | Roughness Repeatability Test (SR Panel)
NX15 | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX15 | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX15 | Roughness Repeatability Test (SiNx Panel)
NX15 | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX15 | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX15 | Roughness Repeatability Test (ABF Panel)
NX15 | Lift Pin Stroke Min Check in DB
NX15 | Lift Pin Stroke Max Check in DB
NX15 | Limit range
NX-TSH | Z Linearity by Z scanner position
NX-TSH | Full range of Z Detector during Z Scan sweep
NX-TSH | Z Height calibration error
NX-TSH | Z Drive calibration error
NX-TSH | Z Linearity by 4 different sample height
NX-TSH | Full range of Z Height during Z Scan sweep
NX-TSH | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH | Z Height noise
NX-TSH | Wafer placement(Loading) error on the load port1 (X offset)
NX-TSH | Wafer placement(Loading) error on the load port1 (Y offset)
NX-TSH | Wafer placement(Loading) error on the load port2 (X offset)
NX-TSH | Wafer placement(Loading) error on the load port2 (Y offset)
NX-TSH | OpticalAxisShiftPerFocusX
NX-TSH | OpticalAxisShiftPerFocusY
NX-TSH | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-TSH | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-TSH | Lift pins synchronization offset check (X offset)
NX-TSH | Lift pins synchronization offset check (Y offset)
NX-TSH | Profiler noise @200um
NX-TSH | Roll motion to X+ full scan
NX-TSH | Position repeatability(10repeats, Position Based)
NX-TSH | OPM Repeatability @5mm
NX-TSH | OPM Repeatability @1mm
NX-TSH | Max OPM @5mm
NX-TSH | Max OPM @1mm
NX-TSH | Torque Profile (10mm, 200um/s)
NX-TSH | Torque Profile (10mm, 20um/s)
NX-TSH | Lateral(C-D) sensitivity
NX-TSH | Lateral(C-D) noise
NX-TSH | Min. Vibration Transmissbility (~100Hz)
NX-TSH | RMS of ZeroScan
NX-TSH | LightStrengthGain
NX-TSH | Full range of X Detector during X Scan sweep
NX-TSH | Full range of X1 Detector during X Scan sweep
NX-TSH | Full range of X2 Detector during X Scan sweep
NX-TSH | Full range of Y Detector during Y Scan sweep
NX-TSH | Full range of Y1 Detector during Y Scan sweep
NX-TSH | Full range of Y2 Detector during Y Scan sweep
NX-TSH | Full range of X1 Detector Raw during X Scan sweep
NX-TSH | Full range of X2 Detector Raw during X Scan sweep
NX-TSH | Full range of Y1 Detector Raw during Y Scan sweep
NX-TSH | Full range of Y2 Detector Raw during Y Scan sweep
NX-TSH | X Detector stroke calibration error (nominal scale error X)
NX-TSH | Y Detector stroke calibration error (nominal scale error Y)
NX-TSH | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-TSH | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-TSH | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-TSH | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-TSH | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-TSH | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-TSH | X Non-linearity error
NX-TSH | Y Non-linearity error
NX-TSH | X axis OPM
NX-TSH | Y axis OPM
NX-TSH | X Ringing(Rpv)
NX-TSH | Y Ringing(Rpv)
NX-TSH | X Detector noise (XY Servo on)
NX-TSH | Y Detector noise (XY Servo on)
NX-TSH | X1 Detector noise (XY Servo on)
NX-TSH | Y1 Detector noise (XY Servo on)
NX-TSH | X2 Detector noise (XY Servo on)
NX-TSH | Y2 Detector noise (XY Servo on)
NX-TSH | RMS error after Stage Calibration
NX-TSH | Image offset after moving from -X to +X (X scan)
NX-TSH | Image offset after moving from -Y to +Y (Y scan)
NX-TSH | Image offset after moving from -X to +X (Y scan)
NX-TSH | Image offset after moving from -Y to +Y (X scan)
NX-TSH | Image offset after moving from +X to -X (X scan)
NX-TSH | Image offset after moving from +Y to -Y (Y scan)
NX-TSH | Image offset after moving from +X to -X (Y scan)
NX-TSH | Image offset after moving from +Y to -Y (X scan)
NX-TSH | XY Stage positioning X offset (Vision Pattern Matching)
NX-TSH | XY Stage positioning Y offset (Vision Pattern Matching)
NX-TSH | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-TSH | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-TSH | XY Stage positioning X offset (In-Die Align)
NX-TSH | XY Stage positioning Y offset (In-Die Align)
NX-TSH | XY Stage positioning X 1σ (In-Die Align)
NX-TSH | XY Stage positioning Y 1σ (In-Die Align)
NX-TSH | XY Stage positioning X offset (LLC Translation)
NX-TSH | XY Stage positioning Y offset (LLC Translation)
NX-TSH | XY Stage positioning X 1σ (LLC Translation)
NX-TSH | XY Stage positioning Y 1σ (LLC Translation)
NX-TSH | XY Stage positioning X offset (Global Align)
NX-TSH | XY Stage positioning Y offset (Global Align)
NX-TSH | XY Stage positioning X 1σ (Global Align)
NX-TSH | XY Stage positioning Y 1σ (Global Align)
NX-TSH | Z Drift
NX-TSH | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | Cycle time for 15-point measurement
NX-TSH | Short-term stability using Step Height Sample (3σ)
NX-TSH | Roughness Repeatability 1x1um(RSD%)
NX-TSH | Roughness Repeatability 1x1um(1σ)
NX-TSH | Roughness Repeatability 80x80um(RSD%)
NX-TSH | Roughness Repeatability 80x80um(1σ)
NX-TSH | Z Height calibrtaion error
NX-TSH | Z Drive calibrtaion error
NX-TSH | Z Linearity by 3 different sample height
NX-TSH | Full range in voltage during Z Scan sweep
NX-TSH | Full range in um during Z Scan sweep
NX-TSH | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH | Lateral(C-D) sensitivity
NX-TSH | Lateral(C-D) noise
NX-TSH | A-B noise
NX-TSH | NCM Amplitude noise
NX-TSH | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-TSH | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-TSH | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-TSH | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-TSH | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-TSH | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-TSH | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-TSH | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-TSH | RMS of Z Height of Wafer ZeroScan (0,0)
NX-TSH | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-TSH | XY Orthogonality
NX-TSH | MicrometerPerAdc of Z Height (Z Height calibration)
NX-TSH | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-TSH | Z linearity by sample height
NX-TSH | 1) Z Scanner position at +3μm
NX-TSH | 3) Z Scanner position at -3μm
NX-TSH | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-TSH | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-TSH | X Scan Stroke
NX-TSH | Y Scan Stroke
NX-TSH | Full range of X Detector Raw during X Scan sweep
NX-TSH | Full range of Y Detector Raw during Y Scan sweep
NX-TSH | X Detector noise (XY Servo off)
NX-TSH | Y Detector noise (XY Servo off)
NX-TSH | X axis OPM
NX-TSH | Y axis OPM
NX-TSH | X Non-linearity
NX-TSH | Y Non-linearity
NX-TSH | OPM Repeatability @25mm
NX-TSH | OPM Repeatability @10mm
NX-TSH | Max OPM @25mm
NX-TSH | Max OPM @10mm
NX-TSH | Torque Profile (50mm, 200um/s)
NX-TSH | Torque Profile (50mm, 20um/s)
NX-TSH | SLD current limit check
NX-TSH | SLD beam power at the end of the lens
NX-TSH | 2) of 20nm
NX-TSH | X Detector noise (XYServo off)
NX-TSH | Y Detector noise (XYServo off)
NX-TSH | X axis OPM raw Rpv
NX-TSH | Y axis OPM raw Rpv
NX-TSH | StageSafetyXMax
NX-TSH | StageSafetyXMin
NX-TSH | StageSafetyY
NX-TSH | StageSafetyYMax
NX-TSH | StageSafetyYMin
NX-TSH | Check X axis upper limit sensor function
NX-TSH | Check X axis lower limit sensor function
NX-TSH | Check Y axis upper limit sensor function
NX-TSH | Check Y axis lower limit sensor function
NX-TSH | X axis Angle Range
NX-TSH | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-TSH | X detector noise (XY Servo on)
NX-TSH | Y detector noise (XY Servo on)
NX-TSH | Full range in um during X Scan sweep (XY Servo off)
NX-TSH | Full range in um during Y Scan sweep (XY Servo off)
NX-TSH | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-TSH | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-TSH | X axis, Repeatability @1st position
NX-TSH | X axis, Repeatability @ 2nd position
NX-TSH | X axis, Repeatability @ 3rd position
NX-TSH | X axis, Repeatability @ 4th position
NX-TSH | X axis, Accuracy @1st position
NX-TSH | X axis, Accuracy @ 2nd position
NX-TSH | X axis, Accuracy @ 3rd position
NX-TSH | X axis, Accuracy @ 4th position
NX-TSH | Y axis, Repeatability @1st position
NX-TSH | Y axis, Repeatability @ 2nd position
NX-TSH | Y axis, Repeatability @ 3rd position
NX-TSH | Y axis, Repeatability @ 4th position
NX-TSH | Y axis, Accuracy @1st position
NX-TSH | Y axis, Accuracy @ 2nd position
NX-TSH | Y axis, Accuracy @ 3rd position
NX-TSH | Y axis, Accuracy @ 4th position
NX-TSH | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-TSH | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-TSH | Z-XY Orthogonality (X direction-Left)
NX-TSH | Z-XY Orthogonality (X direction-Right)
NX-TSH | Z-XY Orthogonality (Y direction-Left)
NX-TSH | Z-XY Orthogonality (Y direction-Right)
NX-TSH | X-Scan Repeatability (10 repeats, RSD%)
NX-TSH | Y-Scan Repeatability (10 repeats, RSD%)
NX-TSH | Mechanical Align Accuracy
NX-TSH | Check the Multi Peattern Upload
NX-TSH | X Detector noise (XY servo off)
NX-TSH | Y Detector noise (XY servo off)
NX-TSH | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH | Z Detector Max
NX-TSH | Z Detector Min
NX-TSH | WLI_Scale(Z Calibration)
NX-TSH | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-TSH | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-TSH | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-TSH | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-TSH | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-TSH | Laser intensity with AC160TS cantilever
NX-TSH | Check if the driving % changes proportionally as the NCM amplitude changes
NX-TSH | NCM driving % at an amplitude of 1 nm
NX-TSH | NCM driving % at an amplitude of 20 nm
NX-TSH | NCM driving % at an amplitude of 100 nm
NX-TSH | - 940nm Step Height Sample (10 Repeats)
NX-TSH | - 200~500nm Step Height Sample (10 Repeats)
NX-TSH | - 50~200nm Step Height Sample (10 Repeats)
NX-TSH | - Under 50nm Step Height Sample (10 Repeats)
NX-TSH | A-B after Auto PSPD Alignment
NX-TSH | NCM driving% with AC160TS cantilever
NX-TSH | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-TSH | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-TSH | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-TSH | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-TSH | X Detector stroke calibration error (nominal scale error X)
NX-TSH | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | Resonance Frequency in DB
NX-TSH | Servo Cut-off Frequency in DB
NX-TSH | Check chuck surface no damage
NX-TSH | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-TSH | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-TSH | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-TSH | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-TSH | X1 Detector noise (XY servo off)
NX-TSH | X2 Detector noise (XY servo off)
NX-TSH | Y1 Detector noise (XY servo off)
NX-TSH | Y2 Detector noise (XY servo off)
NX-TSH | X Detector noise (XY servo on)
NX-TSH | X1 Detector noise (XY servo on)
NX-TSH | X2 Detector noise (XY servo on)
NX-TSH | Y Detector noise (XY servo on)
NX-TSH | Y1 Detector noise (XY servo on)
NX-TSH | Y2 Detector noise (XY servo on)
NX-TSH | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-TSH | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-TSH | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH | Z Detector offset
NX-TSH | Check the full stroke image(servo on)
NX-TSH | ResonanceFrequencyHz
NX-TSH | ServoCutoffFrequencyHz
NX-TSH | FullPixelHeight
NX-TSH | Serial Port#
NX-TSH | Off-axis Camera X Position(ΔX)
NX-TSH | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-TSH | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-TSH | Z Scanner position at -3μm
NX-TSH | Z Scanner position at +3μm
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH | Scan offset accuracy (Precision)
NX-TSH | Aligned stage accuracy (X offset-Point Align1)
NX-TSH | Aligned stage accuracy (Y offset-Point Align1)
NX-TSH | Aligned stage accuracy (X offset-Point Align2)
NX-TSH | Aligned stage accuracy (Y offset-Point Align2)
NX-TSH | Aligned stage accuracy (X offset-Point Align3)
NX-TSH | Aligned stage accuracy (Y offset-Point Align3)
NX-TSH | Aligned stage accuracy (X offset)
NX-TSH | Aligned stage accuracy (Y offset)
NX-TSH | Imag offset after EFEM loading (X scan)
NX-TSH | Imag offset after EFEM loading (Y scan)
NX-TSH | Imag offset after Lock stage (X scan)
NX-TSH | Imag offset after Lock stage (Y scan)
NX-TSH | Mask Roughness Repeatability (10Repeats, RSD%)
NX-TSH | Mask Roughness Repeatability (10Repeats, σ )
NX-TSH | Mask exchange time
NX-TSH | WLI_Scale(C Calibration)
NX-TSH | WLI_Scale(B Calibration)
NX-TSH | WLI_Scale(A Calibration)
NX-TSH | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-TSH | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-TSH | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-TSH | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-TSH | Port 1 lens(to OnAxis) Y Offset
NX-TSH | Port 1 lens(to OnAxis) X Offset
NX-TSH | P Gain
NX-TSH | I Gain
NX-TSH | D Gain
NX-TSH | Check the Port 2 lens light strength
NX-TSH | Check the Port 1 lens light strength
NX-TSH | Check the focus deviation between the lenses(Δf)
NX-TSH | Check the 25mm Moving – Y Axis Offset Check
NX-TSH | C Scanner Detector Min
NX-TSH | C Scanner Detector Max
NX-TSH | B Scanner Detector Min
NX-TSH | B Scanner Detector Max
NX-TSH | Angle align between Port 2 lens and VLSI SHS: Y
NX-TSH | Angle align between Port 2 lens and VLSI SHS: X
NX-TSH | Angle align between Port 2 lens and Sample chuck: Y
NX-TSH | Angle align between Port 2 lens and Sample chuck: X
NX-TSH | Angle align between Port 1 lens and Sample chuck: Y
NX-TSH | Angle align between Port 1 lens and Sample chuck: X
NX-TSH | A Scanner Detector Min
NX-TSH | A Scanner Detector Max
NX-TSH | Y2 Detector noise (XY Servo off)
NX-TSH | X2 Detector noise (XY Servo off)
NX-TSH | Y axis OPM (-52000,-52000)
NX-TSH | Y axis OPM (-52000,52000)
NX-TSH | Y axis OPM (52000,-52000)
NX-TSH | Y axis OPM (52000,52000)
NX-TSH | Y axis OPM (0,0)
NX-TSH | XY Orthogonality (-52000,-52000)
NX-TSH | XY Orthogonality (-52000,52000)
NX-TSH | XY Orthogonality (52000,-52000)
NX-TSH | XY Orthogonality (52000,52000)
NX-TSH | XY Orthogonality (0,0)
NX-TSH | X axis OPM (-52000,-52000)
NX-TSH | X axis OPM (-52000,52000)
NX-TSH | X axis OPM (52000,-52000)
NX-TSH | X axis OPM (52000,52000)
NX-TSH | X axis OPM (0,0)
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-TSH | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-TSH | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-TSH | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-TSH | XY stage calibration
NX-TSH | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-TSH | LD beam power at the end of the lens
NX-TSH | Y Non-lineairty
NX-TSH | X Non-lineairty
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-TSH | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-TSH | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-TSH | UpperLimit
NX-TSH | Under 50nm Step Height Sample (10 Repeats)
NX-TSH | Overlay Camera X Position(ΔX)
NX-TSH | LowerLimit
NX-TSH | Initial Z
NX-TSH | 940nm Step Height Sample (10 Repeats)
NX-TSH | 50~200nm Step Height Sample (10 Repeats)
NX-TSH | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-TSH | StrokeYMin
NX-TSH | StrokeYMax
NX-TSH | StrokeY
NX-TSH | StrokeXMin
NX-TSH | StrokeXMax
NX-TSH | StrokeX
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-TSH | Z Drift (Short Traveling)
NX-TSH | 200~500nm Step Height Sample (10 Repeats)
NX-TSH | XY orthogonality
NX-TSH | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-TSH | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-TSH | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-TSH | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-TSH | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-TSH | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-TSH | Y1 Detector noise (XY Servo off)
NX-TSH | X1 Detector noise (XY Servo off)
NX-TSH | Maximum Y Scan range (Full Stroke)
NX-TSH | Maximum X Scan range (Full Stroke)
NX-TSH | Full range of Y2 Detector raw
NX-TSH | Full range of Y1 Detector raw
NX-TSH | Full range of X2 Detector raw
NX-TSH | Full range of X1 Detector raw
NX-TSH | Z Linearity by Z Scanner position
NX-TSH | -Y Direction software Limit
NX-TSH | -X Direction software Limit
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-TSH | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-TSH | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-TSH | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-TSH | %GRR w/ roughness of non-patterened wafer
NX-TSH | Z Scanner position at -6μm
NX-TSH | Z Scanner position at +6μm
NX-TSH | Approach position at -5μm
NX-TSH | Approach position at +5μm
NX-TSH | * 1800nm Step Height Sample (10 Repeats)
NX-TSH | Full range of Z Detector raw
NX-TSH | Approach position at -3μm
NX-TSH | Approach position at +3μm
NX-TSH | Under 50nm Step Height Sample (10 Repeats)
NX-TSH | 50~200nm Step Height Sample (10 Repeats)
NX-TSH | 200~500nm Step Height Sample (10 Repeats)
NX-TSH | PosMin
NX-TSH | PosMax
NX-TSH | MaxLength
NX-TSH | FullRecoveryProfileLength
NX-TSH | Full range of Y Detector raw
NX-TSH | Full range of X Detector raw
NX-TSH | EFM Phase : For the positive bias
NX-TSH | EFM Phase : For the negative bias
NX-TSH | Z Linearity by 5 different sample height
NX-TSH | Laser intensity with AC160 cantilever
NX-TSH | A－B difference during 3 times run test
NX-TSH | A＋B difference during 3 times run test
NX-TSH | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-TSH | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-TSH | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-TSH | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-TSH | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-TSH | Allowable max warpage, w/ SiC sample chuck
NX-TSH | RMS of Z Height of Wafer Zeroscan (Top)
NX-TSH | RMS of Z Height of Wafer Zeroscan (Right)
NX-TSH | RMS of Z Height of Wafer Zeroscan (Left)
NX-TSH | RMS of Z Height of Wafer ZeroScan (Center)
NX-TSH | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-TSH | PosMin
NX-TSH | PosMax
NX-TSH | MaxLength
NX-TSH | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-TSH | Check sample surface mapping
NX-TSH | Z Direction, Accuracy Test (Bias not significant)
NX-TSH | Y Direction, Accuracy Test (Bias not significant)
NX-TSH | X Direction, Accuracy Test (Bias not significant)
NX-TSH | Roughness Reproducibility Test (SR Panel)
NX-TSH | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-TSH | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-TSH | Roughness Reproducibility Test (SiNx Panel)
NX-TSH | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-TSH | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-TSH | Roughness Reproducibility Test (ABF Panel)
NX-TSH | Roughness Repeatability Test (SR Panel)
NX-TSH | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-TSH | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-TSH | Roughness Repeatability Test (SiNx Panel)
NX-TSH | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-TSH | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-TSH | Roughness Repeatability Test (ABF Panel)
NX-TSH | Lift Pin Stroke Min Check in DB
NX-TSH | Lift Pin Stroke Max Check in DB
NX-TSH | Limit range
NX-Wafer Basic | Z Linearity by Z scanner position
NX-Wafer Basic | Full range of Z Detector during Z Scan sweep
NX-Wafer Basic | Z Height calibration error
NX-Wafer Basic | Z Drive calibration error
NX-Wafer Basic | Z Linearity by 4 different sample height
NX-Wafer Basic | Full range of Z Height during Z Scan sweep
NX-Wafer Basic | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer Basic | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer Basic | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer Basic | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer Basic | Z Height noise
NX-Wafer Basic | Wafer placement(Loading) error on the load port1 (X offset)
NX-Wafer Basic | Wafer placement(Loading) error on the load port1 (Y offset)
NX-Wafer Basic | Wafer placement(Loading) error on the load port2 (X offset)
NX-Wafer Basic | Wafer placement(Loading) error on the load port2 (Y offset)
NX-Wafer Basic | OpticalAxisShiftPerFocusX
NX-Wafer Basic | OpticalAxisShiftPerFocusY
NX-Wafer Basic | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ)
NX-Wafer Basic | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ)
NX-Wafer Basic | Lift pins synchronization offset check (X offset)
NX-Wafer Basic | Lift pins synchronization offset check (Y offset)
NX-Wafer Basic | Profiler noise @200um
NX-Wafer Basic | Roll motion to X+ full scan
NX-Wafer Basic | Position repeatability(10repeats, Position Based)
NX-Wafer Basic | OPM Repeatability @5mm
NX-Wafer Basic | OPM Repeatability @1mm
NX-Wafer Basic | Max OPM @5mm
NX-Wafer Basic | Max OPM @1mm
NX-Wafer Basic | Torque Profile (10mm, 200um/s)
NX-Wafer Basic | Torque Profile (10mm, 20um/s)
NX-Wafer Basic | Lateral(C-D) sensitivity
NX-Wafer Basic | Lateral(C-D) noise
NX-Wafer Basic | Min. Vibration Transmissbility (~100Hz)
NX-Wafer Basic | RMS of ZeroScan
NX-Wafer Basic | LightStrengthGain
NX-Wafer Basic | Full range of X Detector during X Scan sweep
NX-Wafer Basic | Full range of X1 Detector during X Scan sweep
NX-Wafer Basic | Full range of X2 Detector during X Scan sweep
NX-Wafer Basic | Full range of Y Detector during Y Scan sweep
NX-Wafer Basic | Full range of Y1 Detector during Y Scan sweep
NX-Wafer Basic | Full range of Y2 Detector during Y Scan sweep
NX-Wafer Basic | Full range of X1 Detector Raw during X Scan sweep
NX-Wafer Basic | Full range of X2 Detector Raw during X Scan sweep
NX-Wafer Basic | Full range of Y1 Detector Raw during Y Scan sweep
NX-Wafer Basic | Full range of Y2 Detector Raw during Y Scan sweep
NX-Wafer Basic | X Detector stroke calibration error (nominal scale error X)
NX-Wafer Basic | Y Detector stroke calibration error (nominal scale error Y)
NX-Wafer Basic | X1 Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Wafer Basic | X2 Detector calibration value (AuxMicrometerPerAdc in XDetector DB)
NX-Wafer Basic | Y1 Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Wafer Basic | Y2 Detector calibration value (AuxMicrometerPerAdc in YDetector DB)
NX-Wafer Basic | X Scanner calibration value (MicrometerPerDacHigh in XScanner DB)
NX-Wafer Basic | Y Scanner calibration value (MicrometerPerDacHigh in YScanner DB)
NX-Wafer Basic | X Non-linearity error
NX-Wafer Basic | Y Non-linearity error
NX-Wafer Basic | X axis OPM
NX-Wafer Basic | Y axis OPM
NX-Wafer Basic | X Ringing(Rpv)
NX-Wafer Basic | Y Ringing(Rpv)
NX-Wafer Basic | X Detector noise (XY Servo on)
NX-Wafer Basic | Y Detector noise (XY Servo on)
NX-Wafer Basic | X1 Detector noise (XY Servo on)
NX-Wafer Basic | Y1 Detector noise (XY Servo on)
NX-Wafer Basic | X2 Detector noise (XY Servo on)
NX-Wafer Basic | Y2 Detector noise (XY Servo on)
NX-Wafer Basic | RMS error after Stage Calibration
NX-Wafer Basic | Image offset after moving from -X to +X (X scan)
NX-Wafer Basic | Image offset after moving from -Y to +Y (Y scan)
NX-Wafer Basic | Image offset after moving from -X to +X (Y scan)
NX-Wafer Basic | Image offset after moving from -Y to +Y (X scan)
NX-Wafer Basic | Image offset after moving from +X to -X (X scan)
NX-Wafer Basic | Image offset after moving from +Y to -Y (Y scan)
NX-Wafer Basic | Image offset after moving from +X to -X (Y scan)
NX-Wafer Basic | Image offset after moving from +Y to -Y (X scan)
NX-Wafer Basic | XY Stage positioning X offset (Vision Pattern Matching)
NX-Wafer Basic | XY Stage positioning Y offset (Vision Pattern Matching)
NX-Wafer Basic | XY Stage positioning X 1σ (Vision Pattern Matching)
NX-Wafer Basic | XY Stage positioning Y 1σ (Vision Pattern Matching)
NX-Wafer Basic | XY Stage positioning X offset (In-Die Align)
NX-Wafer Basic | XY Stage positioning Y offset (In-Die Align)
NX-Wafer Basic | XY Stage positioning X 1σ (In-Die Align)
NX-Wafer Basic | XY Stage positioning Y 1σ (In-Die Align)
NX-Wafer Basic | XY Stage positioning X offset (LLC Translation)
NX-Wafer Basic | XY Stage positioning Y offset (LLC Translation)
NX-Wafer Basic | XY Stage positioning X 1σ (LLC Translation)
NX-Wafer Basic | XY Stage positioning Y 1σ (LLC Translation)
NX-Wafer Basic | XY Stage positioning X offset (Global Align)
NX-Wafer Basic | XY Stage positioning Y offset (Global Align)
NX-Wafer Basic | XY Stage positioning X 1σ (Global Align)
NX-Wafer Basic | XY Stage positioning Y 1σ (Global Align)
NX-Wafer Basic | Z Drift
NX-Wafer Basic | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | Cycle time for 15-point measurement
NX-Wafer Basic | Short-term stability using Step Height Sample (3σ)
NX-Wafer Basic | Roughness Repeatability 1x1um(RSD%)
NX-Wafer Basic | Roughness Repeatability 1x1um(1σ)
NX-Wafer Basic | Roughness Repeatability 80x80um(RSD%)
NX-Wafer Basic | Roughness Repeatability 80x80um(1σ)
NX-Wafer Basic | Z Height calibrtaion error
NX-Wafer Basic | Z Drive calibrtaion error
NX-Wafer Basic | Z Linearity by 3 different sample height
NX-Wafer Basic | Full range in voltage during Z Scan sweep
NX-Wafer Basic | Full range in um during Z Scan sweep
NX-Wafer Basic | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer Basic | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer Basic | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer Basic | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer Basic | Lateral(C-D) sensitivity
NX-Wafer Basic | Lateral(C-D) noise
NX-Wafer Basic | A-B noise
NX-Wafer Basic | NCM Amplitude noise
NX-Wafer Basic | RMS of Z Height of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Wafer Basic | RMS of Z Drive of ZeroScan (XY/Z Range 100%, XY Servo on)
NX-Wafer Basic | RMS of Z Height of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Wafer Basic | RMS of Z Drive of ZeroScan (XY/Z Range 20%, XY Servo off)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(0,120000)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(-120000,0)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(0,-120000)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(120000,0)
NX-Wafer Basic | RMS of Z Height of Wafer ZeroScan (0,0)
NX-Wafer Basic | Max Wafer Zeroscan-Min Wafer Zeroscan
NX-Wafer Basic | XY Orthogonality
NX-Wafer Basic | MicrometerPerAdc of Z Height (Z Height calibration)
NX-Wafer Basic | MicrometerPerDacHigh of Z Scanner (Z Scanner calibration)
NX-Wafer Basic | Z linearity by sample height
NX-Wafer Basic | 1) Z Scanner position at +3μm
NX-Wafer Basic | 3) Z Scanner position at -3μm
NX-Wafer Basic | MicrometerPerAdc of X Detector at Large Mode (X Detector calibration)
NX-Wafer Basic | MicrometerPerAdc of Y Detector at Large Mode (Y Detector calibration)
NX-Wafer Basic | X Scan Stroke
NX-Wafer Basic | Y Scan Stroke
NX-Wafer Basic | Full range of X Detector Raw during X Scan sweep
NX-Wafer Basic | Full range of Y Detector Raw during Y Scan sweep
NX-Wafer Basic | X Detector noise (XY Servo off)
NX-Wafer Basic | Y Detector noise (XY Servo off)
NX-Wafer Basic | X axis OPM
NX-Wafer Basic | Y axis OPM
NX-Wafer Basic | X Non-linearity
NX-Wafer Basic | Y Non-linearity
NX-Wafer Basic | OPM Repeatability @25mm
NX-Wafer Basic | OPM Repeatability @10mm
NX-Wafer Basic | Max OPM @25mm
NX-Wafer Basic | Max OPM @10mm
NX-Wafer Basic | Torque Profile (50mm, 200um/s)
NX-Wafer Basic | Torque Profile (50mm, 20um/s)
NX-Wafer Basic | SLD current limit check
NX-Wafer Basic | SLD beam power at the end of the lens
NX-Wafer Basic | 2) of 20nm
NX-Wafer Basic | X Detector noise (XYServo off)
NX-Wafer Basic | Y Detector noise (XYServo off)
NX-Wafer Basic | X axis OPM raw Rpv
NX-Wafer Basic | Y axis OPM raw Rpv
NX-Wafer Basic | StageSafetyXMax
NX-Wafer Basic | StageSafetyXMin
NX-Wafer Basic | StageSafetyY
NX-Wafer Basic | StageSafetyYMax
NX-Wafer Basic | StageSafetyYMin
NX-Wafer Basic | Check X axis upper limit sensor function
NX-Wafer Basic | Check X axis lower limit sensor function
NX-Wafer Basic | Check Y axis upper limit sensor function
NX-Wafer Basic | Check Y axis lower limit sensor function
NX-Wafer Basic | X axis Angle Range
NX-Wafer Basic | Short-term stability w/ 450nm Step Height Sample (3σ)
NX-Wafer Basic | X detector noise (XY Servo on)
NX-Wafer Basic | Y detector noise (XY Servo on)
NX-Wafer Basic | Full range in um during X Scan sweep (XY Servo off)
NX-Wafer Basic | Full range in um during Y Scan sweep (XY Servo off)
NX-Wafer Basic | X Detector calibration value (MicrometerPerAdc in XDetector DB)
NX-Wafer Basic | Y Detector calibration value (MicrometerPerAdc in YDetector DB)
NX-Wafer Basic | X axis, Repeatability @1st position
NX-Wafer Basic | X axis, Repeatability @ 2nd position
NX-Wafer Basic | X axis, Repeatability @ 3rd position
NX-Wafer Basic | X axis, Repeatability @ 4th position
NX-Wafer Basic | X axis, Accuracy @1st position
NX-Wafer Basic | X axis, Accuracy @ 2nd position
NX-Wafer Basic | X axis, Accuracy @ 3rd position
NX-Wafer Basic | X axis, Accuracy @ 4th position
NX-Wafer Basic | Y axis, Repeatability @1st position
NX-Wafer Basic | Y axis, Repeatability @ 2nd position
NX-Wafer Basic | Y axis, Repeatability @ 3rd position
NX-Wafer Basic | Y axis, Repeatability @ 4th position
NX-Wafer Basic | Y axis, Accuracy @1st position
NX-Wafer Basic | Y axis, Accuracy @ 2nd position
NX-Wafer Basic | Y axis, Accuracy @ 3rd position
NX-Wafer Basic | Y axis, Accuracy @ 4th position
NX-Wafer Basic | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, 3σ)
NX-Wafer Basic | AFM, Static Repeatability w/ 1800nm Step Height Sample (10 repeats, RSD%)
NX-Wafer Basic | Z-XY Orthogonality (X direction-Left)
NX-Wafer Basic | Z-XY Orthogonality (X direction-Right)
NX-Wafer Basic | Z-XY Orthogonality (Y direction-Left)
NX-Wafer Basic | Z-XY Orthogonality (Y direction-Right)
NX-Wafer Basic | X-Scan Repeatability (10 repeats, RSD%)
NX-Wafer Basic | Y-Scan Repeatability (10 repeats, RSD%)
NX-Wafer Basic | Mechanical Align Accuracy
NX-Wafer Basic | Check the Multi Peattern Upload
NX-Wafer Basic | X Detector noise (XY servo off)
NX-Wafer Basic | Y Detector noise (XY servo off)
NX-Wafer Basic | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer Basic | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer Basic | Z Detector Max
NX-Wafer Basic | Z Detector Min
NX-Wafer Basic | WLI_Scale(Z Calibration)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(0,90000)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(-90000,0)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(0,-90000)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(90000,0)
NX-Wafer Basic | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | Short-term stability w/ 180nm Step Height Sample (3σ)
NX-Wafer Basic | Laser intensity with AC160TS cantilever
NX-Wafer Basic | Check if the driving % changes proportionally as the NCM amplitude changes
NX-Wafer Basic | NCM driving % at an amplitude of 1 nm
NX-Wafer Basic | NCM driving % at an amplitude of 20 nm
NX-Wafer Basic | NCM driving % at an amplitude of 100 nm
NX-Wafer Basic | - 940nm Step Height Sample (10 Repeats)
NX-Wafer Basic | - 200~500nm Step Height Sample (10 Repeats)
NX-Wafer Basic | - 50~200nm Step Height Sample (10 Repeats)
NX-Wafer Basic | - Under 50nm Step Height Sample (10 Repeats)
NX-Wafer Basic | A-B after Auto PSPD Alignment
NX-Wafer Basic | NCM driving% with AC160TS cantilever
NX-Wafer Basic | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB)
NX-Wafer Basic | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB)
NX-Wafer Basic | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB)
NX-Wafer Basic | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB)
NX-Wafer Basic | X Detector stroke calibration error (nominal scale error X)
NX-Wafer Basic | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | Resonance Frequency in DB
NX-Wafer Basic | Servo Cut-off Frequency in DB
NX-Wafer Basic | Check chuck surface no damage
NX-Wafer Basic | MicrometerPerAdc of X1 Detector at Large Mode (X1 Detector calibration)
NX-Wafer Basic | MicrometerPerAdc of X2 Detector at Large Mode (X2 Detector calibration)
NX-Wafer Basic | MicrometerPerAdc of Y1 Detector at Large Mode (Y1 Detector calibration)
NX-Wafer Basic | MicrometerPerAdc of Y2 Detector at Large Mode (Y2 Detector calibration)
NX-Wafer Basic | X1 Detector noise (XY servo off)
NX-Wafer Basic | X2 Detector noise (XY servo off)
NX-Wafer Basic | Y1 Detector noise (XY servo off)
NX-Wafer Basic | Y2 Detector noise (XY servo off)
NX-Wafer Basic | X Detector noise (XY servo on)
NX-Wafer Basic | X1 Detector noise (XY servo on)
NX-Wafer Basic | X2 Detector noise (XY servo on)
NX-Wafer Basic | Y Detector noise (XY servo on)
NX-Wafer Basic | Y1 Detector noise (XY servo on)
NX-Wafer Basic | Y2 Detector noise (XY servo on)
NX-Wafer Basic | AFP, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ)
NX-Wafer Basic | AFP, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%)
NX-Wafer Basic | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer Basic | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer Basic | Z Detector offset
NX-Wafer Basic | Check the full stroke image(servo on)
NX-Wafer Basic | ResonanceFrequencyHz
NX-Wafer Basic | ServoCutoffFrequencyHz
NX-Wafer Basic | FullPixelHeight
NX-Wafer Basic | Serial Port#
NX-Wafer Basic | Off-axis Camera X Position(ΔX)
NX-Wafer Basic | Roughness Repeatability 1x1um(10Repeats, RSD%)
NX-Wafer Basic | Roughness Repeatability 80x80um(10Repeats, RSD%)
NX-Wafer Basic | Z Scanner position at -3μm
NX-Wafer Basic | Z Scanner position at +3μm
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer Basic | Scan offset accuracy (Precision)
NX-Wafer Basic | Aligned stage accuracy (X offset-Point Align1)
NX-Wafer Basic | Aligned stage accuracy (Y offset-Point Align1)
NX-Wafer Basic | Aligned stage accuracy (X offset-Point Align2)
NX-Wafer Basic | Aligned stage accuracy (Y offset-Point Align2)
NX-Wafer Basic | Aligned stage accuracy (X offset-Point Align3)
NX-Wafer Basic | Aligned stage accuracy (Y offset-Point Align3)
NX-Wafer Basic | Aligned stage accuracy (X offset)
NX-Wafer Basic | Aligned stage accuracy (Y offset)
NX-Wafer Basic | Imag offset after EFEM loading (X scan)
NX-Wafer Basic | Imag offset after EFEM loading (Y scan)
NX-Wafer Basic | Imag offset after Lock stage (X scan)
NX-Wafer Basic | Imag offset after Lock stage (Y scan)
NX-Wafer Basic | Mask Roughness Repeatability (10Repeats, RSD%)
NX-Wafer Basic | Mask Roughness Repeatability (10Repeats, σ )
NX-Wafer Basic | Mask exchange time
NX-Wafer Basic | WLI_Scale(C Calibration)
NX-Wafer Basic | WLI_Scale(B Calibration)
NX-Wafer Basic | WLI_Scale(A Calibration)
NX-Wafer Basic | Short-term Static repeatability (RSD): WLI Mode (White light Interferometer)
NX-Wafer Basic | Short-term Static repeatability (3σ): WLI Mode (White light Interferometer)
NX-Wafer Basic | Port2) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Wafer Basic | Port1) Specification of Step Height Sample for WLI Repeatability test : 1σ, 10 repeat
NX-Wafer Basic | Port 1 lens(to OnAxis) Y Offset
NX-Wafer Basic | Port 1 lens(to OnAxis) X Offset
NX-Wafer Basic | P Gain
NX-Wafer Basic | I Gain
NX-Wafer Basic | D Gain
NX-Wafer Basic | Check the Port 2 lens light strength
NX-Wafer Basic | Check the Port 1 lens light strength
NX-Wafer Basic | Check the focus deviation between the lenses(Δf)
NX-Wafer Basic | Check the 25mm Moving – Y Axis Offset Check
NX-Wafer Basic | C Scanner Detector Min
NX-Wafer Basic | C Scanner Detector Max
NX-Wafer Basic | B Scanner Detector Min
NX-Wafer Basic | B Scanner Detector Max
NX-Wafer Basic | Angle align between Port 2 lens and VLSI SHS: Y
NX-Wafer Basic | Angle align between Port 2 lens and VLSI SHS: X
NX-Wafer Basic | Angle align between Port 2 lens and Sample chuck: Y
NX-Wafer Basic | Angle align between Port 2 lens and Sample chuck: X
NX-Wafer Basic | Angle align between Port 1 lens and Sample chuck: Y
NX-Wafer Basic | Angle align between Port 1 lens and Sample chuck: X
NX-Wafer Basic | A Scanner Detector Min
NX-Wafer Basic | A Scanner Detector Max
NX-Wafer Basic | Y2 Detector noise (XY Servo off)
NX-Wafer Basic | X2 Detector noise (XY Servo off)
NX-Wafer Basic | Y axis OPM (-52000,-52000)
NX-Wafer Basic | Y axis OPM (-52000,52000)
NX-Wafer Basic | Y axis OPM (52000,-52000)
NX-Wafer Basic | Y axis OPM (52000,52000)
NX-Wafer Basic | Y axis OPM (0,0)
NX-Wafer Basic | XY Orthogonality (-52000,-52000)
NX-Wafer Basic | XY Orthogonality (-52000,52000)
NX-Wafer Basic | XY Orthogonality (52000,-52000)
NX-Wafer Basic | XY Orthogonality (52000,52000)
NX-Wafer Basic | XY Orthogonality (0,0)
NX-Wafer Basic | X axis OPM (-52000,-52000)
NX-Wafer Basic | X axis OPM (-52000,52000)
NX-Wafer Basic | X axis OPM (52000,-52000)
NX-Wafer Basic | X axis OPM (52000,52000)
NX-Wafer Basic | X axis OPM (0,0)
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(-60000,0)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(60000,0)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(0,-60000)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan(0,60000)
NX-Wafer Basic | XY stage calibration
NX-Wafer Basic | Roughness Repeatability 1x1um(10 Repeats, RSD% )
NX-Wafer Basic | LD beam power at the end of the lens
NX-Wafer Basic | Y Non-lineairty
NX-Wafer Basic | X Non-lineairty
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(-60000, -60000)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(-60000, 60000)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(60000, -60000)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(60000, 60000)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(-60000, 0)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(60000, 0)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(0, -60000)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(0, 60000)
NX-Wafer Basic | RMS of Z Height of Mask Sample Zeroscan(0, 0)
NX-Wafer Basic | Max Wafer Zeroscan-Min Mask Sample Zeroscan
NX-Wafer Basic | UpperLimit
NX-Wafer Basic | Under 50nm Step Height Sample (10 Repeats)
NX-Wafer Basic | Overlay Camera X Position(ΔX)
NX-Wafer Basic | LowerLimit
NX-Wafer Basic | Initial Z
NX-Wafer Basic | 940nm Step Height Sample (10 Repeats)
NX-Wafer Basic | 50~200nm Step Height Sample (10 Repeats)
NX-Wafer Basic | RMS of Z Height of C122 Zeroscan(0,-90000)
NX-Wafer Basic | StrokeYMin
NX-Wafer Basic | StrokeYMax
NX-Wafer Basic | StrokeY
NX-Wafer Basic | StrokeXMin
NX-Wafer Basic | StrokeXMax
NX-Wafer Basic | StrokeX
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB)
NX-Wafer Basic | Z Drift (Short Traveling)
NX-Wafer Basic | 200~500nm Step Height Sample (10 Repeats)
NX-Wafer Basic | XY orthogonality
NX-Wafer Basic | Spring constant calibration by Thermal Tune Method with 3XC-NA-500DC
NX-Wafer Basic | Spring constant calibration by Thermal Tune Method with 3XC-NA-240AC
NX-Wafer Basic | Spring constant calibration by Thermal Tune Method with 3XC-NA-200AC
NX-Wafer Basic | Spring constant calibration by Sader Method with 3XC-NA-500DC
NX-Wafer Basic | Spring constant calibration by Sader Method with 3XC-NA-240AC
NX-Wafer Basic | Spring constant calibration by Sader Method with 3XC-NA-200AC
NX-Wafer Basic | Y1 Detector noise (XY Servo off)
NX-Wafer Basic | X1 Detector noise (XY Servo off)
NX-Wafer Basic | Maximum Y Scan range (Full Stroke)
NX-Wafer Basic | Maximum X Scan range (Full Stroke)
NX-Wafer Basic | Full range of Y2 Detector raw
NX-Wafer Basic | Full range of Y1 Detector raw
NX-Wafer Basic | Full range of X2 Detector raw
NX-Wafer Basic | Full range of X1 Detector raw
NX-Wafer Basic | Z Linearity by Z Scanner position
NX-Wafer Basic | -Y Direction software Limit
NX-Wafer Basic | -X Direction software Limit
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in YScanner DB)
NX-Wafer Basic | Closed loop scanning range (ServoScanRangeMicrometer in XScanner DB)
NX-Wafer Basic | %GRR w/ roughness of Precision/Tolerance (P/T) Reproducibility
NX-Wafer Basic | %GRR w/ roughness of Precision/Tolerance (P/T) Repeatability
NX-Wafer Basic | %GRR w/ roughness of non-patterened wafer
NX-Wafer Basic | Z Scanner position at -6μm
NX-Wafer Basic | Z Scanner position at +6μm
NX-Wafer Basic | Approach position at -5μm
NX-Wafer Basic | Approach position at +5μm
NX-Wafer Basic | * 1800nm Step Height Sample (10 Repeats)
NX-Wafer Basic | Full range of Z Detector raw
NX-Wafer Basic | Approach position at -3μm
NX-Wafer Basic | Approach position at +3μm
NX-Wafer Basic | Under 50nm Step Height Sample (10 Repeats)
NX-Wafer Basic | 50~200nm Step Height Sample (10 Repeats)
NX-Wafer Basic | 200~500nm Step Height Sample (10 Repeats)
NX-Wafer Basic | PosMin
NX-Wafer Basic | PosMax
NX-Wafer Basic | MaxLength
NX-Wafer Basic | FullRecoveryProfileLength
NX-Wafer Basic | Full range of Y Detector raw
NX-Wafer Basic | Full range of X Detector raw
NX-Wafer Basic | EFM Phase : For the positive bias
NX-Wafer Basic | EFM Phase : For the negative bias
NX-Wafer Basic | Z Linearity by 5 different sample height
NX-Wafer Basic | Laser intensity with AC160 cantilever
NX-Wafer Basic | A－B difference during 3 times run test
NX-Wafer Basic | A＋B difference during 3 times run test
NX-Wafer Basic | RMS of Z Drive of Wafer Zeroscan(-60000,0)
NX-Wafer Basic | RMS of Z Drive of Wafer Zeroscan(60000,0)
NX-Wafer Basic | RMS of Z Drive of Wafer Zeroscan(0,-60000)
NX-Wafer Basic | RMS of Z Drive of Wafer Zeroscan(0,60000)
NX-Wafer Basic | RMS of Z Drive of Wafer ZeroScan (0,0)
NX-Wafer Basic | Allowable max warpage, w/ SiC sample chuck
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan (Top)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan (Right)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan (Left)
NX-Wafer Basic | RMS of Z Height of Wafer ZeroScan (Center)
NX-Wafer Basic | RMS of Z Height of Wafer Zeroscan (Bottom)
NX-Wafer Basic | PosMin
NX-Wafer Basic | PosMax
NX-Wafer Basic | MaxLength
NX-Wafer Basic | Allowable max warpage, w/ SiC Anti-warpage sample chuck (Crying type)
NX-Wafer Basic | Check sample surface mapping
NX-Wafer Basic | Z Direction, Accuracy Test (Bias not significant)
NX-Wafer Basic | Y Direction, Accuracy Test (Bias not significant)
NX-Wafer Basic | X Direction, Accuracy Test (Bias not significant)
NX-Wafer Basic | Roughness Reproducibility Test (SR Panel)
NX-Wafer Basic | Roughness Reproducibility Test (Smooth Cu Panel, Trace)
NX-Wafer Basic | Roughness Reproducibility Test (Smooth Cu Panel, Plane)
NX-Wafer Basic | Roughness Reproducibility Test (SiNx Panel)
NX-Wafer Basic | Roughness Reproducibility Test (Rough Cu Panel, Trace)
NX-Wafer Basic | Roughness Reproducibility Test (Rough Cu Panel, Plane)
NX-Wafer Basic | Roughness Reproducibility Test (ABF Panel)
NX-Wafer Basic | Roughness Repeatability Test (SR Panel)
NX-Wafer Basic | Roughness Repeatability Test (Smooth Cu Panel, Trace)
NX-Wafer Basic | Roughness Repeatability Test (Smooth Cu Panel, Plane)
NX-Wafer Basic | Roughness Repeatability Test (SiNx Panel)
NX-Wafer Basic | Roughness Repeatability Test (Rough Cu Panel, Trace)
NX-Wafer Basic | Roughness Repeatability Test (Rough Cu Panel, Plane)
NX-Wafer Basic | Roughness Repeatability Test (ABF Panel)
NX-Wafer Basic | Lift Pin Stroke Min Check in DB
NX-Wafer Basic | Lift Pin Stroke Max Check in DB
NX-Wafer Basic | Limit range